33876 | 1220648-2023-03825 | 17884400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884400&pc=OZD | 10/20/2021 | 09/18/2024 |  | 10/06/2023 |  |  | 10/20/2021 | 10/06/2023 |  |  | 07/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022119333 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/20/2021 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a stroke. as a result of the noted condition, the decision was made to withdraw care. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported stroke/cva has been completed since the original report was filed. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | B2: date of death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03825 was provided.
33877 | 1220648-2023-03826 | 17884422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884422&pc=OZD | 10/21/2021 | 01/26/2024 |  | 10/06/2023 |  |  | 10/21/2021 | 10/06/2023 |  |  | 07/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022119995 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2021 |  |  |  |  | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the implanted 5.5 impella pump was found to be stuck in the papillary muscle. as a result of the position, support could not be raised above p4 without triggering suction alarms. repositioning was recommended, but the staff declined and stated that the pump would be pulled later that night. | [5 empty] | The investigation for the reported positioning issues has been completed. the impella was not returned for evaluation. the root cause of the positioning issue was most likely use related, echo showed 5.5 in the papillary muscle unable to go higher as per the clinical description provided and there was low pump flow with no spike/trend in the motor current in the data logs. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33878 | 1220648-2023-03827 | 17884499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884499&pc=OZD | 10/22/2021 | 01/12/2024 |  | 10/06/2023 |  |  | 10/22/2021 | 10/06/2023 |  |  | 04/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022084103 |  | 0550-008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2021 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced oozing/bleeding from their access site during impella support. the dressing was changed and that patient was given an unspecified quantity of blood products and platelets to counteract the condition. support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the introducer damage ¿ mechanical has been completed. the product was not returned. as per the clinical information provided, the root cause of the access site bleeding for the patient was patient condition. it was reported that the oozing of blood was observed at the access site and the low platelet count is the cause for oozing that identifies as the potential cause for oozing. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33879 | 1220648-2023-03828 | 17884597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884597&pc=OZD | 11/04/2021 | 03/01/2024 |  | 10/06/2023 |  |  | 11/04/2021 | 10/06/2023 |  |  | 08/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022123482 |  | 0550-008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2021 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that support, the patient experienced a post-operative stroke with left sided hemiparesis. the patient was alert and oriented after this event. | [5 empty] | The investigation into the stroke/cva has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33880 | 1220648-2023-03829 | 17884685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884685&pc=OZD | 12/27/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 12/27/2021 | 10/06/2023 |  |  | 09/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143564 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2021 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed heparin-induced thrombocytopenia during impella support. as a result of the noted condition, the decision was made to explant the pump. the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombocytopenia issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33881 | 1220648-2023-03830 | 17884722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884722&pc=OZD | 01/08/2022 | 01/10/2024 |  | 10/06/2023 |  |  | 01/08/2022 | 10/06/2023 |  |  | 10/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022145487 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a male patient noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump came out of position when the patient pulled on the line during support. after several failed attempts to reposition the pump, the decision was made to explant the device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was patient movement, as the patient pulled at the impella catheter and the device lost position as per the clinical description provided.
33882 | 1220648-2023-03831 | 17884810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884810&pc=OZD | 12/01/2022 | 10/06/2023 |  | 10/06/2023 |  |  | 12/01/2022 | 10/06/2023 |  |  | 10/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023263201 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 42yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a bleed from a suture line from a newly implanted device. the bleeding was managed with a suture. the patient received a blood transfusion due to renal issue. there was no lasting patient harm reported. | [5 empty] | The investigation for the access site bleed has been completed. the root cause of the access site bleeding was unable to be determined as limited clinical information and no product was provided for evaluation. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33883 | 1220648-2023-03832 | 17884825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884825&pc=OZD | 02/10/2023 | 10/06/2023 |  | 10/06/2023 |  |  | 02/10/2023 | 10/06/2023 |  |  | 08/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022138988 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2023 |  |  |  |  | The user facility reported that a 44-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows despite running the pump at high p-levels. additionally, the aortic (ao) and left ventricular (lv) signals decoupled. if the p-level was increased above 5, suction alarms would be present. upon pump explant, a large clot was visualized on the inlet of the device. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. analysis of the data logs did not reveal any performance issues. the device was returned. it was confirmed there were suction alarms present based on the rt logs. the product evaluation showed that there was a large clot on the inlet of the device. the root cause of the low pump flow is most likely biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33884 | 1220648-2023-03833 | 17884841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884841&pc=OZD | 12/27/2022 | 10/06/2023 |  | 10/06/2023 |  |  | 12/27/2022 | 10/06/2023 |  |  | 10/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023272171 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2022 |  |  |  |  | The user facility reported that a 39-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed a tear in the sterile sleeve on the 18th day of support. of note, the sterile sleeve was reported to have been torn after multiple pump repositions. there was no reported harm to the patient. | [5 empty] | The investigation into the reported sterile sleeve damage was completed. the device was not returned for evaluation. based on the available information the root cause of the sterile sleeve damage was determined to be most likely due to excessive force from repositioning the pump multiple times. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33885 | 1220648-2023-03834 | 17884856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884856&pc=OZD | 02/26/2022 | 03/10/2025 |  | 10/06/2023 |  |  | 02/26/2022 | 10/06/2023 |  |  | 10/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022155778 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2022 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding at their graft anastomosis during patient support. bleeding was first seen when trimming the graft and the clamp was moved closer to the anastomosis to control the condition. upon further observation of the bleed, the sutures were reinforced and pressure was held. the patient was given two units of packed red blood cells to counteract the blood loss. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was anticoagulation management since act were 253 which cross the range of 160-180. this report is being filed as part of a retrospective review of historical records.
33886 | 1220648-2023-03835 | 17884866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884866&pc=OZD | 12/18/2022 | 10/06/2023 |  | 10/06/2023 |  |  | 12/18/2022 | 10/06/2023 |  |  | 11/30/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023275534 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/18/2022 |  |  |  |  | The user facility reported that a 44-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed a hole in the catheter. the blood began to leak from the red impella plug following closure of right axillary pocket. the pump was subsequently replaced the next day. there was no reported harm to the patient. | [5 empty] | The investigation into the reported hole in catheter was completed. the device was returned for evaluation. visual inspection of the returned product revealed blood leaking from the red handle. a catheter puncture was also found and identified as the entry point for the blood. based on the available information, the root cause of the catheter puncture was likely due to a needle used by the operator while securing the pump. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33887 | 1220648-2023-03836 | 17884877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884877&pc=OZD | 07/12/2023 | 10/06/2023 |  | 10/06/2023 |  |  | 07/12/2023 | 10/06/2023 |  |  | 02/28/2025 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311616 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/12/2023 |  |  |  |  | The user facility reported that a 57-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced thrombus. the patients lactate dehydrogenase (ldh) levels were rising, and the device was not in the optimal position. the physician had concerns for thrombus developing and the patient still needed another two weeks of support. the physician was unable to reposition the device due to the new locking mechanism. as a result, the device was exchanged. there was no reported harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. visual inspection of the device revealed a clot. based on the available, the root cause of the thrombus was due to the physician's inability to reposition the device. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33888 | 1220648-2023-03837 | 17884904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884904&pc=OZD | 11/02/2022 | 11/19/2024 |  | 10/06/2023 |  |  | 11/02/2022 | 10/06/2023 |  |  | 04/30/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023230537 |  | 0550-0006 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The user facility in the eu reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. pump stop after 33 days and exchange by new 5.5 pump. patient stable. | [5 empty] | The cause of the pump stop was not determined since the product, data logs, and sufficient clinical information were not provided. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the pump stop was not determined since the product, data logs, and sufficient clinical information were not provided. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33889 | 1220648-2023-03838 | 17884909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884909&pc=OZD | 04/25/2022 | 01/26/2024 |  | 10/06/2023 |  |  | 04/25/2022 | 10/06/2023 |  |  | 01/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023185162 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support pulled out the impella pump and was sent back to the operating room where a new impella was placed. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the positioning issue was most likely patient movement as it was reported that patient pulled out impella. | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33890 | 1220648-2023-03839 | 17884925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17884925&pc=OZD | 04/20/2023 | 10/06/2023 |  | 10/06/2023 |  |  | 04/20/2023 | 10/06/2023 |  |  | 12/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293239 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2023 |  |  |  |  | The user facility reported that a 59-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced sudden low pump flows. the physician had been concerned about possible thrombus ingestion related to acute suction events the day before. during the patient's explant procedure, a clot was noticed in the inflow. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. analysis of the data logs revealed one suction alarm which resolved two minutes later. there were no issues with the flow for entire case. based on clinical description & data review, the root cause of low pump flows and thrombus was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33891 | 1220648-2023-03840 | 17885326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885326&pc=OZD | 08/28/2021 | 10/21/2025 |  | 10/06/2023 |  |  | 08/28/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022102416 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/28/2021 |  |  |  |  | The user facility reported a 48-year-old male undergoing aortic valve surgey was implanted with an impella 5.5 device for mechanical circulatory support. the patient had an echocardiogram (ultrasound) performed after 13 days of support and showed a large thrombus in left ventricle just above the impella. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03840 was provided.
33892 | 1220648-2023-03841 | 17885349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885349&pc=OZD | 08/29/2021 | 01/11/2024 |  | 10/06/2023 |  |  | 08/29/2021 | 10/06/2023 |  |  | 05/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022084161 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2021 |  |  |  |  | The user facility reported a 57-year-old female was implanted with an impella 5.5 device for mechanical circulatory support in care escalation from a balloon pump (iabp). the patient was being worked up for either heart transplant or lvad device. the medical team had difficulty delivering the impella pump due to the patient's axillary artery which appeared to be too small. four days after the implant the patient experienced bleeding from the access site which required the patient to receive one unit of blood and go to the or for a washout. nine days after that, the patient experienced a fluid leak from the purge line. device later explanted. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding and purge leak has been completed since the original report was filed. access site bleeding- the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation. no product was returned. the root cause of the purge leak-pump issue was not determined since product was not returned, as there was a leak observed but not enough information is provided in the clinical description provided.
33893 | 1220648-2023-03842 | 17885420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885420&pc=OZD | 08/24/2021 | 10/28/2025 |  | 10/06/2023 |  |  | 08/24/2021 | 10/06/2023 |  |  | 05/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022086841 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Ventricular Fibrillation | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/24/2021 |  |  |  |  | The user facility reported a 82-year-old male was implanted with an impella 5.5 device for mechanical circulatory support, after care being escalated from the impella cp. the medical team had difficulty delivering the impella pump and after several attempts they were finally able to get it in the right position. the patient went to ventricular fibrillation and was given cardiopulmonary resuscitation however that did not restore the heart rate and decision was made to withdraw care. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that the medical team had difficulty delivering the impella pump and after several attempts they were finally able to get it in the right position. the patient went to vfib and was given cpr however that did not restore the heart rate, and decision was made to withdraw care. patient passed. |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03842 was provided in sections b2, b5, h1, and h6. | Delivery issue: the root cause of the delivery issue was most likely due to the use issue. vt/vfib: the root cause of the vf/vt was most likely patient condition related because the timing of the event correlated with placement/repositioning and any device contacting the heart wall in conjunction with a poor patient condition could result in vt/vf.
33894 | 1220648-2023-03843 | 17885530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885530&pc=OZD | 08/24/2021 | 01/26/2024 |  | 10/06/2023 |  |  | 08/24/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022102431 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2021 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had a chest x-ray performed which identified a hematoma at the access site. this required the patient to receive 2 units packed red blood cells given. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hematoma has been completed. without the impella device nor sufficient clinical data, the root cause of the access site bleeding was unable to be determined.
33895 | 1220648-2023-03844 | 17885567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885567&pc=OZD | 09/20/2021 | 03/04/2024 |  | 10/06/2023 |  |  | 09/20/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022109112 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2021 |  |  |  |  | The user facility reported a 64-year-old male admitted in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a subarachnoid hemorrhage during impella support. although there were no plans to treat the condition, the patient was switched to bicarb for the purge line. support with the implanted pump continued following the event, for just under 10 days. | [5 empty] | The investigation into the reported stroke/cva has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the stroke/cva was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33896 | 1220648-2023-03845 | 17885606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885606&pc=OZD | 09/23/2021 | 09/20/2024 |  | 10/06/2023 |  |  | 09/23/2021 | 10/06/2023 |  |  | 02/28/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022045681 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/23/2021 |  |  |  |  | The user facility reported a 73-year-old female with cardiogenic shock and cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump began triggering a suction alarm during patient support. an echocardiogram was completed and it was discovered that the pump was posterior in position. re-positioning was attempted, but the medical doctor was unable to get the pump mid-ventricle. the patient coded during the troubleshooting and subsequently passed away as they were do not resuscitate (dnr). a definitive relation between the pump and patient's death could not be ascertained. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. as per the clinical information provided, the pump was mispositioned and it was unable to deposition due to patient movement. the cause of the positioning issue was due to patient movement based on the clinical review. the failure modes will be monitored and trended.
33897 | 1220648-2023-03846 | 17885675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885675&pc=OZD | 08/30/2021 | 01/11/2024 |  | 10/06/2023 |  |  | 08/30/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022096670 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2021 |  |  |  |  | The user facility reported a 55-year-old male with multiple cardiac comorbidities was implanted with an impella 5.5 device for mechanical circulatory support while decision made for either angioplasty or bypass surgery. the patient on impella support experienced bleeding from the access site and surgery was required for the patient to reach hemostasis. the pump supported successfully for over 72 hours til wean and explanted. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the product was not returned. as per the clinical information provided, the root cause of the patient injury is patient condition. it was reported that there was bleeding all over that has occurred and there was no additional information that was provided that identifies as the potential cause for bleeding. bleeding from access sites doesn't relate to impella 5.5 since the access site is not exactly identified. hence, the root cause was patient condition and unknown relation to impella. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33898 | 1220648-2023-03847 | 17885719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885719&pc=OZD | 09/01/2021 | 03/01/2024 |  | 10/06/2023 |  |  | 09/01/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022092832 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2021 |  |  |  |  | The user facility reported a 64-year-old male with ischemic cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a hemorrhagic stroke. systemic heparin was discontinued and protamine given. the device explanted after 21 days of successful support. | [5 empty] | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33899 | 1220648-2023-03848 | 17885829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885829&pc=OZD | 09/07/2021 | 10/06/2023 |  | 10/06/2023 |  |  | 09/07/2021 | 10/06/2023 |  |  | 07/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113573 |  | 0550-008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Unknown |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/07/2021 |  |  |  |  | The user facility reported a 40-year-old patient suffering from covid 19, endocarditis, and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the medical staff had difficulty delivering the device via sponge plugs, but was eventually successful when placing via the axillary. the pump supported the patient for just under 10 days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. only the clinical report was used to determine root cause. the cause of the delivery issue was use related. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33900 | 1220648-2023-03849 | 17885883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885883&pc=OZD | 09/08/2021 | 10/28/2025 |  | 10/06/2023 |  |  | 09/08/2021 | 10/06/2023 |  |  |  |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022105419 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2021 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery/bypass (cabg) was implanted with an impella 5.5 device for mechanical circulatory support. placement completed without issue. the patient was taken to the or due to excessive bleeding in chest post cabg. during case wrap up, console began alarming that purge flow stopped. the staff checked all connections when it was discovered that the pressure reservoir developed a crack and fluid was leaking from device. or staff was able to successfully troubleshoot. upon arrival to icu unit, impella motor failed. device was restarted multiple times with no success. it was then decided to explanation device as patient's vitals were stable. | The complainant reported that a patient underwent axillary impella 5.5 placement for cabg scheduled for friday. placement completed without issue. the patient was taken to the or due to excessive bleeding in chest post cabg. during case wrap up, console began alarming that purge flow stopped. the staff checked all connections when it was discovered that the pressure reservoir developed a crack and fluid was leaking from device. or staff was able to successfully cut off the reservoir, insert a 19g needle into the side arm tubing, and connect the purge line directly to the needle. after securing the set up, impella was running with no alarms and patient was transported back to cvicu. upon arrival to unit, impella motor failed. device was restarted multiple times with no success. it was then decided to explanation device as patient's vitals were stable. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | Product damage - the cause of product damage cannot be determined since the product was not returned. pump stop - the cause of the pump stop is biomaterial since motor current above 1200 was observed in the lt logs.
33901 | 1220648-2023-03850 | 17885923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885923&pc=OZD | 09/08/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 09/08/2021 | 10/06/2023 |  |  | 04/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022072689 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2021 |  |  |  |  | The user facility reported a 48-year-old female with multiple cardiac and systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a low pump flow due to allegedly a kink in the line. the pump remained on for over 52 hours til weaned and explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. from the aic data, we can see that the reduction in the purge flow resulted in decline in the flow. there were no low flow alarms that were raised during the case and the pump was still functioning at the given limits. it can be clearly noted that the issue was resolved, and the pump functioning was stable later till the end of the case. the root cause was of the issue is use issue, since we can clearly see that the pump was stable and no it was not clearly not related to functioning or the patient condition that caused the issue. hence the root cause is use issue. this report is being filed as part of a retrospective review of historical records.
33902 | 1220648-2023-03851 | 17885995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17885995&pc=OZD | 09/15/2021 | 09/13/2024 |  | 10/06/2023 |  |  | 09/15/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022109115 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Death | Other |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/15/2021 |  |  |  |  | The user facility reported a 59-year-old male with multiple cardiac and systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed a large cerebral hemorrhage during impella support. as a result of the noted condition, the patient's family decided to withdraw care after 6+ days of support. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported stroke/cva has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the stroke/cva was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03851 was provided.
33903 | 1220648-2023-03852 | 17886028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17886028&pc=OZD | 09/27/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 09/27/2021 | 10/06/2023 |  |  | 01/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022034360 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2021 |  |  |  |  | The user facility reported a 58-year-old male had a prior impella 5.5 explanted in preparation for the heart transplant or lvad was implanted with an impella 5.5 device and ecmo for mechanical circulatory support when transplant and lvad were denied. a fluid leak was observed from the purge sidearm of the impella 5.5 pump during patient support, after a month. as the leak had not yet affected the functionality of the pump, the decision was made to monitor the issue and perform a purge sidearm bypass as needed. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported purge leak issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33904 | 1220648-2023-03853 | 17886082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17886082&pc=OZD | 10/01/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 10/01/2021 | 10/06/2023 |  |  | 07/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113191 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | IMPELLA RP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2021 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella 5.5 and impella rp devices for mechanical circulatory support needed biventricularly. the complainant reported that a patient on impella support was found to have thrombus when was being weaned off the impella 5.5 pump and a washout had to be performed. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33905 | 1220648-2023-03854 | 17886134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17886134&pc=OZD | 09/28/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 09/28/2021 | 10/06/2023 |  |  | 04/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022070081 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2021 |  |  |  |  | The user facility reported a 58-year-old male suffering from cardiogenic shock and cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that their bedside monitor/defibrillator had not been properly secured and fell from the tray onto the impella cable. the impella plug disconnected and resulted in a pump stop. the patient was subsequently taken to the operating room where the pump was removed and replaced with another impella 5.5 pump. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the pump stop issue was use related, as the impella cable got pulled apart causing the pump to stop as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
33906 | 1220648-2023-03855 | 17886292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17886292&pc=OZD | 09/09/2021 | 10/29/2025 |  | 10/06/2023 |  |  | 09/09/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098115 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/09/2021 |  |  |  |  | The user facility reported a 73-year-old male with biventricular failure was implanted with an impella 5.5 and rp for mechanical circulatory support. the patient on impella support experienced a low pump flow on the 5.5 pump, and imaging confirmed the presence of a thrombus in the left ventricle. device was explanted after 6 days. | Additional information noted care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03855 was provided. | D6a and d6b revised from the initial manufacturer device report 1220648-2023-03855 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03855 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03855 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03855. h6 investigation conclusions revised from the initial submission in accordance with updated procedures.
33907 | 1220648-2023-03856 | 17886354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17886354&pc=OZD | 09/09/2021 | 10/27/2025 |  | 10/06/2023 |  |  | 09/09/2021 | 10/06/2023 |  |  | 02/28/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022052197 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/09/2021 |  |  |  |  | The user facility reported a 65-year-old male had care escalated to the impella 5.5 from the cp due to cardiogenic shock and need for continued mechanical circulatory support. the medical team experienced difficulty delivering the pump to the correct position. after delivery the patient experienced low pump flow. patient had ecmo considered, but instead the care was withdrawn, per family choice after over 5 days of 5.5 support. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03856 was provided. | Positioning issue - the root cause of the positioning issue was most likely patient condition related, because it was mentioned that aorta did not have adequate space above valve and it was difficult to get optimal positioning even with transesophageal echocardiography guidance even when inserting from axillary artery. it was mentioned that patient had stenosis of innominate artery and axillary artery. low pump flow - the root cause of the issue was not determined because there were no sufficient details that were reported and the product and logs were not returned to determine the root cause. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
33908 | 1220648-2023-03857 | 17886485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17886485&pc=OZD | 09/15/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 09/15/2021 | 10/06/2023 |  |  | 03/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022058098 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Unspecified Nervous System Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2021 |  |  |  |  | The user facility reported a 65-year-old male undergoing a high risk surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had the plastic purge clip missing and side arm was stabilized with tongue depressor, gauze and tape. the same patient the following day experienced ventricular tachycardia¿¿and later was found to have a subdural bleed. the patient had a history of stroke and multiple falls. device later explanted after over 7 days of support. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vf/vt/af/at and packaging issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33909 | 1220648-2023-03858 | 17886520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17886520&pc=OZD | 10/05/2021 | 10/06/2023 |  | 10/06/2023 |  |  | 10/05/2021 | 10/06/2023 |  |  | 04/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022078126 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/05/2021 |  |  |  |  | The user facility reported a 64-year-old male who had decompensated post mitra-clip procedure was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 met resistance in the clavicular area during attempted delivery. multiple attempts were made until it was discovered that the right axillary/subclavian artery had a dissection. it was unknown when the dissection occurred. due to the noted issue, the delivery was aborted and the site was closed. no further intervention was completed for the dissection as the medical doctor hoped the blood flow would propagate the dissection. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. the root cause was determined to be vessel condition. instructions for use for the related event are as follows: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation...¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
33910 | 1220648-2023-03859 | 17887192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17887192&pc=OZD | 09/09/2023 | 09/25/2025 |  | 10/06/2023 |  |  | 09/09/2023 | 10/06/2023 |  |  | 02/28/2025 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Retroperitoneal Hemorrhage | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient supported by an impella cp had aic alarmed on/off impella in ventricle and impella in aorta alarms. placement verified by transthoracic echocardiogram (tte) two times and chest radiography. staff gave prbc and albumin. given pulsatile motor current, good flows, clear urine, and confirmation of impella with imaging, staff disabled the alarms and impella was running at p1 for lv vent. additionally, there were concerns for continued hemolysis via lab results and ongoing thrombocytopenia. the patient was on va-ecmo treatment so it is difficult to determine the cause of hemolysis. furthermore, the patient had retroperitoneal bleed (rp) bleeding, but it is also difficult to determine what specifically caused this as the patient had ecmo placed, an impella placed, and an hour of chest compressions. the patient remained on support for 9 days til family made choice to withdraw care rather than persue more intensive therapies. | [5 empty] | The investigation into the reported hemolysis and vascular damage has been completed since the original report was filed. no product was returned for evaluation. as per clinical review, several false impella in aorta and ventricle alarms were observed with good position verified. the patient was suspected to have hemolysis with rising ldh and the impella was pulled back. the patient had a gastrointestinal bleeding which was unknown if it was the impella related with extracorporeal membrane oxygenation (ecmo) being used. data logs were reviewed and several false impella position in aorta alarm, impella unknown alarms were found. the 5s parameters were variable with motor current spike and suction alarms observed due to likely biomaterial ingestion. several suction alarms were before the reported hemolysis event. saturated flows were also observed after motor current spike event. the cause of the hemolysis issue was most likely biomaterial with motor current spike and several suction alarms observed prior to reported hemolysis event per log and clinical review. the cause of the vascular damage - peripheral vessels was most likely patient condition related due to gastrointestinal bleeding observed and unknown relation to impella per clinical. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the cause of the optical signal issue was most likely biomaterial with motor current spike observed before several false "impella position in aorta" alarm per log and clinical review. all pertinent information available to abiomed has been submitted at this time. h6 updated to reflect optical signal findings. d10 concomitant medical products and therapy dates.
33911 | 1220648-2023-03861 | 17887470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17887470&pc=OZD | 04/03/2023 | 10/06/2023 |  | 10/06/2023 |  |  | 04/03/2023 | 10/06/2023 |  |  | 12/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024287866 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/03/2023 |  |  |  |  | The user facility reported that a 46-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows, and suction alarms with a low central venous pressure (cvp). echo results showed device was too deep. the physician repositioned however, flows remained low. the device was explanted after 90hours. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. analysis of the data logs confirmed low flows and several suction alarms after repositioning of the device. based on the available information, the root cause of the low flow issue was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33912 | 1220648-2023-03862 | 17887510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17887510&pc=OZD | 03/29/2023 | 10/06/2023 |  | 10/06/2023 |  |  | 03/29/2023 | 10/06/2023 |  |  | 01/31/2025 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024297587 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2023 |  |  |  |  | The user facility reported that a 48-year-old female patient implanted with the impella 5.5 for mechanical circulatory support developed thrombus. transthoracic echocardiogram (tte) revealed the distal end of device had clot formation. activated clotting time (act) was >300 and 1 dose of protamine was administered. repeat act 2 hours later was 160 so another dose of protamine was administered. the physician subsequently, weaned the device to p3 and monitored hemodynamics until explant. no further information was provided. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation. analysis of the data logs did not reveal any performance issues. based on the available information, the root cause of the reported event was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33913 | 1220648-2023-03860 | 17887516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17887516&pc=OZD | 09/09/2023 | 04/01/2024 |  | 10/06/2023 |  |  | 09/09/2023 | 10/06/2023 |  |  | 07/31/2025 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375612 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an impella 5.5 support ongoing for 16 days of the caridiomyopathy admission. the team noted on day 16 that there was high purge/purge system blocked. additionally, they noted prior there was a purge leak. there was no harm and the pump remained on for another 2 days despite the purge alarms. the patient then was explanted as bridged successfully to transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure. reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressure and the purge leak ¿ pump has been completed since the initial report was submitted. the product was not returned for investigation. the data log does not indicate a prior event of a purge leak. according to clinical records, the cause of the high purge pressure issue was most likely kinked purge line, based on data log review.
33914 | 1220648-2023-03863 | 17887865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17887865&pc=OZD | 03/04/2023 | 10/06/2023 |  | 10/06/2023 |  |  | 03/04/2023 | 10/06/2023 |  |  | 12/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024285018 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2023 |  |  |  |  | The user facility reported that a 50-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the impella catheter was still able to be moved despite the tuohy-borst being locked. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported positioning issue was completed. the device was not returned for evaluation. based on the available information, the root cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33915 | 1220648-2023-03864 | 17888056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17888056&pc=OZD | 03/02/2023 | 09/15/2024 |  | 10/06/2023 |  |  | 03/02/2023 | 10/06/2023 |  |  | 10/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283804 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported patient mitral valve disease was implanted with an impella cp device for mechanical circulatory support. the surgeon pulled back the impella into the aorta during bypass surgery and mitral valve repair. the physician was unable to advance impella across valve after repair was completed. decision was therefore made to take out impella cp and replace with a new impella. | It was reported that after meeting resistance while trying to reposition the impella after the mitral repair, the impella cp was removed and the patient was moved to comfort care only measures, and after implanted with a new impella cp. it was noted the patient expired while on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported positioning issues has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the positioning issue was patient condition related due to anatomical complication with resistance felt across aortic valve after mitral valve repair. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03864 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
33916 | 1220648-2023-03865 | 17888195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17888195&pc=OZD | 03/03/2023 | 09/18/2024 |  | 10/06/2023 |  |  | 03/03/2023 | 10/06/2023 |  |  | 12/31/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024286881 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Material Puncture/Hole | Material Integrity Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | The user facility reported that an 82-year-old male patient implanted with the impella 5.5 for mechanical circulatory support. the graft had been placed without the 23 fr sheath bevel being inserted and the device was unable to pass through the axillary artery. as a result, the physician decided to cannulate the graft with an 8 fr sheath to balloon the vessel which was unsuccessful. the impella sheath was removed, and a 24 fr gore dry sheath was inserted instead, and the device was successfully placed. the end of the gore sheath remained outside the graft. later that day, blood leaking from red plug shortly after patient arrived in intensive care unit, medical team thought likely the leakage was due to a difficult insertion with possible damage to catheter. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported catheter damage issue was completed. the device was returned for evaluation. visual inspection of the catheter revealed several catheter kinks and two holes at the 30 cm mark. no other abnormalities were observed. based on the available information, the root cause of the reported event was blood ingress due to a hole in the catheter resulting from improper use. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
33917 | 1220648-2023-03866 | 17888344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17888344&pc=OZD | 10/11/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 10/11/2021 | 10/06/2023 |  |  | 01/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022038055 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2021 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was found to have a thrombosis and had to have it surgically removed. after the procedure the patient returned to the icu on impella support. the support lasted for just under 6 days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33918 | 1220648-2023-03867 | 17888529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17888529&pc=OZD | 10/20/2021 | 01/26/2024 |  | 10/06/2023 |  |  | 10/20/2021 | 10/06/2023 |  |  | 07/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022120024 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2021 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella support experienced multiple suction alarms and they were resolved after multiple units of blood were administered. the patient was alleged to be actively bleeding. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major and low pump flow has been completed. the device was not returned for investigation. per the provided clinical information, the root cause for access site bleeding was not determined. it was mentioned that they are not concerned with the bleeding anymore. no sufficient clinical details were submitted, nor the product was returned to investigate and determine the root cause. also, the root cause of the low pump flow issue was not determined since product was not returned, there was no motor current spike or trend observed in the data logs and not enough information is provided as per the clinical description.
33919 | 1220648-2023-03868 | 17888707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17888707&pc=OZD | 10/21/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 10/21/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022106310 |  | 0550-008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2021 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support was bleeding from surgical placement site and required six units of blood, 2 platelets, 2 plasma, and 2 cryo. the impella pump remained on for support for another 8 days, til weaned and explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33920 | 1220648-2023-03869 | 17888800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17888800&pc=OZD | 10/23/2021 | 01/26/2024 |  | 10/06/2023 |  |  | 10/23/2021 | 10/06/2023 |  |  | 07/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118216 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2021 |  |  |  |  | The user facility reported a 82-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support after the bypass/cabg surgery. the patient on impella support experienced bleeding resulting in suction alarms. the patient was given 5 units of blood and the suction alarms stopped. pump remained on for support for 5 days and was weaned and explanted. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
33921 | 1220648-2023-03870 | 17889071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17889071&pc=OZD | 10/13/2021 | 01/26/2024 |  | 10/06/2023 |  |  | 10/13/2021 | 10/06/2023 |  |  | 08/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022124025 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2021 |  |  |  |  | The user facility reported a 46year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that impella was implanted per ifu protocol but the position was not ideal. when attempting to reposition, impella became dislodged. md elected to remove and run in basin before re-inserting. the pump ran successfully for just under 20 days. | [5 empty] | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was most likely use related since the pump position was not ideal as per the clinical description provided. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33922 | 1220648-2023-03871 | 17889219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17889219&pc=OZD | 03/07/2023 | 10/06/2023 |  | 10/06/2023 |  |  | 03/07/2023 | 10/06/2023 |  |  | 09/30/2024 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273955 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | Vascular Dissection | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant reported s patient was implanted with an impella cp device for mechanical circulatory support due to severe mitral regurgitation. following implant, impella malposition into aorta was noted. impella was placed on p-2 and return of spontaneous circulation was achieved. patient was placed on bypass machine and the repair to the mitral valve was completed. the impella was removed successfully. | [5 empty] | The investigation into the reported positioning issues has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause of the positioning issue was determined to be patient movement with impella mal-positioning into aorta after patient transfer to operating room.the evaluation revealed that the root cause of the vascular damage issue was determined to be patient condition as the patient had a pre-existing flail mitral valve leaflets and severe mitral regurgitation and was unrelated to impella based on the clinical information. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
33923 | 1220648-2023-03872 | 17890262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17890262&pc=OZD | 10/14/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 10/14/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022099196 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2021 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had the pump inspected due to negative volume and noticed the presence of a thrombus in the right atrium. the patient was returned to the or for a washout and flows were improved. low flows did not prompt pump explant. pump remained on for support for another 7 days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was discarded by the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33924 | 1220648-2023-03873 | 17890333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17890333&pc=OZD | 10/28/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 10/28/2021 | 10/06/2023 |  |  | 08/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128768 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pain | Thrombocytopenia | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2021 |  |  |  |  | The user facility reported a 38-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed pain at their access site coinciding with right arm swelling and a lack of sensation in their right arm. it was verified via ultrasound that there was a lack of flow. furthermore, the patient was being given integrilin for heparin induced thrombocytopenia. the patient began to show signs of recovery the following evening and support with the pump continued. in total the pump supported for 7 days, was weaned, and explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombocytopenia and limb ischemia issue has been completed. the device passed all post-sterile inspection checks. the impella device was discarded by the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33925 | 1220648-2023-03874 | 17890386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17890386&pc=OZD | 10/26/2021 | 01/27/2024 |  | 10/06/2023 |  |  | 10/26/2021 | 10/06/2023 |  |  | 06/30/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022104847 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2021 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient was in the or for a double valve surgery with underlying endocarditis. the complainant reported that a patient on impella support experienced coagulopathy and bleeding. multiple units of blood were given. the pump remained on for support for 76 hours and was then explanted. | [5 empty] | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. no product was returned. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33926 | 1220648-2023-03875 | 17890445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17890445&pc=OZD | 11/12/2021 | 01/26/2024 |  | 10/06/2023 |  |  | 11/12/2021 | 10/06/2023 |  |  | 07/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022117792 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2021 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient inadvertently pulled on the pump which caused it to change position. the patient was taken to the or for a pump exchange. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the positioning issue was most likely patient movement. the pump was pulled by the patient into the aortic arch as per the clinical description provided. | The medical center discarded the impella pump at the time of explant. as such no benchtop analysis is possible. upon completion of the investigation a final report will be filed. this report is being filed as part of a retrospective review of historical records.
33927 | 1220648-2023-03876 | 17890467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17890467&pc=OZD | 11/04/2021 | 03/10/2025 |  | 10/06/2023 |  |  | 11/04/2021 | 10/06/2023 |  |  | 01/31/2023 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022040611 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2021 |  |  |  |  | The user facility reported a 74-year-old female was implanted with an impella 5.5 device for mechanical circulatory support, as care was escalated from an impella 5.0. the complainant reported four suction events prior to the pump signaling that the impella had come back across the valve into the aorta and the patient coded. when return of spontaneous circulation (rosc) was achieved, the echocardiogram (echo) revealed that the catheter was in the aorta. the team made several attempts to re-advance the catheter across the valve into the ventricle with no success. the patient coded and expired before any other intervention could be performed. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. from the clinical details, it was mentioned that they made several attempts to cross the valve but with no success. it is related to the patient condition because it was also noted that the rv also contracting and the sudden decomposition associated with advanced acidosis later the event. from the data logs, the suction alarms appeared for that instance when they have changed the p- level and doesn't support the cause of the positioning issue. hence the root cause was determined to be patient condition. this report is being filed as part of a retrospective review of historical records.
33928 | 1220648-2023-03877 | 17892141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892141&pc=OZD | 09/06/2023 | 12/01/2025 |  | 10/06/2023 |  |  | 09/06/2023 | 10/06/2023 |  |  |  |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1205185 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the console was in for a pm service and the control error/alarm and battery failure alert appeared on start up. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Analysis: before the complaint date on (b)(6) 2023, the last time the console was booted up was on (b)(6) 2022. the console booted up with battery failure event #360 and failed battery communication events #352 and 353. batttest and ltpower logs show both batteries were depleted and were locked out. replacing the batteries resolved the failure. root cause: the cause of the battery alarms were depleted batteries. h3 updated. g1 reporting contact fax number updated. | G1 mdr reporting contact email. d2 device name has been updated.
33929 | 1220648-2023-03879 | 17892142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892142&pc=OZD | 09/06/2023 | 04/15/2024 |  | 10/06/2023 |  |  | 09/06/2023 | 10/06/2023 |  |  | 03/31/2025 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024321772 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 49-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella 5.5 pump support ongoing when there was elevated purge alarms and "purge system blocked". this prompted the team to troubleshoot the alarms by adding tpa to the purge. the team did not need to replace any product, neither pump nor pc. the team allowed the pump to remain on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g.,infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported high purge pressure has been completed. the product was returned for investigation. there was biomaterial in the purge gap and in the cannula. the pump and purge cassette were purging for 10 minutes and the high purge pressure did not reproduce. in the logs, there was a gradual rise in purge pressure over 9 minutes along with a spike in motor current, decrease in purge flow and slightly saturated pump flow. the root cause of the high purge pressure is due to biomaterial ingestion. added- d9 device return date, h6 impact code 4644 , h6 component code 925 h6 medical device problem code should be 1491 and not 1250.
33930 | 1220648-2023-03878 | 17892143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892143&pc=OZD | 09/06/2023 | 01/29/2024 |  | 10/06/2023 |  |  | 09/06/2023 | 10/06/2023 |  |  |  |  | 10/06/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1143010 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Component Missing | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) is missing the purge disc sensor. there was no harm to the patient. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was inspected and evaluated on an engineering lab bench. the console was found to function properly with attached pump and purge line. the reported missing purge disc detection sensor was not found upon inspection and could not be reproduced. the device functioned/operated to specification. therefore the root cause of the missing component could not be determined. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33931 | 1220648-2023-03881 | 17892455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892455&pc=OZD | 12/16/2019 | 10/06/2023 |  | 10/06/2023 |  |  | 09/06/2023 | 10/06/2023 |  |  |  |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN | [7 empty] | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Movement | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2023 |  |  |  |  | A maude database search of mdr reporting against tavr products involving the impella device was completed on (b)(6) 2023. this complaint indicates the impella device may have been a factor the migration of the patient's tavr, and may have occurred during impella delivery: "approximately 10 days post implant of a 29mm sapien 3 valve in the aortic position, the valve migrated and severe paravalvular leak (pvl) was noted. a decision was made to implant a 2nd sapien 3 valve. the initial valve was implanted was reported as very low approximately (30:70) aortic/ventricle, however only trace pvl was noted. the gradient decreased from 18. 6mmhg to zero post gradient. no post dilation was required. the patient was stable when transferred for post management care. as reported, while the patient was in icu the patient condition declined and required chest compressions and an impella device insertion. per report, the physician indicated that the valve migration was most-likely due to the chest compression and the placement of the impella device." | [5 empty] | The impella pump investigation is not possible, and further details are not forwarded. no account can be followed up with for clinical details by the physician/team.
33932 | 1220648-2023-03882 | 17892496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892496&pc=OZD | 09/07/2023 | 01/11/2024 |  | 10/06/2023 |  |  | 09/07/2023 | 10/06/2023 |  |  | 05/31/2025 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352352 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 was placed without difficulty but upon three-point fixation, blood was noticed to be leaking from red impella plug. no current issues with impella flow, and the patient was hemodynamically stable. a decision was made to exchange the impella. | [5 empty] | The investigation for the reported product damage has been completed. the impella device was returned for evaluation. upon visual inspection, blood was found in the pump plug and a hole was found in catheter at mark 30. no other issues were found on the rest of the pump. data logs were reviewed and did not reveal any pump function issues. therefore, the root cause of blood in the red pump plug was due to a hole in catheter. the cause of the hole in catheter was use-related. the instructions for use states the following: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33933 | 1220648-2023-03883 | 17892498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892498&pc=OZD | 09/07/2023 | 02/01/2024 |  | 10/06/2023 |  |  | 09/07/2023 | 10/06/2023 |  |  |  |  | 10/06/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1363891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/07/2023 |  |  |  |  | The user facility's biomedical engineer reported the purge disk broke off their automated impella controller (aic). the aic to be returned for investigation and repair. there was no patient involvement. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into aic - purge disc/flag issues has been completed since the original report was filed. the device was returned for investigation. the console was inspected after arriving. the reported issue with the purge disk breaking off due to a broken purge retainer was confirmed. before returning this console to the customer, field service was instructed to replace the purge retainer and flag, validated sensor accuracy and perform a full functional check on the console and optical system. the issues detecting the purge disc were due to a broken purge pressure sensor retainer. | Section :h6 type of investigation code 10 was inadvertently left off the previous manufacturer device report submitted. section :h6: investigation findings code was inadvertently reported incorrect in the previous manufacturer device report submitted, which has been corrected in this report.
33934 | 1220648-2023-03884 | 17892499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892499&pc=OZD | 09/07/2023 | 01/29/2024 |  | 10/06/2023 |  |  | 09/07/2023 | 10/06/2023 |  |  | 04/30/2025 |  | 10/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333101 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 52-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had an impella 5.5 support ongoing when the patient experienced pain from the access site hematoma. the hematoma was treated after a week of support with a washout. the pump remains on for support for the patient. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for benchtop investigation. according to the clinical notes, the patient had a calcified vessel condition, and a washout procedure was performed for the hematoma noted at the impella site. the cause of the access site bleeding was most likely calcified vessel patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
33935 | 1220648-2023-03752 | 17875228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875228&pc=OZD | 09/05/2023 | 10/05/2023 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  |  |  | 10/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1420044 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fracture | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella controller (aic) power cable had become damaged. the console was replaced without issue. upon investigation of the returned device, it was further noted that the purge disc holder was broken. there was no patient harm as a result of the noted damage. | [5 empty] | The aic was returned and an investigation into the reported 'aic - a/c power issues" has been completed. the device was visually inspected and, although there was no verifiable damage to the power cord, the purge disc retainer was found to be broken (broken blue disc retainer). there was speculation that the biomedical team at the account had repaired the plug before sending the console back for investigation. before returning this console back to the customer, fs was instructed to replace the purge disc retainer, validate sensor accuracy via section 12 of the pm procedure, perform electrical safety test and perform a full functional check on the console and optical system. there were no supporting evidence in logs found. the report of a damaged power cable could not be confirmed and a cause could not be established. the fractured purge disc holder was attributed to a component failure. should additional relevant information become available, a supplemental report will be submitted.
33936 | 1220648-2023-03753 | 17875269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875269&pc=OZD | 09/04/2023 | 01/19/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  | 07/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024369960 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Leak/Splash | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 48-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, there was blood oozing from the red plug of the device. the physician noted "there is a base in the red plug, and the event that blood entered the relevant part was an event that should not have occurred." the impella was replaced with a new device. | [5 empty] | The investigation into product damage (hole in catheter) has been completed. product was returned and investigated. per the investigation, the root cause of the product damage issue was use related to improper technique verified by the hole in catheter and blood in the red handle. h1. corrected from serious injury to malfunction. h6 corrected health effect-clinical code from 1888 to 4582. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33937 | 1220648-2023-03754 | 17875347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875347&pc=PYX | 09/05/2023 | 07/24/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  | 01/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024310632 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 47-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella rp device for mechanical circulatory support. one day after implant, it was noted that the pump stopped unexpectedly. the pump restarted, however, the following day, the patient showed signs of hemolysis. as a result of the failure and the noted condition, the decision was made to explant the pump. the patient remained supported with an already implanted impella 5.5 pump. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The product was returned for investigation. data logs show a motor current spike while running on p-level 7, followed by a pump stop. high purge pressure was noted during the pump stop event. low pump flow was also reported during the case run. no physical abnormalities were reported on the returned product. the pump successfully passed electrical testing and ran as expected during the test run. the motor housing was torn down for further investigation. no evidence of damage or foreign material injection was found. the cause of the hemolysis issue was most likely biomaterial based on data log review. the cause of the temporary pump stop issue was not determined. d6a/d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03754: g1 reporting contact fax number missing.
33938 | 1220648-2023-03755 | 17875382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875382&pc=OZD | 09/05/2023 | 05/15/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  | 05/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365418 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during the impella insertion, the catheter was placed across the aortic valve and, due to tortuous anatomy, the catheter kinked below the outflow when attempting to advance further into the ventricle. once the kink was identified the catheter was pulled back but partial kink remained. the decision was made to leave in place and monitor flows. the patient was stable throughout the case. | [5 empty] | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for investigation. based on the clinical details, the failure was identified as a delivery issue. the root cause of the delivery issue was most likely due to the patient¿s tortuous anatomy causing a kink in the catheter. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33939 | 1220648-2023-03756 | 17875412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875412&pc=OZD | 09/05/2023 | 07/24/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  | 05/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366768 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during support, the patient developed ischemia prompting pump explant. the patient was considered stable following the device explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ d6a/d6b added. h6 component code added 3038.
33940 | 1220648-2023-03757 | 17875426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875426&pc=OZD | 09/04/2023 | 01/18/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  | 06/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373893 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 67-year male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient's groin access site oozed and multiple dressings were changes. hemostasis was achieved with fem-stop and the patient was transfused with packed red blood cells. the physician noted that the patient is doing well and was transferred out of intensive care unit (icu). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the access site bleeding issue was most likely patient condition related due to calcified vessel condition.
33941 | 1220648-2023-03758 | 17875452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875452&pc=OZD | 09/05/2023 | 07/25/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  | 04/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding and small hematoma at the access site. bleeding was managed with manual pressure x 20 minutes and purse-string suture given around site and proper angle matching. support was maintained with the implanted pump. | It was further reported that the patient expired after care was withdrawn. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleed is use issue as it was resolved after doing proper angle matching and additional sutures. the instructions for use for the impella cp with smart assist system states: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-b5 mso review, h6 component code 925, added-f6/f8 should have been left blank in the initial report.
33942 | 1220648-2023-03759 | 17875478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875478&pc=OZD | 09/06/2023 | 01/25/2024 |  | 10/05/2023 |  |  | 09/06/2023 | 10/05/2023 |  |  | 05/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024367258 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Positioning Failure | Migration | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2023 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during cpr, the impella cp ,including pigtail, migrated back across the aortic valve. the pump was not able to be repositioned therefore, the pump was removed, and a new pump was placed. there was no patient harm related to the event. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for benchtop investigation. data logs showed about 1 hour and 13 minutes into support, pump was in aorta that triggered alarm impella position in aorta. pump then was disconnected from the console. based on clinical description and data review, the root cause of positioning issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33943 | 1220648-2023-03760 | 17875511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875511&pc=OZD | 09/06/2023 | 02/02/2024 |  | 10/05/2023 |  |  | 09/06/2023 | 10/05/2023 |  |  | 04/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024341796 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2023 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to left ventricle assist device. patient placement signal was low compared to arterial line. the placement signal is 44 points lower on the aortic mean arterial pressure (map) than the patient¿s radial art line map pressure. the left ventricle waveform was also dampened but not decoupled. whenever there was an attempt to increase p-level from 4 to 5 suction alarm occurs and flows drop. patient has adequate volume and echo shows in a good position with inlet away from any walls of ventricle on multiple echo views. patient has adequate urine output with yellow color. it was further reported that the impella was removed due to a clot found in the inlet. the patient also has been off anticoagulant for five days due to hematoma. another impella 5.5 was placed. | [5 empty] | The investigation for the reported low pump flow, thrombus, and access site bleeding-major has been completed. the impella device was returned for evaluation. upon visual inspection, biomaterial was found inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. data logs confirmed the low flow upon initial pump start and remained the rest of the case. the low flow issue triggered suction response and suction alarms. the root cause of low flow was due to biomaterial ingestion. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33944 | 1220648-2023-03761 | 17875608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875608&pc=OZD | 03/14/2023 | 10/05/2023 |  | 10/05/2023 |  |  | 03/14/2023 | 10/05/2023 |  |  | 11/30/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023279612 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/14/2023 |  |  |  |  | The user facility reported a 24yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to heart transplant. it was reported that there were purge pressure high/purge system blocked alarms requiring administration of tissue plasminogen activator (tpa). the initial alarm occurred with patient standing up for therapy and could be reproduced or aggravated with position changes. of note, catheter also has a possible issue with the optical cable or sensor with intermittent placement signal unreliable alarms. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure alarm. the cause of the issue was not determined since product/data were not returned. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33945 | 1220648-2023-03762 | 17875622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875622&pc=OZD | 03/15/2023 | 10/05/2023 |  | 10/05/2023 |  |  | 03/15/2023 | 10/05/2023 |  |  | 12/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024305783 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2023 |  |  |  |  | The user facility reported a 55yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were purge pressure high/blocked alarms. the purge solution was changed to tissue plasminogen activator (tpa). it is unknown if the high purge pressures resolved. the pump was eventually removed. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure has been completed. without sufficient clinical information and without the sample device, the root cause of the reported issue could not be determined. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33946 | 1220648-2023-03763 | 17875639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875639&pc=OZD | 03/16/2023 | 10/05/2023 |  | 10/05/2023 |  |  | 03/16/2023 | 10/05/2023 |  |  | 12/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024305729 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2023 |  |  |  |  | The user facility reported a 43yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after the pump was implanted, there was a high purge pressure. the purge solution was changed to tissue plasminogen activator (tpa) for a suspected clot. the patient remained on support. no further issues were reported. no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. about 6 hours into support, the "purge pressure high" alarm triggered after a purge pressure rise of less than 5 minutes. the high pure pressure lasted for a total of 45 minutes before resolving. the motor current pulsatility remained low and the flow was saturated at this time. after reviewing the clinical information, it was determined that the cause of the high purge pressure was most likely ingested biomaterial occluding the purge gap. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33947 | 1220648-2023-03764 | 17875644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17875644&pc=OZD | 09/06/2023 | 10/20/2025 |  | 10/05/2023 |  |  | 09/06/2023 | 10/05/2023 |  |  | 12/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315427 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a 63-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was documented that upon placing the impella 5.5 and removing the 0.018 wire, the impella was dislodged from the left ventricle and into the aorta. once the device was removed and pigtail and wire were reinserted the original 0.018 wire was examined and appeared to be kinked. a replacement 0.018 wire was used and impella was successfully placed. there was no patient harm as a result of the noted damage and support continued with the implanted pump. | The customer reported that the device exhibited an optical signal issue. |  |  |  |  | An impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the reported issues has been completed since the original report was filed. the device was returned for investigation and was tested upon return. data logs confirmed the failure mode. logs showed after pump started, ps spiked to 3000 mmhg that triggered alarm ps not reliable & remained to the rest of the case. visual inspection found kinked catheter below endcap. optical continuity test was performed & failed as laser light could not pass the kinked catheter to the sensor face. 5s parameters were out of specification on aic indicating a damaged optical fiber line due to kinked catheter. the root cause of optical signal issue was a damaged fiber line. the cause of the damaged fiber line was use related. the root cause of the guide wire damage was most likely use related. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the investigation for optical signal issue and product damage (guidewire damage) has been completed. the device was returned. data logs confirmed the failure mode. logs showed after pump started, ps spiked to 3000 mmhg that triggered alarm ps not reliable & remained to the rest of the case. upon device evaluation, visual inspection found kinked catheter below endcap. optical continuity test was performed & failed as laser light could not pass the kinked catheter to the sensor face. 5s parameters were out of specification on aic indicating a damaged optical fiber line due to kinked catheter. no other issues were found on the rest of the pump.gw was not returned for analysis. the root cause of optical signal issue was a damaged fiber line. the cause of the damaged fiber line was use related. the root cause of the guide wire damage was most likely use related. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. corrections to the initial mfr report: g1 mdr reporting contact email updated h3 to device evaluation completed h6 type of investigation code added. added optical signal issue.
33948 | 1220648-2023-03776 | 17876276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876276&pc=OZD | 08/07/2021 | 01/11/2024 |  | 10/05/2023 |  |  | 08/07/2021 | 10/05/2023 |  |  | 05/31/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088842 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2021 |  |  |  |  | The user facility reported a 58-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the impella had migrated out of the left ventricle (lv) and another time it migrated into the aorta. in both instances, impella needed to be repositioned. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ minor and the positioning issues has been completed. the product was not returned. as per the clinical information provided, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. the root cause of the positioning issue was most likely patient movement.
33949 | 1220648-2023-03777 | 17876332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876332&pc=OZD | 08/18/2023 | 09/19/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  | 06/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024364387 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility |  |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Ischemia Stroke | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | A notification via the long-term outcome and quality indicator (loqi) impella registry was received regarding a 60 year old male patient that presented with acute myocardial infarction and cardiogenic shock for hemodynamic support via the impella 5.5 device. a possible relationship between sepsis and stroke and the impella device was made. the patient was noted to have been confirmed covid-19 positive. a ct scan revealed a "small volume subarachnoid hemorrhage along right frontal lobe. multiple acute/subacute infarcts to right cerebral hemisphere." the patient was ultimately withdrawn from care and expired, however, there was no allegation the patient's death was related. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | B2: date of patient death is unknown. the investigation for the reported infection and stroke has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the infection and stroke and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03777 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
33950 | 1220648-2023-03778 | 17876501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876501&pc=OZD | 08/19/2021 | 01/18/2024 |  | 10/05/2023 |  |  | 08/19/2021 | 10/05/2023 |  |  | 03/31/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022084061 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2021 |  |  |  |  | The user facility reported a 62-year-old male with multiple cardiac comorbidities was implanted with an impella 5.5 device for mechanical circulatory support for an electrophysiology procedure. the patient on impella support experienced bleeding which lead to 200 cc being drained in the jackson-pratt drain. the patient had to be taken to the or to control the bleeding. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, the root cause of the issue was not determined since product was not returned and sufficient clinical information was not provided.
33951 | 1220648-2023-03779 | 17876731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876731&pc=OZD | 08/19/2021 | 09/16/2024 |  | 10/05/2023 |  |  | 08/19/2021 | 10/05/2023 |  |  | 05/31/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022088098 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Death |  |  |  |  | PROTEK DUO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2021 |  |  |  |  | The complainant reported that a patient on impella 5.5 support experienced an embolic stroke. the patient was also on right sided support with a protek duo. the medical team and the family decided to withdraw care and the patient ultimately expired after being made a dnr. the pump support had been for 8 days. | [5 empty] | The medical center did not return for benchtop analysis the impella pump. no imaging was shared. the investigation into the stroke is ongoing at this time. upon completion of the investigation the final will be forthcoming. this report is being filed as part of a retrospective review of historical records. of note the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03779 was provided.
33952 | 1220648-2023-03780 | 17876792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876792&pc=OZD | 08/28/2021 | 03/10/2025 |  | 10/05/2023 |  |  | 08/28/2021 | 10/05/2023 |  |  | 01/13/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032055 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2021 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the care had been escalated for the cardiomyopathy patient awaiting a vad or heart transplant. the patient on impella support had experienced a leak from the connection from the purge line to the yellow connection on the side arm, which was loose and d5 was leaking from the connection into the sidearm retainer. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir. clean the connector cable with 70% isopropyl alcohol. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported purge leak issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was use related since it was loose and then resolved by tightening. this report is being filed as part of a retrospective review of historical records.
33953 | 1220648-2023-03781 | 17876832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876832&pc=OZD | 02/08/2023 | 07/25/2024 |  | 10/05/2023 |  |  | 09/09/2023 | 10/05/2023 |  |  | 04/30/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022078653 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2023 |  |  |  |  | We received a notification from the long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis & thrombocytopenia with a possible relationship to the impella device. one unit of prbc was delivered in response. | [5 empty] | The impella device was not received by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for hemolysis and thrombocytopenia has been completed. the impella device was not returned for evaluation. data logs show no signs related to positioning issue, placement signal trend or motor current spike during the case run. without additional clinical details, the root cause of the reported issues could not be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h6 component code added 925 the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03781: d4 model number f6/f8-should have been left blank g1 reporting contact facility name and fax number missing g2 report source h5 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
33954 | 1220648-2023-03782 | 17876866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876866&pc=OZD | 08/27/2021 | 03/10/2025 |  | 10/05/2023 |  |  | 08/27/2021 | 10/05/2023 |  |  | 06/30/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022102548 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2021 |  |  |  |  | The user facility reported a 39-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the patient was geting worked up for a durable vad. the patient on impella support had experienced fungemia which was confirmed by the blood cultures being positive. the fungemia was treated by antibiotics. device was explanted after 45 days and durable vad placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported infection/sepsis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33955 | 1220648-2023-03783 | 17876922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876922&pc=OZD | 08/28/2021 | 07/18/2024 |  | 10/05/2023 |  |  | 08/28/2021 | 10/05/2023 |  |  | 06/30/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022106337 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2021 |  |  |  |  | The user facility reported a 32-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was taken to the or for impella replacement for extreme hemolysis. in the or the surgeon placed a new graft with end to end technique onto the right axillary. impella placed per ifu protocol. proper position was verified by fluoroscopy and trans-"esophageal" echo, md pleased with position. the pump explantation and crrt dialysis were the treatments for the hemolysis. the patient was supported in total for just under 28 days and was bridged to lvad successfully. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis has been completed. the root cause of patient injury (hemolysis) was not determined. it was mentioned that they have observed hemolysis during crrt (continuous renal replacement therapy) and they have declined the assistance needed for hemolysis. patient was stable till the end of the case after the issue and no other concerns related with impella. the product and the data were not returned to investigate further. therefore, the cause of hemolysis was not determined and unknown relation to impella. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in the following sections: section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿
33956 | 1220648-2023-03784 | 17876976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17876976&pc=OZD | 09/06/2023 | 01/25/2024 |  | 10/05/2023 |  |  | 09/06/2023 | 10/05/2023 |  |  | 06/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368829 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2023 |  |  |  |  | The user facility reported a 47-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient's right axillary was noted to have a large hematoma. the patient was stable. all anti-coagulation was put on hold and the patient received packed red blood cells. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. according to the clinical notes, all anticoagulation medications were put on hold after reported bleeding from the right axillary artery. the cause of the access site bleeding issue was most likely related to high act, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33957 | 1220648-2023-03785 | 17877070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877070&pc=OZD | 09/11/2023 | 01/25/2024 |  | 10/05/2023 |  |  | 09/11/2023 | 10/05/2023 |  |  | 03/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024334152 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2023 |  |  |  |  | The user facility in the eu reported a 66-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had massive bleeding in the right groin. spontaneous activated partial thromboplastin time of >250 due to liver failure. after removing the peel-away sheath, there was excessive bleeding and was unable to stop. the team removed the impella cp and placed a second cp in the other groin. the team additionally infused protamine and novosevenrt for coagulation. the team left the 2nd peel-away sheath in for the 2nd pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. based on clinical description, the root cause of access site bleeding was most likely due to patient condition.
33958 | 1220648-2023-03786 | 17877176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877176&pc=OZD | 09/11/2023 | 12/01/2025 |  | 10/05/2023 |  |  | 09/11/2023 | 10/05/2023 |  |  |  |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105122 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/11/2023 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the team observed a console issue after the replacement of the purge cassette. the cable broke off in the console and the team was unable to remove a piece of it. there was no patient harm related to this event. aic will be returned for investigation and repair. | [5 empty] | The impella console was returned for investigation upon the benchtop. the investigation is ongoing at this time. upon investigation closure, a supplemental mdr will be filed. | D.2 common device name was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03786 was submitted. d.4 model number was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03786 was submitted. revised lot number as it was reported incorrectly on the initial manufacturer device report number 1220648-2023-03786. e.1 fax number was inadvertently left off the previously submitted manufacturer device report number 1220648-2023-03786 and was added. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-03786 and should not have been. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-03786 in accordance with updated procedures. h.6 codes 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2023-03786 in accordance with updated procedures. h.8 revised as selection was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-03786.
33959 | 1220648-2023-03787 | 17877307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877307&pc=OZD | 09/20/2023 | 03/10/2025 |  | 10/05/2023 |  |  | 09/20/2023 | 10/05/2023 |  |  | 12/31/2022 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022025282 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed a bacteremia during impella support. although the source of the infection was unknown, the decision was made to explant the impella 5.5 and any other existing lines from the patient. the patient was stable following the event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported infection/sepsis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33960 | 1220648-2023-03788 | 17877343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877343&pc=OZD | 05/30/2022 | 01/26/2024 |  | 10/05/2023 |  |  | 05/30/2022 | 10/05/2023 |  |  | 01/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023175763 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient was non-compliant and pulled impella from left ventricle into aorta during postoperative delirium. impella repositioning failed and the pump was surgically removed. no further escalation planned. | [5 empty] | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issue has been completed since the original report was filed. accused product was not returned. the root cause of the positioning issue was most likely due to patient movement.
33961 | 1220648-2023-03789 | 17877366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877366&pc=OZD | 07/01/2022 | 10/21/2025 |  | 10/05/2023 |  |  | 07/01/2022 | 10/05/2023 |  |  | 02/29/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023193686 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Difficult to Advance | [6 empty] | Vascular Dissection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team was having difficulty delivering an impella pump due to the patient's tortuous vasculature. during the attempt the innominate artery was dissected and the vascular surgeon was called. ultimately, the pump was inserted successfully and the decision was made to repair any dissections upon pump explant. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
33962 | 1220648-2023-03791 | 17877450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877450&pc=OZD | 06/11/2022 | 03/10/2025 |  | 10/05/2023 |  |  | 06/11/2022 | 10/05/2023 |  |  | 12/31/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023171552 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2022 |  |  |  |  | The user facility reported a 21-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the patient developed a hematoma on their back right side. the patient was taken to the operating room, where the hematoma was treated. support continued with the implanted pump, however, the patient later developed an infection around their access site. the patient was given unspecified antibiotics to treat the infection. support was maintained throughout the event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding, infection/sepsis and placement signal issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the infection was patient condition related and the relation to impella is unknown. the cause of the positioning issue was use issue since the false alarm occurred during the manipulation of the impella graft during procedure. the root cause of the access site bleeding issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
33963 | 1220648-2023-03790 | 17877455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877455&pc=OZD | 01/11/2022 | 03/10/2025 |  | 10/05/2023 |  |  | 01/11/2022 | 10/05/2023 |  |  | 07/31/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113074 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with dilated cardiomyopathy, was implanted with an impella 5.5 device for mechanical circulatory support via the axillary site. a day post implantation, the patient had no pulsatility in left arm and severe hemolysis that required multiple packet red blood cell (prbcs), platelets, and albumin. the patient returned to the or to be placed on ecmo and explant the impella 5.5. after placing the patient on ecmo, the surgeon discovered pocket of blood impeding flow to artery. pocket was therefore drained and a drain was placed while leaving impella in for left ventricle vent /exit strategy. impella 5.5 was repositioned from 6.5 to 5.3cm mid lv. however, bedside echo noted impella to be flipped and diving into apex. the impella 5.5 was ultimately explanted in lieu of intra-aortic balloon pump (iabp). upon explant, impella was noted to have a fair amount of clot built up at outlet area. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported hemolysis, access site bleeding, limb ischemia and thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33964 | 1220648-2023-03792 | 17877519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877519&pc=OZD | 05/19/2022 | 05/23/2025 |  | 10/05/2023 |  |  | 05/19/2022 | 10/05/2023 |  |  | 03/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023199494 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHARMACIST | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/19/2022 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient had impella 5.5 inserted directly via aortic approach with no issues. while closing the chest, the patient decompensated. under echo it was found out that the patient had a large dissection/hematoma thought to be from aortic cannulation site post removal. the patient was recannulated and impella placed in surgical mode. physician to repair aorta and replace impella. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03792 was provided. | No product was returned. the root cause of the access site bleeding issue was patient condition related, as the patient had a large dissection/hematoma as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
33965 | 1220648-2023-03793 | 17877571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17877571&pc=OZD | 04/07/2022 | 05/23/2025 |  | 10/05/2023 |  |  | 04/07/2022 | 10/05/2023 |  |  | 01/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023185224 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Disability | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/07/2022 |  |  |  |  | The user facility reported a 34-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced a low pump flow due to a possible clot. the medical team was concerned for decreased perfusion to right arm and the hand specialist team and ortho consulted and present. decision was made to perform right arm fasciotomy in cath lab at the same time left heart catheterization (arterial catheterization) is performed. during the fasciotomy performed large hematoma and muscle death noted. no distal flow past right iliac was found on angio. family made decision to withdraw care after evidence of multiple brain clots and poor prognosis. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The root cause of the thrombus and high purge pressure could not be determined as insufficient clinical details were provided. the root cause of the limb ischemic reversible issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
33966 | 1220648-2023-03794 | 17878009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878009&pc=OZD | 06/24/2022 | 01/16/2024 |  | 10/05/2023 |  |  | 06/24/2022 | 10/05/2023 |  |  | 02/29/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194793 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the femoral access site. as a result of the condition coinciding with the need for prolonged therapy until a scheduled bypass surgery/cabg, the decision was made to explant the impella cp and place an impella 5.5 pump via the axillary. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. according to the clinical information provided, the root cause of the access site bleeding issue was most likely low patient act values.
33967 | 1220648-2023-03795 | 17878038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878038&pc=OZD | 06/14/2022 | 03/10/2025 |  | 10/05/2023 |  |  | 06/14/2022 | 10/05/2023 |  |  | 02/29/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023190547 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support and management to bridge the patient to lvad. two weeks post implant, the patient was going to the or for possible debridement of impella site to rule out infection. patient was noted to have had positive blood cultures. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation for the reported infection/sepsis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33968 | 1220648-2023-03796 | 17878203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878203&pc=OZD | 07/21/2022 | 01/26/2024 |  | 10/05/2023 |  |  | 07/21/2022 | 10/05/2023 |  |  | 04/30/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211211 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced persistent bleeding at their impella access site during patient support. multiple units of packed red blood cells (prbcs) were given and a pressure dressing was placed to treat the condition. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
33969 | 1220648-2023-03797 | 17878230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878230&pc=OZD | 07/21/2022 | 03/04/2024 |  | 10/05/2023 |  |  | 07/21/2022 | 10/05/2023 |  |  | 03/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023209914 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 73-year-old female patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient lost radial and ulnar pulses during impella support. due to the noted condition, the pump was explanted. the patient's condition was reported to be stable and their pulses were noted to have returned following device removal. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
33970 | 1220648-2023-03798 | 17878250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878250&pc=OZD | 10/11/2022 | 01/11/2024 |  | 10/05/2023 |  |  | 10/11/2022 | 10/05/2023 |  |  | 05/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023224573 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 79-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient pulled their impella 5.5 pump back out of position by ten centimeters during support. the patient was without support for over two hours, but remained hemodynamically stable. due to the patient's mental condition and their stability, the decision was made to explant the pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was most likely patient movement, pt pulled his own impella back 10cm as per clinical description provided. the failure modes will be monitored and trended.
33971 | 1220648-2023-03799 | 17878288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878288&pc=OZD | 10/04/2022 | 03/10/2025 |  | 10/05/2023 |  |  | 10/04/2022 | 10/05/2023 |  |  | 06/30/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247451 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 66-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site at the right subclavian cut down. the site was packed and sutured shut to control the bleeding. support was maintained with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the bleeding was determined to be use issue. it was reported that the patient had copious amount of bleeding from right subclavian even before the introduction of impella and the site was packed and sutured shut to control the bleeding. later, the impella was successfully implanted. the functioning was stable, and no issues were reported with the impella, and no issue of bleeding was reported. hence, the root cause is determined to be patient condition. this report is being filed as part of a retrospective review of historical records.
33972 | 1220648-2023-03800 | 17878314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878314&pc=OZD | 07/20/2022 | 01/26/2024 |  | 10/05/2023 |  |  | 07/20/2022 | 10/05/2023 |  |  | 03/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023207673 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 60-year-old female patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella 5.5 pump dislodged overnight while the patient was ambulating. the pump was subsequently explanted and the patient was considered stable. there was no patient harm. | [5 empty] | The investigation into the reported positioning issues has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was patient movement, the patient was dislodged overnight while ambulating as per the clinical description provided. | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33973 | 1220648-2023-03801 | 17878435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878435&pc=OZD | 07/08/2022 | 09/19/2024 |  | 10/05/2023 |  |  | 07/08/2022 | 10/05/2023 |  |  | 01/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182491 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 55-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient became febrile and lactic levels went from 2.9 to 49 within a 24 hour period during impella support. the patient was given antibiotics and a washout of the impella site was performed to counteract the infection. the patient showed signs of recovery and support with the implanted pump continued. | Additional information noted that patient was suffering from multi-organ failure and family decided to withdraw care. impella was turned off and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon conclusion of the investigation, it was deduced that the cause for the issue was attributed to a patient related condition. should additional relevant information become available, a supplemental report will be submitted. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03801 was provided in sections b2, b5,d6, h1, and h6. | The investigation into the infection/sepsis issue has been completed since the original report was filed. no product was returned. based off the clinical information available, the root cause of the infection/sepsis issue was patient condition related, as there was no clear indication of false procedure and the relation to the impella is unknown.
33974 | 1220648-2023-03802 | 17878470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878470&pc=OZD | 07/05/2022 | 01/19/2024 |  | 10/05/2023 |  |  | 07/05/2022 | 10/05/2023 |  |  | 04/30/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023214638 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 61-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed an impella 5.5 access site bleed during patient support. it was noted by the surgeon that they had difficulty in preventing the anastomosis from bleeding. thrombin, manual pressure, and sutures were added to address the bleed. furthermore, the platelet levels were low and the patient received brilinta. support continued with the implanted pump following the intervention. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related, as the patient had very low platelets and blood supplements were also given as per the clinical description provided. | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33975 | 1220648-2023-03803 | 17878513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878513&pc=OZD | 07/03/2022 | 03/10/2025 |  | 10/05/2023 |  |  | 07/03/2022 | 10/05/2023 |  |  | 04/30/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210785 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/03/2022 |  |  |  |  | The user facility reported a 51-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that they were unable to get adequate flow from an impella 5.5 pump during patient support. at support level p-6, flow rates were only reaching 3.2 l/min. re-positioning was advised; however, the pump initially remained untouched. suction was reported to be frequent whenever support levels were raised above p-5. there was no patient harm and support continued with the implanted pump despite the noted issues. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33976 | 1220648-2023-03804 | 17878630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878630&pc=OZD | 06/30/2022 | 01/24/2024 |  | 10/05/2023 |  |  | 06/30/2022 | 10/05/2023 |  |  | 04/30/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211232 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that following 5.5 implant, the patient experienced some oozing at their axillary access site. after removing the peel-away sheath, a pressure dressing and sandbag were put in place to treat the bleed. the patient was given an unspecified number of blood products and support continued with the implanted pump. | [5 empty] | The investigation for the access site bleed has been completed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33977 | 1220648-2023-03805 | 17878850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878850&pc=OZD | 06/24/2022 | 03/10/2025 |  | 10/05/2023 |  |  | 06/24/2022 | 10/05/2023 |  |  | 10/31/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022152032 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed a large apical left ventricle thrombus during impella support. after the thrombus was identified via echocardiogram (echo), the decision was made to explant the pump and place the patient on veno-arterial extracorporeal membrane oxygenation (ecmo). the patient was stable following the explant. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33978 | 1220648-2023-03806 | 17878924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17878924&pc=OZD | 09/08/2023 | 01/11/2024 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  | 10/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309157 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and impella was repositioned. the pump remained on support for 3 days. | It was further reported that upon initiation of support, flows started in auto and immediately the left ventricle (lv) completely decompressed with reduced flows. echo at bedside showing severe lv dysfunction and moderate-severe rv dysfunction. eventually, the p-level support was increased to 7. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported hemolysis and low pump flow has been completed. the impella device was not returned for evaluation. clinical information provided was sufficient to determine the root cause. as per clinical details, the patient had suction and low flows at the start of support due to left ventricle decompression due to severe dysfunction and later p-level was able to be increased. hemolysis issue experienced with hematuria and low urine output after impella position observed too deep with several aorta alarms. data logs were reviewed and suction alarms with p-level decrease to p-2 and later flow improved with increase to p-7. several impella position in aorta and ventricle alarms seen with down trending placement signal. the cause of the hemolysis issue was most likely improper positioning of the device due to hematuria developed after several improper position alarms per clinical and log review. the cause of the low pump flow issue was most likely patient condition related with lv decompressed due to severe dysfunction per and suction alarm with low flows at start seen per log and clinical review.
33979 | 1220648-2023-03807 | 17879130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17879130&pc=OZD | 06/29/2022 | 10/21/2025 |  | 10/05/2023 |  |  | 06/29/2022 | 10/05/2023 |  |  | 03/31/2023 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022056283 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was unable to be increased more than p6 due to suction. however, there was no reported patient injury. the pump was explanted the next day. the pump had remained on for support of 6 days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
33980 | 1220648-2023-03808 | 17879200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17879200&pc=OZD | 09/08/2023 | 03/27/2025 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  |  |  | 10/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1439042 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility in japan reported had an automated impella controller (aic) with a purge disc issue. the purge cassette was replaced and found the aic to be malfunctioning and so the team replaced the aic. the 5.5 support proceeded on the second aic without any harm or injury from the interruption in purge. this aic issue was noted on day 11 of the support which continued. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported issue has been completed. data analysis: event logs confirm that the complaint by presence of purge disc not detected alarm. device analysis: the automated controller unit arrived at the tokyo office. it was confirm that the position sensing lever of the purge pressure sensor unit is damaged, from visual inspection. conclusion: the root cause for the purge issues in the complaint was a damaged purge pressure flag.
33981 | 1220648-2023-03809 | 17879231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17879231&pc=OZD | 09/12/2023 | 08/02/2024 |  | 10/05/2023 |  |  | 09/12/2023 | 10/05/2023 |  |  | 04/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349466 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2023 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. support was ongoing when there was bleeding from the femoral site. the access site bleeding prompted the team to troubleshoot by infusing 2 units of blood and prematurely explanting the pump. the pump had allowed for the 5+ hour of support. the pump was explanted and the groin closed with a manta. the patient decompensated post and became hypotensive and had pulseless electrical activity which prompted cpr. the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. per medical safety officer review use of the device cannot conclusively be associated with the outcome. d6a and d6b added. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03809: f6/f8-should have been left blank.
33982 | 1220648-2023-03810 | 17879268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17879268&pc=PYX | 09/08/2023 | 01/25/2024 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  | 03/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP EU |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023205705 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu reported a 54-year-old male patient has been treated for right heart failure with acute myocardial infarction and atrial fibrillation. support has been provided by va-ecmo, an impella 5.0 heart pump and an impella rp heart pump. it has been reported that hemolysis has been observed. the impella rp device has been removed from the right ventricle and the patient has been treated with continuous hemodialysis. within 24 hours of impella rp removal, the hemolysis signs resolved. the patient has been stable afterwards, on continued va-ecmo and impella 5.0 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis has been completed. the device was not returned for benchtop investigation. as per clinical, the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
33983 | 1220648-2023-03811 | 17879378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17879378&pc=OZD | 09/08/2023 | 01/25/2024 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  |  |  | 10/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had an error message (battery communication failure). the staff followed the prompts and exchanged the aic for a new one that worked properly. | [5 empty] | The investigation into the reported aic - battery comm. failure (yellow alarm) has been completed since the original report was filed. imc log from the complaint date shows a controller error (opm comm crc invalid) [optical pump connected] ¿ alarm, event set #1045, which confirms the reported failure. ¿controller error¿ recovered by itself. when the controller error occurred, the rt log recorded the snr optical signal dropped to ¿-16384,¿ which points to the failure mode transient loss of opm data. the console was tested on an engineering lab bench. the failure mode was not reproduced by running the console with the optical pump. there was no issue with the console hardware related to the reported failure. the root cause for the controller error was not determined due to the inability to reproduce. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33984 | 1220648-2023-03812 | 17879584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17879584&pc=OZD | 09/08/2023 | 12/06/2023 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  | 06/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024369189 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown race/age/gender received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. after one day of impella support, a pump stop occurred. the impella cp has been removed. the patient was transferred to another hospital and a new impella cp heart pump was inserted. the patient was subsequently stable. | [5 empty] | The investigation into the pump stop has been completed. the impella ump was not returned for investigation from the eu. data logs were not provided. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency..
33985 | 1220648-2023-03813 | 17880122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17880122&pc=OZD | 05/06/2023 | 10/05/2023 |  | 10/05/2023 |  |  | 05/06/2023 | 10/05/2023 |  |  | 01/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024310006 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/06/2023 |  |  |  |  | The user facility reported a 51yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day 3 of support there was a diagnosis made of a cerebrovascular accident (cva). the patient remained on for support despite the cva. the cva occurring during the support caused right upper limb issues. the team also adjusted the anticoagulation. the pump remained on for support for 24 days until explant. | [5 empty] | The investigation for the reported stroke has been completed. the device was not returned for evaluation. however, the clinical information provided sufficient details to determine the root cause. the cause of the stroke event was most likely patient condition related since no heparin was given during the case. the instructions for use lists stroke as a potential adverse event. it states :¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33986 | 1220648-2023-03814 | 17880450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17880450&pc=OZD | 09/12/2023 | 01/29/2024 |  | 10/05/2023 |  |  | 09/12/2023 | 10/05/2023 |  |  | 12/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024301966 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2023 |  |  |  |  | The user facility in japan reported a male (age unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was removed as a suction alarm would not resolve and re-positioning the impella cp were unsuccessful. an impella 5.5 was placed and the cp explanted. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the product and the data logs were not returned for review. based on clinical description, the root cause of positioning issue was most likely due to patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33987 | 1220648-2023-03815 | 17880597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17880597&pc=OZD | 09/08/2023 | 05/22/2024 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  | 04/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024335641 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male in st elevated myocardial infarction was implanted with an impella cp and ecmo for mechanical circulatory support. the complainant had a patient on impella cp support when the patient expired from hemorrhagic complications. the relationship between the impella and the death is unknown. the patient had dic, bladder bleeding, gi bleed, as well as access site bleeding. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the vascular damage - peripheral vessels cannot be determined as limited clinical information was provided. access site bleeding was also reported. the root cause of the access site bleeding is most likely due to patient condition as bleeding was occurring from multiple locations. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33988 | 1220648-2023-03816 | 17882714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17882714&pc=OZD | 09/05/2023 | 01/23/2024 |  | 10/05/2023 |  |  | 09/05/2023 | 10/05/2023 |  |  |  |  | 10/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1211689 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) system self-check failed when aic started for in-service. the aic was removed. | [5 empty] | The investigation into the aic - system self-check error has been completed since the original report was filed. the device was returned for investigation. it was reported that the cause of the aic issue was contamination. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33989 | 1220648-2023-03817 | 17882720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17882720&pc=OZD | 09/05/2023 | 01/25/2024 |  | 10/05/2023 |  |  |  | 10/05/2023 |  |  | 04/30/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355033 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 45-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the cardiovascular intensive care unit (cvicu) did not initiate systemic heparin nor switch the impella cp purge fluid over to the recommended d5w 25meq/l bicarb purge overnight on this patient. the abiomed clinical support consultant was present for the case and transferred the patient to the icu for check in on 9/4. act was greater than 250 upon leaving the cath lab so its not uncommon for the cvicu to wait a few hours before initiating systemic anti coagulation of some kind. the field representative rounded on this patient this morning 9/5 and asked about anti coagulation and why the purge bag was not changed. cvicu stated that the patient was not started on any anti coagulation over night and that he wasn¿t sure why the purge fluid wasn¿t changed. education provided concerning this matter. patient was scheduled for explant today. cp was explanted with out issue. upon explant of the cp, a right common femoral artery (cfa) thrombus greater than 90% was noted on angio while floating the perclose wire. the left cfa accessed for penumbra thrombectomy of right cfa. thrombectomy to the right cfa successful. right distal angio performed and peripheral vascular disease (pvd) noted throughout the patients lower right extremity. the right cfa was also ballooned to ensure a patent artery. distal flow was restored. left common iliac was stented x1 for pvd and angio of the left lower extremity displaced pvd similar to the right lower extremity. bil groins were perclosed at the conclusion of the case and patient returned to cvicu with bil distal pulses intact. the physician was made aware of the cvicu not initiating systemic anti coagulation overnight and not changing out the d5w to our recommended purge fluid. the physician stated that this in not an impella issue, rather a system issue of why anti coagulation was not started on his patient. plavix was given in the cath lab on arrival and this owning prior to explant. the patient was not harmed, in neurologically intact and following commands and will most likely be extubated this afternoon in cvicu. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. as per clinical, thrombus observed on right common femoral artery (cfa) upon explant of cp pump and thrombectomy done to the right cfa to successfully evacuate the clot. the cause of the thrombus issue was patient condition related due to thrombus observed in right femoral artery during explant and thrombectomy performed to resolve complication per clinical.
33990 | 1220648-2023-03818 | 17882721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17882721&pc=OZD | 09/05/2023 | 06/25/2024 |  | 10/05/2023 |  |  |  | 10/05/2023 |  |  | 05/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | There is not enough information to associate use of the device with the outcome. | The user facility reported a 71-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced a large hematoma in the impella groin. manual pressure was held, and additional sutures were placed by the surgeon. the patient is stable. |  |  |  |  | The investigation into the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to patient movement. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
33991 | 1220648-2023-03819 | 17882729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17882729&pc=OZD | 09/05/2023 | 01/29/2024 |  | 10/05/2023 |  |  |  | 10/05/2023 |  |  | 05/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358660 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 63-year-old female/male noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the pci was completed and the patient tolerated well. the team then observed the pump needed to be explanted due to limb ischemia in the impella accessed leg. the pulses were lost. the physician noted the patient appears to be doing well. headed to the icu for post care. | [5 empty] | The investigation into the limb ischemia ¿ irreversible has been completed. the device was not returned for investigation. the cause of the limb ischemia issue was most likely patient condition related since device removal resolved the limb ischemia. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33992 | 1220648-2023-03820 | 17883014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17883014&pc=OZD | 12/12/2022 | 10/05/2023 |  | 10/05/2023 |  |  | 12/12/2022 | 10/05/2023 |  |  | 08/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259594 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The user facility in japan reported a 51-year-old male admitted in cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant in japan had a 5.5 support of over 24 days. the team noted an infection that could have been due to the use of impella but, the causality is not known. | [5 empty] | The data logs were returned and show a slow increase in purge pressure and decrease in purge flow over multiple days/ purge system blocked and high purge pressure alarms were triggered. motor current remained stable throughout the case. the root cause of the high purge pressure was most likely biomaterial in the purge gap because the increase in purge pressure was gradual, it had no effect of the motor current, and it was able to be resolved. no product was returned for this investigation. the root cause of the infection was most likely patient condition as the infection did not occur until day 25 of support and the pump was confirmed to be sterilized under validated conditions. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ . ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ . this report is being filed as part of a retrospective review of historical records.
33993 | 1220648-2023-03821 | 17883016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17883016&pc=OZD | 01/19/2023 | 10/05/2023 |  | 10/05/2023 |  |  | 01/19/2023 | 10/05/2023 |  |  | 10/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023269608 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/19/2023 |  |  |  |  | The user facility reported that a 46-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the catheter was reported to be stuck in the blue suture hub/t-lock housing. as a result, the device could not advance, retract, or torque. the physician had to torque graft and suture the hub, and catheter en bloc in order to reposition. a blue suture pad was sewed on inverted (graft twisted 3 times). there was no reported harm to the patient. | [5 empty] | The investigation into the reported positioning issues was completed. the device was not returned for evaluation. based on the available information, the root cause of the positioning and suture issues could not be determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33994 | 1220648-2023-03822 | 17883023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17883023&pc=OZD | 01/16/2023 | 10/05/2023 |  | 10/05/2023 |  |  | 01/16/2023 | 10/05/2023 |  |  | 05/31/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217463 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/16/2023 |  |  |  |  | The user facility reported that a 75-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the device rotated toward the mitral valve during insertion and afterwards, echo revealed that this would occur immediately after the tuory bourst (tb) was not tightened. device repositioned many times with same result. all torque was removed from the catheter during insertion and during repositioning. the pump remained in place with orientation toward mitral valve as pump performance did not appear to be affected. there was no reported harm to the patient. | [5 empty] | The investigation into the reported positioning issues was completed. the device was not returned for evaluation. based on the available information, the root cause of the positioning issues was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33995 | 1220648-2023-03823 | 17883080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17883080&pc=OZD | 08/28/2023 | 07/25/2024 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  | 05/31/2025 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024348568 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a successful cp support explanted after two days of support. hemodynamics had stabilized. at the time of explant, there was thrombus observed. additional intervention was not required. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." a.3 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2023-03823. b.3 date of event was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-03823 and has been updated. d.4 catalog number was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-03823 and has been updated. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-03823 in accordance with updated procedures. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-03823 was submitted. h.6 code 925 was added since the initial submission of manufacturer device report number 1220648-2023-03823 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-03823 in accordance with updated procedures.
33996 | 1220648-2023-03824 | 17883091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17883091&pc=OZD | 09/08/2023 | 02/26/2024 |  | 10/05/2023 |  |  | 09/08/2023 | 10/05/2023 |  |  | 11/30/2024 |  | 10/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308525 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The family member of a (b)(6) 2023 supported impella cp patient contacted abiomed to report upon the patient death and an allegation that the use of impella caused the patient outcome. the patient had multiple comorbidities (inclusive of cgs, ef 15%, copd, tbi, tobacco, meth use and sob) and had a twenty-five hour support. the pump was then explanted without any harm or injury noted. the family member noted that after explant the patient transferred out to another hospital and then nine days later the patient expired from cardiogenic shock (cgs). | [5 empty] | The investigation into the embolism issue has been completed. the device was not returned for investigation only a usb drive was sent. the case log data was returned but did not show any device malfunction; the device functioned within specification. the root cause of the embolism issue was not determined since product was not returned. without further information, it is not possible to determine the exact root cause of the patient¿s subsequent death in another hospital and establish a definitive association with impella use. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33997 | 1220648-2023-03670 | 17867744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17867744&pc=OZD | 04/06/2023 | 10/04/2023 |  | 10/04/2023 |  |  | 04/06/2023 | 10/04/2023 |  |  | 12/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024315450 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/06/2023 |  |  |  |  | The user facility reported a male patient was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after removing the soft clamps used to cross clamp the aorta, high pressure alarms/low purge flows were alarming. the pump was removed and replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure alarms has been completed. the data log was reviewed and upon pump activation, suction, gradually increasing purge pressure, and position in aorta alarms were observed. the cause of the issue was not determined since product was not returned. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33998 | 1220648-2023-03671 | 17867783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17867783&pc=OZD | 04/17/2023 | 10/04/2023 |  | 10/04/2023 |  |  | 04/17/2023 | 10/04/2023 |  |  | 12/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024318646 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/17/2023 |  |  |  |  | The user facility reported 69yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day 4 of the 5.5 support there was high purge pressure alarms that was troubleshot by the use of tissue plasminogen activator (tpa). the purge pressure and purge flows alarmed the following day as well and so they troubleshot with purge cassette (pc) replacement, as the pc was seen to be leaking. the new cassette was placed and support proceeded till wean and explant. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. the datalogs and the returned pump was visually inspected and biomaterial was observed in the outflow. the leaking purge cassettes were not returned for review. the high purge pressure alarm was reproduced. a kink was observed in the catheter. the catheter was trimmed and no kink in the purge line was observed. the catheter was then cut and purge was confirmed to exit the catheter purge line with a low purge pressure alarm resulting. the cause of the high purge pressure was biomaterial. the instructions for use includes troubleshooting for high purge pressure alarms. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33999 | 1220648-2023-03676 | 17867927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17867927&pc=OZD | 09/07/2023 | 02/09/2024 |  | 10/04/2023 |  |  | 09/07/2023 | 10/04/2023 |  |  | 12/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 38-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was found to have left ventricle thrombus and thought to be causing the suction alarms on the impella. the physician repositioned the impella and gave iv fluids throughout the day without suction alarms resolving. the impella was removed, and a clot was found in the cannula. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issues has been completed. the device was not returned for investigation. data logs confirmed the failure mode & clinical narrative. logs showed after pump started, mc spiked followed low flow that triggered auto suction response & suction alarms. based on clinical description & data review, the root cause of low flow was most likely due to biomaterial ingestion.
34000 | 1220648-2023-03677 | 17867968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17867968&pc=OZD | 09/07/2023 | 09/18/2024 |  | 10/04/2023 |  |  | 09/07/2023 | 10/04/2023 |  |  | 04/30/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024345194 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 94 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A 94-year-old patient with many underlying comorbidities received hemodynamic support from an impella cp pump due to acute myocardial infarction/cardiogenic shock. three days after insertion of the impella cp, the patient had recovered sufficiently that removal of the impella heart pump was attempted. when the impella was pulled from the left ventricle into the aorta, ventricular fibrillation occurred. the patient required cardiopulmonary resuscitation (cpr) followed by a high dose of catecholamines. the impella heart pump was able to be reinserted into the left ventricle, but the patient did not survive the event. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ventricular fibrillation has been completed. the device nor sufficient clinical information was returned for evaluation. the relationship between the arrhythmia and death was not specified. therefore, the root cause of the ventricular fibrillation could not be determined. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ added: b2 selected death and required intervention to prevent permanent damage. h6-impact code 1802 h1-updated to death. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
34001 | 1220648-2023-03678 | 17867993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17867993&pc=OZD | 05/13/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 09/28/2023 | 10/04/2023 |  |  | 01/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186135 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient on impella 5.5 support was found to be hit+ and had concern for clots in graft. physician removed the impella and graft completely, and attached new graft. patient was supported with milrinone while off impella support and a new impella 5.5 inserted. | Upon review of the noted case, it was documented that leading up to explant, persistent high purge pressures were encountered for which tpa was administered. the pump was subsequently explanted along with the graft due to concerns of a potential clot. patient was supported with milrinone while off impella support and a new impella 5.5 and graft were inserted. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information added to section b, g4, and h6 pertaining to the report of a high purge pressure. upon investigation completion, a supplemental report will be filed. | The investigation for the reported thrombocytopenia and high purge pressure issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34002 | 1220648-2023-03679 | 17868036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868036&pc=OZD | 09/07/2023 | 12/17/2024 |  | 10/04/2023 |  |  | 09/08/2023 | 10/04/2023 |  |  | 05/31/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024356122 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2023 |  |  |  |  | The user facility in japan reported an 83-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the impella had high purge pressure alarms and a kink in the catheter was noted. the impella cp was replaced. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into high purge pressure and product damage (cannula kink) has been completed. the product was returned for investigation. full data logs were not provided for this case. the provided data logs were missing the case start activity, and purge pressure was observed to be high. no physical abnormalities were observed on the returned product. the cannula also showed no noticeable kinks. the pump was successfully primed, and high purge pressure was reproduced during the case. the pump was cut near the motor housing, and purge was seen coming out of the cut catheter. further investigation revealed the biomaterial in the purge gap. as per the clinical details, high purge pressure was noted at the start of the case. the purge cassette replacement did not resolve the issue. the cannula kink failure mode was ruled out after investigating the product, and the clinical team did not report any damage on the cannula. the cause of the high purge pressure issue was most likely biomaterial, based on returned product investigation. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03679. d6a and d6b revised from the initial submission in accordance with updated procedures.
34003 | 1220648-2023-03680 | 17868086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868086&pc=OZD | 09/28/2022 | 09/13/2024 |  | 10/04/2023 |  |  | 09/28/2022 | 10/04/2023 |  |  | 05/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023228124 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient on impella support experienced a stroke/bleed and passed away. the pump had supported for just under 7 days. the patient had been admitted in cardiogenic shock and with kidney function issues. | Procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | B2: date of death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03631 was provided.
34004 | 1220648-2023-03681 | 17868134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868134&pc=OZD | 12/13/2021 | 09/20/2024 |  | 10/04/2023 |  |  | 12/13/2021 | 10/04/2023 |  |  | 08/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022137423 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2021 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced ventricular tachycardia (vt) and fibrillation (vf) as a result of suction on impella 5.5. the patient was shocked 5 times. the impella was repositioned, pulled back 1 cm, vt temporarily calm down. | Additional information noted that care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03681 was provided in sections b2, b5, d6a, d6b, h1, and h6.
34005 | 1220648-2023-03683 | 17868202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868202&pc=OZD | 11/23/2021 | 09/26/2024 |  | 10/04/2023 |  |  | 11/23/2021 | 10/04/2023 |  |  | 08/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130666 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2021 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support, via the left axillary access. the complainant reported that the patient experienced clots in the distal extremities, left upper axillary and brachial arteries, along with pulmonary embolism (pe) in left and right lobes. six days after this event the patient condition had deteriorated and decision was made to withdraw care. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34006 | 1220648-2023-03684 | 17868241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868241&pc=OZD | 09/15/2021 | 10/04/2023 |  | 10/04/2023 |  |  | 09/15/2021 | 10/04/2023 |  |  | 09/30/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1505139 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/15/2021 |  |  |  |  | The user facility reported a 39-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed an acute thrombus during the course of impella support. the thrombus was discovered via transesophageal echocardiogram (tee) imaging and the decision was made to explant the pump. extracorporeal membrane oxygenation (ecmo) was subsequently placed for ongoing patient support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the clinical report was reviewed and the cause of the thrombus was patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34007 | 1035166-2023-00100 | 17868244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868244&pc=DYB | 10/06/2022 | 10/04/2023 |  | 10/04/2023 |  |  | 09/12/2023 | 10/04/2023 |  |  | 12/01/2023 |  | 10/04/2023 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | C1-19947 |  | 0052-3006 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | We have been notified of a complaint involving an oscor 23 fr introducer (material #0052-3006, lot # c1-19947). during impella 5.5 implant, the introducer cracked and began leaking. the introducer was replaced and the subsequent impella 5.5 implant was successful. there was no patient harm and no intervention. additional devices used during the procedure include impella 5.5 pump, (b)(6), lot # 2022043492, product number 0550-0008. there was minimal delay. swapped introducer and implant was subsequently successful. standard technique used for insertion. additional information received 09/14/2023. the impella 5.5 pump was the only device inserted through the sheath at the time of the event. the procedure was completed with model number and lot number 0052-3006, c1-19943. no issue reported regarding patients' anatomy. product was discarded; therefore, will not be returned. | [5 empty] | No product was provided for analysis as product was discarded by customer. there was no patient harm and no intervention reported. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. no product was returned to oscor inc. for evaluation; however, a complaint notification was provided. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. the instructions for use (ifu): never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. based on the investigation, a capa is not required. the event will be re-evaluated if additional information becomes available. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
34008 | 1220648-2023-03682 | 17868263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868263&pc=OZD | 11/26/2021 | 01/11/2024 |  | 10/04/2023 |  |  | 11/26/2021 | 10/04/2023 |  |  | 08/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139026 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2021 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that heparin was on hold due to bleeding issues that happened overnight and partial thromboplastin time (ptt) was over 100. it was also noted that due to a hematoma at the access site, the patient was going to the operation room for exploration. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. the root cause of the access site bleeding could not be determined since the product was not returned and enough clinical information was not provided.
34009 | 1220648-2023-03685 | 17868385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868385&pc=OZD | 04/02/2021 | 09/15/2024 |  | 10/04/2023 |  |  | 04/02/2021 | 10/04/2023 |  |  | 02/28/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022044408 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/02/2021 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the impella 5.5 pump was triggering persistent positioning alarms during patient support. the alarm was subsequently disabled after correct positioning was verified via echocardiogram. after three to four days of support, it was documented that the patient experienced a brain bleed. a craniotomy was performed following discovery of the brain bleed. a definitive cause for the brain bleed could not be determined. support with the pump continued until discussions with the patient's family and patient was allowed to expire. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. it is reported that was a hemorrhage type of stroke which caused the bleeding in the brain. they have done the craniotomy the following morning but subsequently the patient expired. it was noted that heparin was used in the purge stream. there is no evidence that relates the use of heparin has caused the stroke. hence, the root cause of the stroke is patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the investigation was made with only clinical details furnished by the medical center. the root cause was patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03685 was provided.
34010 | 1220648-2023-03686 | 17868422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868422&pc=OZD | 09/20/2021 | 09/17/2024 |  | 10/04/2023 |  |  | 09/20/2021 | 10/04/2023 |  |  |  |  | 10/04/2023 |  | AUTOMATED IMPELLA CONTROLLER¿(AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084503 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Computer Operating System Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/20/2021 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the nurse had issues with air in the purge line in the unit and was unable to de-air effectively. shortly thereafter the nurse noticed an automated impella controller (aic) failure notification on the console, and so she had to change it out for a new aic unit. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the aic - purge issue ¿ other has been completed. the root cause of the purge issues was due to a defective purge assembly. the exact subcomponent of the purge assembly was not determined. | The impella console was not received from the customer, therefore, investigation of the device was not possible. a supplemental mdr will be filed at the conclusion of the investigation. this report is being filed as part of a retrospective review of historical records. | B2: date of death unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03686 was provided.
34011 | 1220648-2023-03687 | 17868439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868439&pc=OZD | 09/24/2021 | 01/24/2024 |  | 10/04/2023 |  |  | 09/24/2021 | 10/04/2023 |  |  | 07/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113195 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2021 |  |  |  |  | The user facility reported a 30-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. there was a large hematoma at the insertion site. a total of 3 units of packed red blood cells (prbc) were given. it was further reported that during the impella placement blood products were also given to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details.
34012 | 1220648-2023-03688 | 17868472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868472&pc=OZD | 04/20/2021 | 01/24/2024 |  | 10/04/2023 |  |  | 04/20/2021 | 10/04/2023 |  |  | 01/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022043261 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2021 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella 5.5 pump came out of position during patient support. as the patient had stabilized, the pump was explanted with no plans for insertion of another device. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the positioning issue was most likely use related. impella was found dislocated after 8 days as per the clinical description. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34013 | 1220648-2023-03689 | 17868524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868524&pc=OZD | 12/16/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 12/16/2022 | 10/04/2023 |  |  | 06/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247645 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2022 |  |  |  |  | The complainant reported a 67-year-old male patient experiencing post cardiotomy cardiogenic shock was presented for impella 5.5 placement. upon completions of support, it was noted that the doctor encountered resistance when attempting to explant the pump. after multiple unsuccessful attempts, the patient was reintubated and the sternum was reopened to visualize the aortic anastomosis site. it was observed that the impella was lodged within the anastomosis location. the pump was subsequently withdrawn and the sites were closed appropriately. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported removal issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34014 | 1220648-2023-03691 | 17868563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868563&pc=OZD | 05/03/2021 | 02/21/2025 |  | 10/04/2023 |  |  | 05/03/2021 | 10/04/2023 |  |  | 10/31/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022034204 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 12 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2021 |  |  |  |  | The user facility in the eu reported a 12-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a thrombus developed on the inlet of the impella 5.5 pump during patient support. the pump was explanted as a result of the noted issue. a new impella 5.5 pump was subsequently implanted for ongoing support, along with ecmo. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
34015 | 1220648-2023-03690 | 17868607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868607&pc=OZD | 04/29/2021 | 01/23/2024 |  | 10/04/2023 |  |  | 04/29/2021 | 10/04/2023 |  |  | 11/30/2021 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1466760 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2021 |  |  |  |  | The user facility reported a 47-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a clot in the apex of their left ventricle during impella 5.5 support. as a result of the clot, the decision was made to explant the pump. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the cause of the thrombus was patient condition since a clot was discovered in the apex of left ventricle (lv). | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34016 | 1220648-2023-03692 | 17868611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868611&pc=OZD | 05/28/2021 | 01/11/2024 |  | 10/04/2023 |  |  | 05/28/2021 | 10/04/2023 |  |  | 03/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060397 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2021 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support, and patient was experiencing oozing from the insertion site. a drain was placed. thought in the field was that the patient vomited potentially breaking open small capillary. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. per the provided clinical information, the cause for access site bleeding is patient movement since the patient vomited causing to break open small capillary.
34017 | 1220648-2023-03693 | 17868632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868632&pc=OZD | 12/10/2022 | 09/29/2025 |  | 10/04/2023 |  |  | 12/10/2022 | 10/04/2023 |  |  | 06/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234690 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Material Twisted/Bent | [5 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2022 |  |  |  |  | The user facility reported a 58-year-old female experiencing cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, the pump triggered high purge pressure alarms and upon inspection it was determined that there was some bending on the proximal side of the red plug. the rn used a 3 point fixation to prevent this issue from occurring again. the patient was found to have a hematoma at the access site which appeared to be improving after starting heparin. in addition, the patient was found to have issues with hemolysis and the pump had to be explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03693 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03693 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03693 in accordance with updated procedures.
34018 | 1220648-2023-03694 | 17868670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868670&pc=OZD | 05/19/2021 | 10/09/2025 |  | 10/04/2023 |  |  | 05/19/2021 | 10/04/2023 |  |  | 02/28/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022047462 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/19/2021 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support the patient experienced suction issues which resolved when impella was turned down to p-8. | The procedural outcome noted that the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | The complainant reported that while a patient was on impella support, the patient experienced continuous suction issues which resolved when impella was turned down to p-8. however, eventually the patient expired. |  |  |  | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
34019 | 1220648-2023-03695 | 17868754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868754&pc=OZD | 05/28/2021 | 03/10/2025 |  | 10/04/2023 |  |  | 05/28/2021 | 10/04/2023 |  |  | 03/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022065486 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2021 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. multiple issues were encountered by the patient during attempted impella use. while attempting to implant the device, resistance was met proximal to the anastomosis and the pump had to be removed. flouro and contrast was subsequently utilized and it was noted that the subclavian artery had disease. the patient's right arm was manipulated and placement was successfully completed. however, upon starting the pump and increasing support levels, suction was encountered, and the patient's mean arterial pressure (map) dropped to the 30s. volume was provided, however the issue persisted. the position was checked, and it was discovered that the mitral valve was being impeded by the pump. the device was repositioned, and the issue resolved. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow, positioning and delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue- anatomy was most likely due to patient condition. the root cause of the positioning issue was determined to be use/ technique. low pump flow was caused due to the improper positioning and from the clinical that repositioning improved the flows. hence, the root cause is determined to be improper use/ technique. this report is being filed as part of a retrospective review of historical records.
34020 | 1220648-2023-03696 | 17868826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868826&pc=OZD | 06/20/2021 | 01/24/2024 |  | 10/04/2023 |  |  | 06/20/2021 | 10/04/2023 |  |  | 03/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064501 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2021 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient on impella support experienced a tear in the sterile sleeve. there was no harm reported as a result of this incident. | [5 empty] | The investigation into the reported sterile sleeve damage has been completed since the original report was filed. the cause of the tear cannot be determined as the product cannot be retrieved. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34021 | 1220648-2023-03697 | 17868835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868835&pc=OZD | 11/18/2022 | 01/15/2024 |  | 10/04/2023 |  |  | 11/18/2022 | 10/04/2023 |  |  | 02/29/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196466 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that complications were encountered while attempting to implant an impella 5.5 pump for escalated support. it was documented that after obtaining access to the left ventricle with the .018 guidewire, the doctor tried unsuccessfully to implant the pump for roughly an hour. due to the noted difficulty, the patient suffered bleeding from the site and required an unspecified amount of blood products. the implant was subsequently aborted as a result of the blood loss and the time that it was taking to implant the device. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported delivery issue and access site bleeding has been completed since the original report was filed. no product was returned. delivery use- the root cause of the delivery issue - anatomy was most likely due to the patient condition access site bleeding - no product was returned. the root cause of the access site bleeding issue was patient condition related, as there was loss of blood at access site with blood products been given due to the difficult inserting attempts taking approximately an hour as per the clinical description provided.
34022 | 1220648-2023-03698 | 17868866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868866&pc=OZD | 05/28/2021 | 07/29/2025 |  | 10/04/2023 |  |  | 05/28/2021 | 10/04/2023 |  |  | 03/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064497 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2021 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support, and that the device¿s sidearm luer cracked, after 15 days of support. the device was replaced with another 5.5 without any issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | D6a and d6b revised from the on the initial manufacturer device report number 1220648-2023-03698 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03698. g1 reporting contact email revised on manufacturer device report 1220648-2023-03698 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03698.
34023 | 1220648-2023-03699 | 17868892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868892&pc=OZD | 06/03/2021 | 01/26/2024 |  | 10/04/2023 |  |  | 06/03/2021 | 10/04/2023 |  |  | 03/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064512 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2021 |  |  |  |  | The user facility reported a 50-year-old patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient was transported to the operating room due to the catheter migrating into the aorta. the issue was corrected, and the patient did well. it was further reported that the patient was taken again to the operating room for a wash out and evacuation of a hematoma at the site. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning and access site bleeding. the device was not returned for evaluation. however, the clinical information was sufficient to determine root cause of the positioning. the root cause of the positioning issue was use related, due to the catheter migrating into aorta as per the clinical description. the clinical information was not sufficient to determine the root cause of the access site bleeding. therefore the root cause of the access site could not be determined.
34024 | 1220648-2023-03700 | 17868935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868935&pc=OZD | 12/02/2022 | 10/28/2025 |  | 10/04/2023 |  |  | 12/02/2022 | 10/04/2023 |  |  | 06/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023240247 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that during placement of the impella device, there was extreme difficulty passing wires and catheters through the anastamosis. as a result, the device was removed and the graft was un-tunneled. another attempt was made for placement but there was still difficulty crossing the valve. propofol was used to lubricate and the impella was placed without issue. | The complainant reported a 67-year-old male was presented for impella 5.5 placement. following implant, it was documented that the patient experienced significant ectopy as support levels were increase to higher levels. the pump was verified to be in an optimal position, but the issue persisted. the patient was given epinephrine, vasopressin, methylene blue and calcium throughout the procedure. a manageable support level was eventually established with the implanted pump. |  |  |  |  | The investigation for the reported delivery issues and ectopy has been completed. the device was not returned, however, clinical information was sufficient to determine the root cause. according to clinical details, there was difficulty crossing the anastamosis and the valve, but once in optimal positioning, there was still some ectopy. the root cause of the delivery issue and the ectopy was due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | In order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
34025 | 1220648-2023-03701 | 17868936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868936&pc=OZD | 06/07/2021 | 03/10/2025 |  | 10/04/2023 |  |  | 06/07/2021 | 10/04/2023 |  |  | 12/31/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022029476 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2021 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. medical staff had difficulty inserting the impella pump due to the patient's anatomy. after a few failed attempts the patient was taken to the catheterization lab where impella was implanted successfully. five days after the impella implant the computed tomography scan showed that the patient had possibly experienced a stroke. seven days after that the patient was shown to have bloody bowel which was suspected to be from hemolysis and several blood units had to be administered. the md decided to slowly wean the patient off impella and finally decision was made to completely withdraw care. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes. may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis, stroke/cva and delivery issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was determined to be the patient condition. since it was mentioned that the patient has high grade stenosis in the right femoral artery and a heavily calcified left femoral, so impella placement was attempted through right innominate artery in or but was unsuccessful pass the impella 5.5. it was noted that it was a hemorrhage type of stroke which caused the bleeding in the brain. it was noted that heparin was used in the purge stream. there is no evidence that relates to the use of heparin has caused the stroke. therefore, it is an unknown relation to heparin that caused the stroke and the root cause of the stroke is determined to be patient condition. the root cause of the bloody bowel moment (hemolysis) was not determined. there was not enough information that hemolysis was caused by impella and the product was not returned to inspect the root cause and hence, the root cause is undetermined. this report is being filed as part of a retrospective review of historical records.
34026 | 1220648-2023-03702 | 17868983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17868983&pc=OZD | 07/12/2021 | 03/10/2025 |  | 10/04/2023 |  |  | 07/12/2021 | 10/04/2023 |  |  | 02/28/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054965 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2021 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a low purge alarm which was later resolved after changing the purge bag. the patient also experienced ventricular tachycardia allegedly caused by the impella being directed slightly posterior. device eventually explanted after 82 days of support. the patient bridged to transplant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported purge leak and vf/vt/af/at issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34027 | 1220648-2023-03704 | 17869025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869025&pc=OZD | 09/27/2021 | 01/26/2024 |  | 10/04/2023 |  |  | 09/27/2021 | 10/04/2023 |  |  | 02/28/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022046209 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2021 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed oozing from their access site coinciding with ecchymosis. the patient underwent two washouts and a jackson-pratt drain pulled 500cc of fluid. to counteract the blood loss, an unspecified quantity of packed red blood cells and platelets were given to the patient. the pump was explanted the following day. | [5 empty] | The investigation for the reported access site bleeding has been completed. without the impella device or sufficient clinical details, the root cause of the access site bleeding was unable to be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34028 | 1220648-2023-03705 | 17869061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869061&pc=OZD | 05/25/2021 | 09/17/2024 |  | 10/04/2023 |  |  | 05/25/2021 | 10/04/2023 |  |  | 12/03/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060317 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/25/2021 |  |  |  |  | The user facility reported a 39-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support and was unable to maintain flows about p-4 without suction. the patient also developed septic shock. decision made to withdraw care. device was explanted. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03705 was provided. | The investigation into the reported infection/sepsis and low pump flow has been completed since the original report was filed. the device was not returned by the medical facility. low pump flow - the root cause for low pump flow was patient condition as the rv failure had developed as flow could not be maintained above p-4 without suction. infection/ sepsis - the cause for the septic shock cannot be determined as there is not enough information available and product evaluation is not possible since the product cannot be retrieved. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
34029 | 1220648-2023-03707 | 17869146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869146&pc=OZD | 11/08/2022 | 01/26/2024 |  | 10/04/2023 |  |  | 11/08/2022 | 10/04/2023 |  |  | 06/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232903 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Difficult to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following a transcatheter aortic valve replacement (tavr), the 92-year old female began to decompensate, and a decision was made to place patient on impella 5.5 support. during the procedure, the medical staff had difficulty delivering an impella pump due to the patient's anatomical limitations. after several attempts, bleeding was found at the groin that had to be repaired. after such complications, the procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the delivery issue and access site bleeding - major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the delivery issue - anatomy was most likely due to the patient condition. however, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34030 | 1220648-2023-03706 | 17869148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869148&pc=OZD | 05/24/2021 | 10/30/2025 |  | 10/04/2023 |  |  | 05/24/2021 | 10/04/2023 |  |  | 02/28/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022052505 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/24/2021 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support and experienced suction alarms which were resolved after the physician repositioned the device. there was no harm to the patient as a result of this malpositioning. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical (patient had pre-existing ventricular septal defect). |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03706 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03706 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03706 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03706. b2 death and date of death on manufacturer device report number 1220648-2023-03706-1 h1 type of report was erroneously selected on manufacturer device report number 1220648-2023-03706-1.
34031 | 1220648-2023-03708 | 17869189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869189&pc=OZD | 07/04/2021 | 03/10/2025 |  | 10/04/2023 |  |  | 07/04/2021 | 10/04/2023 |  |  | 04/30/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022074306 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/04/2021 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a pump stop alarm. the pump was restarted and it was running normal up to p-8. at p-9 it was getting flow saturation higher than normal and the pump failed. a new impella 5.5 was placed after into the same graft. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34032 | 1220648-2023-03709 | 17869256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869256&pc=OZD | 07/16/2021 | 01/23/2024 |  | 10/04/2023 |  |  | 07/16/2021 | 10/04/2023 |  |  | 12/30/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022026880 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Pressure Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2021 |  |  |  |  | The user facility reported a 28-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was found to have a clot found on tricuspid valve and three days after that experienced a purge system blocked alarm. device was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported thrombus and high purge pressure has been completed since the original report was filed. the device was not returned by the medical facility. thrombus - the cause of the thrombus was patient condition since there is no relation between impella 5.5 and clot found on tricuspid valve which is in rv. high purge pressure - the cause of high purge pressure was most likely due to the kink was in the purge cassette tubing as switching the purge cassette resolved the issue.
34033 | 1220648-2023-03710 | 17869295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869295&pc=OZD | 07/05/2021 | 01/11/2024 |  | 10/04/2023 |  |  | 07/05/2021 | 10/04/2023 |  |  | 04/30/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071393 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2021 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella was positioned against mitral valve and after multiple attempts to reposition over wire md removed impella and reaccessed left ventricle again. the second time impella positioned nicely away from mitral into the free space of the left ventricle. there was no harm to the patient as a result of this event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the product was not returned. clinical notes detail an access site bleed upon insertion of impella cp; the patient had high act. the cause of the access site bleeding was high patient act.
34034 | 1220648-2023-03711 | 17869396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869396&pc=OZD | 07/06/2021 | 10/04/2023 |  | 10/04/2023 |  |  | 07/06/2021 | 10/04/2023 |  |  | 05/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088304 |  | 0550-0008 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/06/2021 |  |  |  |  | The user facility reported a 72-year-old males undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had a clot develop and observed in the left ventricle. bicarb was ordered for purge and transfusions protocol ordered. ecmo was placed when right ventricle failed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the clinical report alone was analyzed. the cause for the thrombus is patient condition since the right ventricle fail and clot was found in left ventricle. the failure mode will be monitored and trended. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34035 | 1220648-2023-03712 | 17869403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869403&pc=OZD | 11/04/2022 | 01/25/2024 |  | 10/04/2023 |  |  | 11/04/2022 | 10/04/2023 |  |  | 02/29/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196025 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. following placement of an impella 5.5, the patient was returned to the or for ongoing bleeding. anastomosis of one the grafts was found to be the issue and was repaired by the surgeon. support with the implanted pump was maintained throughout the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. data logs were returned and investigated. per the clinical information provided, the root cause of the access site bleeding issue was most likely patient condition related to patient anatomy. data logs show that there was a small motor current spike and then high purge pressure, and purge system was blocked alarms. multiple bag changes and were then done. the issue was resolved with a decrease in the p-level of the pump and starting systemic heparin. the root cause of the high purge pressure issue was most likely biomaterial.
34036 | 1220648-2023-03713 | 17869556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869556&pc=OZD | 08/06/2021 | 01/11/2024 |  | 10/04/2023 |  |  | 08/06/2021 | 10/04/2023 |  |  | 10/31/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021006136 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/06/2021 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient developed a right side hematoma during impella support. the patient was taken to the operating room and that venous bleeding was the source for the hematoma. unspecified intervention was performed to treat the condition and support with the pump continued. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was not determined. it was mentioned that hematoma was developed and the patient was taken to or. the cause of hematoma was unknown and no clinical data was reported to support the cause was related to impella and the relation to impella was unknown. hence, the root cause is undetermined.
34037 | 1220648-2023-03714 | 17869607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869607&pc=OZD | 10/18/2022 | 09/19/2024 |  | 10/04/2023 |  |  | 10/18/2022 | 10/04/2023 |  |  | 05/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217151 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 40-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that complications were encountered during delivery of an impella 5.5 pump. while advancing the pump through the patient's vasculature, it was noted that the tortuous graft tunneling was causing the wire and catheter to kink. despite the pump eventually making its way into the left ventricle, the wire could not be removed and both devices had to be removed simultaneously. a .035 j wire was then placed and the device was loaded and implanted successfully. support with the 5.5 was initiated and the existing cp pump was removed. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the delivery issues- use and the product damage (guidewire damage - great vessel) issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issues is most likely due to the patient's condition of diseased/tortuous vessels causing the guidewire and catheter to kink. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03714 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34038 | 1220648-2023-03716 | 17869676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869676&pc=OZD | 10/21/2022 | 01/25/2024 |  | 10/04/2023 |  |  | 10/21/2022 | 10/04/2023 |  |  | 06/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247452 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced a hematoma and pain at the access site, signs of bleeding. heparin was held off and the site was going to be evaluated by surgery. support was maintained throughout the event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
34039 | 1220648-2023-03717 | 17869977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17869977&pc=OZD | 07/11/2021 | 03/04/2024 |  | 10/04/2023 |  |  | 07/11/2021 | 10/04/2023 |  |  |  |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2021 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support showed to have questionable neurological status due to suspected stroke with seizure. cerebral vascular accident (cva) ruled out via computed tomography (ct) scan twice. although there was no confirmation for the stroke by the ct scan. the device eventually was explanted 4 days after this event. patient survived event. | [5 empty] | The investigation into the reported stroke/cva has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the stroke/cva was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34040 | 1220648-2023-03718 | 17870404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870404&pc=OZD | 10/11/2022 | 10/10/2023 |  | 10/04/2023 |  |  | 10/11/2022 | 10/04/2023 |  |  | 07/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248856 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 17-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed an access site bleed during impella 5.5 support. the patient was given an unspecified amount of blood products to treat the condition. support continued with the implanted pump following the intervention. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, the root cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the failure mode will be monitored and trended. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34041 | 1220648-2023-03719 | 17870430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870430&pc=OZD | 10/24/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 10/24/2022 | 10/04/2023 |  |  | 07/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023250263 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient had developed hemolysis during impella 5.5 support. the hemolysis was said to have compromised the patient's kidneys and the pump was explanted after four days of support. additional details pertaining to the hemolysis and/or potential intervention were not provided. the patient was reported to be stable following explant. | [5 empty] | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34042 | 1220648-2023-03720 | 17870463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870463&pc=OZD | 10/12/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 10/12/2022 | 10/04/2023 |  |  | 06/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023239090 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2022 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced saturated flows when powered on during patient support. as a result of the noted issue, the pump was removed and replaced with another impella 5.5 pump. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported saturated flow issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34043 | 1220648-2023-03721 | 17870578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870578&pc=OZD | 07/25/2022 | 09/22/2025 |  | 10/04/2023 |  |  | 07/25/2022 | 10/04/2023 |  |  | 03/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208304 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at their access site during impella support. the patient was taken to the operating room for exploration and evacuation of the hematoma. this resulted in the removal of serous fluid and the evacuation of blood from the pectoral muscle. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03721 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03721 in accordance with updated procedures. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03721. g1 reporting contact email revised on manufacturer device report 1220648-2023-03721 in accordance with updated procedures. h6 health effect - clinical code 1930 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03721. h6 component code revised from the initial submission in accordance with updated procedures.
34044 | 1220648-2023-03722 | 17870641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870641&pc=OZD | 08/05/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 08/05/2022 | 10/04/2023 |  |  | 02/29/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023189492 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2022 |  |  |  |  | The user facility reported a 35-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced persistent suction during patient support. volume was given and the pump was manipulated to be in a shallow position in an attempt to avoid suction and ectopy. as a result of the shallow positioning, the pump came back across the valve when the patient was moved. the patient remained hemodynamically stable despite the lack of support and the decision was made to explant the pump. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34045 | 1220648-2023-03724 | 17870676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870676&pc=OZD | 09/25/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 09/25/2022 | 10/04/2023 |  |  | 05/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023226480 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2022 |  |  |  |  | The user facility reported a 43-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during the attempted delivery the graft failed causing an artery dissection. a bypass intervention was required and the procedure was abandoned. | [5 empty] | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vascular damage - great vessel and delivery issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported vascular damage issue could not be determined. the cause of the delivery issue was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
34046 | 1220648-2023-03723 | 17870718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870718&pc=OZD | 08/04/2022 | 01/25/2024 |  | 10/04/2023 |  |  | 08/04/2022 | 10/04/2023 |  |  | 03/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203510 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma under their right arm and access site bleeding during impella support. a washout was performed and a jp drain was placed to treat the hematoma. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
34047 | 1220648-2023-03725 | 17870783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870783&pc=OZD | 09/14/2022 | 09/22/2024 |  | 10/04/2023 |  |  | 09/14/2022 | 10/04/2023 |  |  | 05/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023228231 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a purge door open alarm and a leak coming out of the yellow luer lock connector of the purge sidearm. the md wanted to wean the patient out of impella but decision was made not to change the pump that night and would only do so if the pump stopped. the alarm was resolved by following the normal bypass instructions. five days after this event the patient condition had not improved and decision was made to withdraw care. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir". ¿clean the connector cable with 70% isopropyl alcohol¿. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03725 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34048 | 1220648-2023-03726 | 17870847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870847&pc=OZD | 07/12/2022 | 10/21/2025 |  | 10/04/2023 |  |  | 07/12/2022 | 10/04/2023 |  |  | 03/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210365 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed hemolysis during impella support. with the resulting low hgb levels, the patient was given packed red blood cells (prbc) to treat the condition. discussions were had to reposition the pump, but the decision was ultimately made to explant the pump a couple days later. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the cause of the hemolysis can not be determined since the product was not returned and data showed no mc spike or positioning issue.
34049 | 1220648-2023-03727 | 17870887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870887&pc=OZD | 08/15/2022 | 09/30/2025 |  | 10/04/2023 |  |  | 08/15/2022 | 10/04/2023 |  |  | 04/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023213171 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced high purge pressure coinciding with a low purge flow during impella support. after three days of increased purge pressure, the staff was informed of the risks involved with the persistent issue. the decision was subsequently made to explant the pump and plans were made to place a durable lvad for ongoing support. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03727 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03727 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03727 in accordance with updated procedures. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03727. h6 medical device problem code 1248 was erroneously entered on the initial manufacturer device report number 1220648-2023-03727.
34050 | 1220648-2023-03728 | 17870980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17870980&pc=OZD | 07/14/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 07/14/2022 | 10/04/2023 |  |  | 04/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211990 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella patient was losing high volumes and appeared to be bleeding around the suture line at the graft anastomosis. further observation noted that the vessel was heavily damaged. the bleeding was controlled after the artery was surgically repaired with sutures and vein graft pledgets. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vascular damage- peripheral vessel issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the vascular damage was most likely due to patient condition since the vessel appeared to be disintegrating with lot of blood loss. this report is being filed as part of a retrospective review of historical records.
34051 | 1220648-2023-03729 | 17871015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17871015&pc=OZD | 07/17/2022 | 09/29/2025 |  | 10/04/2023 |  |  | 07/17/2022 | 10/04/2023 |  |  | 04/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210782 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the patient suffered a stroke. the patient required increased inotropic support and they were eventually weaned off of support. | [5 empty] | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03729 in accordance with updated procedures. f6/f8. should have been left blank on the initial manufacturer device report number 1220648-2023-03729 in accordance with updated procedures. g1. reporting contact email revised on manufacturer device report 1220648-2023-03729 in accordance with updated procedures. g1. reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03729. h6. component code revised from the initial submission in accordance with updated procedures.
34052 | 1220648-2023-03730 | 17871209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17871209&pc=OZD | 07/29/2022 | 02/18/2025 |  | 10/04/2023 |  |  | 07/29/2022 | 10/04/2023 |  |  | 10/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023161345 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a 36-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a gastrointestinal bleed during impella support. the patient was temporarily taken off heparin/anti-coagulants as a result of the condition. impella support was maintained throughout the event and the patient was considered stable. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation...¿ cp and 5.5 ifu: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the reported issue was unable to be determined as no product was returned for investigation.
34053 | 1220648-2023-03731 | 17871307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17871307&pc=OZD | 06/29/2022 | 03/08/2024 |  | 10/04/2023 |  |  | 06/29/2022 | 10/04/2023 |  |  | 03/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210223 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced stroke like symptoms during impella support. despite impella flows remaining unimpeded, a computed tomography(ct) scan confirmed the presence of a right proximal carotid thrombus. the pump was explanted as a result of the condition and a balloon pump was placed for ongoing support. the fogarty balloon verified clearance of any thrombus from the graft and right subclavian between the vessel loops. the patient was considered to be stable following the intervention. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34054 | 1220648-2023-03732 | 17871338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17871338&pc=OZD | 06/18/2022 | 01/26/2024 |  | 10/04/2023 |  |  | 06/18/2022 | 10/04/2023 |  |  | 01/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182478 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was documented that the patient experienced bleeding from their access site . the patient was taken to the operating room where an unspecified procedure was undertaken to resolve the bleed. the source of the bleeding was unknown. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34055 | 1220648-2023-03733 | 17871494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17871494&pc=OZD | 06/30/2022 | 10/04/2023 |  | 10/04/2023 |  |  | 06/30/2022 | 10/04/2023 |  |  | 09/30/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143604 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that during attempted delivery of an impella 5.5 pump, the md was unable to control the bleeding at the cut down/anastomosis despite the use of several sealants and clotting agents. as a result, the 5.5 implant was aborted and support continued with the already in-place impella cp pump. | [5 empty] | The impella device was not by the customer and therefore, an evaluation of the device was not possible. an investigation was completed and, based on the provided information, it was advised that the noted bleed was attributed to the patient's condition. should additional relevant information become available, a supplemental report will be submitted. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34056 | 1220648-2023-03734 | 17872116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872116&pc=OZD | 07/08/2022 | 10/09/2025 |  | 10/04/2023 |  |  | 07/08/2022 | 10/04/2023 |  |  | 01/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023179483 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced a gradual bleed from multiple sites, including the impella access site, during pump support. blood transfusions were performed as needed based on the patient's hemoglobin levels. it was noted that a coagulation abnormality was present in the patient at the time of the noted condition. support continued with the implanted pump throughout the event. | The complainant reported that the patient experienced a gradual bleed from multiple sites, including the impella access site, during pump support. blood transfusions were performed as needed based on the patient's hemoglobin levels. it was noted that a coagulation abnormality was present in the patient at the time of the noted condition. support continued with the implanted pump throughout the event. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. no product was returned. the root cause of the access site bleeding issue was most likely patient condition related as the clinical details mention that the patient had a coagulation abnormality.
34057 | 1220648-2023-03735 | 17872279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872279&pc=OZD | 06/29/2021 | 01/22/2024 |  | 10/04/2023 |  |  | 06/29/2021 | 10/04/2023 |  |  | 12/31/2022 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022027989 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2021 |  |  |  |  | The user facility reported a 14-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team had difficulty inserting impella through the graft and after several attempts aborted the procedure for concerns for the integrity of the catheter shaft. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. the device was not returned by the medical facility. the cause of the delivery issue was patient condition since narrowing (stenosis) was observed. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34058 | 1220648-2023-03736 | 17872489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872489&pc=OZD | 07/01/2021 | 09/18/2024 |  | 10/04/2023 |  |  | 07/01/2021 | 10/04/2023 |  |  | 03/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022067247 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/01/2021 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced significant amount of bleeding and was required to receive three units of blood. the patient went into asystole and had cpr performed. the patient survived the event and stayed on impella support for 5 more days. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
34059 | 1220648-2023-03737 | 17872559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872559&pc=OZD | 07/17/2021 | 03/10/2025 |  | 10/04/2023 |  |  | 07/17/2021 | 10/04/2023 |  |  | 02/28/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022050215 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2021 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was found to have a thrombus in the left ventricle (lv). the following day the medical team had a computed tomography (ct) scan performed and showed that the patient had bilateral hematoma and they had to stop heparin and give the patient 2 units of blood. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus and access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. as per clinical review lv thrombus was found to be present in patient and also bilateral hematoma was observed. the cause of the thrombus issue was most likely patient condition based on the clinical information provided. access site bleeding- the root cause of the access site bleeding major issue was not determined since product was not returned and limited clinical details were provided. this report is being filed as part of a retrospective review of historical records.
34060 | 1220648-2023-03738 | 17872597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872597&pc=OZD | 04/06/2023 | 10/04/2023 |  | 10/04/2023 |  |  | 04/06/2023 | 10/04/2023 |  |  | 01/31/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024294774 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  |  | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/06/2023 |  |  |  |  | The user facility reported a 37yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a high purge pressure alarm coinciding with a saturated motor current during impella support. the purge cassette was replaced in an attempt to troubleshoot the issue, however the failure remained. the pump was eventually replaced with another impella 5.5 pump for continued support. there was no patient harm reported. | [5 empty] | The investigation for the reported high pressure issue has been completed. the data log was reviewed and prior to a gradual 4 hour increase in purge pressure, suction alarms were observed. no motor current spikes or placement signal trend was observed. the prior suction is indicative of biomaterial rather than a kink/blockage in the purge line. the cause of the issue was biomaterial. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34061 | 1220648-2023-03739 | 17872657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872657&pc=OZD | 09/10/2023 | 01/26/2024 |  | 10/04/2023 |  |  | 09/10/2023 | 10/04/2023 |  |  | 05/31/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024382975 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2023 |  |  |  |  | The user facility in the eu reported a 64year-old male (unknown race) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the team noted that mild bleeding was present but not at the puncture site. the patient's leg appeared to be mottled and the foot pulses were mismatched. after consultation with the physician, a fem-fem bypass was planned. the pump remained on for support of over 6 days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of limb ischemia was most likely due to patient condition.
34062 | 1220648-2023-03740 | 17872711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872711&pc=OZD | 09/10/2023 | 09/17/2024 |  | 10/04/2023 |  |  | 09/10/2023 | 10/04/2023 |  |  | 06/30/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368917 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/10/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support post procedure. the impella 5.5 inserted via the direct access. the pump was positioned and experienced suction that prompted blood products. the team then observed a condition deterioration that prompted the team to treat the patient for left ventricle (lv) perforation. the lv perforation was due to the myocardial tissue being weakened. the patient survived. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported ventricle perforation has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on the note from the physician, the patient's tissue was noted to be too fragile to handle a device in the left ventricle. therefore, the cause of the perforation was due to patient condition. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
34063 | 1220648-2023-03741 | 17872762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872762&pc=OZD | 10/22/2022 | 10/04/2023 |  | 10/04/2023 |  |  | 10/22/2022 | 10/04/2023 |  |  | 07/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023236849 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/22/2022 |  |  |  |  | The user facility in the eu reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. when the nurse repositioned the console, he noticed that the pump stopped as a result. the patient remained hemodynamically stable, and the pump could be restarted without any problems. no harm done to the patient. | [5 empty] | The root cause of the optical signal issue was unable to be determined as no product or logs were returned for analysis. the root cause of the pump stop was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
34064 | 1220648-2023-03580 | 17872804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872804&pc=OZD | 09/05/2023 | 01/17/2024 |  | 10/04/2023 |  |  | 09/05/2023 | 10/04/2023 |  |  | 03/31/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024333669 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella support in the eu was enrolled in the protect iv study. the information from protect iv study: a patient of unknown age and gender underwent a high-risk pci procedure with hemodynamic support from an impella cp. severe anemia of uncertain origin has been reported, with patient¿s hemoglobin level on hospital admission of 92 g/l (reference 127 - 163 g/l). one day after pci and impella removal, the hemoglobin level was 84 g/l, two days after pci and impella removal the hemoglobin level was 69 g/l. pre-pci the patient received one unit of erythrocyte concentrate and post-pci two units of erythrocyte concentrate. four days post pci, the hemoglobin level was 94 g/l. the event required prolonged hospitalization. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported anemia/hemolysis has been completed since the original report was filed. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and no sufficient clinical data provided. no data logs to review. the patient had low hemoglobin level pre existing in hospital.
34065 | 1220648-2023-03577 | 17872810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872810&pc=OZD | 04/10/2022 | 01/26/2024 |  | 10/04/2023 |  |  | 04/10/2022 | 10/04/2023 |  |  | 12/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023171950 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 34 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2022 |  |  |  |  | The user facility reported a 34-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support had to be taken in the or to remove a hematoma from the access site. in addition the patient had to receive 2 units of blood. support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34066 | 1220648-2023-03578 | 17872815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872815&pc=OZD | 04/05/2022 | 09/18/2024 |  | 10/04/2023 |  |  | 04/05/2022 | 10/04/2023 |  |  | 12/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023172006 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/05/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. while on impella support, the patient experienced bleeding from the access site which prompted a blood transfusion and an access site washout. as support continued, the medical team also noted that the patient had cold feet and there were no pulses. care was subsequently withdrawn a few days later and the pump was explanted. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding and limb ischemia has been completed since the original report was filed. the cause of the access site bleeding could not be determined since the product was not returned and enough clinical information was not provided. the cause of the limb ischemia was patient condition related since both feet were ischemic. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
34067 | 1220648-2023-03581 | 17872816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872816&pc=OZD | 03/18/2022 | 03/10/2025 |  | 10/04/2023 |  |  | 03/18/2022 | 10/04/2023 |  |  | 10/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154316 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The user facility reported a 32-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the patient suffered ventricular tachycardia on multiple occasions due to pump positioning. the pump was repositioned after each occurrence and the ventricular tachycardia(vt) resolved. support was maintained throughout the events. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vf/vt/af/at and optical signal issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
34068 | 1220648-2023-03582 | 17872821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872821&pc=OZD | 06/02/2023 | 10/03/2023 |  | 10/04/2023 |  |  | 06/02/2023 | 10/04/2023 |  |  | 12/31/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024315854 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2023 |  |  |  |  | The user facility reported 70yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after one day on support, the while trying to reposition the pump, they were unable to pull the pump across the aorta. as a result, the pump was replaced. the pump's graft was seen to be thrombosed and had to be re-sewn. there was no lasting patient harm reported. no further intervention was deemed necessary. | [5 empty] | The investigation for the positioning issue and pump thrombosis has been completed. the product was returned for evaluation. it was determined that the root cause of the pump pulling out was most likely patient movement related and the root cause of the thrombus on on the graft after explant was also most likely patient condition related. the instructions for use states provides information on the prevention of thrombus by stating ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34069 | 1220648-2023-03579 | 17872825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872825&pc=OZD | 05/03/2023 | 10/03/2023 |  | 10/04/2023 |  |  | 05/03/2023 | 10/04/2023 |  |  | 03/31/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024327619 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2023 |  |  |  |  | The user facility reported 20yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a low purge flow / high purge pressure coinciding with a saturated pump flow occurred during patient support. both the console and cassette were swapped during troubleshooting, however, the issue remained. the pump was subsequently swapped as a result of the ongoing failure. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure and flow saturation issues has been completed. the data log from the first console shows high purge pressure alarms within a couple of minutes of the start of the case. the data log from the second console which the pump was switched to shows immediate high purge pressure alarms. the pump was investigated and the first time it was purged the high purge pressure was not reproduced. the second time it was purged the high purge pressure reproduced after a slow rise in purge pressure over the course of several minutes. the third time it was purged the high purge pressure again did not reproduce. the pump was run in a bucket of water while purging with sterile water and there was saturated flow. the saturated flow reproduced each time the pump was purged and run. the red handle was opened and no abnormalities were observed. no kinks were observed anywhere on the pump. no biomaterial was observed in the purge gap. the cause of the high purge pressure was most likely biomaterial due to the saturated flow reproducing each time. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34070 | 1220648-2023-03576 | 17872828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17872828&pc=OZD | 01/31/2023 | 10/03/2023 |  | 10/04/2023 |  |  | 01/31/2023 | 10/04/2023 |  |  | 11/30/2024 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023276316 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2023 |  |  |  |  | The user facility reported 39yo female was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to heart and kidney transplant. it was reported that there was a leak from both the pump and purge cassette during support. the cassette was replaced, however, the catheter continued to leak. no harm to the patient and support was able to proceed. there was no patient harm reported. | [5 empty] | The investigation for the reported purge leaks has been completed. the device was returned for evaluation. the catheter was analyzed to observe the location of the purge leak. the purge leak was found in the purge tubing near the motor housing. the cause of the pump purge leak was damaged purge tubing near the motor housing. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34071 | 1220648-2023-03743 | 17873373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873373&pc=OZD | 09/04/2023 | 07/18/2024 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  |  |  | 10/04/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1441400 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Display/Image | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | During impella procedure, the automated impella controller (aic) screen turned black. | Additional information states that there was no adverse events related to this product malfunction. |  |  |  |  | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) display issue has been completed. data logs were reviewed and were unable to confirm the reported failure. the console was returned for evaluation. functional testing was performed and was unable to reproduce the reported failure by wiggling the groups of wires that connect the display to the 4-in-1 connectors x26 and x51. no loose contact was intact. during visual inspection, one of the power supply screws was missing, and the other screw was floating inside the aic, which could have made any short circuit in the pbm board and likely could have resulted in an intermittent black screen. unable to find any damaged components or discoloration on both sides of the pbm. the root cause for the black screen was not determined due to the inability to reproduce. added- b5-no patient harm reported. d2-corrected common device name
34072 | 1220648-2023-03742 | 17873376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873376&pc=OZD | 09/04/2023 | 02/06/2024 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 06/30/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024358647 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 38-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was inserted with good placement. overnight the patient had clotted off the continuous renal replacement therapy (crrt) circuit and then had a suction alarm with sudden drop in flows to 0 l/min. the clinician decreased flow to p-4. aic showed flat motor current and when attempting to ramp up flows pump would have suction alarms and flows dropped again to 0 l/min. sedation turned off and the patient followed commands and moved all extremities. the patient was taken back to the operation room. trans-esophageal echo (tee) performed in the operation room reveals a very large mobile thrombus on the mitral valve (greater than 5cm). the pump was removed without incident. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. product & data were not returned for review. the root cause of low flow was most likely due to biomaterial ingestion. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup. unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the purge pressure high alarm message. could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.
34073 | 1220648-2023-03744 | 17873653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873653&pc=OZD | 01/20/2023 | 09/19/2024 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 04/30/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022092605 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2023 |  |  |  |  | The user facility in the eu reported a 48-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. on the 9th day of impella 5.5 support a purge leak was reported at the filter. no purge alarms were triggered during support. the sidearm retainer was not used and it is unknown if excessive force was applied. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | No product or logs were returned for analysis. the root cause of the purge leak was unable to be determined as the product was not returned so the exact location of the leak could not be determined. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34074 | 1220648-2023-03745 | 17873658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873658&pc=OZD | 08/10/2023 | 11/09/2023 |  | 10/04/2023 |  |  | 08/11/2023 | 10/04/2023 |  |  | 04/30/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337455 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2023 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support awaiting heart transplant. suction alarms occurred for 24 hours during support. staff was unable to increase flows over p4. echo showed adequate placement and volume being given. staff continued support and the patient eventually had a heart transplant. | [5 empty] | The investigation into the low pump flow issue and the optical signal issue has been completed since the original report was filed. the device was returned for investigation. it was determined that the cause of the low pump flow issue was most likely patient condition related with several suction alarms observed at higher p-levels and p-level automatically dropping to p-4 resolved the suction event. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34075 | 1220648-2023-03746 | 17873702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873702&pc=OZD | 05/16/2022 | 10/04/2023 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 12/31/2023 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023171786 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a 48-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the medical team in the eu had a 5.5 placed to replace one that had supported the patient for 2 months. on day 10 of the support the team observed thrombus burden in the left ventricle¿¿(lv) at the 5.5 pump. there was no harm or injury and the patient remained on support despite the 5.5 pump thrombosis until she recieved her heart transplant. of note the patient was on extracorporeal membrane oxygenation (ecmo) support additionally due to right sided failure. | [5 empty] | Two purge cassettes were returned for review. the pump was not returned for review. the cause of the thrombus was most likely high patient act values as reported in the clinical report. the failure mode will be monitored and trended.
34076 | 1220648-2023-03747 | 17873712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873712&pc=OZD | 09/07/2023 | 06/18/2024 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 05/31/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2023 |  |  |  |  | The patient's outcome of death was reviewed by the abiomed's medical safety officer/post market surveillance team. it was determined that the initial reportable event type of hemolysis has now been updated to death. | The user facility reported a 74-year-old female undergoing in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during cardiac surgery. the patient was exhibiting signs of hemolysis. the pump was attempted to be repositioned and came out of left ventricle position into the aorta. the team withdrew care and the patient expired. no noted treatment done for the malposition or hemolysis. |  |  |  |  | The investigation for the reported hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details and log analysis, the root cause of the positioning issue is most likely due to use as the pump was pulled into the aorta during repositioning of the device. the root cause of the hemolysis could not be determined due to insufficient clinical information and no product returned. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. h.6 code 4641 was reported incorrectly on the previously submitted report. a new code has been added to health effect - impact codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34077 | 1220648-2023-03748 | 17873718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873718&pc=OZD | 09/08/2023 | 02/02/2024 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 04/30/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350732 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2023 |  |  |  |  | The user facility reported a 48-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while implanting the impella, the sheath was peeled away and reposition sheath was advanced. physician reported significant bleeding requiring placement of a femstop. the abiomed sales representative confirmed that the patient did not receive blood products, but anticoagulants were not initiated that night due to the bleeding, but they were initiated in the afternoon of the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
34078 | 1220648-2023-03749 | 17873725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873725&pc=OZD | 09/08/2023 | 02/26/2024 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 07/31/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379834 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2023 |  |  |  |  | The user facility reported an 86-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. support was ongoing with suction and pulses lost on the impella accessed limb. the team sent the patient to interventional radiology for evaluation of the leg flow and had fem-fem bypass placed with regained flow and pulses by doppler. the pump remained on for support for one day and then care was withdrawn per family request. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34079 | 1220648-2023-03750 | 17873753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873753&pc=OZD | 09/08/2023 | 11/03/2023 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 06/30/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356781 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2023 |  |  |  |  | The user facility reported a 53-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. when the single-access technique was used to place the guide catheter, automated impella controller (aic) motor current waveform and impella flows dropped. staff was unable to regain flows, and there were recurrent suction alarms. the team troubleshot, checked pump position, infused fluid bolus, and explanted the pump due to concerns of clot. | [5 empty] | The investigation into the low pump flow and the thrombus issues has been completed since the original report was filed. product was returned for evaluation. pump was returned for investigation. no physical abnormalities were observed. pump ran expected during test. no biomaterial was observed on pump. data logs shows the low flow from the beginning of the case with reported suction alarm. based on clinical, patient had small left ventricle and blood product was given during treatment. the root cause of the low flow issue was patient condition related based on return product evaluation and provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34080 | 1220648-2023-03751 | 17873762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17873762&pc=OZD | 09/08/2023 | 07/23/2024 |  | 10/04/2023 |  |  |  | 10/04/2023 |  |  | 06/30/2025 |  | 10/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359623 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2023 |  |  |  |  | The user facility reported a 55-year-old male was implanted with an impella 5.5 for circulatory support after heart transplant. the patient developed a hematoma at the insertion site. patient received two units of red blood cells and had axillary site's hematoma washed out. the patient was noted as stable post interventions. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a and d6b added f6/f8-should have been left blank in the initial report.
34081 | 1220648-2023-03572 | 17859681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17859681&pc=OZD | 10/29/2021 | 02/14/2025 |  | 10/03/2023 |  |  | 10/29/2021 | 10/03/2023 |  |  | 07/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113032 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/29/2021 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed a subarachnoid bleed coinciding with a gastrointestinal bleed during impella 5.5 support. due to the noted bleeding, heparin was removed from the purge line. the patient had melena stools and required an unspecified quantity of blood products in response to the bleeding. the decision was subsequently made to withdraw care on account of the patient's overall condition. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03572 was provided in sections b2, b5, h1, and h6. | Due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34082 | 1220648-2023-03573 | 17859740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17859740&pc=OZD | 07/04/2022 | 01/19/2024 |  | 10/03/2023 |  |  | 07/04/2022 | 10/03/2023 |  |  |  |  | 10/03/2023 |  | AUTOMATED IMPELLA CONTROLLER¿ (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1462509 |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | No Display/Image | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/04/2022 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced a ¿controller error¿ alarm during the night. the rn stated that the automated impella controller¿(aic) screen went blank for approximately 1 min but denies loss of support or change in hemodynamics. there was no harm to the patient reported as a result of this incident. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation has been completed. no product was returned. the cause of the issue was not determined since the issue could not be reproduced.
34083 | 1220648-2023-03574 | 17859933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17859933&pc=OZD | 04/23/2023 | 10/03/2023 |  | 10/03/2023 |  |  | 04/23/2023 | 10/03/2023 |  |  | 12/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024318693 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/23/2023 |  |  |  |  | The user facility reported that a 56-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a pump did not start issue. when the device connected to automated impella controller (aic), the ¿defective impella catheter¿ alarm sounded during set up, prior to insertion. reconnecting the pump resolved the issue. there was no reported harm to the patient. | [5 empty] | The investigation into the reported pump did not start issue was completed. the device was returned for evaluation. based on the available information, the root cause of the impella defective alarm was most likely use-related due to the cable not being plugged in fully. reconnection resolved the issue. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34084 | 1220648-2023-03575 | 17859979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17859979&pc=OZD | 05/09/2023 | 10/03/2023 |  | 10/03/2023 |  |  | 05/09/2023 | 10/03/2023 |  |  | 03/31/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024320254 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Suction Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2023 |  |  |  |  | The user facility reported that a 32-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the device was inserted without issue but upon ramping up, the device had suction alarms at levels of p-7 or higher. transesophageal echocardiogram (tee) showed haziness around the inlet. as a result, the physician opted to removed and replace the device with a new one. the second device ran without issue. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. analysis of the data logs confirmed the low flow within about five minutes of the pump starting. visual inspection found no defects on the pump and functional testing passed. based on the available information, the root cause of the low flow was most likely due to patient condition as no issues were found on the pump & the low flow could not be reproduced. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34085 | 1220648-2023-03583 | 17860238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860238&pc=OZD | 05/23/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 05/23/2022 | 10/03/2023 |  |  | 12/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023170266 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The complainant reported a 40-year-old male patient suffering from cardiomyopathy was presented for impella 5.5 placement. during support, it was documented that high purge pressure was experienced by the pump. the purge cassette was initially replaced to troubleshoot the issue, however the flow remained low. tissue plasminogen activator (tpa) was subsequently administered and the purge flow began to increase accordingly. the purge pressure and flows were to be monitored closely and instructions for next steps were provided based on the results. there was no patient harm and support continued with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The (automated impella controller) aic were swapped most likely because after the cassette change, the cassette or disc was not properly fitted into the first console and instead of checking if it was fitted properly, the aic were replaced in a hurry. the cause of the high purge pressure was kinked purged line as the product was not returned and log indicate spikes/trend in purge pressure. this report is being filed as part of a retrospective review of historical records.
34086 | 1220648-2023-03584 | 17860317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860317&pc=OZD | 05/11/2022 | 03/28/2025 |  | 10/03/2023 |  |  | 05/11/2022 | 10/03/2023 |  |  | 02/29/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023194837 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that complications were encountered during implant of of an impella 5.5 pump. when advanced into place, the .018" guidewire was seen under echocardiogram (echo) to be located under the papillary muscle. the decision was made to remove the pump and guidewire for re-insertion; however, the guidewire was found to be kinked. the wire was swapped and the pump was implanted without further issue. after implant, some oozing was noted at the axillary access site, but intervention was not required. there was no patient harm and support with the implanted pump continued as planned. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Access site bleeding- the root cause of the access site bleed was unable to be determined as insufficient clinical information and product was provided for the investigation. positioning issue- no product was returned. the root cause of the issue was most likely use related. the pump was positioned under anterior mitral valve leaflet on echo as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
34087 | 1220648-2023-03585 | 17860446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860446&pc=OZD | 06/28/2021 | 01/24/2024 |  | 10/03/2023 |  |  | 06/28/2021 | 10/03/2023 |  |  | 04/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022069630 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Hemorrhagic Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2021 |  |  |  |  | The medical center reported that a patient on impella 5.5 support had a fever and the blood cultures were positive after 2 weeks of support. the following day the patient experienced a stroke. the stroke was treated by neurology team and a drain was placed for the hemorrhage. the patient was transitioned to palliative care where eventually expired. of note the patient had known heart failure and was on the list awaiting transplant. | [5 empty] | The investigation is ongoing at this time. upon completion of the investigation a final mdr will be forthcoming. per the ifu: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ "infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen." | The investigation into the reported stroke and infection has been completed since the original report was filed. the device was not returned by the medical facility. stroke/ cva - the cause for hemorrhagic stroke/cva is patient condition since the patient had acute head bleed and an increase in size of intercranial hemorrhage was observed. infection/ sepsis - the cause of the infection was patient condition since enough clinical info was not provided and the relation between infection and impella 5.5 is unknown. per the ifu: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ "infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen."
34088 | 1220648-2023-03586 | 17860497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860497&pc=OZD | 06/20/2021 | 10/03/2023 |  | 10/03/2023 |  |  | 06/20/2021 | 10/03/2023 |  |  |  |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022061927 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/20/2021 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was found to have a thrombus adjacent to the impella. the patient was shown to have bilateral upper extremity edema present. vascular surgery has been consulted for possible removal of clot. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. only the clinical report was avaialable for investigation. the cause of this issue is patient condition as edema and clot was discovered in the patient's body. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34089 | 1220648-2023-03587 | 17860594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860594&pc=OZD | 06/17/2021 | 01/23/2024 |  | 10/03/2023 |  |  | 06/17/2021 | 10/03/2023 |  |  | 04/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071791 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | Sepsis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2021 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced impella in the aorta alarms. an echo was performed and it showed that the device was in the correct position. on a later date and incorrect placement signal occurred. impella had migrated further into the ventricle. echo was performed and the device was pulled back to reposition it in the right place. the alarm stopped. the patient later had a noted infection of unknown cause. antibiotics were given. causality to the impella is unknown. the pump remained on for support for another 29 days after infection noted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the optical signal issue and the infection/sepsis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the root cause of the optical signal issue was patient condition related; however, the root cause of the infection/sepsis was not determined since product was not returned and enough information was not provided. the failure modes will be monitored and trended.
34090 | 1220648-2023-03588 | 17860666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860666&pc=OZD | 06/24/2021 | 01/24/2024 |  | 10/03/2023 |  |  | 06/24/2021 | 10/03/2023 |  |  | 02/28/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022043462 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2021 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support, via the right axillary. the patient on impella support experienced a decreased right radial pulse. the arm was painful and the ultrasound did not reveal any deep vein thrombosis. device was explanted 5 days later. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported limb ischemia has been completed since the original report was filed. the device was not returned by the medical facility. the cause for limb ischemia was most likely patient condition related. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34091 | 1220648-2023-03589 | 17860707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860707&pc=OZD | 06/18/2021 | 01/24/2024 |  | 10/03/2023 |  |  | 06/18/2021 | 10/03/2023 |  |  | 04/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022080684 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2021 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team attempting to implant an impella 5.5 made numerous attempts to get stable position of the pump but they were unable to due to patient's anatomy. decision was made to explant the impella 5.5 and re-implant a 5.0 pump due to ability to pigtail to stabilize in apex. no lasting harm or injury to the patient for the delay in support in initiation. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issue has been completed since the original report was filed. the device was not returned by the medical facility. the cause of the positioning issue is patient condition as due to the anatomy of the patient, md could not position the pump correctly.
34092 | 1220648-2023-03590 | 17860736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860736&pc=OZD | 06/23/2022 | 03/07/2024 |  | 10/03/2023 |  |  | 06/23/2022 | 10/03/2023 |  |  | 02/29/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197482 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient suffered an ischemic stroke during impella support. the patient was monitored following the event to ascertain the severity of the stroke. support continued with the implanted pump throughout the event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva issue has been completed since the original report was filed. the device was not returned for investigation. as no clinical information was provided, the root cause of the stroke/cva could not be determined.
34093 | 1220648-2023-03591 | 17860750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860750&pc=OZD | 06/20/2021 | 05/06/2024 |  | 10/03/2023 |  |  | 06/20/2021 | 10/03/2023 |  |  | 11/30/2022 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021020026 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2021 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had to be taken in the or for an emergent tamponade removal one day after implant. the pump remained on for 6 more days afterwards. the patient survived event. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
34094 | 1220648-2023-03592 | 17860780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860780&pc=OZD | 02/27/2022 | 01/11/2024 |  | 10/03/2023 |  |  | 02/27/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141838 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2022 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site which required the patient to receive 3 units of blood. the systemic heparin was stopped and purge heparin had to be decreased. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. the root cause of the access site bleeding issue was [low/high] patient act values as per the clinical description provided. the failure modes will be monitored and trended.
34095 | 1220648-2023-03593 | 17860818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860818&pc=OZD | 03/17/2022 | 10/27/2025 |  | 10/03/2023 |  |  | 03/17/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022146591 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Unspecified Infection | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2022 |  |  |  |  | The user facility reported a 41-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was alleged to have a possible leak from the red plug. an inspection of that plug showed it to be dry and intact. the patient had also developed fever of unknown origin which subsided in after initiating antibiotics. the pump remained on for support for another 5 days til removed for the placement of centrimag device. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen¿. this report is being filed as part of a retrospective review of historical records.
34096 | 1220648-2023-03594 | 17860819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860819&pc=OZD | 06/27/2022 | 01/25/2024 |  | 10/03/2023 |  |  | 06/27/2022 | 10/03/2023 |  |  | 02/29/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196085 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the pump was accidentally pulled back across the valve. the pump was subsequently explanted as the patient had increased pulsatility and increased blood pressure without pressor support. there was no patient harm as a result of the noted issue. | [5 empty] | The impella device was not returned by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
34097 | 1220648-2023-03595 | 17860843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860843&pc=OZD | 04/10/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 04/10/2022 | 10/03/2023 |  |  | 12/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023175364 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2022 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that while removing the impella 5.5 through the supra clavicular site the aorta tore. the chest was immediately opened and the bleed was repaired. patient was given 2 units of blood. the pump had supported the patient for over 12 days, and survived the adverse event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | The cause of the vascular damage in great vessel was use issue since the graft tore most likely due to excessive force. this report is being filed as part of a retrospective review of historical records.
34098 | 1220648-2023-03596 | 17860912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860912&pc=OZD | 03/11/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 03/11/2022 | 10/03/2023 |  |  | 10/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022156736 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient had the impella implanted via direct access developed bleeding in their left pleural and mediastinal spaces. the patient was taken to the operating room for a washout due to the bleeding in the chest. positioning was verified via echocardiogram (echo); however, the ventricles were underfilled and the patient was given 2 units prbcs, 2 units ffp, and 2 units of platelets before the chest could be closed. support with the implanted pump was maintained throughout the intervention. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the bleeding was not determined. it was reported that the patient experiencing bleeding in left pleural and mediastinal spaces and the patient was taken for wash out. later the wash out, there were no patient updated concerning the bleeding and the issue seemed to be resolved. there were no clinical details submitted that concerns the impella in relation to the cause of the bleeding. hence the root cause of the issue not determined. but, the pump functioning was good till the end of the case.
34099 | 1220648-2023-03597 | 17860949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860949&pc=OZD | 06/11/2022 | 02/20/2024 |  | 10/03/2023 |  |  | 06/11/2022 | 10/03/2023 |  |  | 01/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182439 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Inflammation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced bleeding from the impella access site. the patient was given one unit of blood to counteract the bleed. however, as support continued, the site became inflamed and the patient developed a spiking fever. furthermore, blood entered the repositioning sleeve and the decision was ultimately made to explant the pump. | [5 empty] | The investigation into the access site bleeding major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34100 | 1220648-2023-03598 | 17860951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860951&pc=OZD | 03/25/2022 | 03/28/2025 |  | 10/03/2023 |  |  | 03/25/2022 | 10/03/2023 |  |  | 12/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023174115 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced two small strokes after 2 weeks of support. transesophageal echocardiography (b)(6) revealed no visible clot on device, in apex or appendage. decision was made to explant the device, however, the physician had difficulty passing the motor housing freely out. device was removed, graph and chest closed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. stroke - the cause of the stroke was patient condition related since the product was not returned. removal issue - the cause of the removal issue was use issue since the angle of the graph under the sternum caused the removal issue. this report is being filed as part of a retrospective review of historical records.
34101 | 1220648-2023-03599 | 17860971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860971&pc=OZD | 06/08/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 06/08/2022 | 10/03/2023 |  |  | 02/29/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196416 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that following impella insertion, the patient experienced notable bleeding from their access site. a jackson-pratt drain (jp drain) had not been placed at the time of insertion. the patient received one unit of platelets, 2 units of fresh frozen plasma, and 1 unit of packed red blood cells to counteract the blood loss. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. no product was returned. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34102 | 1220648-2023-03600 | 17860985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17860985&pc=OZD | 02/17/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 02/17/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022137387 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2022 |  |  |  |  | The patient experienced intermittent hemolysis and supraventricular tachycardia as a result of impella 5.5 positioning during support. although the pump was repositioned to initially resolve the issue, the angle of the catheter combined with the patient's anatomy was causing the device to consistently migrate towards the mitral valve. the patient was placed on continuous renal replacement therapy following the bout of hemolysis. the patient was monitored following the events and support continued with the implanted pump for 13 days til explant and bridge to lvad device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the vt was use issue since echo confirmed the pump out of position in mitral valve. the cause of the hemolysis was due to improper positioning since tte revealed impella facing mitral valve and cr confirmed it. this report is being filed as part of a retrospective review of historical records.
34103 | 1220648-2023-03601 | 17861011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861011&pc=OZD | 02/20/2022 | 10/03/2023 |  | 10/03/2023 |  |  | 02/20/2022 | 10/03/2023 |  |  | 10/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023159648 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/20/2022 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced an "impella position in aorta" alarm. the medical team attempted to reposition the pump but opted to replace it instead. there was no harm to the patient reported. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34104 | 1220648-2023-03602 | 17861096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861096&pc=OZD | 06/23/2022 | 01/12/2024 |  | 10/03/2023 |  |  | 06/23/2022 | 10/03/2023 |  |  | 01/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023181536 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a 59-year-old male suffering from acute myocardial infarction and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced access site bleeding during impella support. the patient was given an unspecified number of blood products due to the complication and heparin was stopped. support was withdrawn the following day after discussions with the family on the patient's existing condition. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34105 | 1220648-2023-03603 | 17861109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861109&pc=OZD | 03/08/2022 | 03/01/2024 |  | 10/03/2023 |  |  | 03/08/2022 | 10/03/2023 |  |  | 11/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166273 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The user facility reported a 41-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella cp was transitioned to impella 5.5. when the new pump was inserted then there were no flows at the superficial femoral artery. surgical repair of the artery was performed to restore flow to the limb, and jp drains were placed in the access site. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the limb ischemia ¿ reversible event has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
34106 | 1220648-2023-03604 | 17861177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861177&pc=OZD | 06/07/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 06/07/2022 | 10/03/2023 |  |  | 01/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023180543 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Swelling/ Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The complainant reported a 17-year-old male patient suffering from heart transplant rejection was presented for impella 5.5 placement. during support, it was documented that the patient experienced bleeding from their access site after ambulating. the dressing was changed and it was discovered that a hematoma had formed at the site causing swelling. the patient was taken to the operating room where the pocket was explored for any active bleeding and the hematoma was evacuate. support with the pump continued following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34107 | 1220648-2023-03605 | 17861195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861195&pc=OZD | 01/28/2022 | 03/28/2025 |  | 10/03/2023 |  |  | 01/28/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149900 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2022 |  |  |  |  | The user facility reported a 52-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced access site bleeding during the procedure. the or time was extended attempting to tamponade bleeding next to graft site. the physicians had to be scrubbed back in to assist with controlling bleeding, and made the choice to have chest left open. patient returned to or for washout and partial closure. the pump was successfully weaned and explanted after just under 6 days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the cause of the access site bleeding was patient condition as patient had various heart problems and valve replacements prior to the implant. this report is being filed as part of a retrospective review of historical records.
34108 | 1220648-2023-03606 | 17861214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861214&pc=OZD | 08/10/2022 | 02/18/2025 |  | 10/03/2023 |  |  | 08/10/2022 | 10/03/2023 |  |  | 01/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186096 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 55-year-old male suffering from cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced a purge leak at the yellow luer during patient support. a purge bypass was initially completed to resolve the issue, but the decision was ultimately made to replace the pump. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. the device history record review confirmed passed all post-sterile inspection checks. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation was not determined.
34109 | 1220648-2023-03607 | 17861230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861230&pc=OZD | 03/10/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 03/10/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141243 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2022 |  |  |  |  | The user facility reported a 67-year-old patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced ventricular tachycardia (vt). the patient had to be shocked 3 times, went to full code and had to be intubated. two days after the patient condition had improved and decision was made to extubate and explant impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
34110 | 1220648-2023-03608 | 17861319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861319&pc=OZD | 03/08/2022 | 01/23/2024 |  | 10/03/2023 |  |  | 03/08/2022 | 10/03/2023 |  |  | 12/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023174109 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding from their access site of the impella. it was noted that direct access was utilized after it was observed that a right axillary insertion would not be possible due to heavy calcification in the vessel. once the graft was trimmed and secured around the repositioning sheath, some bleeding was observed while trying to close the chest. the patient subsequently received four units of packed red blood cells to counteract the blood loss. support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the patient's access site bleeding has been completed. the impella products were not returned for evaluation. the clinical investigation follows: patient has calcified aorta. cutdown was performed to right axillary and the artery was found to be heavily calcified. decision then made to place graft directly onto aorta and graft tunneled to right chest. omni flush used to cross av. 5.5 inserted with ease into lv, initially deep, but then pulled back to ~5cm. patient was found to have some bleeding while closing the chest. the cause of the access site bleeding was patient condition since patient had heavily calcified aorta causing many problems. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34111 | 1220648-2023-03609 | 17861341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861341&pc=OZD | 02/04/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 02/04/2022 | 10/03/2023 |  |  | 06/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022100765 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2022 |  |  |  |  | The user facility reported a 20-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient developed a hematoma at their axillary access site during impella support. the site was evacuated and the patient was given one unit of packed red blood cells to treat the condition. support continued following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding - major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34112 | 1220648-2023-03610 | 17861371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861371&pc=OZD | 01/26/2022 | 09/22/2025 |  | 10/03/2023 |  |  | 01/26/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022136275 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team had some difficulty delivering the impella pump and failed on the first attempt. after another attempt they were able to successfully implant the device, however, significant bleeding noted from open chest and chest tubes requiring multiple blood products. chest was closed and impella secured using three point fixation. there was no clear source of the bleeding. the causal relationship to the use of impella could not be denied. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number. d6a and d6b revised from the initial manufacturer device report in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number. h6 medical device problem code 1157 was inadvertently omitted on the initial manufacturer device report number. h6 component code revised from the initial submission in accordance with updated procedures.
34113 | 1220648-2023-03611 | 17861377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861377&pc=OZD | 07/25/2023 | 10/27/2025 |  | 10/03/2023 |  |  | 07/25/2023 | 10/03/2023 |  |  | 12/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023176816 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2023 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the pump had to be replaced due to a suspected clot within the device. the pump was removed and another impella 5.5 pump was implanted for ongoing support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. instructions for use for the related event are as follows: ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. d6a/d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated. h6 updated.
34114 | 1220648-2023-03612 | 17861397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861397&pc=OZD | 01/30/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 01/30/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022123447 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2022 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced an "impella position in aorta" alarm. an echocardiogram (echo) was performed, and it showed that the catheter was deep in ventricle with the inlet appearing to be involved with the mitral valve and its structures. the pump was repositioned with echo guidance performed under fluoroscopy in cath lab. there was no harm to the patient reported. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issues has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was most likely use related, as the position of catheter was deep in ventricle with the inlet appearing to be involved with the mitral valve and its structures as per the clinical description provided.
34115 | 1220648-2023-03613 | 17861435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861435&pc=OZD | 05/29/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 05/29/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022137356 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at the impella access site during pump support. although discussions were had on potentially explanting the pump, the staff feared the patient may decompensate and opted on monitoring the patient's condition. the patient's arm was swollen as a result of the hematoma and the arm was subsequently wrapped and elevated to prevent the condition from worsening. the hematoma, despite not worsening, remained present until the pump was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34116 | 1220648-2023-03614 | 17861437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861437&pc=OZD | 02/22/2022 | 09/18/2024 |  | 10/03/2023 |  |  | 02/22/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143126 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Ventricular Fibrillation | [11 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2022 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced ventricular fibrillation (vf) and it was alleged that patient had sepsis but there was no confirmation of that diagnosis. the medical team planned to swap out the lines due to sepsis, but the patient expired. there was no allegation that impella was a factor in the patient expiry. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03614 was provided in sections b2, b5, h1, and h6.
34117 | 1220648-2023-03615 | 17861509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861509&pc=OZD | 02/16/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 02/16/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022146540 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding from the access site which required the patient to receive blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product were returned for investigation.
34118 | 1220648-2023-03616 | 17861531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861531&pc=OZD | 05/03/2022 | 02/21/2025 |  | 10/03/2023 |  |  | 05/03/2022 | 10/03/2023 |  |  | 12/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023172887 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The user facility reported a male suffering from cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced low purge pressures resulting in a purge system open alarm during patient support. during troubleshooting, it was identified that the cause for the noted issues was a fluid leak in the purge system. a purge bypass was completed twice before the issue resolved. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the purge leak ¿ pump could not be determined as insufficient clinical details were provided.
34119 | 1220648-2023-03617 | 17861556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861556&pc=OZD | 01/29/2022 | 01/25/2024 |  | 10/03/2023 |  |  | 01/29/2022 | 10/03/2023 |  |  | 10/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154705 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced a hematoma at the access site. incision site drainage was performed and the patient had to receive one unit of blood. the medical staff was considering taking the patient in the or for a washout. no further information provided. patient survived event and had impella support for just under 15 days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the access site bleed has been completed. however, with limited clinical details nor the impella device, the root cause of the reported issue could not be determined.
34120 | 1220648-2023-03618 | 17861575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861575&pc=OZD | 04/20/2022 | 01/12/2024 |  | 10/03/2023 |  |  | 04/20/2022 | 10/03/2023 |  |  | 11/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169807 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2022 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that resistance was met during attempted delivery of an impella 5.5 pump. it was noted that existing stenosis of the innominate and axillary artery made advancing the pump difficult during implant. the device subsequently kinked during manipulation and was no longer usable. the pump was removed and another impella 5.5 pump was utilized for implant. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the delivery issues- use has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the delivery issue was most likely patient condition related as the patient had stenosis.
34121 | 1220648-2023-03619 | 17861594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17861594&pc=OZD | 01/30/2022 | 05/03/2024 |  | 10/03/2023 |  |  | 01/30/2022 | 10/03/2023 |  |  | 10/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154719 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2022 |  |  |  |  | The user facility reported a 62-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump had a fluid leak from the purge sidearm during patient support. a purge bypass was initially completed to resolve the leak; however, due to the patient's critical status, the decision was made to replace the pump. there was no patient harm as a result of the purge leak. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the purge leak was unable to be determined as the pump was not returned for investigation. d6a, d6b. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
34122 | 1220648-2023-03620 | 17862235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862235&pc=OZD | 03/01/2022 | 09/19/2024 |  | 10/03/2023 |  |  | 03/01/2022 | 10/03/2023 |  |  | 10/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023159328 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced arrhythmia as a result of the positioning of the impella pump. it was noted during support that the pump was unintentionally pushed out of the left ventricle during an operation to close the patient's chest. the patient began experiencing ventricular tachycardia and the pump was quickly removed and properly re-inserted. it was documented that during repositioning, the pump shifted and flipped into the mitral apparatus and caused arrhythmia. the pump was placed in a manner to avoid further complications and support continued. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34123 | 1220648-2023-03621 | 17862264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862264&pc=OZD | 01/21/2022 | 01/24/2024 |  | 10/03/2023 |  |  | 01/21/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022133326 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump stopped functioning following a high motor current during patient support. the pump was subsequently explanted as a result of the noted failure. the pump had supported for 124 hours. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported pump stop has been completed since the original report was filed. the device was not returned for evaluation, therefore a root cause could not be positively identified. the most likely cause of the pump stop was most likely a kinked purge line resulting in a biomaterial ingestion pump stop. sudden purge pressure had risen indicating a kink which lead to high purge pressure. pump had been stopped per high mc for about an hour. logs indicate high mc above 1200 along with purge system blocked alarm. device was not returned for analysis. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34124 | 1220648-2023-03622 | 17862299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862299&pc=OZD | 02/28/2022 | 03/05/2024 |  | 10/03/2023 |  |  | 02/28/2022 | 10/03/2023 |  |  | 02/28/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022049824 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. they encountered difficulty while placing the impella 5.5. during the initial delivery, the doctor was reported to have incorrectly positioned the wire for guided delivery. as a result, the pump could not be torqued into the correct position. the pump was removed twice from the body and rinsed with saline before interventional cardiology was consulted to place the device. with the noted assistance, the pump was able to recross the aortic valve and support was initiated. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the delivery issues- use has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue was due to use as the wire was placed incorrectly.
34125 | 1220648-2023-03623 | 17862370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862370&pc=OZD | 02/03/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 02/03/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022137382 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2022 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that their impella 5.5 pump experienced a loss in flows coinciding with an increase in motor current and suction alarms during patient support. it was noted that the events occurred when support levels were raised above p4/p5. due to the noted issue, the pump was replaced and it was discovered that a clot was lodged in the inlet of the pump. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34126 | 1220648-2023-03624 | 17862429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862429&pc=OZD | 01/11/2022 | 03/15/2024 |  | 10/03/2023 |  |  | 01/11/2022 | 10/03/2023 |  |  | 10/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151952 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a right thalamic stroke during impella support. the patient had symptoms of left sided weakness and hemiplegia due to the stroke. no neuro intervention was planned for the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva issue has been completed since the original report was filed. the device was not returned for investigation. as no clinical information was provided, the root cause of the stroke/cva could not be determined.
34127 | 1220648-2023-03625 | 17862476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862476&pc=OZD | 01/17/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 01/17/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151090 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2022 |  |  |  |  | The user facility reported a 41-year-old male undergoing cardiac surgery, and on transplant list, was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support required right axillary site washout in a return to the or. there were no complications intra or post operation. the pump support ran for over 49 days til wean and explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. with limited clinical details nor the impella device for evaluation, the root cause of the reported access site bleeding was unable to be determined.
34128 | 1220648-2023-03626 | 17862518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862518&pc=OZD | 01/14/2022 | 03/28/2025 |  | 10/03/2023 |  |  | 01/14/2022 | 10/03/2023 |  |  | 10/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153885 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2022 |  |  |  |  | The user facility reported a 18-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient had the labs drawn and the results indicated that the patient had hemolysis. the physician repositioned the pump in attempt to resolve hemolysis, however, the symptoms did not improve. the physician was unable to get the pump to a better position due to fear of completely removing it out of the left ventricle (lv). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. positioning issue - the cause of the positioning issue was patient movement as patient was independently ambulating prior to the issue. hemolysis - the cause of the hemolysis was improper positioning since the position of the pump got shifted. this report is being filed as part of a retrospective review of historical records.
34129 | 1220648-2023-03627 | 17862611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862611&pc=OZD | 01/03/2022 | 10/03/2023 |  | 10/03/2023 |  |  | 01/03/2022 | 10/03/2023 |  |  | 07/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022119316 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2022 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that patient was experiencing some suction issues from a suspected thrombus formation. this required the medical team to open the access site to remove the thrombus and replace the pump which caused some generalized bleeding to occur. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the clinical report was analyzed. the data logs showed suction. the root cause of the low flow was not determined. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34130 | 1220648-2023-03628 | 17862709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862709&pc=OZD | 12/27/2021 | 01/12/2024 |  | 10/03/2023 |  |  | 12/27/2021 | 10/03/2023 |  |  | 03/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060321 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2021 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that difficulty was met when attempting to place an impella 5.5 pump for patient support. it was noted that the pump was unable to make the turn from the axillary into the aorta. while attempting to manipulate the device, the graft anastomosis tore and the patient began to bleed. multiple units of packed red blood cells (prbc) were given and surgical intervention was utilized to treat the bleed. the implant was subsequently aborted and an intra aortic balloon (iabp) was placed for patient support. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the delivery issue and vascular damage has been completed. the product was not returned. as per the clinical information provided, the root cause of the delivery issue- anatomy was most likely due to patient condition. the cause of the vascular damage was also patient condition since anastomosis tore while attempting to navigate the patient's difficult anatomy.
34131 | 1220648-2023-03629 | 17862790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862790&pc=OZD | 12/23/2021 | 01/26/2024 |  | 10/03/2023 |  |  | 12/23/2021 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022121462 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2021 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced persistent bleeding/hematoma issues from the axillary access site during impella support. following implant, oozing was noted from the access site. the site was redressed and an extra stitch was added in an attempt to address the bleed. a hematoma formed at the site and a jackson-pratt drain (jp drain) was also placed to manage the condition. after nine days of support, some bleeding was still evident and the patient was given blood products due to dropping hemoglobin levels. the hematoma and drainage from the jp drain were noted for roughly nine more days throughout the remainder of support. the patient was stable at the time of planned explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
34132 | 1220648-2023-03630 | 17862826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862826&pc=OZD | 12/22/2021 | 03/28/2025 |  | 10/03/2023 |  |  | 12/22/2021 | 10/03/2023 |  |  | 03/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071469 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2021 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced multiple issues with clotting and bleeding during impella 5.5 support. as a result of the noted conditions, the decision was made to explant the pump and leave extracorporeal membrane oxygenation (ecmo) in place for ongoing support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. low pump flow - the cause of the low pump flow was patient condition as patient was found to have multiple conditions. access site bleeding - the cause of the access site bleeding cannot be determined since the product was not returned and no other clinical info was provided. thrombus - the cause of the thrombus was patient condition as low blood flow was present in the patient's body. this report is being filed as part of a retrospective review of historical records.
34133 | 1220648-2023-03631 | 17862841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862841&pc=OZD | 05/24/2022 | 05/23/2025 |  | 10/03/2023 |  |  | 05/24/2022 | 10/03/2023 |  |  | 02/29/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023191623 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/24/2022 |  |  |  |  | Additional information was received regarding the outcome of the patient. despite a successful perforation repair, the patient still struggled after coming off bypass. the patient's condition remained poor the following day, and the family decided to withdraw care after making the patient dnr. the patient passed away following explant. a definitive relationship between the perforation and patient's passing was not established. | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced positioning issues during impella support that coincided with a left ventricle perforation. after placing the pump across the aortic valve during two consecutive delivery attempts, the pump flipped towards the mitral valve. following the second attempt, it was noted that the left ventricle had been perforated. a covered stent and balloon tamponade were utilized to treat the perforation. once the perforation was treated, the pump was re-advanced over the guidewire and placed successfully for patient support. |  |  |  |  | Additional information added to b2, b5, h1, and h6 to coincide with the report of the patient's passing. the investigation remains ongoing. upon completion of the investigation, a supplemental report will be submitted. | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34134 | 1220648-2023-03632 | 17862867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862867&pc=OZD | 05/27/2022 | 02/06/2024 |  | 10/03/2023 |  |  | 05/27/2022 | 10/03/2023 |  |  | 03/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023202586 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2022 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed hemolysis during impella support. lactate dehydrogenase (ldh) levels were measured to be 1000 u/l and platelet levels dropped to 27,000. as a result, the pump was repositioned, and the patient was switched to bivalirudin. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the hemolysis issue has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the hemolysis cannot be determined since product and logs were not returned and there was insufficient clinical information provided.
34135 | 1220648-2023-03633 | 17862922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17862922&pc=OZD | 12/22/2021 | 03/28/2025 |  | 10/03/2023 |  |  | 12/22/2021 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141354 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2021 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient experienced multiple issues of bleeding and clotting. the impella was explanted. the patient remained on ecmo support alone. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. thrombus - the cause of the thrombus was patient condition as patient had low blood flow present in the body. access site bleeding - the cause of the access site bleeding was patient condition since patient was extremely volume sensitive and was given multiple blood units. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34136 | 1220648-2023-03634 | 17863006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863006&pc=OZD | 01/07/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 01/07/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022144158 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/07/2022 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a purge leak was identified at the yellow luer following a scheduled purge cassette change. a purge bypass was completed utilizing a 19ga needle and support continued with the implanted pump. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the purge leak was unable to be determined as the pump was not returned for investigation. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34137 | 1220648-2023-03635 | 17863045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863045&pc=OZD | 11/30/2021 | 10/03/2023 |  | 10/03/2023 |  |  | 11/30/2021 | 10/03/2023 |  |  | 07/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118213 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2021 |  |  |  |  | The user facility reported a male patient of unknown age undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a hematoma at the impella site which required washout. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. no data logs were returned for analysis. the clinical details shared were not sufficient to determine any root cause. the root cause can not be determined. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34138 | 1220648-2023-03636 | 17863060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863060&pc=OZD | 12/20/2021 | 01/11/2024 |  | 10/03/2023 |  |  | 12/20/2021 | 10/03/2023 |  |  | 07/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128703 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2021 |  |  |  |  | The user facility reported a 80-year-old undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance during attempted delivery due to the patient's anatomy. as a result of the complication, the patient experienced bleeding at their access site and an unspecified number of blood products were given to the patient. the pump was removed and the site was closed. an intra-aortic balloon pump (iabp) remained in place for support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed and the delivery issues has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause of the delivery issues. according to clinical details, the doctor was unable to advance the implant due to anatomy, leading to pump removal. therefore the root cause of the delivery issue was due to patient conditions. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
34139 | 1220648-2023-03637 | 17863087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863087&pc=OZD | 04/25/2022 | 01/25/2024 |  | 10/03/2023 |  |  | 04/25/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149849 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a positioning issue involving an impella 5.5 pump during patient support. it was documented that the pump migrated back across the aortic valve and triggered an impella in aorta alarm. attempts were made to reposition the device, but the surgeon was unable to re-advance the pump across the valve. due to the noted issue, the decision was made to explant the pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
34140 | 1220648-2023-03638 | 17863090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863090&pc=OZD | 01/01/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 01/01/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022132632 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2022 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella 5.5 pump met resistance during attempted delivery. it was documented that the device was unable to pass the anastomosis due to the presence of a thrombus in the graft. upon removal of the pump, a thrombus was identified in the inlet of the cage and the decision was made to abort the implant. a balloon pump was subsequently inserted for ongoing patient support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | No product was returned. the root cause of the thrombus issue was patient condition related, there was severe thrombus observed in the graft as per the clinical description provided. the root cause of the delivery anatomy issue was patient condition related, as the pump was unable to pass through the anastomosis as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
34141 | 1220648-2023-03639 | 17863102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863102&pc=OZD | 12/26/2021 | 01/24/2024 |  | 10/03/2023 |  |  | 12/26/2021 | 10/03/2023 |  |  | 06/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098192 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2021 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that following impella implant, significant bleeding was seen from the anastomosis. multiple blood products were given to counteract the blood loss and the bleeding subsequently resolved. as a precaution, heparin administration in the purge line was delayed until the following day to verify that the bleeding had stopped. support with the implanted pump proceeded as planned. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34142 | 1220648-2023-03640 | 17863184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863184&pc=OZD | 08/11/2022 | 09/29/2025 |  | 10/03/2023 |  |  | 08/11/2022 | 10/03/2023 |  |  | 05/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023225622 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 47-year-old patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported the patient had asystole and then ventricular tachycardia/ ventricular fibrillation (vt/vf) and was shocked. the rhythm was corrected and echo at bedside showed that impella was against the wall and orientation not toward apex. attempted to torque but unsuccessful. pump had to be pulled back to 3cm. pump remained on for support for over 4 days and was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03640 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03640 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03640 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03640. h6 component code revised from the initial submission in accordance with updated procedures.
34143 | 1220648-2023-03641 | 17863225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863225&pc=OZD | 10/21/2022 | 01/25/2024 |  | 10/03/2023 |  |  | 10/21/2022 | 10/03/2023 |  |  | 04/30/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023211968 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an access site bleed prompting vascular intervention. the source of the bleed was found to be from the patient's pectoral muscle. intervention was taken via cauterization and the bleed subsided. the 5.5 pump had been inserted via the axillary artery and graft. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34144 | 1220648-2023-03642 | 17863329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863329&pc=PYX | 09/07/2023 | 11/03/2023 |  | 10/03/2023 |  |  | 09/07/2023 | 10/03/2023 |  |  | 02/28/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024346415 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was prepped and drapped in normal sterile fashion for delivery of an impella rp. the physician had difficulty delivery product to the right heart. every time physician tried to advance the impella rp flex device over any guidewire it would kink the wire. multiple wires by multiple manufacturers were utilized. wire kept coming back every time rp flex was advanced or rp flex catheter would collapse in on itself. physician decided to abort procedure and put in a protek suo with no issues. same access site as impella rp flex and same lunderquist wire used for placement. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the root cause of the delivery issue is most likely due to the improper technique that was used to place the guidewire to try to advance the rp flex..
34145 | 1220648-2023-03643 | 17863568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17863568&pc=OZD | 01/31/2022 | 09/26/2024 |  | 10/03/2023 |  |  | 01/31/2022 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141805 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2022 |  |  |  |  | The user facility in the eu reported a 69-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a purge side arm leak that developed during use of the device. a bypass was performed and support continued with no harm to the patient. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | No product was returned for evaluation and no data logs were available for analysis. the root cause of the purge side arm leak (between the reservoir and the filter) was unable to be determined. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported purge leak has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the purge leak and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03643 was provided in sections b2, b5, h1, and h6.
34146 | 1220648-2023-03644 | 17864208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864208&pc=OZD | 10/07/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 10/07/2022 | 10/03/2023 |  |  | 04/30/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023215095 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. positioning issues were encountered with an impella 5.5 pump following placement in the left ventricle. it was documented that once placed in the left ventricle, the pump pulled towards the sub-annular structures of the mitral valve when flows were increased. the pump was removed and reinserted, but the physician was unable to direct the pump away from the sub-annular structures. this resulted in suction when the pump was once again turned on. the pump was again removed and re-inserted; however, the physician could not torque the pump into the desired position and the decision was made to abort the placement. the patient remained supported with an intra-aortic balloon pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | No product was returned. the root cause of the positioning issue was most likely patient condition related, as flows were increased impella pulled toward sub-annular structures of the mv as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
34147 | 1220648-2023-03645 | 17864351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864351&pc=OZD | 10/26/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 10/26/2022 | 10/03/2023 |  |  | 01/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022038052 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 69-year-old male experiencing cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at their impella access site during support. the patient was subsequently taken to the operating room where the hematoma was evacuated. although the site was sore and tender following the intervention, support with the impella pump was able to continue. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34148 | 1220648-2023-03654 | 17864450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864450&pc=OZD | 08/08/2022 | 09/30/2025 |  | 10/03/2023 |  |  | 08/08/2022 | 10/03/2023 |  |  | 05/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023220833 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Difficult to Advance | Material Deformation | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. difficulty was met while attempting to implant an impella 5.5 pump for patient support. multiple instruments were utilized including a jr4 guide catheter and buddy wire to allow for the pump to traverse the extremely tortuous subclavian/aorta. despite the additional techniques used, multiple attempts were still required before the pump was successfully delivered in the left ventricle. as a result of the reported difficulty, it was noted that the initial guidewire was damaged and needed to be replaced before the pump could be placed. | It was reported that the patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03654. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03654. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03654. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03654 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03654 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03654 in accordance with updated procedures. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03654. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2023-03654. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2023-03654. h6 health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number 1220648-2023-03654. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03654.
34149 | 1220648-2023-03648 | 17864524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864524&pc=OZD | 09/07/2023 | 04/16/2024 |  | 10/03/2023 |  |  | 09/07/2023 | 10/03/2023 |  |  | 06/30/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024359658 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2023 |  |  |  |  | The user facility in japan reported an 86-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. during the support a thrombus was discovered in the patient's left atrium that had adhered to the inlet of the impella affecting the flows. ecmo was added. the field team was made aware that the patient's cardiac function was improving, however, after the removal of the impella 5.5 and ecmo, the patient had a cerebral infarction and expired. the cause of death was determined to be cerebral infraction. the relationship to the impella and the cause of death is unknown, but cannot be denied. the patient was heparin free for two days after the impella insertion that potentially led to the thrombus and cerebral infarction. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus and the stroke/cva has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and the stroke/cva was not determined.
34150 | 1220648-2023-03647 | 17864551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864551&pc=OZD | 08/19/2022 | 01/25/2024 |  | 10/03/2023 |  |  | 08/19/2022 | 10/03/2023 |  |  | 06/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022102411 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced positioning issues during patient support. it was documented that when the patient was taken to the operating room for scheduled coronary artery bypass grafting (cabg), the motor housing was too close to the innominate artery for a cross clamp to be placed on the aortic arch. the pump was measured at 3.5cm from the aortic annulus at the time of the noted issue. repositioning was attempted, but the surgeon was afraid that a perforation was possible if heart manipulation was needed during the bypass. the decision was subsequently made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
34151 | 1220648-2023-03650 | 17864683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864683&pc=OZD | 08/21/2022 | 03/15/2024 |  | 10/03/2023 |  |  | 08/21/2022 | 10/03/2023 |  |  | 12/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170251 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The user facility reported a 16-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced bouts of non-sustained ventricular tachycardia and ectopy during impella 5.5 support. the condition was attributed to the positioning of the pump and the device was pulled back roughly 1 centimeter the following morning. support continued with the implanted pump following the event. | [5 empty] | The investigation into the ventricular tachycardia (vt/vf) issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34152 | 1220648-2023-03655 | 17864701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864701&pc=OZD | 08/13/2022 | 09/17/2024 |  | 10/03/2023 |  |  | 08/13/2022 | 10/03/2023 |  |  | 05/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222354 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient suffered an unspecified stroke during impella support. support with the pump initially continued, but due to the patient's poor overall prognosis secondary to the stroke, the decision was made to withdraw care. | Additional information noted care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
34153 | 1220648-2023-03653 | 17864703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864703&pc=OZD | 09/28/2023 | 03/11/2025 |  | 10/03/2023 |  |  | 09/28/2023 | 10/03/2023 |  |  | 03/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023199535 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during explant of the pump, the graft anastomosis became compromised and the md had to regain proximal hemostatic control. the pump was successfully explanted and a patch was used on the axillary artery where the anastomosis was. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34154 | 1220648-2023-03656 | 17864706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864706&pc=OZD | 09/05/2023 | 07/17/2024 |  | 10/03/2023 |  |  | 09/05/2023 | 10/03/2023 |  |  | 05/31/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345457 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 39-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support for a bridge to transplant. during support, the patient developed a hematoma at the insertion site. the patient received two units of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Section h6 has been updated, as the investigation has been completed. according to clinical, the patient developed a hematoma at the insertion site and consequently received two units of packed red blood cells within 24 hours. the impella 5.5 device was not returned, and therefore, an evaluation of the device was not possible. product history review was completed, and the pump passed all post sterile investigating checks. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ correction was made to the following section(s): in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6: date captured to the initial report should be disregarded; this section is not applicable. f8: date captured to the initial report should be disregarded; this section is not applicable.
34155 | 1220648-2023-03651 | 17864716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864716&pc=OZD | 09/07/2023 | 07/19/2024 |  | 10/03/2023 |  |  | 09/07/2023 | 10/03/2023 |  |  | 05/31/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366456 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2023 |  |  |  |  | The user facility reported an 82-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's groin was bleeding and manual pressure being held. patient apparently threw his leg up over night and was moving the impella access site. groin started oozing at this time. patient did not have a knee immobilizer. patient again slightly moved leg after shift change and groin started bleeding heavily. manual pressure was held and bleeding would stop with manual pressure. every time manual pressure was released, bleeding would resume. patients heparin was discontinued, blood and fluids were given. after one hour, still no hemostasis. the physician decided he wanted to remove the pump at this time and trial weaned patient at p2. after two hours, patient clotted and access site maintained hemostasis. the physician still wanted to remove even though hematoma was absent and patient on multiple drips. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleed was most likely patient movement as bleeding occurred in conjunction with the patient moving. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank in the initial report.
34156 | 1220648-2023-03657 | 17864737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864737&pc=PYX | 08/27/2023 | 10/09/2024 |  | 10/03/2023 |  |  | 09/06/2023 | 10/03/2023 |  |  | 07/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023281610 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2023 |  |  |  |  | The user facility reported a 41-year-old male with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. before implant, a self-resolving pump stop alarm occurred. the case continued after the event and the impella rp was successfully inserted for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the pump stop has been completed. according to the clinical, a high motor current pump stop occurred. the pump was able to be restarted and self-resolved. no product was returned for a visual inspection. data log analysis confirmed a high motor current pump stop around 23 hours into support while on p-6. prior to the pump stop there were multiple motor current spikes with speed deviations independent of p-level. after the motor current spikes, purge pressure, purge flow, placement signal, and motor current were variable for 5 hours before eventually stabilizing. the cause of the temporary pump stop was not determined because no product was returned and not enough clinical information was given. the instructions for use for the impella rp with smartassist system states the following: section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ corrections to the initial mfr report: added-d6a/d6b, f6/f8 should have been left blank, h6 med dev prob code 1502 changed to 1503.
34157 | 1220648-2023-03659 | 17864746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864746&pc=OZD | 10/14/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 10/14/2022 | 10/03/2023 |  |  | 05/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023224030 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2022 |  |  |  |  | The user facility reported a 42-year-old female experiencing cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed an access site bleed/ooze during impella support. silver nitrate and manual pressure were applied to stop the bleed. furthermore, heparin was discontinued and removed from the purge. the oozing subsequently resolved and patient support continued with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34158 | 1220648-2023-03658 | 17864751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864751&pc=OZD | 08/09/2022 | 01/26/2024 |  | 10/03/2023 |  |  | 08/09/2022 | 10/03/2023 |  |  | 05/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022086848 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The complainant reported a 79-year-old male patient suffering from acute myocardial infarction and cardiogenic shock was presented for impella 5.5 placement. during support, it was documented that the patient developed a hematoma at the impella access site. the site was evacuated and systemic heparin was halted to treat the noted condition. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34159 | 1220648-2023-03662 | 17864780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864780&pc=OZD | 09/08/2023 | 07/16/2024 |  | 10/03/2023 |  |  | 09/08/2023 | 10/03/2023 |  |  | 03/31/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024326086 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2023 |  |  |  |  | The user facility in japan reported a 78-year-old female post cardiotomy cardiogenic shock and with low cardiac output syndrome undergoing aortic valve replacement (avr) surgery was scheduled to be switched to an impella cp when weaned from cardiopulmonary bypass after surgery. after impella cp was inserted, there was severe bleeding near the avr anastomosis, and although the condition was observed for several hours, the amount of bleeding did not change. the impella was removed and an intra-aortic balloon pump was inserted. the doctor stated that the impella entered deeply, and at that time, the flow at the outlet site may have interfered with the aortic valve anastomosis site. a blood transfusion was performed. on (b)(6) 2023, abiomed was made aware of an update/new clinical report, that the bleeding was from the graft anastomosis. additionally, before the cardiopulmonary bypass, the patient's chest had been opened; bleeding was able to be confirmed visually. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported vascular damage has been completed. according to the clinical, shortly after placing the impella cp device, bleeding was observed near the avr anastomosis. the physician noted that during insertion there was a situation where the pump entered deeply. this bleeding persisted for several hours until the pump was removed and replaced with iabp in an attempt to resolve the issue. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the great vessel vascular damage was use related. the instructions for use for the impella cp with smartassist in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 clinical code 1888 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
34160 | 1220648-2023-03661 | 17864802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864802&pc=OZD | 09/03/2023 | 02/02/2024 |  | 10/03/2023 |  |  | 09/03/2023 | 10/03/2023 |  |  | 09/30/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, that a hematoma was noted in the right femoral artery, the impella access site. manual pressure was applied for 45 minutes. protamine was given and heparin was put on hold. support was maintained throughout the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported , access site bleeding has been completed since the original report was filed. product was not returned for review. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
34161 | 1220648-2023-03646 | 17864828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864828&pc=OZD | 10/22/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 10/22/2022 | 10/03/2023 |  |  | 05/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023224590 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2022 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced thrombocytopenia and required a transfusion of platelets. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the thrombocytopenia issue was not determined due to insufficient clinical details. this report is being filed as part of a retrospective review of historical records.
34162 | 1220648-2023-03660 | 17864840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864840&pc=OZD | 08/03/2022 | 09/22/2024 |  | 10/03/2023 |  |  | 08/03/2022 | 10/03/2023 |  |  | 04/30/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023213175 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced high purge pressure during patient support. the pressure increases coincided with tissue plasminogen activator infusion. the staff was advised on the risks of unresolved high purge pressures and their effect on the pump and purge system. care was withdrawn a couple days later and the pump was explanted. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03660 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34163 | 1220648-2023-03664 | 17864864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864864&pc=OZD | 11/22/2021 | 09/17/2024 |  | 10/03/2023 |  |  | 11/22/2021 | 10/03/2023 |  |  | 06/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022106325 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/22/2021 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery, for a mitral valve replacement, was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a stroke and went to cardiac arrest. this occurred after just under 5 days of support. patient later expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03664 was provided.
34164 | 1220648-2023-03666 | 17864926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864926&pc=OZD | 12/14/2021 | 01/26/2024 |  | 10/03/2023 |  |  | 12/14/2021 | 10/03/2023 |  |  | 09/30/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022144198 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that difficulty was encountered while attempting to place an impella 5.5 pump for impella support, due to the patient's anatomy. manipulation was attempted, but the decision was made to use another impella 5.5 pump for implant. coinciding with the delivery issues, bleeding complications were noted around the anastomosis of the graft and aorta. an unknown quantity of blood products were given and the site was surgically repaired to treat the bleed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported delivery issues and the access site bleeding has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause of the delivery issues. according to the clinical details, md had difficulty crossing ao valve resulting the impella to bend in 2 locations. therefore the root cause of the delivery issues was most likely user-related. without sufficient details nor a returned device, the root cause of the access site bleeding was unable to be determined.
34165 | 1220648-2023-03665 | 17864928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864928&pc=OZD | 12/31/2021 | 02/06/2024 |  | 10/03/2023 |  |  | 12/31/2021 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022123560 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/31/2021 |  |  |  |  | The user facility reported a 31-year-old black undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a purge leak was noted from an impella 5.5 sidearm during patient support. a visual inspection identified that there was a crack in the purge line. bone wax was in place to seal the leak and instructions were provided to the staff on how to perform a purge bypass if needed. support with the implanted pump was maintained. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol¿. this report is being filed as part of a retrospective review of historical records. | The investigation into the purge leak ¿ pump issue has been completed. based on the clinical details, the root cause of the purge leak - pump is most likely due to a damaged sidearm. product was not returned, and insufficient clinical details were provided. the location on the sidearm or cause of the leak is unknown. | G.3 the date received by the manufacturer on the previously submitted manufacturer device report (mdr) should have been 28-jan-2024 and not 27-jan-2024. h.3 device not returned to manufacturer for evaluation did not need to be marked on the previously submitted mdr as it was already selected on the original mdr that was submitted.
34166 | 1220648-2023-03649 | 17864950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864950&pc=OZD | 09/06/2023 | 12/07/2023 |  | 10/03/2023 |  |  | 09/06/2023 | 10/03/2023 |  |  | 07/31/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375099 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male noted as high risk percutaneous cardiac intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. the impella was removed without issue and the right axillary site was washed out. there was some thrombus present, so thrombectomy was performed post impella removal on right axillary to continued flow to right arm. additionally, the patient was noted to have had a mild stroke and is currently recovering. | [5 empty] | The investigation into patient's thrombus and stroke/cva has been completed. the product was not returned for evaluation. the data logs showed that there are no motor current spikes that can be indicative of where the stroke could have occurred, and no clear sign of ingestion. there are insufficient details provided to determine when the stroke occurred. it was noted that thrombus was seen upon explant. the relationship of the impella to the stroke is unknown. the root cause of the stroke and thrombus is most likely due to the patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34167 | 1220648-2023-03652 | 17864967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17864967&pc=OZD | 08/22/2022 | 03/11/2025 |  | 10/03/2023 |  |  | 08/22/2022 | 10/03/2023 |  |  | 05/31/2024 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023225943 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient had transcatheter aortic valve replacement (tavr) and the medical staff was unable to snare 5.5 back across valve. the patient was stable after this event, however, impella had to be explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34168 | 1220648-2023-03667 | 17865060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17865060&pc=OZD | 08/01/2023 | 01/16/2024 |  | 10/03/2023 |  |  | 08/01/2023 | 10/03/2023 |  |  | 03/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LEFT DIRECT |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2023188894 |  | 005082 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a failure to deliver the impella device due to patient's calcified vasculature. insertion was aborted due to potential damaged endured to the device. a new impella was used without further issues. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported delivery issues has been completed. the device was not returned for evaluation. however, clinical details stated that the patient had calcification of the axillary artery. therefore, the root cause of the delivery issue was determined to be patient condition.
34169 | 1220648-2023-03669 | 17865104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17865104&pc=OZD | 09/03/2023 | 10/03/2023 |  | 10/03/2023 |  |  | 09/03/2023 | 10/03/2023 |  |  |  |  | 10/03/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1160666 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported had brought a console (automated impella controller (aic)) in for the impella support. the aic was found to have a broken purge disc. the team removed and replaced the console with another per ifu recommendations. | [5 empty] | The product was returned for investigation. the device was visually inspected and found the console was tested after arriving in fs. the purge disc retainer and flag were confirmed to be broken. data analysis found that (b)(6) had nothing in the imc or rt logs indicating the purge flag and retainer had been dislodged. as per clinical investigation, the complainant had brought a console (aic) in for the impella support. the aic was found to have a broken purge disc. the team removed and replaced the console with another per ifu recommendations. the investigation determined that the cause of the issue was a missing purge retainer and flag.
34170 | 1220648-2023-03663 | 17865287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17865287&pc=OZD | 01/01/2022 | 01/11/2024 |  | 10/03/2023 |  |  | 01/01/2022 | 10/03/2023 |  |  | 08/31/2023 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022136209 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2022 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the pump alarmed for/experienced placement signal low alarms. it was alleged that when the patient moved up the chair the pump appeared to have dove deeper. when laying the patient back the position corrects. impella was assessed for repositioning via echo guidance due to being too deep and getting placement signal low alarms. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was most likely patient movement that resulted in placement signal alarms. the failure modes will be monitored and trended.
34171 | 1220648-2023-03668 | 17865344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17865344&pc=OZD | 09/03/2023 | 01/24/2024 |  | 10/03/2023 |  |  | 09/03/2023 | 10/03/2023 |  |  | 06/30/2025 |  | 10/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373494 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella was placed and started but no flows showed on the pump. the motor current was also flat. the pump was repositioned with no change in the flow or motor current. the pump was replaced. the pump was run in a bowl of water as a test and still no raise in motor current or flows. there were only flows in the bowl with no visible clot. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup¿. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop¿. | The investigation into the low pump flow issue has been completed. the device was not returned for benchtop investigation. data confirmed the low flow with multiple spiked mc when pump started & remained to the rest of the case that triggered auto suction response & suction alarms. based on clinical description & data review, the root cause of low flow was most likely due to biomaterial ingestion.
34172 | 1220648-2023-03522 | 17851006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851006&pc=OZD | 07/06/2021 | 01/26/2024 |  | 10/02/2023 |  |  | 07/06/2021 | 10/02/2023 |  |  | 04/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071822 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2021 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a large hematoma at the access site and required 1 unit of blood. support with the implanted pump was maintained following the event. | [5 empty] | The impella device was not returned by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hematoma has been completed. the device was not returned for evaluation nor was clinical information provided sufficient. therefore, the root cause of the reported issue could not be determined.
34173 | 1220648-2023-03523 | 17851082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851082&pc=OZD | 07/09/2021 | 03/11/2025 |  | 10/02/2023 |  |  | 07/09/2021 | 10/02/2023 |  |  | 04/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022078136 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/09/2021 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed sepsis during impella support for which the impella was suspect. the impella was successfully explanted and cultures of unknown results were taken from the site. additional information was not provided regarding the event. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | An infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician physical examination findings results of special investigations : laboratory test results & imaging studies microbiological culture and sensitivity results response to previous treatments and therapies any relevant epidemiological information or exposure history concomitant devices in use preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. this report is being filed as part of a retrospective review of historical records.
34174 | 1220648-2023-03524 | 17851100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851100&pc=OZD | 07/09/2021 | 10/02/2023 |  | 10/02/2023 |  |  | 07/09/2021 | 10/02/2023 |  |  | 05/31/2025 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088857 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/09/2021 |  |  |  |  | The user facility reported a 86-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient had heavy access site oozing during impella support. the bandages were removed and an additional stitch was placed to control the bleeding. support with the implanted pump continued following the event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34175 | 1220648-2023-03525 | 17851135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851135&pc=OZD | 07/09/2021 | 01/24/2024 |  | 10/02/2023 |  |  | 07/09/2021 | 10/02/2023 |  |  | 04/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022081784 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/09/2021 |  |  |  |  | The user facility reported a 38-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed oozing at their impella access site during pump support. a jackson-pratt drain was placed and the patient was taken to the operating room to stitch the source of the bleed at the axillary site. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation.
34176 | 1220648-2023-03526 | 17851285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851285&pc=OZD | 07/11/2021 | 01/26/2024 |  | 10/02/2023 |  |  | 07/11/2021 | 10/02/2023 |  |  | 04/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022095900 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2021 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding at the axillary graft anastomosis requiring blood transfusion. decision was made to take the impella out and insert peripheral extracorporeal membrane oxygenation (ecmo) on lower flow setting. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34177 | 1220648-2023-03527 | 17851373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851373&pc=OZD | 07/11/2021 | 11/19/2024 |  | 10/02/2023 |  |  | 07/11/2021 | 10/02/2023 |  |  | 04/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022074322 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2021 |  |  |  |  | The user facility reported a 19-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed an infection around their impella access site during pump support. no noted treatment was reported for the infection and support continued with the implanted pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the patient injury is patient condition. it was mentioned that there were concerns of infection in the impella site and the medical team planned to culture and determine whether there was an infection at impella site. later on the date of reporting the concern of infection, there was no further update regarding the concern of infection and the pump continued to be in use till the end of the case and no concerns were reported that related to the pump. hence, the root cause is patient condition. hence, the relationship to impella is unknown. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34178 | 1220648-2023-03528 | 17851406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851406&pc=OZD | 06/15/2021 | 01/19/2024 |  | 10/02/2023 |  |  | 06/15/2021 | 10/02/2023 |  |  | 02/28/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022052235 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2021 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced generalized bleeding from an unknown source during impella support. the patient went to the operating room and was transfused with 4 units red blood cell, 2 units of platelets and 4 units of fresh frozen plasma to counteract the bleed. support with the implanted pump continued following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. the device was not returned by the medical facility, only the clinical report was evaluated. the root cause of the patient injury is patient condition. it was reported that patient was taken to or for generalized bleeding and there was no additional information that was provided that identifies as the potential cause for bleeding. it was mentioned that there was no clear source that was identified and there were no details that can be verified that related impella 5.5 for the cause of bleeding.
34179 | 1220648-2023-03529 | 17851540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851540&pc=OZD | 07/21/2021 | 01/24/2024 |  | 10/02/2023 |  |  | 07/21/2021 | 10/02/2023 |  |  | 02/28/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054946 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Swelling/ Edema | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2021 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed swelling at their impella access site during support. the patient was provided an unspecified quantity of packed red blood cells (prbc) and support continued with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation.
34180 | 1220648-2023-03530 | 17851620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851620&pc=OZD | 10/18/2021 | 09/18/2024 |  | 10/02/2023 |  |  | 10/18/2021 | 10/02/2023 |  |  | 04/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071374 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/18/2021 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a right sided stroke during impella support. it was noted that the right hemisphere of the patient's brain was infarcted and edemic. due to the noted condition, the decision was made by the patient's family to withdraw care. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva has been completed since the original report was filed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03530 was provided.
34181 | 1220648-2023-03531 | 17851670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851670&pc=OZD | 10/16/2021 | 03/11/2025 |  | 10/02/2023 |  |  | 10/16/2021 | 10/02/2023 |  |  | 04/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071875 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2021 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a thrombus was discovered in the apex of the left ventricle during impella support. due to the thrombus, a discussion was had with the family and the decision was made to continue utilizing the implanted pump despite the condition. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency) relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined" this report is being filed as part of a retrospective review of historical records.
34182 | 1220648-2023-03532 | 17851722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17851722&pc=OZD | 10/15/2021 | 09/18/2024 |  | 10/02/2023 |  |  | 10/15/2021 | 10/02/2023 |  |  | 02/28/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022043420 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/15/2021 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient showed a neurological change overnight during impella support. upon reviewing the patient's condition, an acute subdural hematoma was identified coinciding with watershed stroke lesions. as a result of the condition, the patient's family opted to withdraw care. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva has been completed since the original report was filed. the device was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03532 was provided.
34183 | 1220648-2023-03533 | 17852132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17852132&pc=OZD | 10/14/2021 | 03/01/2024 |  | 10/02/2023 |  |  | 10/14/2021 | 10/02/2023 |  |  | 03/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022056281 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2021 |  |  |  |  | The user facility reported a 12-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a hemorrhagic stroke during impella support. only purge heparin was being administered at the time of the noted stroke. the patient was noted to have stabilized following the stroke and support continued with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined.
34184 | 1220648-2023-03534 | 17852716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17852716&pc=OZD | 09/04/2023 | 02/05/2025 |  | 10/02/2023 |  |  | 09/04/2023 | 10/02/2023 |  |  |  |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565140 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/04/2023 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the impella automated controller (aic) displayed a controller failure alarm. the aic was replaced and support continued without harm or injury to the patient. | [5 empty] | The impella console has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. according to the device and data analysis the cause of the aic issue was an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets. a5 should have been left blank on manufacturer device report 1220648-2023-03534, since the ethnicity was reported as unknown. d2: common device name revised on manufacturer device report 1220648-2023-03534 in accordance with updated procedures. d4: model number was inadvertently reported incorrectly manufacturer device report 1220648-2023-03534 f6/f8-should have been left blank on manufacturer device report 1220648-2023-03534. g1 reporting contact email revised on manufacturer device report 1220648-2023-03534 in accordance with updated procedures.
34185 | 1220648-2023-03535 | 17852721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17852721&pc=OZD | 09/04/2023 | 09/20/2024 |  | 10/02/2023 |  |  | 09/04/2023 | 10/02/2023 |  |  | 09/30/2024 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2023261041 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility, in japan, reported a 62-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced a cerebral infarction and expired. the cause of cerebral infarction cannot be determined. | It was noted that the patient expired while on support. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/cva has been completed since the original report was filed. since enough clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03535 was provided in sections b2, b5, d6, h1, and h6.
34186 | 1220648-2023-03536 | 17852751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17852751&pc=OZD | 09/06/2023 | 08/16/2024 |  | 10/02/2023 |  |  | 09/06/2023 | 10/02/2023 |  |  | 05/31/2025 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355548 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Malposition of Device | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2023 |  |  |  |  | The user facility reported a 52-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient was on support for over 8 days when the team had the pump come out of placement. the mal-positioning occurred when the patient pulled upon the pump cord. the pump failed to reposition and the pump stopped. the stop occurred after the md believed the pump ingested clot at the mal-positioning. decision made not to replace the impella. the patient was successfully weaned and the site surgically closed. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. of note the ifu states: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump positioning and pump stop has been completed. the impella pump was returned for evaluation. upon visual inspection, the three electrical wires were found to be broken inside the red handle suggesting that excessive force was applied. clinical details stated that the patient had tugged on the wire in the morning which led to the impella being pulled into the aorta and "pump stop" alarm. therefore, the root cause of the pump stop issue was an open electrical line due to excessive force exerted by the patient while tugging the impella. the root cause of the positioning issues was likely patient movement as the pump was in the graft after the patient had pulled on the pump. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen. impella 5.5 with smartassist : section: troubleshooting the purge system. unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. section: impella stopped. if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency. impella 5.5 with smartassist for use during cardiogenic shock: section: warnings, cautions, and precautions. to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2023-03536. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-03536 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-03536 and should not have been. h.6 code 3009 was reported incorrectly on the initial manufacturer device report number 1220648-2023-03536. codes 4627 and 4628 were inadvertently omitted from the initial manufacturer device report number 1220648-2023-03536 and were added to this report. a new code has been added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-03536 in accordance with updated procedures.
34187 | 1220648-2023-03537 | 17852856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17852856&pc=OZD | 10/22/2021 | 03/11/2025 |  | 10/02/2023 |  |  | 10/22/2021 | 10/02/2023 |  |  | 05/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088307 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2021 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella 5.5 was unable to make the turn into the aorta during attempted delivery. multiple attempts were made until a retrograde dissection was noted via angiogram. it was unknown when the dissection occurred. due to the noted issue, the delivery was aborted and the site was closed. no further intervention was completed for the dissection. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | Vascular damage - the cause of the vascular damage was patient condition related since a retrograde dissection was discovered prior to the implant of pump. delivery issue - the root cause of the delivery issue- anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
34188 | 1220648-2023-03538 | 17852864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17852864&pc=OZD | 09/05/2023 | 01/22/2024 |  | 10/02/2023 |  |  | 09/05/2023 | 10/02/2023 |  |  |  |  | 10/02/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1251852 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility biomedical engineer reported their automated impella controller (aic) was being set up for a procedure when the team inserted the cassette and discovered the gray piece that the purge discs plug in to came off. it separated from the wire behind it. the instructions for use were followed and a back up aic was used. there was no patient involvement. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issue has been completed. the device and the data logs were returned for evaluation. the data logs showed that upon inserting the purge cassette, the console received the case start: purge fluid not detected error, which indicates the blue purge retainer was most likely broken and came off from the aic, later case start cancelation has been confirmed. this aligns with the reported failure. the console was tested on an engineering lab bench. upon visual inspection, it was confirmed that the blue purge retainer and the flag were broken. therefore, the root cause for the broken blue purge retainer was due to the defective component.
34189 | 1220648-2023-03539 | 17852886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17852886&pc=OZD | 10/21/2021 | 05/03/2024 |  | 10/02/2023 |  |  | 10/21/2021 | 10/02/2023 |  |  | 06/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022109087 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2021 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced some some suction alarms, potentially related to not enough blood being returned by perfusion. decision was made to explant impella. | [5 empty] | The impella device was not returned by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the cause of the low pump flow was patient condition since enough blood supply was not present. d6a, d6b.
34190 | 1220648-2023-03540 | 17853182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853182&pc=OZD | 08/13/2021 | 10/02/2023 |  | 10/02/2023 |  |  | 08/13/2021 | 10/02/2023 |  |  | 01/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022037358 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2021 |  |  |  |  | The complainant reported a 65-year-old male patient suffering from cardiomyopathy was presented for impella 5.5 placement. during support, it was documented that a fluid leak was seen from the purge sidearm. a purge bypass was performed, and the purge pressure/purge flow stabilized. there was no patient harm and support with the implanted pump was maintained. | [5 empty] | The device was returned and an evaluation was completed. the purge sidearm filter section was returned attached to the bypass that had been put in place to troubleshoot the leak. unfortunately, the blue pressure reservoir was not returned, and reproduction testing could not be completed due to the condition of the returned product. a visual inspection noted no signs of physical damage to the filter section of the device. a clinical review also observed that sodium bicarbonate had been used in the purge system during support. with the limited information. it was advised that the cause for the noted leak was a potential bond failure. this report is being filed as part of a retrospective review of historical records.
34191 | 1220648-2023-03541 | 17853214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853214&pc=OZD | 10/22/2021 | 11/19/2024 |  | 10/02/2023 |  |  | 10/22/2021 | 10/02/2023 |  |  | 06/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098176 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2021 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at their access site during impella support. as the patient's hemodynamics were stable, the decision was made to explant the pump and evacuate the hematoma in the operating room. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the access site bleeding cannot be determined since enough clinical info is not provided. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34192 | 1220648-2023-03542 | 17853461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853461&pc=OZD | 10/22/2021 | 10/02/2023 |  | 10/02/2023 |  |  | 10/22/2021 | 10/02/2023 |  |  | 08/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022123555 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/22/2021 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. following the placement of an impella 5.5 pump, a clot was noted on the device. as a result of the biomaterial, the decision was made to explant the pump. the patient was stable following the event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation of the provided information, it was advised that the patient condition caused or contributed to the noted event. should additional relevant information become available, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34193 | 1220648-2023-03543 | 17853465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853465&pc=OZD | 09/04/2023 | 01/25/2024 |  | 10/02/2023 |  |  | 09/04/2023 | 10/02/2023 |  |  |  |  | 10/02/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1386344 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Charge | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility in the eu reported that the battery does not charge on the automated impella controller. red battery failure alarm displayed. there was not any noted patient use at the time, and as such no harm or injury. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - battery failure (red alarm) has been completed since the original report was filed. the console was inspected and evaluated on an engineering lab bench. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure, a known pattern failure.
34194 | 1220648-2023-03544 | 17853689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853689&pc=OZD | 05/03/2023 | 10/02/2023 |  | 10/02/2023 |  |  | 05/03/2023 | 10/02/2023 |  |  | 03/31/2025 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024327620 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2023 |  |  |  |  | The user facility reported 20yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the complainant reported abnormal placement signal readings and high purge pressure during support. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the high purge pressure and the placement signal issues has been completed. the cause of the high purge pressure was not determined since the issue could not be reproduced, no abnormalities were found on the returned product, and no logs were returned. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34195 | 1220648-2023-03545 | 17853702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853702&pc=OZD | 09/05/2023 | 01/25/2024 |  | 10/02/2023 |  |  | 09/05/2023 | 10/02/2023 |  |  |  |  | 10/02/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1107791 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility's biomedical engineer reported that the automated impella controller (aic) has a broken purge clip. the aic will be returned. no patient involvement. | [5 empty] | The impella console has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the console aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and the purge disc retainer was confirmed to be broken. the cause of the issue was a broken purge pressure disk retainer.
34196 | 1220648-2023-03546 | 17853857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853857&pc=OZD | 10/28/2021 | 03/11/2025 |  | 10/02/2023 |  |  | 10/28/2021 | 10/02/2023 |  |  | 06/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022099214 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2021 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed an infection around their right axillary access site during impella support. due to the noted condition, the pump was explanted and a new impella 5.5 pump was implanted via left axillary access for ongoing support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | An infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician physical examination findings results of special investigations: laboratory test results & imaging studies microbiological culture and sensitivity results response to previous treatments and therapies any relevant epidemiological information or exposure history concomitant devices in use preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. approved this report is being filed as part of a retrospective review of historical records.
34197 | 1220648-2023-03408 | 17853904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853904&pc=OZD | 04/25/2023 | 09/26/2023 |  | 10/02/2023 |  |  | 04/25/2023 | 10/02/2023 |  |  | 03/31/2025 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024324192 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/25/2023 |  |  |  |  | The user facility reported 60 yo male was implanted with an impella 5.5 device for mechanical circulatory support.. it was noted that the impella flows were seemingly out of range and an impella in aota alarm was triggered. after troubleshooting, the readings showed potential saturated flow and the decision was made to transfer the pump to another console. once on the new console, the alarms resolved. | [5 empty] | The investigation for the reported saturated flow issue has been completed. data log showed that about 2 minutes after starting the pump, there several motor current spikes in quick succession, accompanied by a simultaneous drop in motor speed and elevated lv waveforms. immediately afterwards, significant flow saturation is observed. additionally, the "impella position in aorta" alarm triggers even though the placement signal does not become aortic. in conclusion, it was determined that the cause of saturated flow was most likely ingested biomaterial. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34198 | 1220648-2023-03547 | 17853914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17853914&pc=OZD | 10/20/2021 | 10/02/2023 |  | 10/02/2023 |  |  | 10/20/2021 | 10/02/2023 |  |  | 03/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022065615 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/20/2021 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient's surgical graft had become infected in the area in which the impella had been inserted prior. the site was treated surgically, the infection was removed and patched. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the reported infection was confirmed unrelated to impella by the medical office and is likely related to surgical technique. this report is being filed as part of a retrospective review of historical records.
34199 | 1220648-2023-03548 | 17854031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17854031&pc=OZD | 11/23/2021 | 09/17/2024 |  | 10/02/2023 |  |  | 11/23/2021 | 10/02/2023 |  |  | 08/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022133335 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/23/2021 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a right sided stroke during impella support. after speaking with the family regarding the noted condition, the decision was made to withdraw care. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva issue has been completed. the device was not returned for benchtop investigation. however, the data logs were returned, and no motor current spikes were observed. per the given clinical information and data logs, the root cause of the stroke issue was most likely patient condition related and the relation to impella is unknown. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34200 | 1220648-2023-03549 | 17854080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17854080&pc=OZD | 09/04/2023 | 08/09/2024 |  | 10/02/2023 |  |  | 09/04/2023 | 10/02/2023 |  |  | 05/31/2025 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the heart team made the decision for escalation to an impella 5.5 due to ischemic leg and need for ongoing mechanical circulatory support (mcs). | The patient did not survive the procedure. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information d6a and d6b added g1 reporting contact fax number missing on manufacturer device report 1220648-2023-03549.
34201 | 1220648-2023-03550 | 17854777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17854777&pc=OZD | 11/18/2021 | 10/29/2025 |  | 10/02/2023 |  |  | 11/18/2021 | 10/02/2023 |  |  | 08/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022124061 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2021 |  |  |  |  | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that they experienced difficulty during the delivery of an impella 5.5 pump. upon insertion and manipulation of the 5.5 pump to pass through the anastomosis and innominate, the peel away came out of the graft and the patient lost a unit of blood (500ml). the pump was flushed and then successfully placed across the aortic valve. however, upon powering up the device, the position of the pump changed towards the mitral valve and the device torqued into the mid-left ventricle and closer to the septum. the pump stopped momentarily until the torque produced during repositioning was slackened. support continued with the implanted pump following the troubleshooting. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
34202 | 1220648-2023-03551 | 17854998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17854998&pc=OZD | 11/09/2021 | 03/05/2024 |  | 10/02/2023 |  |  | 11/09/2021 | 10/02/2023 |  |  | 08/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022125943 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2021 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced a stroke on day 2 in icu. the patient was hemodynamically doing very well and the impella was weaned and explanted. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
34203 | 1220648-2023-03552 | 17855278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855278&pc=OZD | 11/06/2021 | 03/11/2025 |  | 10/02/2023 |  |  | 11/06/2021 | 10/02/2023 |  |  | 06/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022096668 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2021 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced positioning issues and required the doctor to reposition the pump in the or. about 20 days later the same patient was reported to have suboptimal position and the patient experienced persistent ventricular tachycardia(vt) leading the doctor to explant the pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Positioning issue - no product was returned. the root cause of the positioning issue was most likely use related as per the clinical description provided. vt - the cause of the vt was use issue since the position of the device was poorly placed and repositioning had be attempted multiple times. high purge pressure - the cause of the high purge pressure was purged kinked line due to patient movement. this report is being filed as part of a retrospective review of historical records.
34204 | 1220648-2023-03553 | 17855305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855305&pc=OZD | 11/04/2021 | 01/24/2024 |  | 10/02/2023 |  |  | 11/04/2021 | 10/02/2023 |  |  | 08/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022132641 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2021 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced significant oozing/bleeding from their axillary site. an additional suture was added and manual pressure was applied to treat the condition. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the access site bleed has been completed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product was returned for investigation.
34205 | 1220648-2023-03554 | 17855322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855322&pc=PYX | 09/05/2023 | 02/06/2024 |  | 10/02/2023 |  |  | 09/05/2023 | 10/02/2023 |  |  | 07/31/2024 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023229839 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Tamponade | Pericardial Effusion | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 65-year-old male in non-st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. right internal jugular percutaneous access obtained using micropuncture kit under ultrasound guidance. rp sheath and dilator were used with subsequent perforation of inferior vena cava/right atrium junction resulting in large pericardial effusion and tamponade, requiring conversion to emergent sternotomy and cannulation for cardiopulmonary bypass. | [5 empty] | The impella rp device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined.
34206 | 1220648-2023-03556 | 17855374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855374&pc=OZD | 09/04/2023 | 01/25/2024 |  | 10/02/2023 |  |  | 09/04/2023 | 10/02/2023 |  |  | 12/31/2024 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024300753 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility in japan reported that there was difficulty in adjusting the position of the impella cp. while in the icu the pump shifted to the aorta. the patient¿s vitals were stable so impella continued to be used until the next day when it was explanted. no harm or injury noted from the malpositioning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. product and data were not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement as it was out of position after moving the patient to the icu.
34207 | 1220648-2023-03555 | 17855501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855501&pc=OZD | 11/03/2021 | 03/11/2025 |  | 10/02/2023 |  |  | 11/03/2021 | 10/02/2023 |  |  | 08/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128821 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2021 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was experiencing high purge pressure during patient support. the cassette was replaced in an attempt to troubleshoot the high pressure, but the pressure remained high (1065 mmhg) with a purge flow of 1.2 ml/hr. with an upper ideal pressure limit of 1100mmhg, the staff was advised that tissue plasminogen activator (tpa) was needed to prevent a potential purge system blocked alarm. the staff declined the usage of tpa as the patient 's chest was still open and they were scheduled for surgery. support continued under close monitoring and the pump was explanted a couple days later. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34208 | 1220648-2023-03557 | 17855536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855536&pc=OZD | 10/02/2021 | 03/28/2025 |  | 10/02/2023 |  |  | 10/02/2021 | 10/02/2023 |  |  | 10/31/2022 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022033242 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Female |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility in the eu reported a 21-year-old female noted as high risk percutaneous cardiac intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 had suction alarms with higher flows during patient support. the physician suspected thrombus in inlet area. despite the alarms, the pump remained on for support for a total of 80 days, was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The cause of the low flow was patient condition since right ventricle was weak and volume had to be given. this report is being filed as part of a retrospective review of historical records.
34209 | 1220648-2023-03558 | 17855537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855537&pc=OZD | 03/28/2022 | 03/08/2024 |  | 10/02/2023 |  |  | 03/28/2022 | 10/02/2023 |  |  | 10/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154712 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient suffered an unspecified stroke on their right side during impella support. the patient was subsequently monitored to determine the extent of the damage and support continued with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
34210 | 1220648-2023-03559 | 17855610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855610&pc=OZD | 03/21/2022 | 01/11/2024 |  | 10/02/2023 |  |  | 03/21/2022 | 10/02/2023 |  |  | 12/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023175372 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced positioning issues during patient support. it was documented that at one point during support, the pump had shifted and pointed towards the mitral. this resulted in persistent suction alarms for which multiple unsuccessful attempts were made to reposition the device. a placement signal not reliable error was subsequently triggered and the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
34211 | 1220648-2023-03560 | 17855667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855667&pc=OZD | 03/20/2022 | 03/11/2025 |  | 10/02/2023 |  |  | 03/20/2022 | 10/02/2023 |  |  | 11/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166241 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2022 |  |  |  |  | The user facility reported a 41-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a thrombus in the middle cerebral artery during impella support. it was noted that the patient had stenosis at the location and a thrombectomy was subsequently performed. the pump was explanted the same day as the noted intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Analysis summary the failure mode of "thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
34212 | 1220648-2023-03561 | 17855771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855771&pc=OZD | 03/16/2022 | 01/11/2024 |  | 10/02/2023 |  |  | 03/16/2022 | 10/02/2023 |  |  | 11/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166256 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2022 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the aic displayed an impella in aorta alarm coinciding with a flat motor current during patient support. multiple attempts were made unsuccessfully to reposition the pump and the device ended up in the descending aorta. as the patient was reliant on impella support, emergent venoarterial extracorporeal membrane oxygenation (va ecmo) was initiated and the patient was intubated. the pump was removed and a new impella 5.5 pump was placed for support. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the positioning issues has been completed. the product was not returned. per the clinical information provided, the root cause of the positioning issue was most likely use related, as there was limited information provided in the clinic description.
34213 | 1220648-2023-03562 | 17855816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855816&pc=OZD | 03/10/2022 | 03/28/2025 |  | 10/02/2023 |  |  | 03/10/2022 | 10/02/2023 |  |  | 05/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022085236 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2022 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding and oozing from their anastomosis site during impella support. it was noted that difficulty had been met when initially inserting the impella which resulted in a tear in the anastomosis and part of the vessel. the patient required several hours of stitching combined with 5 units packed red blood cells, 3 units of fresh frozen plasma, 1 unit cryoprecipitate, 500 colloid, and 5 l crystalloids to treat the condition. the site was dressed following the intervention and support with the implanted pump continued. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. vascular damage- peripheral vessel- the root cause of the vascular damage was use issue. at the anastomosis site due to initial resistance and push through causing a tear in the anastomosis and part of vessel. vascular damage has occurred because of the excessive fore. hence, the root cause is use issue. access site bleeding- the root cause of the access site bleeding was determined to be use issue. the bleeding has occurred due the initial tear in the anastomosis. hence, the root cause is determined to be use issue.
34214 | 1220648-2023-03563 | 17855840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855840&pc=OZD | 03/09/2022 | 03/28/2025 |  | 10/02/2023 |  |  | 03/09/2022 | 10/02/2023 |  |  | 12/31/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023174069 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2022 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced massive bleeding from the access site which required the patient to be taken in the or for repair and have the bleeding stop. five days after this event. the patient was found to have a clot in the brain. six days after the clot was found decision was made by the medical team to withdraw care and the patient expired. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. access site bleeding- the root cause of the access site bleed was unable to be determined as insufficient clinical information and product was provided for the investigation. thrombus- the root cause of the thrombus is determined to be patient condition. it was mentioned that the clot was found in brain and has no details to determine the cause was related to impella. hence the root cause is patient condition. this complaint has been identified as part of a retrospective review.
34215 | 1220648-2023-03564 | 17855868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855868&pc=OZD | 03/03/2022 | 01/26/2024 |  | 10/02/2023 |  |  | 03/03/2022 | 10/02/2023 |  |  | 09/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149783 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2022 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella had come back across the valve during patient support. the patient was taken to the operating room where the pump was removed and an impella 5.0 was placed for continued support. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
34216 | 1220648-2023-03566 | 17855916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17855916&pc=OZD | 09/08/2022 | 03/11/2025 |  | 10/02/2023 |  |  | 09/08/2022 | 10/02/2023 |  |  | 06/30/2024 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237514 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that following five to ten minutes of support, high purge pressure alarms were triggered. attempts to troubleshoot the catheter were unsuccessful and the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34217 | 1220648-2023-03565 | 17856089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17856089&pc=OZD | 09/07/2022 | 01/26/2024 |  | 10/02/2023 |  |  | 09/07/2022 | 10/02/2023 |  |  | 05/31/2024 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229017 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was unable to pass through the patient's tortuous anatomy during attempted delivery. as a result of the noted resistance, the pump kinked and the implant was subsequently aborted. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the delivery issue has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the delivery issue is patient anatomy, and the anatomy was most likely due to patient condition.
34218 | 1220648-2023-03567 | 17856520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17856520&pc=OZD | 05/23/2022 | 01/10/2024 |  | 10/02/2023 |  |  | 05/23/2022 | 10/02/2023 |  |  | 11/30/2023 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169679 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a 36-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella 5.5 pump experienced multiple positional issues during patient support. as a result, the impella 5.5 pump was explanted and the patient was placed on ecmo. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related, as there were multiple position changes.
34219 | 1220648-2023-03568 | 17857813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17857813&pc=OZD | 09/05/2023 | 11/08/2024 |  | 10/02/2023 |  |  |  | 10/02/2023 |  |  |  |  | 10/02/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1338418 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) ((b)(6)) was replaced with aic ((b)(6)) it was reported that during the removal of the purge disc, the blue retainer broke on the new console ((b)(6)). the impella 5.5 support proceeded without any harm or injury from the interruption. the retainer breakage occurred on day 12 of support. later the family withdrew care. patient outcome not due to the aic malfunction. | Additional information noted that the care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The automated impella controller (aic) device was received from the customer, and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | Corrected information for the initial mfr 1220648-2023-03568 was provided in sections b1, b2, b5, h1, and h6.
34220 | 1220648-2023-03569 | 17857822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17857822&pc=OZD | 09/05/2023 | 08/19/2024 |  | 10/02/2023 |  |  |  | 10/02/2023 |  |  |  |  | 10/02/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a battery/red failure. the instructions for use were followed and a backup aic was used. no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. follow up with the user facility biomedical engineer has not been completed. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) battery red alarm has been completed. the console was returned for evaluation. upon testing of the console, it was confirmed that the batteries were operational and non-defective. data logs were reviewed which confirmed the battery temperature high alarm (event set #49). imc logs also showed that the battery temperature high alarm coincided with the ac power simultaneously being disconnected (event set #109), and the alarm resolved itself approximately 10 seconds later. data analysis indicates that the alarm was solely due to a miscommunication between the pbm and the 4-in-1, causing a false alarm and a misreporting of battery temperature. previous investigations have revealed that the removal of ac power can trigger the pbm to incorrectly report an out-of-range battery temperature. added- d9 device return date, h6 impact code 4629 d2-corrected common device name d4-corrected model number. f6/f8 should have been left blank in the initial report.
34221 | 1220648-2023-03570 | 17858121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17858121&pc=OZD | 09/02/2023 | 02/26/2024 |  | 10/02/2023 |  |  | 09/02/2023 | 10/02/2023 |  |  | 04/30/2025 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347835 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2023 |  |  |  |  | The user facility reported an 86-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a cold and mottled leg on the impella inserted side. the team successfully weaned the patient and explanted the impella cp on (b)(6) 2023. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
34222 | 1220648-2023-03571 | 17858123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17858123&pc=OZD | 09/02/2023 | 01/15/2024 |  | 10/02/2023 |  |  | 09/02/2023 | 10/02/2023 |  |  | 04/30/2025 |  | 10/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient became restless causing the access site to bleed substantially. the patient became hypotensive requiring chemical support. ns bolus was given and manual pressure held followed by fem-stop placement controlling bleeding. patient was also given one unit of blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the access site bleeding issue was patient movement reported during the time of bleeding.
34223 | 1220648-2023-03513 | 17849408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849408&pc=OZD | 01/30/2023 | 10/01/2023 |  | 10/01/2023 |  |  | 01/30/2023 | 10/01/2023 |  |  | 10/31/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023280064 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Transient Ischemic Attack | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/30/2023 |  |  |  |  | The user facility reported that a 46-year-old male patient implanted with the impella 5.5 for mechanical circulatory support displayed facial paralysis. the physician had previously performed an echocardiogram and suspected bubbles in the ascending aorta. a transient ischemic attack (tia) was diagnosed, which resolved two days later. the hospital staff reported that patient also had displayed diplopia on 1/24 but it had since resolved. the facial paralysis resolved as well. computerized tomography (ct) imaging was performed twice along with a ct angiography (cta). all were negative. | [5 empty] | The investigation into the reported embolism was completed. the device was returned for evaluation. analysis of the logs revealed there were no motor current spikes or air in purge alarms. there was no evidence to indicate the embolic stroke was related to the device. sufficient product to perform an evaluation was not returned. the catheter was cut and only the pump cannula was returned. based on the available information, the root cause of the issue was not determined since sufficient product was returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34224 | 1220648-2023-03514 | 17849409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849409&pc=OZD | 04/03/2023 | 10/01/2023 |  | 10/01/2023 |  |  | 04/03/2023 | 10/01/2023 |  |  | 05/31/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023223538 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/03/2023 |  |  |  |  | The user facility reported that a 55-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a leak from the purge side. visual inspection revealed that the purge fluid was leaking from the luer lock of the extension tubing. sodium bicarbonate was in the purge solution and there was no evidence of a leak from the pump luers or purge tubing. the leak resolved when the extension tubing was removed. there was no reported harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34225 | 1220648-2023-03515 | 17849410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849410&pc=OZD | 05/01/2023 | 10/01/2023 |  | 10/01/2023 |  |  | 05/01/2023 | 10/01/2023 |  |  | 11/30/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278485 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/01/2023 |  |  |  |  | The user facility reported that a 23-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the device was unable to function above p4 for a long period of time. the pump was repositioned, and the patient given additional volume; however, the issue persisted. the staff reported that it had been difficult to manage the patient's volume status. an inspection of the pump was requested. the pump was eventually able to function at p7. the patient had pump explanted and received a new heart. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the pump, purge cassette, and data stick were returned for evaluation. analysis of the data logs revealed multiple suction alarms. visual inspection of the pump did not reveal any evidence of biomaterial. based on the available information, the root cause of the low pump flow was determined to be patient's low volume status. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34226 | 1220648-2023-03516 | 17849413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849413&pc=OZD | 08/03/2022 | 10/01/2023 |  | 10/01/2023 |  |  | 08/03/2022 | 10/01/2023 |  |  | 03/31/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023209649 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported that a 53-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge cassette leak coming from the cassette five days into support. the purge cassette was replaced without issue. on the 27th day of support a purge leak was reported to be coming from the yellow luer. staff observed that the purge cassette may have been misaligned when threaded onto the yellow luer. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported purge leak and purge cassette leak was completed. the pump and purge cassette were returned for evaluation. visual inspection of the pump confirmed a crack in the yellow luer. based on the available information, the root cause of the purge leak was determined to be yellow luer damage most likely as a result of use. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34227 | 1220648-2023-03517 | 17849414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849414&pc=OZD | 04/12/2023 | 10/01/2023 |  | 10/01/2023 |  |  | 04/12/2023 | 10/01/2023 |  |  | 12/31/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318778 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Infusion or Flow Problem | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the decision was made to place left sided impella cp with possible right sided support also with and impella rp placement. the cp pulled out into aorta when attempted to reposition. the impella cp was removed from the body and physician made the decision to run in water and use same pump for reinsertion. flows when restarted were 1.2 and continuous suction alarms. after looking at left ventricle (lv) signal, it was suggested to check position again and at that time noted a kink in the impella cannula. discussion between cv surgery and interventional cardiology was held, and the decision was made to pull impella and place iabp. | [5 empty] | The investigation into the reported "positioning issues and low pump flow" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the positioning issue was use related due to cannula kink when pump was malpositioned into aorta during repositioning. data logs were reviewed, and positioning issue and motor current spike alarms were observed. so, the low pump flow was likely due to cannula kink with improper positioning. the root cause of the low pump flow issue was most likely due to kinked cannula based on clinical information and data log review. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34228 | 1220648-2023-03518 | 17849420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849420&pc=OZD | 07/13/2022 | 09/22/2024 |  | 10/01/2023 |  |  | 07/13/2022 | 10/01/2023 |  |  | 03/31/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203520 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Fluid/Blood Leak | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2022 |  |  |  |  | The user facility reported that a 46-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced sudden high purge pressure and a purge leak. the patient initially received a purge blocked/purge pressure high alarm that began noted on the impella connect. the purge cassette was exchanged and the issue troubleshooted without resolution. the cardiologist subsequently ordered cardiologists 2mg tissue plasminogen activator (tpa) in 50ml sterile water. after about 1 hour of tpa, purge flows increased to ~2.2 ml/hr with purge pressure <1000 mmhg. it was additionally reported that the purge sidearm leaked during support, which required a purge bypass to resolve. sodium bicarbonate was used in the purge solution. no further information was provided. the device was still in use at the time the report was written. there was no reported harm to the patient. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported high purge pressure and purge leak was completed. the device was not returned for evaluation. analysis of the data logs confirmed a sudden spike in purge pressure, accompanied by the "purge system blocked" and "purge pressure high" alarms hours before gradually resolving. after reviewing the clinical information, it was determined that the cause of the high purge pressure was most likely biomaterial buildup in the purge gap since the administration of tpa resolved the issue. there is additional evidence that the purge line may have been kinked as well. however, the root cause of the purge leak could not be determined since the device was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03518 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34229 | 1220648-2023-03519 | 17849421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849421&pc=OZD | 04/22/2023 | 10/01/2023 |  | 10/01/2023 |  |  | 04/22/2023 | 10/01/2023 |  |  | 02/28/2025 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024313473 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/22/2023 |  |  |  |  | The user facility reported that a 68-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the physician attempted to reposition the device but was unable to do so successfully. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. the root cause of the positioning issue was not determined since product was not returned and there is insufficient clinical information. his report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34230 | 1220648-2023-03520 | 17849426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849426&pc=OZD | 10/15/2022 | 10/01/2023 |  | 10/01/2023 |  |  | 10/15/2022 | 10/01/2023 |  |  | 06/30/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234679 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/15/2022 |  |  |  |  | The user facility reported that a 47-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the patient had suction when the device ran above p-5. multiple experienced physicians attempted to reposition multiple times under echo and transesophageal echocardiogram (tee) but were unable to without getting suction above p-5. the patient needed to be decannulated from ecmo and remain only on 5.5 support. therefore, the device was replaced to achieve p-9 flows. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the pump and purge cassette were returned for evaluation. analysis of the data logs showed "suction" alarms throughout support with the p-level unable to exceed p-4. no damage was noted on the pump upon visual inspection. functional testing of the pump passed. based on the available information, the root cause of the low pump flow was related to patient condition, specifically to the patient's volume with ecmo use. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34231 | 1220648-2023-03521 | 17849427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17849427&pc=OZD | 07/06/2023 | 10/01/2023 |  | 10/01/2023 |  |  | 07/06/2023 | 10/01/2023 |  |  | 12/31/2024 |  | 10/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315401 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported that a 63-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flow. upon implant of the device, the p level was slowly ramped up to p4. once the p level reached p5 suction was immediately obtained with resolution at p4.transesophageal echocardiogram (tee) showed the device placement to be appropriate. volume given as physician thought the issue could be volume related however, the low pump flows persisted. 10,000 u of heparin was administered prior to implant so clot was not considered an issue. the device was eventually explanted and replaced with an impella cp. impella cp was flowing p7 and 3.2l with no suction. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. visual inspection of the pump revealed no issues. functional testing of the pump did not observe any suction alarms and flow were in the specific flow rate at each p-level. analysis of the data logs confirmed many suction alarms occurred. based on the available information, the root cause of the low pump flow issue most likely was patient condition related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34232 | 1220648-2023-03473 | 17842490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842490&pc=OZD | 08/30/2023 | 01/17/2024 |  | 09/29/2023 |  |  | 08/30/2023 | 09/29/2023 |  |  | 06/30/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360218 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Break | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2023 |  |  |  |  | The user facility reported a 63-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the 14 fr impella short sheath with dilator was placed. the clear piece twisted to remove dilator and the valve and orange piece came out with dilator. there was bleeding significantly out of sheath. broken impella sheath was then exchanged for 16fr check-flo sheath. impella then placed with ease. during support with the implanted pump, it was noted that the left leg became mottled and a doppler pulse could not be found. after the device was explanted, the left leg color and temperature improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleed and limb ischemia has been completed. the device was not returned. the cause of the access site bleeding issue was not determined since location of damage was unable be confirmed without return product. the cause of the limb ischemia issue was most likely patient condition related to small vessels since the impella removal resolved the issue, based on given clinical details.
34233 | 1220648-2023-03474 | 17842514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842514&pc=OZD | 08/30/2023 | 01/25/2024 |  | 09/29/2023 |  |  | 08/30/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369430 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2023 |  |  |  |  | The user facility reported an 82-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. prior to impella explant, post procedure, the patient was noted to have large hematoma at left femoral artery (lfa). interventional cardiology (ic) gained wire access and explanted the pump while leaving sheath. attempts at per-close x 2 failed and plan for dry close and covered stent. the patient had large blood loss and required packed red blood cells. ic successfully placed covered stent to lfa impella access site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the root cause of the access site bleeding-major cannot be determined due to insufficient clinical information and lack of product returned.
34234 | 1220648-2023-03475 | 17842587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842587&pc=OZD | 08/30/2023 | 02/06/2024 |  | 09/29/2023 |  |  | 08/30/2023 | 09/29/2023 |  |  | 02/28/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024319700 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 70-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support ongoing when the team observed an access site bleed. the bleed was treated by the team suturing and 6-10 units of blood delivered. the team did not explant the pump. the support continues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for benchtop investigation. per the clinical information provided, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. the cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided.
34235 | 1220648-2023-03476 | 17842694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842694&pc=OZD | 08/30/2023 | 01/18/2024 |  | 09/29/2023 |  |  | 08/30/2023 | 09/29/2023 |  |  | 04/30/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024339692 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 62-year-old patient was treated for nonischemic cardiomyopathy. he received extracorporeal membrane oxygenation (va-ecmo) support and was ventilated. an impella 5.5 smartassist heart pump was then used for hemodynamic support. high purge pressure was reported on day 28 of impella support. the user performed lysis therapy in an attempt to clear the purge lumen of any potential biomaterial. the day after the incident, the patient was successfully weaned from the impella 5.5 heart pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g.,infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressures has been completed. the device nor data logs were returned for evaluation. therefore, the root cause of the high purge pressure was unable to be determined since the product was not returned and limited clinical details were provided.
34236 | 1220648-2023-03477 | 17842749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842749&pc=OZD | 05/27/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 05/27/2023 | 09/29/2023 |  |  | 03/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327094 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/27/2023 |  |  |  |  | The user facility reported that a 56-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the physician attempted to reposition but stated that with the blue button the repositioning was much harder and seems less maneuverable. there was no reported harm to the patient. | [5 empty] | The investigation into the reported positioning issues was completed. the device was not returned for evaluation. based on the available information, the root cause of positioning issue due to the catheter anchor issue was unable to be determined as no product was returned for review. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34237 | 1220648-2023-03478 | 17842760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842760&pc=OZD | 08/31/2023 | 09/26/2024 |  | 09/29/2023 |  |  | 08/31/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024353316 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 67-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had an impella 5.5 placed via the axillary for a patient with biventricular failure. the patient condition is very poor and she is being worked up for bilateral lower limb amputation. the upper extremity with the impella showed signs of ischemia after six days of support. the limb is grey and mottled but, no intervention is being done as palliative care is being discussed. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. the root cause of limb ischemia was most likely due to the patients condition as they had pad/pvd. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03478 was provided in sections b5, d6a, d6b, and h1.
34238 | 1220648-2023-03479 | 17842764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842764&pc=OZD | 05/05/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 05/05/2023 | 09/29/2023 |  |  | 03/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324087 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/05/2023 |  |  |  |  | The user facility reported that a 60-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak a tiny leak was noted at the luer connection even after tightening. the ventricular assist device (vad) engineer added short extension tubing which was successful in stopping the leak. the patient had d5w + 12.5 (u/ml) heparin in the purge. there was no reported harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was most likely use related because the leak resolved with the addition of an extension set. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34239 | 1220648-2023-03480 | 17842792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842792&pc=OZD | 08/31/2023 | 09/20/2024 |  | 09/29/2023 |  |  | 08/31/2023 | 09/29/2023 |  |  | 02/28/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311684 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/31/2023 |  |  |  |  | The user facility reported a 52-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. during insertion of the impella 5.5, the left ventricle was perforated. the issue was repaired with sutures. the patient expired the following day. the field service representative confirmed there is no indication that the cause of death had anything to do with the impella. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, wireless insertion was made since 5.5 was accessed with direct surgical artery and during insertion of 5.5 the left ventricle was perforated. no data logs were returned for review. the cause of the ventricular perforation issue was most likely use related per clinical. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34240 | 1220648-2023-03481 | 17842809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842809&pc=OZD | 05/25/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 05/25/2023 | 09/29/2023 |  |  | 12/31/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024300587 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2023 |  |  |  |  | The user facility reported that a 51-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced limb ischemia. the device was inserted into the right axillary without issues. the following day, the patient complained of numbness in right arm and four days later he could only move his hand and had complete numbness from the right wrist to the right shoulder. neurology was consulted. the access site pocket had minimal bleeding in jackson pratt (jp) drain and had no unusual pocket swelling. | [5 empty] | The investigation is completed, with review of clinical details only. product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34241 | 1220648-2023-03482 | 17842849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842849&pc=OZD | 05/16/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 05/16/2023 | 09/29/2023 |  |  | 02/28/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311172 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/16/2023 |  |  |  |  | The user facility reported that a 34-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and suction alarms. device wouldn¿t go above p4 without suction. after running down different causes, the physician determined that the pump inlet had axillary artery intima stuck to it. as a result, the pump was replaced with a new device. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. analysis of the data logs confirmed the low flow when the pump started that triggered the auto suction response and suction alarms. based on clinical description & data review, the root cause of low flow was most likely due to biomaterial ingestion. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34242 | 1220648-2023-03483 | 17842861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842861&pc=OZD | 09/01/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 09/01/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024342256 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 56-year-old female patient underwent high-risk cardiac surgery and received hemodynamic support with an impella 5.5 smartassist heart pump. on day 7 of impella support, the impella 5.5 pump stopped and the impella was removed. the patient suffered no harm from the incident and did not require further impella support. | [5 empty] | H3: the investigation into the reported "pump stop" has been completed. product was returned for investigation. the device was visually inspected and found three open lines during electrical testing. kinks found in catheter and evidence of excessive force found on red handle at catheter kink protector. the ground wire in the red handle on the catheter side is broken. section of visible motor cable lines in red handle are intact. no data logs were returned for investigation. as per clinical investigation, the impella 5.5 had a pump stop during support. there was an impella defective alarm. the pump reconnection did not resolve the issue and the device was removed. the investigation determined that the root cause of the pump stop issue was an open electrical line due to excessive force as seen by the broken ground wire in the red handle.
34243 | 1220648-2023-03484 | 17842913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842913&pc=OZD | 01/31/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 01/31/2023 | 09/29/2023 |  |  | 10/31/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023274547 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2023 |  |  |  |  | The user facility reported 61yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, the patient developed an ectopic heartbeat and it was noted that the tip of the device was against the mitral. attempts were made to reposition the device, however, the impella locking device was said to have failed and repositioning was unsuccessful. the pump was subsequently removed as a result of the noted issue. there was no harm to the patient. | [5 empty] | The investigation for the pump positioning issue has been completed. the returned device was visually inspected and the inner sleeve of the cath anchor was not within the cath anchor and several kinks in the catheter were observed. the cath anchor locking mechanism was functional. for the cath anchor inner sleeve to detach, the cath anchor halves had to be separated due to excessive force. kinks in the catheter are also indicative of excessive force. the cause of the issue was excessive force on the cath anchor use related. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34244 | 1220648-2023-03485 | 17842949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17842949&pc=OZD | 02/06/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 02/06/2023 | 09/29/2023 |  |  | 11/30/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023282849 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/06/2023 |  |  |  |  | The user facility reported that a 58-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed a hole in the catheter. after the device was inserted, blood was observed to be leaking from the red plug. there were no signs of puncture on anterior part of the catheter, and no blood visible in the sterile sleeve. the pump was subsequently removed. there was no reported harm to the patient. | [5 empty] | The investigation into the reported product damage was completed. the device was returned for evaluation. visual inspection of the catheter was performed, and a hole was found at the 35-centimeter mark. based on the available, the root cause of the hole was likely indicative of a needle stick as a result of improper technique. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34245 | 1220648-2023-03486 | 17843020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843020&pc=OZD | 02/20/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 02/20/2023 | 09/29/2023 |  |  | 12/31/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024286905 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/20/2023 |  |  |  |  | The user facility reported 56yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the device could not be moved despite pressure used when pump was being repositioned by the physician. the pump was eventually placed successfully. there was no patient harm reported. | [5 empty] | The investigation for the positioning issue has been completed. the product was returned and the catheter was able to move without issue when the button on the cath anchor was depressed. the sterile sleeve was torn showing signs that excessive force was used. the root cause of the positioning issue was determined to be cath anchor use issue. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34246 | 1220648-2023-03487 | 17843069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843069&pc=OZD | 03/07/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 03/07/2023 | 09/29/2023 |  |  | 12/31/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024286885 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2023 |  |  |  |  | The user facility reported 52yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump has been moving frequently inside the patient's left ventricle, requiring frequent repositions. patient support continued with the pump despite the noted observation. there was no patient harm reported. | [5 empty] | The investigation for the reported device positioning issue has been completed. the device was not returned for evaluation. additionally, clinical information was not sufficient to determine the root cause. therefore, the exact root cause of the reported positioning issue could not be determined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34247 | 1220648-2023-03488 | 17843075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843075&pc=OZD | 09/01/2023 | 01/24/2024 |  | 09/29/2023 |  |  | 09/01/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369971 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient was also on ecmo support. the patient had pupillary change requiring a head ct scan. the scan was suggestive of a very large head bleed. heparin purge was removed. the patient was transferred to another hospital for escalation of care. after 4 days on support, the family withdrew care for the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke has been completed since the original report was filed. no product was returned. data logs were reviewed, and no motor current spike was observed. as per clinical, head ct scan revealed hemorrhagic stroke with large brain bleed and systemic heparin was stopped. the cause of the hemorrhagic stroke issue was most likely patient condition related and unknown relation to impella per clinical and data log review. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34248 | 1220648-2023-03489 | 17843124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843124&pc=OZD | 01/18/2023 | 09/22/2024 |  | 09/29/2023 |  |  | 01/18/2023 | 09/29/2023 |  |  | 06/30/2023 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023246696 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2023 |  |  |  |  | The user facility reported that a 53-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. previously, an impella cp provided support for three days but was replaced with an impella 5.5 after the patient developed thrombus. the 5.5 pump was subsequently removed, and a small bit of clot was noted on the cannula. ultimately, the decision was made to place the same device back in the patient, and flow at whatever p level possible as the alternative to no device. the device was reinserted again, and entrained clot once started. the patient remained on p-6 until transfer to the icu at about 1.6lpm of flow and no drips. there was no reported harm to the patient. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. analysis of the data logs did not reveal any performance issues. based on the available information the root cause of the low pump flow was most likely biomaterial per clinical account. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03489 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34249 | 1220648-2023-03490 | 17843158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843158&pc=OZD | 03/07/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 03/07/2023 | 09/29/2023 |  |  | 11/30/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023278985 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2023 |  |  |  |  | The user facility reported 75 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during the night, the impella 5.5 dislocated into the aorta. while trying to reposition, crossing the aortic valve proved difficult and during the process the sterile sleeve got damaged. pump removed as it was no longer required at that time. no harm done to the patient. | [5 empty] | The investigation for the positioning issue has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause of the reported event. it was determined that the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34250 | 1220648-2023-03491 | 17843262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843262&pc=OZD | 09/01/2023 | 02/26/2024 |  | 09/29/2023 |  |  | 09/01/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024349885 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported a 62-year-old female was implanted with an impella 5.5 as a bridge to left ventricular assist device (lvad). during support, the patient experienced decreased sensation and mobility in the arm where the impella was placed. distal pulses were still present. no interventions have been done but monitoring will continue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
34251 | 1220648-2023-03492 | 17843284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843284&pc=OZD | 12/14/2022 | 03/11/2025 |  | 09/29/2023 |  |  | 12/14/2022 | 09/29/2023 |  |  | 10/31/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023268368 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the device started to ramp up to p-3, and started alarming impella position in aorta. ultrasound transesophageal echocardiography (tee) showed that lv cavity was being compressed by a large known effusion. after reconfiguring venoarterial extracorporeal membrane oxygenation (va ecmo) cannulation, physician elected to remove 5.5 and leave patient on va ecmo only. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34252 | 1220648-2023-03493 | 17843294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17843294&pc=PYX | 03/13/2023 | 01/16/2024 |  | 09/29/2023 |  |  | 03/13/2023 | 09/29/2023 |  |  | 02/28/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023188626 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella rp device for mechanical circulatory support. the complainant had a rp support ongoing along with ecmo for a patient in right heart failure. on day eight of the support there were concerns for sepsis. the patient was treated with antibiotics for the gram- rods. the impella was not the only support device/access site and causality to sepsis is not known. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported infection/sepsis has been completed since the original report was filed. per the provided clinical details, the root cause of the infection/sepsis issue was patient condition related and the relation to impella is unknown. the clinical details suggest, the root cause of the access site bleeding issue was most likely a introducer sheath kink.
34253 | 1220648-2023-03494 | 17844782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17844782&pc=OZD | 09/02/2023 | 08/01/2024 |  | 09/29/2023 |  |  | 09/02/2023 | 09/29/2023 |  |  |  |  | 09/29/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150569 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/02/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while the patient was on support, the staff went to change the bag and cassette. the automated impella controller (aic) would not recognize the new cassette after multiple ttempts. the team was unable to advance through the purge wizard. the field service representative walked the customer through and aic to aic transfer without issue. no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation automated impella controller (aic) - purge issue has been completed. the console was received and tested. reproduced the issue, unable to pass through step 2 of the case start. upon inserting the purge disk, received the ¿purge disc not detected ¿ alarm. upon visual inspection, it was confirmed that the purge flag was missing/broken. the root cause for not detecting the presence of the purge disc is because of the broken disk detect flag. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03494: d2 common device name. d4 model number. f6/f8-should have been left blank.
34254 | 1220648-2023-03495 | 17844824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17844824&pc=OZD | 05/02/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 05/02/2023 | 09/29/2023 |  |  | 02/28/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024311433 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2023 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported and optical signal failure during preparation of the device. a new impella was opened and setup without further issue. there was no adverse impact to the patient. | [5 empty] | Data logs showed that the placement signal spiked immediately upon connecting the pump to the aic; the optical sensor parameters were characteristic of an optical fiber break. no external damage or abnormalities were found on the returned pump. when the pump was connected to the aic, the ps spiked out of range and the placement signal not reliable alarm reproduced. the optical parameters of the returned pump were out of specification. optical testing revealed optical fiber damage within the red handle. in conclusion, it was determined that the cause of the optical signal issue was an out-of-box damaged optical fiber line. this report is being filed as part of a retrospective review of historical records.
34255 | 1220648-2023-03496 | 17844950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17844950&pc=OZD | 04/10/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 04/10/2023 | 09/29/2023 |  |  | 01/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024297530 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported an optical signal failure after moving the pump from one console to another. there was no harm endured and support continues. | [5 empty] | The team returned the data logs but, not the pump for analysis. the imc logs show an eeprom corruption occurred. the cause of the eeprom corruption was not determined since the pump was not returned.
34256 | 1220648-2023-03497 | 17845030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17845030&pc=OZD | 09/15/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 09/15/2023 | 09/29/2023 |  |  | 03/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339968 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the physician was unable to advance impella past sheath, the physician thought that maybe there were some stent struts getting caught. the patient had known peripheral disease with bilateral iliac stenting. the short sheath was exchanged for 25cm sheath in an attempt to get past stents with the sheath. once again pigtail and wire were crossed into left ventricle, but physician was still unable to advance the impella out of sheath. impella placement was aborted. | [5 empty] | The investigation into the reported "delivery issues" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the difficult inserting pump through introducer was unable to be determined as limited clinical details were provided and no product was returned for review.
34257 | 1220648-2023-03498 | 17845994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17845994&pc=OZD | 09/02/2023 | 11/17/2025 |  | 09/29/2023 |  |  | 09/02/2023 | 09/29/2023 |  |  |  |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133146 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/02/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was transferred to another hospital while on the automated impella controller (aic) took place with no issues. a few hours later, the rn reported that the aic was now reading ¿purge disk not detected.¿ team was advised to go back on the original console since it was still available and there were no other aics. inspection of faulty aic revealed a crack in plastic surrounding the area where the purge disk is inserted. there was no patient harm related to this event. | Patient continued to slowly decompensate, and remained do not resuscitate. procedure outcome was that patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03498 was provided in sections b1, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34258 | 1220648-2023-03500 | 17846777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17846777&pc=OZD | 09/02/2023 | 01/17/2024 |  | 09/29/2023 |  |  | 09/02/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366776 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2023 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cannula kinked during introduction of the repo sheath. when the pump was pulled back, straightening out the cannula, flows were reduced and the physician decided to replace the pump with another impella cp. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported cannula kink has been completed. the cannula was returned for evaluation. upon visual inspection, a kink was found. no other issues were found on the rest of the pump. the pump ran without an issue under the bench test. the root cause was due to a kink in the cannula.
34259 | 1220648-2023-03499 | 17847220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17847220&pc=OZD | 06/20/2022 | 10/21/2025 |  | 09/29/2023 |  |  | 06/20/2022 | 09/29/2023 |  |  | 02/29/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197471 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Cardiac Perforation | No Clinical Signs, Symptoms or Conditions | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical staff had difficulty delivering the impella pump. during the delivery the heart manipulation caused newly patched ventriculotomy to bleed and multiple blood products had to be administered to the patient. the patient was sent to the icu for recovery on the support with ecmo and the impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ this report is being filed as part of a retrospective review of historical records.
34260 | 1220648-2023-03501 | 17848252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848252&pc=OZD | 08/27/2023 | 01/26/2024 |  | 09/29/2023 |  |  | 09/05/2023 | 09/29/2023 |  |  | 11/30/2024 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024302724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility |  |  |  |  | 19 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2023 |  |  |  |  | A long-term outcome and quality indicator (loqi) impella registry notification was received which indicated an impella patient had endured hemolysis while on support. symptoms started on (b)(6) 2023 and v-a ecmo was added (b)(6) 2023, as a result. | [5 empty] | The impella device was not returned and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported hemolysis has been completed since the original report was filed. the device was not returned for investigation there are large placement signal spikes on 8/27. however, there are no positioning alarms or suction present. the clinical states that hemolysis was occurring 24 hours into support and then the pump was repositioned. there was nothing else indicative in the logs for the rest of the case that indicates hemolysis resumed. the root cause of the hemolysis is most likely due to improper positioning.
34261 | 1220648-2023-03502 | 17848332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848332&pc=OZD | 08/30/2023 | 01/24/2024 |  | 09/29/2023 |  |  | 08/30/2023 | 09/29/2023 |  |  | 04/30/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2023 |  |  |  |  | The user facility reported a 79-year-old female/male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was inserted and started on auto with flows of 3.7 l/min. doctor then wired right coronary artery (rca), took multiple wires to gain access. percutaneous transluminal angioplasty (pta) to rca, intravascular ultrasound (ivus) to rca, stent placed to distal/mid rca. another stent placed placed on the proximal rca. after, the doctor noticed perforation to mid/distal rca. echo noted a small effusion. two covered stents were placed on the rca. megatron stent placed to ostial rca with good results, effusion did not getting larger. the peel away introducer was removed. repositioning sheath placed. reaccess port used, contrast flushed to observe distal flow. slow filling noted, doctor elected to do a distal fem-fem bypass with good results. forward tension sutures placed and dressing was applied. additionally, there were no distal pulses noted in right lower extremity which was treated with medication. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the vascular damage - great vessel and the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the cause of the vascular damage issue was most likely patient condition related and unknown relation based on clinical review. the cause of the limb ischemia - reversible issue was patient condition related due to fem-fem bypass done to resolve flow down the leg.
34262 | 1220648-2023-03503 | 17848380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848380&pc=OZD | 09/01/2023 | 08/02/2024 |  | 09/29/2023 |  |  | 09/01/2023 | 09/29/2023 |  |  | 03/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342948 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported a 72-year-old female/male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was on day two of support when coronary artery bypass grafting (cabg) was getting started, right after intubation and while vein from leg was being harvested, the automated impella controller (aic) gave alarm ¿purge system open¿. a team member looked at the purge sidearm and saw fluid leaking from the purge reservoir. shortly after, the pump stopped. the pump was able to restart but it kept stopping and restarting until it would not restart. the surgeon was just getting cannulation for bypass. pa pulled impella out of ventricle and bypass pump was used for about two minutes and the patient was stable during that time. after the cabg surgery, the surgeon was not comfortable implanting the impella. the impella was pulled back all the way, then intra-aortic balloon pump (iabp) started. impella explanting options were discussed with surgeon, but it was decided to explant in sicu. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the pump stop and the purge leak has been completed. the pump was returned for evaluation. product evaluation confirmed the membrane of purge sidearm's reservoir was damaged. while attempting to reproduce the pump stop, the pump could not start under bench test. motor housing was then torn down and blood was found on rear ball bearing. data log analysis confirmed the pump started and ran without issue for 43 hours and 54 minutes, then purge pressure suddenly dropped from 500 mmhg to 40 mmhg and was followed by high motor current and a pump stop. the pump was then restarted two times until it could not be restarted. the cause of the pump stop was due to blood ingress into the motor housing. the most likely cause of the purge leak was due to sidearm reservoir damage. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi states the following: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ added- d6a/d6b. f6/f8 should have been left blank in the initial report.
34263 | 1220648-2023-03504 | 17848793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848793&pc=OZD | 09/01/2023 | 01/19/2024 |  | 09/29/2023 |  |  | 09/01/2023 | 09/29/2023 |  |  | 03/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOO PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED INC. |  |  | US |  | 2024349476 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2023 |  |  |  |  | The us medical center reported that with the micro puncture single access attempt, there was a puncture through the impella cp's peel away sheath. the peel away was removed and the repositioning sheath advanced with oozing, but the doctor felt the vessel was intact and continued the procedure. patient was hemodynamically unstable with dropping hematocrit and hemoglobin levels and requiring volume replacement. coronary intervention aborted due to instability and vascular surgery proceeded with a right femoral vessel repair. the cp was removed to gain control of the vessel damage. staff increased the medical/pharmaceutical hemodynamic support and gave blood transfusion. | [5 empty] | The medical center has not yet returned for analysis any impella product, neither pump that failed delivery nor the introducer sheaths. upon completion of the investigation, a final report will be forthcoming. per the ifu: "potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation into the delivery issues- use and the vascular damage - great vessel has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was patient condition related with pump unable to advance over short sheath due to tortuous vessels. the cause of the vascular damage issue was patient condition related with femoral vessel dissected due to tortuous vessel condition during insertion.
34264 | 1220648-2023-03505 | 17848797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848797&pc=OZD | 06/13/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 06/13/2023 | 09/29/2023 |  |  | 04/30/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024337070 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/13/2023 |  |  |  |  | The user facility reported 59yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were high purge pressures and low purge flows. no additional information was provided. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the clinical report was evaluated and found that after the purge bag change the purge pressure gradually climbed and stayed consistently high. sodium bicarbonate was the purge solution. it took about an hour for the purge pressure to rise and trigger the alarm. the purge bag change was routine and took about 2 minutes. there was no evidence of a purge leak before the bag change. during the bag change, there was most likely biomaterial ingested causing the gradual increase in pressure after it was complete. the root cause of the high purge pressure is most likely due to ingested biomaterial. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34265 | 1220648-2023-03506 | 17848818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848818&pc=OZD | 09/03/2023 | 11/10/2023 |  | 09/29/2023 |  |  | 09/03/2023 | 09/29/2023 |  |  | 07/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024372438 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2023 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the pump came out of position and was unable to be delivered. the pump was replaced with another impella 5.5 due to the malposition. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. it was reported that the pump was accidentally pulled out of the lv during repositioning. the root cause of the positioning issues was most likely a cath anchor use issue.
34266 | 1220648-2023-03507 | 17848822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848822&pc=OZD | 09/01/2023 | 01/22/2024 |  | 09/29/2023 |  |  | 09/01/2023 | 09/29/2023 |  |  |  |  | 09/29/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED INC. |  |  | US |  | 1467028 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/01/2023 |  |  |  |  | The us complainant was setting up a new impella cp support. the team observed a purge disc alarm on the automated impella controller (aic) and trouble shot by aic and pump replacement. the team then successfully started support on new pump and aic. there was no noted harm or injury to the patient for any delay in the support initiation. the impella cp supported for 4 days, was weaned, and explanted. | [5 empty] | The medical center has returned for benchtop analysis the impella controller. the investigation is on-going at this time. upon the completion of the investigation, a final report will be forthcoming. the purge disc is noted in the ifu as follows: "transmits pressure to the controller based on the purge pressure in the purge tubing; a sensor in the controller measures the pressure so that it can be displayed on the screen and used by the purge pressure algorithm to maintain the purge pressure." | The investigation for the reported console disc/flag issues has been completed. the device and data logs were returned for evaluation. the data logs were reviewed which confirmed that the alarms for purge disc not detected and purge system open were issued on the reported occurrence date. the data logs also confirmed an instance where purge pressure was high, but it was not persistent enough to trigger a purge pressure high alarm. the console was inspected and tested after arriving. issues reported were not able to be reproduced under normal operating conditions with know-good pump and purge unit. the cause of the issue is suspected to be use related.
34267 | 1220648-2023-03508 | 17848832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848832&pc=OZD | 09/02/2023 | 05/07/2024 |  | 09/29/2023 |  |  | 09/02/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024366168 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Optical Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an optical sensor system error alarm occurred when the connection cable was connected for an impella cp support. it was confirmed that the cable was firmly stuck in the back, and it was also confirmed that the cable was removed and the light was visible. the controller was restarted several times, but no improvement was seen. the staff was unable to replace the automated impella controller (aic). the pump catheter was replaced and the alarm resolved. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the optical signal issue has been completed. the pump was returned for evaluation and during visual inspection, there were no kinks, no holes in the catheter, and no blood in red handle observed. the red lumen was still in the pump. "placement signal not reliable" alarm was observed as soon as the pump was connected to the aic. the pump passed light continuity test and electrical testing was within the limits. the 5s parameters were normal apart from contrast being a little low. the optical ferrule on the patient cable was cleaned and the pump was connected, and "placement signal not reliable" was observed again. the aic was returned and investigated per 23-18213-2, which revealed a defective optical bench with low snr. the defective bench caused the g0 to be improperly calibrated, resulting in the placement signal not reliable alarm upon pump connection until the pump was recalibrated. there was no problem found with the pump. see investigation 23-18213-2 for more details on the aic investigation. in accordance with updated procedures, section d6 has been revised from the initial submission.
34268 | 1220648-2023-03509 | 17848846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848846&pc=OZD | 09/02/2023 | 04/11/2024 |  | 09/29/2023 |  |  | 09/02/2023 | 09/29/2023 |  |  | 04/30/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348395 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the insertion of the impella cp, the pigtail became entangled in what was assumed to be papillary structures. the physician pulled back on device to free pigtail and then tried to advance catheter again, but the cannula appeared to become kinked on fluoro. the physician attempted to initiate therapy, but a pump stop alarm appeared despite the device initially appearing to ramp up. the physician removed the device without incident and successfully replaced it with a new cp. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the pump "did not start" issue has been completed. data logs were not returned for review. the impella was returned for evaluation and during fuctional testing, the pump could start and run without issue under bench testing. the flow was saturated during bench testing due to dry blood in the purge gap. the kinked cannula was found about 15 mm from bonding site of can & of cage. the pump was caught on the papillary muscle and then attempted to be re-advanced and the cannula became kinked. the root cause of the positioning issues is most likely related to use. the root cause of the pump did not start was most likely related to use as there was a positioning issue and then it was re-advanced, the cannula kinked and then it was attempted to be started. added- h6 medical device problem code 3009, h6 impact code 4628
34269 | 1220648-2023-03510 | 17848850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848850&pc=OZD | 06/19/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 06/19/2023 | 09/29/2023 |  |  | 04/30/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024336855 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/19/2023 |  |  |  |  | The user facility reported 19yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were high purge pressure alarms on the pump. the pump was running sodium bicarbonate in the purge. the purge pressure was steadily rising and the purge flow was falling since the insertion. there were no kinks or occlusions found. the purge solution was changed to tissue plasminogen activator (tpa) to troubleshoot and remedy the pressure alarms. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure alarms has been completed. clinical report was reviewed finding that a purge bag change is completed, for about 4 minutes. at the end of the bag change the purge pressure immediately increases and triggers the high purge pressure alarm but it relieves itself. additionally, at the time of the purge pressure spike, there is no motor speed change or motor current spike. this resembles a kinked purge line that most likely happened right at the end of the bag change. the root cause of the high purge pressure is mostly due to a kinked purge line. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34270 | 1220648-2023-03511 | 17848852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848852&pc=OZD | 07/08/2023 | 09/29/2023 |  | 09/29/2023 |  |  | 07/08/2023 | 09/29/2023 |  |  | 03/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024327491 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/08/2023 |  |  |  |  | The user facility reported 66yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge system blocked. the customer ran tissue plasminogen activator (tpa) through the purge. pump shows signs of flow saturation which has gotten progressively worse since. pump was explanted secondary to pump failure as a result of increased purge pressure. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure issues has been completed. pump was not returned for investigation. data log shows a rise in purge pressure and purge-related alarm after the three minute long purge cassette/bag change procedure. the cause of the high purge pressure issue was use related since the purge cassette/bag change procedure change took more than 2 minutes based on data log review. the recommended purge cassette/bag change time is 90 seconds. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34271 | 1220648-2023-03512 | 17848870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17848870&pc=OZD | 09/03/2023 | 01/25/2024 |  | 09/29/2023 |  |  | 09/03/2023 | 09/29/2023 |  |  | 05/31/2025 |  | 09/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356526 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the doctor attempted to treat the coronary artery disease, but the patient became uncooperative and bleeding started at peel away sheath site of the impella cp pump. there was significant blood loss and the team treated by one unit of blood given in in the cath lab, 250ml of albumin, and 1 additional unit prbc and plasma given by flight crew prior to transfer out to the tertiary medical center. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. based the clinical information provided, the root cause of the access site bleeding issue was most likely patient movement since the bleeding started after performing balloon angioplasty when the patient became uncooperative.
34272 | 1220648-2023-03442 | 17831298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831298&pc=OZD | 08/28/2023 | 02/07/2024 |  | 09/28/2023 |  |  | 08/28/2023 | 09/28/2023 |  |  | 05/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella was placed and the peel away sheath was removed. hemostasis was attempted but pre-close failed, staff elected to place 8fr angioseal, but again failed to achieve adequate closure. the left femoral artery (lfa) was then accessed and uf was place up and over for dsa imaging which revealed perforation, staff gave protamine and proceeded with balloon tamponade to the vessel x3 with 20 minutes intervals. vascular surgery was ultimately consulted for surgical cut down and closure and hemostasis was achieved. | [5 empty] | The investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return the impella device. as per clinical review, the impella was weaned, explanted and peel away sheath was removed. hemostasis was attempted using perclose and angioseal but both had failed attempts and imaging revealed perforation. the cause of the vascular damage issue was most likely use related due to failed attempts of perclose and angioseal device per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34273 | 1220648-2023-03437 | 17831300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831300&pc=OZD | 08/12/2023 | 09/27/2023 |  | 09/28/2023 |  |  | 08/12/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357382 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 34-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient was noted to be net negative five liters post diuresis. the patient's labs hemolyzed sever times. several labs were obtained, including ldh and pfhgb. results are not known, but the patient underwent impella explant. | [5 empty] | (h3) the investigation into the reported limb ischemia and vascular damage has completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that there were over 8,000 occurrences of the "impella position unknown" alarms observed in the data logs. there were no other abnormalities observed in the logs. the patient was noted to be negative 5l post diuresis the day before. the patient's labs showed that the patient hemolyzed several times. in those labs that were taken are ldh and pfhgb. the results were not known but the impella was explanted the following day. the root cause of the suspected hemolysis is most likely due to improper positioning. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34274 | 1220648-2023-03447 | 17831339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831339&pc=OZD | 08/30/2023 | 01/25/2024 |  | 09/28/2023 |  |  | 08/30/2023 | 09/28/2023 |  |  |  |  | 09/28/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1607654 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | There was no report of patient harm. no additional information was provided. | The user facility reported that the purge cassette clip was ripped off during the device explant. |  |  |  |  | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and inspected. the reported issue with the broken purge retainer was confirmed. the root cause of this issue was a broken purge pressure sensor retainer. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
34275 | 1220648-2023-03438 | 17831341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831341&pc=OZD | 08/28/2023 | 01/23/2024 |  | 09/28/2023 |  |  | 08/28/2023 | 09/28/2023 |  |  | 04/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353622 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2023 |  |  |  |  | The user facility reported an 88-year-old male in anterior myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the team made an observation of hemolysis by hemolyzed labs and urine color change. no noted intervention was done, no red blood cells nor dialysis initiated for renal impairment. the pump remains on for support. | [5 empty] | The investigation into the hemolysis has been completed. the device was not returned for benchtop evaluation and investigation. data logs showed about 13 hours into the case corresponding with the time hemolysis occurred in clinical description, the pump position was in the ventricle for three minutes that triggered alarm impella position in ventricle then in aorta for one minute. about 16 hours & 45 minutes, the pump was in the aorta position again for five minutes that triggered alarm impella position in aorta. about 4 hours & 45 minutes later, the pump was stopped manually and disconnected from the console. based on the clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine. hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure. patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34276 | 1220648-2023-03439 | 17831342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831342&pc=OZD | 08/28/2023 | 01/24/2024 |  | 09/28/2023 |  |  | 08/28/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359741 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2023 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, the pump suddenly stopped. the impella 5.5 was replaced with another successfully. there was no patient harm. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, it was reported that the data logs show that alarm #115 was triggered indicating an electrical issue. placement signal goes out at the same time as the pump stop meaning that the optical fiber broke. the root cause of the pump stop was most likely an open electrical line due to excessive force. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34277 | 1220648-2023-03440 | 17831358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831358&pc=OZD | 08/28/2023 | 09/20/2024 |  | 09/28/2023 |  |  | 08/28/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371345 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Ischemia | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/28/2023 |  |  |  |  | Additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while the impella was crossing the aortic valve (av), placement appeared to be under the pap and repositioning was attempted. the impella came back out across the av. attempts to re-cross were unsuccessful. the device was removed from the sheath, rinsed, and re-loaded over the .018 wire with no issues. two days later on (b)(6) 2023, patient limb ischemia was noted. no treatment or blood products were provided for limb ischemia. the patient remained on impella cp support until (b)(6) 2023. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the positioning issues and the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the positioning issues is due to use as the impella was caught under the papillary muscle after it crossed the av. the root cause of the limb ischemia is most likely due to use as they left the large od sheath in instead of replacing it with the smaller od repo sheath. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34278 | 1220648-2023-03446 | 17831362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831362&pc=OZD | 08/30/2023 | 01/24/2024 |  | 09/28/2023 |  |  | 08/30/2023 | 09/28/2023 |  |  | 02/28/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324693 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2023 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the staff unable to get doppler pulse right popliteal, post tib or dorsalis pedis. per cath lab staff report, the patient had worsening bruising and mottling of his right flank extending across the anterior of his abdomen. concern with continued bleeding from the impella access site, so the impella was explanted. | [5 empty] | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed. the device was not returned for benchtop investigation. as per clinical, patient had bleeding at access site with worsening bruising and mottling of right flank extending across the anterior of abdomen. unable to get doppler pulse on right popliteal with patient having pad and calcified vessel condition. the cause of the access site bleeding issue was most likely patient condition related with patient having diseased vessel condition leading to bleeding. the cause of the limb ischemia issue was patient condition related with patient having peripheral artery disease and calcified vessel condition per clinical. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34279 | 1220648-2023-03448 | 17831365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831365&pc=OZD | 08/30/2023 | 01/25/2024 |  | 09/28/2023 |  |  | 08/30/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364880 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient was implanted with the impella 5.5 and after closing the cutdown pocket impella measured slightly deep and was retracted to show inlet 4.98 cm below atrioventricular (av) on transesophageal echocardiography (tee). additionally, the patient had access site bleeding that lead to multiple blood products being given to the patient. | [5 empty] | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. based on the clinical information, the root cause of the access site bleeding issue was most likely patient condition related since bleeding was reported from multiple site and calcified vessels was reported, based on provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34280 | 1220648-2023-03441 | 17831435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831435&pc=OZD | 08/28/2023 | 01/12/2024 |  | 09/28/2023 |  |  | 08/28/2023 | 09/28/2023 |  |  | 04/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353570 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump had low flows which the team assumed was due to improper flushing. when they observed the pump at the time of explant there as a large clot burden within the inlet. a second pump was placed which also failed and was replaced by an iabp. | [5 empty] | The investigation into the low pump flow, positioning issues and thrombus has been completed. the product was not returned. as per the clinical information provided, the motor current spikes right at the start of this case and then the flows are low while suction is present. additionally, the pump was not in a good position as it was accidentally pulled into the aorta and not able to cross the av again. a clot was visualized in the inflow cage. the sheath was not aspirated upon insertion and likely caused the clot ingestion. the root cause of the low flow is most likely due to biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34281 | 1220648-2023-03443 | 17831438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831438&pc=OZD | 08/28/2023 | 01/24/2024 |  | 09/28/2023 |  |  | 08/29/2023 | 09/28/2023 |  |  | 01/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024297554 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Injury | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 68-year-old male patient has been treated for acute myocardial infarction / cardiogenic shock and received hemodynamic support from an impella 5.5 smart assist cardiac pump. it has been reported that one day after impella insertion, high purge pressure has been observed. the user suspected that a thrombus caused the issue and independently performed a lysis therapy through the purge sidearm of the impella 5.5 pump. the purge fluid has been changed from 5% glucose solution with sodium bicarbonate to 5% glucose solution with heparin. no harm has been done to the patient. the pump has not been removed and to our current knowledge is still supporting the patient. | [5 empty] | The investigation into the thrombus issue has been completed. the device was not returned for benchtop investigation. as per the given clinical information, the root cause of the thrombus issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34282 | 1220648-2023-03444 | 17831445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831445&pc=OZD | 08/29/2023 | 09/26/2025 |  | 09/28/2023 |  |  | 08/29/2023 | 09/28/2023 |  |  | 05/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024345494 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pressure Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/29/2023 |  |  |  |  | Optical signal issues were uncovered during investigation. | The user facility reported a 38-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient had significant swelling in right upper chest. previous jp drain was pulled due to low output and pressure dressing applied. the patient was taken for ct which showed moderate amount of blood pocketed in tissue. the decision was made to bring the patient back to the operation room for a washout of the cutdown. approximately 5 day later, saturated flows occurred which prompted explant of the device. a clot was removed from the surgical graft upon removal of the device and a new impella was inserted without further complications. |  |  |  |  | The investigation into the reported access site bleeding and saturated flow has been completed since the original report was filed. impella 5.5 s2 pump was returned with pump cut open at the motor housing end. visually inspected and biomaterial ingestion was found to be present in the outflow cage and impeller of the pump. the cath anchor blue button repositioning unit was visually inspected, and no damages were observed. no other abnormalities were found. the data logs were reviewed and there was motor current spike observed after several suction events. also, several false alarm of impella in aorta alarm was found after motor current spike followed with saturated flow. purge pressure rise was observed in span of 5 minutes post the biomaterial ingestion. device was explanted due to pump being saturated. the cause of the access site bleeding issue was not determined since sufficient clinical information was not provided. the cause of the high purge pressure issue was due to biomaterial ingestion in the purge gap with purge pressure rise in 5 minutes after the motor current spike along and saturated flow. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was biomaterial with false alarm of 'impella in aorta' observed. all pertinent information available to abiomed has been submitted at this time. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34283 | 1220648-2023-03445 | 17831457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17831457&pc=OZD | 08/29/2023 | 01/25/2024 |  | 09/28/2023 |  |  | 08/29/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371810 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/29/2023 |  |  |  |  | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry: the originally reported access site hematoma required multiple units of packed red blood cell transfusion due to a decrease in hgb levels. additionally, the patient exhibited thrombocytopenia that prompted stoppage of heparin. | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. after implantation, the sheath was peeled away. while prepping for icu, flows and pulsatility were reduced on impella and ecmo. hematoma was noted on the left side. it was unclear if it was associated with the left femoral artery or left femoral vein. impella was removed and 16 fr gore dry seal placed. the impella was run under heparinized saline and placed in surgical mode. the valve re-crossed and another impella cp was placed. flows have since returned to normal. |  |  |  |  | B(5): additional information received via abiomed study. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. data logs confirmed the failure mode & clinical narrative. logs showed after pump started, mc spiked followed low flow that triggered auto suction response & suction alarms. pump then was run on p-2 (surgical mode) for about 47 hours. after that, pump was run on p-4 with flow in specification. based on clinical description & data review, the root cause of low flow was most likely due to biomaterial ingestion. the root cause of access site bleeding was most likely a lack of vessel recoil onto the sheath. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
34284 | 1220648-2023-03449 | 17833221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17833221&pc=OZD | 11/10/2022 | 09/28/2023 |  | 09/28/2023 |  |  | 11/10/2022 | 09/28/2023 |  |  | 08/31/2024 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248906 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported that a 57-year-old female patient implanted with the impella 5.5 for mechanical circulatory support had a cannula kink in the patient's body near the outlet. the pump was exchanged as a result. it was noted that there were delivery issues prior to the observation of the kink and "that the patient was large". it was also noted that excessive force was used and there the patient had tortuous vessels. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported kink was completed. the device was not returned for evaluation. an image provided by the clinical representative confirms the cannula and the catheter kinks. the kinks likely occurred during delivery of the pump and were due to the reported tortuosity of the vessels. therefore, the root cause of the delivery issue was patient condition related and related to the patient's anatomy.
34285 | 1220648-2023-03450 | 17833284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17833284&pc=OZD | 02/19/2023 | 09/28/2023 |  | 09/28/2023 |  |  | 02/19/2023 | 09/28/2023 |  |  | 01/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024296585 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/19/2023 |  |  |  |  | The user facility reported 38yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient had right arm pain and numbness. they also had ventricular tachycardia that was treated with amiodarone infusion. the patient was eventually weaned off all inotropes with plans to withdraw care. | [5 empty] | The investigation for the reported limb ischemia and ventricular tachycardia has been completed. no product was returned. the cause of the limb ischemia issue was most likely due to patient condition related with high dosage of epinephrine or norepinephrine. the cause of the ventricular tachycardia issue was due to patient condition related with heart wall damage pre-existing condition through indication of thrombosis observed on left ventricle. the instructions for use states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34286 | 1220648-2023-03451 | 17833336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17833336&pc=OZD | 08/29/2023 | 01/24/2024 |  | 09/28/2023 |  |  | 08/29/2023 | 09/28/2023 |  |  | 05/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360348 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had an access site bleed. the patient developed the hematoma after the reposheath was placed and sutured. the team had to deliver epinephrine (epi) and promtamine. the pump was weaned and explanted with the team placing a femostop. | [5 empty] | The investigation into the access site bleeding issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to a lack of vessel recoil. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34287 | 1220648-2023-03452 | 17833483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17833483&pc=OZD | 08/29/2023 | 01/23/2024 |  | 09/28/2023 |  |  | 08/29/2023 | 09/28/2023 |  |  | 04/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2023 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the impella cp was inserted along with ecmo for a patient and the team placed antegrade catheters for the limb to be perfused. the pump remained on for support with eccymosis and edema in the limb. the team made choice to explant and escalate to the 5.5 pump via the axillary. the cp support was for 19.3 hours. axillary impella 5.5 placed without difficulty. | [5 empty] | The investigation has been completed. the device was not returned for benchtop evaluation and investigation. the root cause of the access site bleeding issue was most likely due to patient movement, as it was reported that the bleed occurred after transfer to the intensive care unit. the root cause of the limb ischemia issue was most likely patient condition related to calcified vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34288 | 1220648-2023-03455 | 17834343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17834343&pc=OZD | 08/31/2023 | 10/14/2024 |  | 09/28/2023 |  |  | 08/31/2023 | 09/28/2023 |  |  | 05/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024348702 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/31/2023 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed sepsis. the sepsis and relationship to the impella could not be proven or determined by the japanese team as there were multiple catheters inserted during the hospitalization. later the patient expired from unrelated to the use of impella causes. the pump was known to support malfunction free, and as such the death will be unrelated. | The patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the patient's sepsis has been completed. product was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the infection/sepsis was not determined. added h6 component code f6 and f8 should have been blank. | Corrections to the initial mfr report #1220648-2023-03455: b2-selected death and added date of patient death b5-it was initially reported that the patient expired and the death was unrelated to use. however, medical safety officer reviewed the patient death. statement provided in b5. added- d6a/d6b. h1-type of reportable event changed to death. h6 impact code 1802.
34289 | 1220648-2023-03454 | 17834528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17834528&pc=OZD | 08/30/2023 | 11/06/2023 |  | 09/28/2023 |  |  | 08/30/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024367969 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 76-year-old patient received extracorporeal membrane oxygenation (va-ecmo) support during high-risk cardiac surgery. in order to wean the patient from ecmo more quickly, hemodynamic support should also be provided by an impella 5.5 smartassist heart pump. when attempting to insert impella 5.5 via the right subclavian artery, a structural change was observed fluoroscopically. the vascular surgeon called in diagnosed a dissection of the right subclavian artery. after surgical reconstruction of the injured artery, an attempt was made to insert the impella 5.5 heart pump via the left subclavian artery. the pump was able to be inserted into the left ventricle, but dissection of the artery was also noted. the impella heart pump was removed and the dissection was repaired using a vascular prosthesis. the patient was transferred to the intensive care unit without impella support. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the vascular damage has been completed since the original report was filed. the device was returned for investigation. as per the clinical information provided, the root cause of the vascular damage issue was patient condition related to small/ diseased vessels.
34290 | 1220648-2023-03453 | 17834545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17834545&pc=PYX | 08/31/2023 | 01/16/2024 |  | 09/28/2023 |  |  | 08/31/2023 | 09/28/2023 |  |  | 01/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024322355 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2023 |  |  |  |  | The user facility reported a 51-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device and an impella 5.5 device for mechanical circulatory support. the doctor attempted to wean the patient from the impella rp. doctor tried to restart the rp and a "motor current high" alarm occurred and the "impella stopped" alarm sounded. staff tried to swap to a different console and rp wouldn't start. the va ecmo was decannulated the day prior. patient was brought to cath lab and the rp was explanted. patient is stable. doctor noted a visible clot in inflow cage of rp when explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the thrombus and the pump stop has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the thrombus is most likely due to the patient¿s condition while the root cause of the pump stop cannot be determined.
34291 | 1220648-2023-03457 | 17834702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17834702&pc=OZD | 08/31/2023 | 08/15/2025 |  | 09/28/2023 |  |  | 08/31/2023 | 09/28/2023 |  |  |  |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124694 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported biomed department filed a complaint for a failed self check. there was no patient use or harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | A.1 added information in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03457 was submitted. a.2 age should have been left blank in accordance with updated procedures on the initial manufacturer device report number 1220648-2023-03457 as no patient was involved. a.3 na should have been selected on the initial manufacturer device report number 1220648-2023-03457 as no patient was involved. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03457 was submitted. d.4 model number, catalog number and serial number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03457 was submitted. e.2 revised selection in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03457 was submitted. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-03457 in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03457 was submitted. h.5 and h.8 revised selection as it was reported incorrectly on the initial manufacturer device report number 1220648-2023-03457. h.6 codes 2199 and 3221 were reported incorrectly on the initial manufacturer device report number in accordance with updated procedures. a new code was added.
34292 | 1220648-2023-03456 | 17835452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17835452&pc=OZD | 03/08/2022 | 09/19/2024 |  | 09/28/2023 |  |  | 03/08/2022 | 09/28/2023 |  |  | 12/31/2023 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023171924 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Other | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced ventricular tachycardia (vt) and four days after that event the patient experienced cardiogenic shock, shock liver, kidney failure, while trying to support left ventricle. eventually care was withdrawn per family request. | Additional information note the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34293 | 1220648-2023-03458 | 17838097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17838097&pc=OZD | 08/31/2023 | 07/08/2024 |  | 09/28/2023 |  |  | 08/31/2023 | 09/28/2023 |  |  | 05/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352073 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2023 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced groin bleeding at impella site despite angle matching and manual pressure. ct of pelvis ruled out retroperitoneal bleeding. bleeding resolved with holding of systemic heparin with bruising present in the groin. one unit of blood given. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values, heparin was stopped to achieve hemostasis as per the clinical description. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations section: entering cardiac output section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b. f6/f8 should have been left blank.
34294 | 1220648-2023-03459 | 17838224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17838224&pc=OZD | 08/31/2023 | 01/16/2024 |  | 09/28/2023 |  |  | 08/31/2023 | 09/28/2023 |  |  | 03/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024334344 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 42-year-old male patient was on impella support for 1,611 hours. the staff stated there was access site bleeding and the patient expired on support. patient went into hemorrhagic shock. the patient expired on impella support. on (b)(6) 2023, abiomed was made aware that there was no relationship between the impella and the cause of death. the doctor stated there was a mistake in puncturing the access site and bleeding occurred. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that there was an error of puncturing the access site which caused the bleeding to occur. therefore, the root cause of the access site bleed was most likely due to use as improper technique was used to puncture the access site.
34295 | 1220648-2023-03460 | 17838396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17838396&pc=OZD | 08/31/2023 | 11/03/2023 |  | 09/28/2023 |  |  | 08/31/2023 | 09/28/2023 |  |  | 04/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333471 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had a 68-year-old female patient supported by the impella 5.5 pump. the patient experienced a cva during the support. this was observed on second day of the support. patient had ct imaging done. the patient remained on support for 5 days when care withdrawn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs showed there was no motor current spike or other relevant abnormal pump metrics on the day of the reported cva. therefore, it was determined that the cause of the stroke/cva was most likely patient condition and its relation to the impella is unknown.
34296 | 1220648-2023-03462 | 17838889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17838889&pc=OZD | 09/01/2023 | 02/08/2024 |  | 09/28/2023 |  |  | 09/01/2023 | 09/28/2023 |  |  | 10/31/2024 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292334 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had an impella cp placed for support of a 72-year-old female patient with angina when after just fourteen hours the team noted a pump stop that was not resolved. the pump stop was remedied with pump explant and placement of an iabp instead. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. the returned pump was visually inspected and no abnormalities were observed. the pump was run in a basin of water and the high motor current pump stop was reproduced. electrical resistances were measured and found to be within specification. the catheter was then trimmed nearest to the motor and no blood was observed in the purge line. the returned purge cassette was attempted to be purge but was unable to complete de-air. kinks in the purge line tubing were observed. the tubing line was then cut to remove the kinks and the purge cassette was able to complete de-air. the pump was manually rotated and no interference with the outflow cage struts were observed. the pump was then torn down and denatured proteins and evidence of biomaterial were observed in the motor. the cause of the issue was biomaterial.
34297 | 1220648-2023-03461 | 17838919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17838919&pc=OZD | 09/01/2023 | 11/10/2023 |  | 09/28/2023 |  |  | 09/01/2023 | 09/28/2023 |  |  | 07/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376260 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the doctor noticed clots in graft before inserting abiomed peel-away introducer. the physician removed all the clots and act is therapeutic.shortly after insertion automated impella controller (aic) began to alarm ¿purge pressure high.¿ even after switching consoles the alarm continued. the staff placed new impella 5.5. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressure has been completed since the original report was filed. the device was returned for investigation. upon investigation, biomaterial was found in the purge gap. data logs were also returned and investigated, and it shows purge pressure rising since the initiation of the pump. also, another data log showed purge pressure at 1000mmhg. the root cause of the high purge pressure issue was most likely biomaterial found in the purge gap.
34298 | 1220648-2023-03463 | 17839190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839190&pc=OZD | 09/09/2022 | 09/28/2023 |  | 09/28/2023 |  |  | 09/09/2022 | 09/28/2023 |  |  | 06/30/2024 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237497 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 36-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported optical signal failing after cross clamp was placed for surgery. the pump was removed/replaced with a new impella, resolving the issue. there was no harm to the patient. | [5 empty] | No product was returned. the pump's data log confirms a placement signal not reliable alarm. the trend in parameters are indicative of a kinked fiber line. the cause of the optical signal issue was most likely a damaged optical fiber line. this report is being filed as part of a retrospective review of historical records.
34299 | 1220648-2023-03465 | 17839406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839406&pc=OZD | 09/01/2023 | 09/28/2023 |  | 09/28/2023 |  |  | 09/01/2023 | 09/28/2023 |  |  | 05/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368284 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported a 56-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician had problems getting the impella cp length to accommodate the very tall patient. the sterile sleeve was adjusted and the locking mechanism was damaged. the physician was able to get the device into position. there was no patient harm. | [5 empty] | The investigation into the product damage has been completed. the user facility did not return the product for evaluation. the clinical review determined the cause of the positioning issue, specifically the tuohy borst not tightening was not determined since product was not returned for investigation. the failure modes will be monitored and trended.
34300 | 1220648-2023-03467 | 17839472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839472&pc=OZD | 05/01/2023 | 09/29/2025 |  | 09/28/2023 |  |  | 05/01/2023 | 09/28/2023 |  |  | 12/31/2024 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024285912 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Optical Problem | Activation Failure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an optical sensor alarm that occurred during setup of the device. a new impella pump was opened and used without further issue. there was no adverse impact to patient support. | [5 empty] | The investigation into the reported optical signal issue has been completed . product was returned for investigation. the device was visually inspected and found no abnormalities were observed. the optical light test was performed. the catheter was connected to first aic and optical sensor error was displayed on the screen. in reviewing the data logs, after going through trouble shooting procedures with first aic a second aic was brought to the or and connected to the catheter. the same optical sensor error screen was displayed and the same trouble shooting steps were followed before deciding to try a new catheter. the second catheter worked without any issue. the investigation determined that the root cause of the issue was a damaged optical fiber line. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
34301 | 1220648-2023-03468 | 17839488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839488&pc=OZD | 08/29/2023 | 01/23/2024 |  | 09/28/2023 |  |  | 08/30/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373017 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was successfully weaned and explanted on (b)(6) 2023. the physician pulled the impella pump at icu bedside, but did not receive any bleed back once pump was explanted. additionally, the physician could not find a distal pulse. a vascular surgeon was consulted, and the patient was brought to or for surgical closure. a fem patch was performed due to a dissected common femoral artery (cfa). the surgery was successful, with palpable distal pulses of the right lower extremity. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia-reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the limb ischemia-reversible is most likely plaque buildup patient condition related.
34302 | 1220648-2023-03469 | 17839532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839532&pc=OZD | 08/30/2023 | 01/11/2024 |  | 09/28/2023 |  |  | 08/30/2023 | 09/28/2023 |  |  | 02/28/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325251 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had ongoing suction overnight. echo revealed aortic annulus clot. the device was explanted and clot was visualized in the inflow tract. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. according to the clinical details, the patient had ongoing suction and an echo revealed an aortic annulus clot. upon explant, clot was observed in the inflow tract of the impella device. additionally, the data logs reviewed showed suction alarms were observed with small fluctuation in flow. the cause of the low pump flow issue was due to biomaterial observed on aortic annulus and on inflow cage of pump after explant per clinical and suction alarms observed per data log review.
34303 | 1220648-2023-03470 | 17839552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839552&pc=OZD | 08/31/2023 | 01/24/2024 |  | 09/28/2023 |  |  | 08/31/2023 | 09/28/2023 |  |  | 06/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374209 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2023 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed via a femoral artery access. the team made an observation made of limb ischemia. the team treated the issue with tissue plasminogen activator infusion. the pulseless leg also had a distal perfusion catheter placed and the team left impella support on for a total of 176 hours until wean and explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the given clinical information, the root cause of the limb ischemia issue was patient condition related to calcified vessels.
34304 | 1220648-2023-03471 | 17839563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839563&pc=OZD | 08/27/2023 | 08/09/2024 |  | 09/28/2023 |  |  | 09/01/2023 | 09/28/2023 |  |  | 04/30/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024346117 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported a 86-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had impella access site bleeding and manual pressure was held however, the bleeding could not be stopped. the bleeding occurred after replacing the repositioning sheath and the sheath was secured with purse-string sutures. eventually, the patient expired after six hours of impella support. cause of death was determined to be bleeding from the impella access site. reportable due to access site bleeding that led to the patient expiring. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.2 date of death and b.3 date of event revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2023-03471. d.6a and d.6b revised from the initial submission of initial manufacturer device report number 1220648-2023-03471 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-03471 and should not have been. e.1 first name revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2023-03471. h.6 added code 925 since submission of the initial manufacturer device report number 1220648-2023-03471 in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2023-03471. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-03471 in accordance with updated procedures. additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2023-03471 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
34305 | 1220648-2023-03472 | 17839572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839572&pc=OZD | 09/01/2023 | 09/28/2023 |  | 09/28/2023 |  |  | 09/01/2023 | 09/28/2023 |  |  |  |  | 09/28/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1676151 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 09/01/2023 |  |  |  |  | The user facility reported their automated impella controller (aic) started up without the self-test and ended with a black window. a backup aic was used by the user facility. there was no patient involvement. | [5 empty] | The investigation into the malfunction of the user facility's automated impella controller (aic) has been completed. the device was returned and evaluated. the cause of the issue was a corrupt compact flash card. the failure modes will be monitored and trended.
34306 | 1220648-2023-03464 | 17839735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839735&pc=OZD | 07/07/2022 | 12/04/2025 |  | 09/28/2023 |  |  | 07/07/2022 | 09/28/2023 |  |  |  |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123949 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient who was just transitioned to impella 5.5 support, during transfer from the surgical table to the bed, had the white power cable became unplugged and pump stopped. initial attempt to re-plug and continue support was unsuccessful and we received an error message stating that the impella was defective and needed to be replaced, then the screen went black. upon restart of the automated impella controller (aic) it acted normal and accepted the pump and resumed support. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34307 | 1220648-2023-03466 | 17839763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17839763&pc=OZD | 08/29/2023 | 01/31/2024 |  | 09/28/2023 |  |  | 08/29/2023 | 09/28/2023 |  |  | 03/31/2025 |  | 09/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344777 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2023 |  |  |  |  | The user facility reported a 71-year-old male (race unknown) with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had a hematoma on the ecmo access site and bleeding at the impella site. the staff halted heparin. the patient was on impella support for 298.95 hours before the family withdrew care and the patient expired. no relationship between the cause of death and the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. the cause of the access site bleeding issue was most likely patient condition related since the bleeding was reported at multiple access site, based on given clinical details.
34308 | 1220648-2023-03424 | 17824460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17824460&pc=OZD | 08/21/2023 | 03/04/2024 |  | 09/27/2023 |  |  | 08/21/2023 | 09/27/2023 |  |  | 04/30/2025 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353454 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported an 83-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had poor limb perfusion after 72 hours on impella support. overtime, the staff noticed the right leg swelling and poor distal flows. the poor distal perfusion resolved with a fasciotomy. it was also noted that patient did have history of distal peripheral vascular disease, but right femoral was clear of any calcium which is why the cp had no difficulty with insertion. with chemical support weaned and turned off, the doctor decided to wean and remove the impella, but the patient did not tolerate well after the removal and an impella 5.5 was implanted for continued support. | [5 empty] | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34309 | 1220648-2023-03425 | 17824529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17824529&pc=OZD | 04/17/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 04/17/2023 | 09/27/2023 |  |  | 10/31/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023276867 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/17/2023 |  |  |  |  | The user facility reported that a 53-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak at the connection site. sodium bicarb was used in purge solution. also, while changing out the cassette, it was noted that there was leaking at the yellow check valve and a crack was visible. a purge bypass was completed with a19ga needle without issue. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. pump was returned.the yellow luer on the sidearm was confirmed to have a leak. a syringe was attached to the sidearm and the fluid leaked at the yellow luer. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34310 | 1220648-2023-03426 | 17824615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17824615&pc=OZD | 03/05/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 03/05/2023 | 09/27/2023 |  |  | 12/31/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024293246 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/05/2023 |  |  |  |  | The user facility reported 67yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that while the patient was getting out of bed and the device was pulled on by either arm or stepped on, the pump stopped and was unable to be restarted. the pump was exchanged for a new impella. there was no patient harm reported. | [5 empty] | The investigation for the pump stop has been completed. the pump was returned and an electrical resistance test was performed. it was found that all three motor cables had open lines. the red handle was opened and necking and damage was seen on the motor cables on the catheter side. the root cause of the pump stop was determined to be an electrical open caused by excessive force that was applied when the patient was getting out of bed and the pump was either pulled or stepped on. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34311 | 1220648-2023-03427 | 17824771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17824771&pc=OZD | 02/27/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 02/27/2023 | 09/27/2023 |  |  | 11/30/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023278026 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/27/2023 |  |  |  |  | The user facility reported 57yo male in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a left ventricle (lv) signal went very negative and teh flows came inaccurate. there was a concern for pump stoppage so decision made to exchange pump in the operating room. there was no patient harm. | [5 empty] | The investigation for the flow saturation issues has been completed. the data log shows an increase in motor current at the end of the case during the time in which there is saturated flow. at the same p-level as before the saturated flow, the motor current is higher during the time of the saturated flow event. biomaterial was observed on the impeller and the saturated flow reproduced in the lab when the pump was run. the cause of the saturated flow was biomaterial on the impeller. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34312 | 1220648-2023-03428 | 17824844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17824844&pc=OZD | 04/19/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 04/19/2023 | 09/27/2023 |  |  | 03/31/2025 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024319091 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/19/2023 |  |  |  |  | The user facility reported that a 77-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak coming from the connection of the purge line and sidearm while in the operating room (or). sodium bicarbonate purge solution was utilized as purge solution from beginning of case. the line was subsequently disconnected and retightened but leaking continue and a "purge pressure low" alarm ensued. plan was to attempt changing the purge cassette, but due to the patient's declining condition the family to withdraw care. no further information was provided. | [5 empty] | The investigation into the reported purge leak was completed. the pump, purge cassette, and data stick were returned for evaluation. visual inspection found a crack on the surface of the purge sidearm filter. functional testing confirmed the leak. the root cause of the purge leak was due to damage of the filter. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34313 | 1220648-2023-03429 | 17824907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17824907&pc=OZD | 08/28/2023 | 03/01/2024 |  | 09/27/2023 |  |  | 08/30/2023 | 09/27/2023 |  |  | 05/31/2025 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363600 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2023 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient's right leg had no pulse and was ischemic. the impella had been implanted on the right side. patient was emergently taken back to the cath lab for impella cp explant and re-implant on the left side. | [5 empty] | The investigation into the limb ischemia ¿ irreversible event has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34314 | 1220648-2023-03430 | 17824935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17824935&pc=OZD | 04/19/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 04/19/2023 | 09/27/2023 |  |  | 12/31/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 202494416 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/19/2023 |  |  |  |  | The user facility reported that a 58-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the physician had just finished ambulating the patient and the purge side arm "around the hub area" starting leaking profusely. purge bypass was subsequently performed to resolve the issue. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak issue was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34315 | 1220648-2023-03431 | 17827411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17827411&pc=OZD | 04/20/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 04/20/2023 | 09/27/2023 |  |  | 12/31/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024315808 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2023 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported an optical sensor alarm after connecting the impella 5.5. a new automated impella controller (aic) was used but still the alarm remained and so the pump was replaced. there was no harm to the patient. | [5 empty] | The root cause of the optical signal issue is due to a broken fiber line found during benchtop analysis of the pump. this report is being filed as part of a retrospective review of historical records.
34316 | 1220648-2023-03432 | 17827513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17827513&pc=OZD | 03/29/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 03/29/2023 | 09/27/2023 |  |  | 04/30/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211998 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2023 |  |  |  |  | The user facility reported a 37-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a loss of placement signal during impella support. a visual inspection was subsequently performed, and it was identified that there was a kink in the catheter near the red impella plug proximal to the pump. the pump was removed and replaced to resolve the failure. there was no patient harm. | [5 empty] | Data logs show the placement signal suddenly spikes, accompanied by the "placement signal not reliable" alarm. optical parameters were characteristic of optical fiber damage. visual inspection of the returned pump revealed a significant catheter kink next to the kink protector. optical continuity test revealed fiber damage at the catheter kink. the placement signal not reliable alarm reproduced when the pump was connected to the aic. in conclusion, it was determined that the cause of the optical signal issue was a damaged optical fiber line stemming from patient ambulation. this report is being filed as part of a retrospective review of historical records.
34317 | 1220648-2023-03433 | 17827655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17827655&pc=OZD | 01/10/2023 | 10/07/2025 |  | 09/27/2023 |  |  | 01/10/2023 | 09/27/2023 |  |  | 11/30/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278011 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | The user facility reported 41yo male was implanted with an impella 5.5 device for mechanical circulatory. it was reported that the pump stopped four times. they were able to restart at p2 but it kept stopping. the pump stop prompted device removal and replacement. there was no patient harm reported. | The investigation uncovered placement signal issues. |  |  |  |  | The investigation for the reported pump stop has been completed. data logs show that the placement signal suddenly spiked out of specification about 18 days into support, accompanied by ¿placement signal not reliable¿. two days later, the pump began to stop intermittently. these pump stops were accompanied by motor current spikes. and were indicative of a pump stop due to an open electrical line. the benchtop analysis showed an open line, in conclusion, it was determined that the cause of the pump stop was an open electrical line and the cause of the optical signal issue was a damaged optical fiber line; both failures were due to the catheter kink near the red pump plug. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. all pertinent information available to abiomed has been submitted at this time. h6. updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-03433. b7 other relevant history, including preexisting medical conditions updated. d4. model number updated. d6.a. / d6.b. updated. d.10. concomitant medical products and therapy dates updated. f6. and f8. should have been left blank in the initial report. g1. reporting contact email and reporting contact fax number updated.
34318 | 1220648-2023-03434 | 17827727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17827727&pc=OZD | 07/23/2022 | 09/27/2023 |  | 09/27/2023 |  |  | 07/23/2022 | 09/27/2023 |  |  | 05/31/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023219723 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/23/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the 5.5 support that had position signal loss. after the loss of position signal the pump was replaced after 3 days. the pump was weaned and explanted. a 2nd 5.5 was placed and it supported successfully. no lasting harm or injury. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
34319 | 1220648-2023-03435 | 17827815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17827815&pc=OZD | 04/20/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 04/20/2023 | 09/27/2023 |  |  | 12/31/2024 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024289672 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2023 |  |  |  |  | The user facility reported that a 59-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the physician mentioned they didn't believe the impella was successfully unloading the patient. positioning was confirmed with fluoroscopy. the patient was noted to have severe right ventricular dysfunction. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. analysis of the data logs revealed a large occurrence of suction alarms but did not show low pump flow. flows were as expected for the p level. visual inspection of the device did not show any visual defects or abnormalities. based on the clinical account, the root cause of the reported event was due to the pump was not successfully unloading the patient because of the patient's severeright ventricular dysfunction. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34320 | 1220648-2023-03436 | 17829749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17829749&pc=OZD | 06/11/2023 | 09/27/2023 |  | 09/27/2023 |  |  | 06/11/2023 | 09/27/2023 |  |  | 03/31/2025 |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024327661 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Alarm | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/11/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that due to the position signal not recognized alarm (peak signal to noise) they decided to change the console. upon exchanging, the alarm was resolved. this procedure was completed successfully using the initial console. no patient harm or injury was reported. | [5 empty] | The console was able to run a pump without any errors but was also able to recreate the alarms by slightly unplugging the pump and having it run. the patient was transferred to ic9735 which ran without reproducing this failure. these results are consistent with an air gap between the optical pump and connector. the cause of the issue was due to the optical pump not being fully plugged in. this report is being filed as part of a retrospective review of historical records.
34321 | MW5146271 | 17836635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17836635&pc=OZD | 09/14/2023 | 09/25/2023 |  | 09/27/2023 |  |  |  | 09/27/2023 |  |  |  |  | 09/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA IMPELLAR CATHETER |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED INC. | [8 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 64 YR |  | Female |  |  |  | Malfunction | No | Yes | [7 empty] | Cardiac Arrest | [12 empty] | Other | [28 empty] | Initial submission | [9 empty] | During a code blue in the cath lab, an impella device was placed by the cardiologist. the impella was placed in p2 and compressions were continued. during compressions, the impella became lodged too far into the left ventricle and subsequently folded in half. in attempting to remove the device, the cardiologist called on a vascular surgeon for assistance as well as the abiomed impella employee. there was worry about whether the device could be safely removed without rupturing vasculature. the device was safely removed by the cardiologist after different maneuvers and tries. the patient did achieve rosc, and the code was run and carried out very well. the patient did code again after this, but the impella was not re-implanted.
34322 | 1220648-2023-03382 | 17817060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817060&pc=OZD | 08/29/2023 | 01/22/2024 |  | 09/26/2023 |  |  | 08/29/2023 | 09/26/2023 |  |  |  |  | 09/26/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1122793 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | There was no patient involvement. no additional information was provided. | The user facility reported the automated impella controller (aic) had a system self check error. |  |  |  |  | The investigation for the reported error message on the console has been completed. the device and the data logs were returned for evaluation. the data logs were reviewed which showed "system self check error". the console was tested on an engineering lab bench with the original parts and configuration. the ¿system self check error¿ was unable to be reproduced, but during testing the continuous alarm, and the log recorded the controller error (loss of comm (pbm1)) - alarm event set # (b)(4). upon visual inspection, pbm corrosion was found. the root cause of the system self check error was due to the defective pbm board. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
34323 | 1220648-2023-03383 | 17817149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817149&pc=OZD | 08/29/2023 | 01/25/2024 |  | 09/26/2023 |  |  | 08/29/2023 | 09/26/2023 |  |  | 05/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365551 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the team observed an access site bleed. the bleed was observed on day five of the seven day support. the bleed was treated by 1 unit of platelets and consult with surgery. the team did not need further intervention. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. as per clinical, patient bleeding from impella site, surgical sutures added, and hemostasis was achieved. improper management of access site was noted with not doing forward tension sutures. the cause of the access site bleeding issue was most likely use related due to improper management of access site per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34324 | 1220648-2023-03384 | 17817368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817368&pc=OZD | 08/27/2023 | 07/30/2024 |  | 09/26/2023 |  |  | 08/27/2023 | 09/26/2023 |  |  | 03/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342441 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2023 |  |  |  |  | The user facility reported a 53-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was experiencing access site bleeding. the surgical team proceeded with manual pressure without any success. the surgeon decided to cut down and place a vascular patch in place and the bleeding resolved. the device was explanted on the same day. the patient remained stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be use issue because the issue was resolved by surgically placing the vascular patch at the site. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34325 | 1220648-2023-03385 | 17817394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817394&pc=OZD | 03/09/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 03/09/2023 | 09/26/2023 |  |  | 12/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024307349 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/09/2023 |  |  |  |  | The user facility reported 61yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were high purge pressure alarms during use. cassette changes were performed but this did not resolve the alarms. eventually, the pump was removed. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure alarms has been completed. the product was returned and a leak was confirmed in the filter of the sidearm. the cracks are consistent with chemical interaction. the sidearm was removed and the purge was reattached. the high purge pressure was then reproduced. blood was visualized in the purge gap. as the pump was run, the biomaterial was ejected and the purge pressure normalized. the data logs were reviewed and show a slow purge increase over an hour. the root cause of the purge leak was determined to be filter leak due to use. the root cause of the high purge pressure was determined to biomaterial. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34326 | 1220648-2023-03386 | 17817442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817442&pc=OZD | 12/30/2022 | 09/26/2023 |  | 09/26/2023 |  |  | 12/30/2022 | 09/26/2023 |  |  | 12/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023282859 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/30/2022 |  |  |  |  | The user facility reported that a 70-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced continued low pump flows resulting in pump removal after attempts to reposition and provided volume failed to correct the issue. no clot was noted on removal. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. analysis of the data log shows multiple motor current spikes about an hour into the case. at the time of these spike, the flow dropped, and the motor speed dropped suggesting that biomaterial was contacting the impeller. based on the available information, the root cause of the low pump flow was most likely biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34327 | 1220648-2023-03387 | 17817477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817477&pc=OZD | 01/30/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 01/30/2023 | 09/26/2023 |  |  | 09/30/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023257463 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/30/2023 |  |  |  |  | The user facility reported that a 55-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the device became dislodged from the left ventricle when the patient abruptly sat up in the bed after being moved off the operating room (or) table. the patient reached to pull out his endotracheal (et) tube and the "impella in aorta" alarmed and then resolved when the patient laid down. the alarm occurred again in the intensive care unit (icu). the physician was unable to recross the valve, so the pump was removed. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. according to the clinical details the pump became dislodged due to patient movement. based on this information, the root cause of the positioning issues was determined to be patient movement. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34328 | 1220648-2023-03388 | 17817508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817508&pc=OZD | 01/30/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 01/30/2023 | 09/26/2023 |  |  | 12/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023284608 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/30/2023 |  |  |  |  | The user facility reported that a 59-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced lower than expected flows and unreliable placement signals (ps). echocardiogram revealed left ventricular thrombus on inlet of the left ventricle. the patient was therapeutic on anticoagulation. the purge solution was dw5 at 25 units/ml. the patient was subsequently transferred to another facility and the decision was made to keep the pump in place as the flows stabilized and the patient was adequately supported. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation and no logs are available from the console for the time of the event. based on the available information the root cause of the low pump flow was most likely biomaterial ingestion. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34329 | 1220648-2023-03389 | 17817534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817534&pc=OZD | 02/13/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 02/13/2023 | 09/26/2023 |  |  | 12/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291878 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/13/2023 |  |  |  |  | The user facility reported that a 44-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the device was placed in direct fashion; subsequently flows became erratic at times despite position confirmation under transesophageal echocardiogram (tee). the physician elected to remove the pump and the patient completed the aortic valve replacement (avr). no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. based on the available information and review of the data logs, the root cause of the low pump flow issue was undetermined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34330 | 1220648-2023-03390 | 17817561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817561&pc=OZD | 02/13/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 02/13/2023 | 09/26/2023 |  |  |  |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291917 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/13/2023 |  |  |  |  | The user facility reported that a 67-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. when repositioning, there was lots of built-up torque and when the device was torqued and pulled back it spun and flipped out of the left ventricular. the pump was unable to recross and was removed as a result. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the available information, the root cause of the positioning issue was determined to be use related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34331 | 1220648-2023-03391 | 17817608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817608&pc=OZD | 08/28/2023 | 11/10/2023 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  | 06/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024361220 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 had purge leak. the 5.5 replacement was aborted and proceeded with insertion of impella cp. no lasting harm or injury. the 5.5 had only supported the patient for 7 minutes. | [5 empty] | The investigation purge leak ¿ pump has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, impella 5.5 had purge leak but the leak location was unknown. the cause of the purge leak issue was not determined since product was not returned and with inconclusive clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿.
34332 | 1220648-2023-03392 | 17817640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817640&pc=OZD | 02/21/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 02/21/2023 | 09/26/2023 |  |  | 02/29/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023193716 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/21/2023 |  |  |  |  | The user facility reported that a 47-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the physician attempted to reposition device by bedside echo imaging and inadvertently pulled the device across the valve and into the aorta. several attempts were made to place pump back across the valve and into proper position but ultimately, they were unsuccessful. no further information was provided. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the available information, the root cause of the positioning issue was use related as the was pump inadvertently pulled out by the physician. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34333 | 1220648-2023-03393 | 17817671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817671&pc=OZD | 02/19/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 02/19/2023 | 09/26/2023 |  |  | 03/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210158 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/19/2023 |  |  |  |  | The user facility reported that a 63-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the dislodged into the aorta during support and needed to be exchanged when it was unable to recross the aortic valve. no further information was provided. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. analysis of the data log shows the "impella position unknown" alarms throughout the case. based on the available information, the root cause of the positioning issue was most likely use related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34334 | 1220648-2023-03394 | 17817689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817689&pc=OZD | 08/28/2023 | 02/19/2024 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  | 05/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356197 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had an impella 5.5 prepped and ready for use when the team observed a purge leak from both purge reservoir and microfilter. the impella 5.5 and purge cassette were replaced and support proceeded with new without any harm or injury from delay. | [5 empty] | The investigation into the purge leak ¿ pump and the purge leak cassette issues has been completed since the original report was filed. the pump and purge cassette were returned for investigation. when purged, the sidearm was found to be leaking from the right side of the filter at the joint. the root cause of the purge leak was determined to be sidearm filter damage most likely due to insufficient adhesive. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g.infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34335 | 1220648-2023-03395 | 17817702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817702&pc=OZD | 08/28/2023 | 09/24/2024 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  | 06/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024362542 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It is noted the patient expired after care was withdrawn. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 79-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed disseminated intravascular coagulation (dic) and bleeding from the insertion site. the bleeding was treated with blood transfusion and pressure hemostasis. it was further reported that a ct scan revealed cerebral hemorrhage. it was decided that additional treatment would not be performed due to cerebral ventricular rupture. the physician stated that the patient being treated with antithrombotic and anticoagulant drugs, and that the blood vessels were already in bad condition, possibly resulting in the cerebral hemorrhage and no direct causal relationship with impella. |  |  |  |  | The investigation into the access site bleeding ¿ major, stroke/cerebral vascular accident (cva) and unrelated death has been completed. the device was not returned for benchtop evaluation and investigation. the root cause of access site bleeding was most likely due to anticoagulation management. the patient had high activated clotting time (act) values. the cause of the stroke was most likely due to patient condition and unknown relation to impella. the cause of the death was most likely due to patient condition and unknown relation to impella. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03395 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34336 | 1220648-2023-03396 | 17817740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817740&pc=OZD | 08/28/2023 | 02/26/2024 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  | 03/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345220 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced ischemia. the impella cp was explanted. the patient had an embolectomy to left femoral artery (lfa) after impella cp was explanted by vascular surgeon. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34337 | 1220648-2023-03397 | 17817754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817754&pc=OZD | 04/07/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 04/07/2023 | 09/26/2023 |  |  | 12/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307994 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a delivery issue involving an impella 5.5 pump. it was noted that the pump failed to advance through the innominate artery and into the ascending aorta during the attempted implant. upon removal, it was observed that the catheter appeared to be slightly deformed proximal to the motor housing. an impella cp pump was subsequently implanted without further issue. there was no patient harm. | [5 empty] | The investigation into the reported "delivery issue" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that as the pump was unable to advance through the patient's vasculature. the device deformation likely occurred while trying to push past the resistance. the root cause for delivery issue is most likely patient condition. this report is being filed as part of a retrospective review of historical records.
34338 | 1220648-2023-03398 | 17817778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817778&pc=OZD | 02/22/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 02/22/2023 | 09/26/2023 |  |  | 12/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024291832 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/22/2023 |  |  |  |  | The user facility reported that a 24-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. as a result, the device lost the ability to support the patient overnight. flows dropped and the patient experienced suction. the issue was unable to be corrected and so was taken back to operating room for pump exchange. no further information was provided. there were reports of harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the pump and data log stick were returned for evaluation. no damage was noted on the pump upon visual inspection. biomaterial was observed around the impella. functional testing of the pump was able to reproduce the low flow issue. based on the available information, the root cause of the low pump flow was biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34339 | 1220648-2023-03399 | 17817801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817801&pc=OZD | 08/27/2023 | 02/16/2024 |  | 09/26/2023 |  |  | 08/27/2023 | 09/26/2023 |  |  |  |  | 09/26/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 2024351025 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Delayed Alarm | Device Alarm System | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/27/2023 |  |  |  |  | The user facility reported a 78-year-old female with heart failure was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a "controller error" alarm. the pressure surge and left ventricle waveforms disappeared approximately 90 seconds before the alarm appeared. impella motor current and flows appeared within normal limits throughout alarm. aic successfully exchanged. there was no patient harm. | [5 empty] | The investigation into the console "aic - controller error (yellow alarm)" has been completed since the original report was filed. impella console was returned. the evaluation found that the yellow alarm is a known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. after analyzing of the logs from the console, it's been determined that the reported issue is a known pattern failure caused by a temporary loss of optical placement signal. abiomed is aware of the issue and is working on appropriate corrective action. this is a low probability event and is very difficult to reproduce. as a result, the root cause of the issue was not determined since the issue could not be reproduced during testing. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34340 | 1220648-2023-03400 | 17817820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817820&pc=OZD | 04/06/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 04/06/2023 | 09/26/2023 |  |  | 07/31/2023 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118210 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a delivery issue involving an impella 5.5 pump. it was documented that the device kinked during attempted implant and was unable to advance through the patient's vasculature. the implant attempt was aborted and patient support continued with the already implanted impella cp, as escalation to the 5.5 was not possible. there was no patient harm. | [5 empty] | The investigation into the reported "delivery issues" has been completed. product was returned for investigation. the evaluation revealed that the root cause of delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
34341 | 1220648-2023-03401 | 17817852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817852&pc=OZD | 02/23/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 02/23/2023 | 09/26/2023 |  |  | 03/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023198327 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/23/2023 |  |  |  |  | The user facility reported that a 62-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. repositioning was attempted. the patient also had a ¿pump in aorta¿ after a coughing fit. ultimately, the physician was unable to readvance the pump and it was exchanged. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the available information, the root cause of the positioning issues was most likely due to wireless re-access. the device is not indicated for wireless re-access per ifu 0550-9002. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34342 | 1220648-2023-03402 | 17817915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817915&pc=OZD | 09/02/2022 | 09/26/2023 |  | 09/26/2023 |  |  | 09/02/2022 | 09/26/2023 |  |  | 03/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023201371 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a no flow issue with an impella device. during patient support, the placement signal was lost and the measured motor current was flat. the patient subsequently crashed that same night and was supported with extracorporeal membrane oxygenation (ecmo). the impella device was removed the following morning to prevent a potential recurrence. | [5 empty] | The data logs confirm the low flow and low pulsatility in the motor current as well as the loss of placement signal after reconnecting the patient cable. the product was returned with a cut catheter so the failure modes were unable to be reproduced. a large thrombus was found in the outflow located on the back of the optical sensor. the root cause of the low flow was determined to be ingested biomaterial. the imc logs were reviewed and show evidence of a communication error between the console and the eeprom while reading the previously calibrated g0 value. the software read this value as zero and the placement signal is lost as a result. this failure indicates corruption on the eeprom. the root cause of the optical signal failure was determined to be a corrupted eeprom.
34343 | 1220648-2023-03403 | 17817945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817945&pc=OZD | 08/22/2023 | 01/26/2024 |  | 09/26/2023 |  |  | 08/22/2023 | 09/26/2023 |  |  | 06/30/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023237448 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2023 |  |  |  |  | The user facility reported a 34-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced oozing form the access site and was required to be given protamine to stop the bleeding. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34344 | 1220648-2023-03404 | 17818597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17818597&pc=OZD | 06/12/2022 | 01/26/2024 |  | 09/26/2023 |  |  | 06/12/2022 | 09/26/2023 |  |  | 01/31/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. access site for impella dressed, but upon scrubbing out and removing the drape the staff noticed a hematoma forming. the physician scrubbed back in and performed a cut down to assess bleeding and establish hemostasis. physician solved the issue, and closed the site. | [5 empty] | The investigation into the high purge pressure and the access site bleeding ¿ major issue has been completed. the product was not returned for benchtop investigation. data logs showed the "purge system failure" alarm was triggered during purge prime. the alarm resolved after the purge cassette was replaced for another. the cause of the priming issue and the access site bleeding issue was not determined since product was not returned for investigation and there were limited clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34345 | 1220648-2023-03405 | 17818761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17818761&pc=OZD | 12/23/2021 | 09/26/2023 |  | 09/26/2023 |  |  | 12/23/2021 | 09/26/2023 |  |  | 09/30/2023 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143638 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2021 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported false "impella in aorta" alarms during support. the decision was made to remove and replace the impella. there was no harm to the patient. | [5 empty] | The returned pump was visually inspected, and biomaterial was observed within the inlet. based on the provided clinical information that the pump was properly positioned, and that the patient had tissue in the left ventricle, the root cause of the false positioning alarms was determined to be related to ingested biomaterial. the presence of biomaterial can reduce the ability of the pump¿s optical sensor to detect pulsatility. no data logs were provided. this report is being filed as part of a retrospective review of historical records.
34346 | 1220648-2023-03406 | 17819145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17819145&pc=OZD | 12/30/2021 | 11/19/2024 |  | 09/26/2023 |  |  | 12/30/2021 | 09/26/2023 |  |  | 07/31/2023 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130589 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported an unreliable placement signal during support. there was no adverse impact to the patient because of this. | [5 empty] | Neither the product nor the data logs were returned for investigation. therefore, the cause of the optical placement signal issue was not determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. neither the product nor the data logs were returned for investigation. therefore, the cause of the suction and of the optical placement signal issue was not determined. this report is being filed as part of a retrospective review of historical records.
34347 | 1220648-2023-03407 | 17819164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17819164&pc=OZD | 08/22/2023 | 01/29/2024 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  | 09/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024365586 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the patient acquired an infection at the device access site. a re-operation, wound revision, wound irrigation, vacuum and dressing were all performed. | [5 empty] | The investigation into the infection/sepsis issue has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, the patient had infection of access site and re-operation of access site was done to resolve the infection, but this issue occurred 12 days after impella device explant. the cause of the infection issue was most likely patient condition related and unrelated to impella based on clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
34348 | 1220648-2023-03409 | 17821465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821465&pc=OZD | 05/25/2022 | 09/26/2023 |  | 09/26/2023 |  |  | 05/25/2022 | 09/26/2023 |  |  | 02/29/2024 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023195774 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported false positioning alarms during support. the pump was ultimately removed and replaced. there was no harm to the patient. | [5 empty] | The returned pump was connected to a console and no placement signal alarms triggers. ps was a steady and operated to specification. the cause of the issue was most likely patient condition related since the patient had heart failure and the returned device operated to specification. this report is being filed as part of a retrospective review of historical records.
34349 | 1220648-2023-03410 | 17821480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821480&pc=OZD | 08/29/2023 | 08/02/2024 |  | 09/26/2023 |  |  | 08/29/2023 | 09/26/2023 |  |  | 05/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2023 |  |  |  |  | The user facility reported a 74-year-old female in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. a small hematoma was noted at the access site at the time of impella removed. perclose closure device was used and manual pressure used for hemostasis in cath lab. additionally, the patient developed hematuria hours after the impella was removed. | [5 empty] | The investigation for the hemolysis has been completed. according to the clinical, a small hematoma was found at the access site. in order to accomplish hemostasis, a perclose was used and manual pressure was applied. additionally, the patient developed hematuria hours after the pump was removed. no product was returned for a visual inspection. data log analysis confirmed that the pump ran without issue for the entire case. no motor current spikes, suction alarms, positioning alarms, or ps trends were observed. the cause of the minor access site bleeding was not determined because no product was returned and not enough clinical information was given. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations section: entering cardiac output section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34350 | 1220648-2023-03411 | 17821496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821496&pc=PYX | 08/30/2023 | 01/25/2024 |  | 09/26/2023 |  |  | 08/30/2023 | 09/26/2023 |  |  | 06/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024363954 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the doctors attempted to advance the impella rp through to the right ventricle without success. patient became asystolic briefly and it resolved with the pump removal. decision was made to abort the rp insertion. | [5 empty] | The investigation into the into the delivery issues- anatomy and the asystolic arrest issues has been completed since the original report was filed. per the given clinical information, the root cause of the delivery issue was patient condition related to small/ diseased vessels. any device contacting the heart wall in conjunction with a poor patient condition could result in vf/vt but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vf/vt is unable to be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34351 | 1220648-2023-03412 | 17821555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821555&pc=OZD | 08/29/2023 | 01/11/2024 |  | 09/26/2023 |  |  | 08/29/2023 | 09/26/2023 |  |  |  |  | 09/26/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1137162 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Inappropriate Audible Prompt/Feedback | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/29/2023 |  |  |  |  | The user facility reported the impella had an impella defective alarm when it connected to the automated impella controller (aic). the impella 5.5 was swapped out, but the issue persisted. upon replacing the aic, the problem was resolved. the was no patient harm. | [5 empty] | The investigation into the aic - pump issue has been completed since the original report was filed. the console was returned for investigation. testing confirmed that the cause of the issue was a faulty impellatronic board, specifically the epm/eep system of the impellatronic board, which communicates with the impella¿s eeprom. | The impella console device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34352 | 1220648-2023-03413 | 17821567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821567&pc=OZD | 08/30/2023 | 09/04/2024 |  | 09/26/2023 |  |  | 08/30/2023 | 09/26/2023 |  |  |  |  | 09/26/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1249363 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was air in purge alarm observered on the impella console. the console would purge the line, but once connected the alarm remained active. the abiomed field team advised swapping out the console. no harm or injury noted. | [5 empty] | Upon further review the engineer determined that the automated imeplla controller (aic) was operating as expected. there were no issues with the aic. the engineer determined that the issue was with the purge cassette, not the aic. a regulatory report is no longer necessary. | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34353 | 1220648-2023-03414 | 17821574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821574&pc=OZD | 08/29/2023 | 01/12/2024 |  | 09/26/2023 |  |  | 08/29/2023 | 09/26/2023 |  |  | 04/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355005 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on the second day of support, the team explanted the device for issues with hemolysis. the hemolysis was observed by signs of: urine color change, hemolyzed labs, and cp in poor position in the lv. the pump was prematurely explanted and the patient was escalated to the impella 5.5 pump. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. the data logs did not record any events related to biomaterial interaction or improper pump position. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34354 | 1220648-2023-03415 | 17821598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821598&pc=OZD | 05/22/2023 | 09/26/2023 |  | 09/26/2023 |  |  | 05/22/2023 | 09/26/2023 |  |  | 04/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024333573 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/22/2023 |  |  |  |  | The user facility reported 55yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after the pump was started, a ventricle alarm was triggered which required the explant of the pump. there were concerns about pump saturation. once removed, the patient was placed on an intra-aortic balloon pump (iabp). there was no patient harm reported. | [5 empty] | The investigation for the reported saturated flows has been completed. the data log shows a motor current spike towards the beginning of the case followed by slightly saturated flow. the returned product was found to have biomaterial around the impeller. the cause of the saturated flow was biomaterial. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34355 | 1220648-2023-03416 | 17821712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821712&pc=OZD | 08/26/2023 | 01/18/2024 |  | 09/26/2023 |  |  | 08/27/2023 | 09/26/2023 |  |  | 04/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330392 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as escalation of care. during support, the impella purge sidearm had minimal amounts of leaking. purge pressures remain stable and in normal range. there was no patient harm. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34356 | 1220648-2023-03417 | 17821734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821734&pc=OZD | 08/27/2023 | 01/16/2024 |  | 09/26/2023 |  |  | 08/27/2023 | 09/26/2023 |  |  | 03/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327525 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2023 |  |  |  |  | The user facility reported a 56-year-old male in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the team noted an access site bleed. the bleed prompted the team to visit the or for a repair. the artery was repaired and the jackson pratt drain was placed. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. per the clinical details provided, the root cause of the access site bleeding issue was use related to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34357 | 1220648-2023-03418 | 17821798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821798&pc=OZD | 08/30/2023 | 03/19/2024 |  | 09/26/2023 |  |  | 08/30/2023 | 09/26/2023 |  |  | 03/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328754 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was attempted to be inserted without success due to groin complications. the impella cp was taken out and shockwave was performed. the impella cp was then reinserted successfully and ran for a few minutes prior to shutting off. the pump was taken out and a new pump was exchanged at this time. | [5 empty] | The investigation into the delivery issues- use and the pump stop issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the delivery issue was patient condition related to calcified vessel condition since device was successfully placed after shockwave procedure. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34358 | 1220648-2023-03419 | 17821819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17821819&pc=OZD | 08/28/2023 | 01/18/2024 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  | 07/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374293 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient was sent to cardiovascular catheterization lab (ccl) for right atrial (ra) pacing lead and possible right coronary artery (rca) percutaneous coronary intervention (pci). the staff shocked patient four times in ccl, and then found clot on inflow cage via tee. the staff had to turn flows down to p4 to break suction and care was withdrawn shortly after. | [5 empty] | The investigation into the thrombus has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the clot was observed on inflow cage via tee and had to turn flows down to p4 to break suction. data logs were reviewed, and motor current spikes were seen a few hours before multiple suction alarms were triggered with low flows at p-8. after reduction of p-level to p-4, low flows were still observed. the cause of the low pump flow issue was most likely biomaterial ingestion with suction alarms per clinical and log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34359 | 1220648-2023-03420 | 17822393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17822393&pc=OZD | 08/28/2023 | 01/23/2024 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  |  |  | 09/26/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1095120 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the purge disk blue holder broke off while the staff was removing the impella connections from the console after completion of a case. as such, there was no patient harm as a result of the noted issue. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the device was returned for investigation. the console was tested after arriving in fs. purge retainer was confirmed to be broken. the cause of the issue was a defective component. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
34360 | 1220648-2023-03421 | 17822409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17822409&pc=OZD | 08/28/2023 | 01/24/2024 |  | 09/26/2023 |  |  | 08/28/2023 | 09/26/2023 |  |  | 05/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366765 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported as the staff was advancing impella over the wire the patient went into third degree atrioventricular (av) block. the impella and wire were removed through peel away sheath, access obtained venous for temporary transvenous pacemaker. the impella was outside of body for 5-10 minutes. the impella was removed and a new impella was implanted successfully. | [5 empty] | The investigation into the valvular insufficiency/regurgitation leading to delivery issue has been completed since the original report was filed. the device was returned for investigation. the cause of the delivery issue was use related since the pump was outside the body for 10 minutes and aborted delivery due to patient safety per clinical review. the cause for the third degree atrioventricular block could not be established, and the pump was swapped as a result of it being outside of the patient for 10 minutes. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34361 | 1220648-2023-03422 | 17822432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17822432&pc=OZD | 08/29/2023 | 07/15/2024 |  | 09/26/2023 |  |  | 08/29/2023 | 09/26/2023 |  |  | 05/31/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352446 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 70-year-old female in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient had frequent runs of ventricular tachycardia (vt), requiring multiple rounds of defibrillation. rhythm was finally broken and amio gtt was initiated along with heavy sedation. an echocardiogram revealed that the pump was against the apex of the left ventricle (lv). care team was attempting to pull the impella back, and it would not move. the rep. advised to undo the three point anchoring and get the impella in line with the insertion angle, but the impella still would not move despite all best practices. the patient was subsequently taken to the operating room where a surgeon was able to reposition the device under echocardiogram. support continued with the implanted pump following the intervention. |  |  |  |  | The investigation for the ventricular tachycardia (vt), thrombus, and positioning issues has been completed. according to the clinical details, the patient suffered frequent runs of vt. imaging confirmed the pump was against the apex of the lv. multiple attempts were made to reposition the pump, but the pump would not move despite all best practices being followed. there was also known thrombus within the graft. the patient continued on support for another three days until the pump was removed. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issue was use related. without further clinical details, the root cause of the vt and the thrombus could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock in the following sections: section: potential adverse events (unites states), which now includes the statement "cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34362 | 1220648-2023-03423 | 17822459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17822459&pc=OZD | 08/29/2023 | 01/25/2024 |  | 09/26/2023 |  |  | 08/29/2023 | 09/26/2023 |  |  | 06/30/2025 |  | 09/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373248 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 49-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient has lower extremities pulses via palpation or doppler, the cardiology aware. the patient bilateral lower extremities skin is pink and cool. lasix drip initiated. per echo there was little left ventricle (lv) wall motion. the plan of care is to consult hf and consult surgery. a distal perfusion sheath was placed on (b)(6) 2023. the patient remained on impella support until (b)(6) 2023. | [5 empty] | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. based on the clinical information, the patient bilateral lower extremities skin were pink and cool and distal perfusion sheaths were placed. the cause of the limb ischemia issue was most likely patient condition related and unknown relation to impella with ischemia observed on both limbs per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34363 | 1220648-2023-03378 | 17813624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17813624&pc=OZD | 08/25/2023 | 01/19/2024 |  | 09/25/2023 |  |  | 08/26/2023 | 09/25/2023 |  |  | 06/30/2025 |  | 09/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2023 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support via axillary. the patient experience right forearm and hand ischemia. there were no doppler signals in right radial or ulnar arteries. brachial artery signal present. antegrade reperfusion sheath was placed in right brachial artery and connected to rewire port of repositioning sheath. field service rep. advised that this was not recommended and that this passive flow would need to be frequently checked for patency. the team did not want to splice into ecmo circuit as leg reperfusion was already connected. vascular surgery did fasciotomy of right forearm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's limb ischemia has been completed. product was not returned for review. the root cause of limb ischemia was most likely due to patient condition.
34364 | 1220648-2023-03379 | 17813629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17813629&pc=OZD | 08/26/2023 | 01/24/2024 |  | 09/25/2023 |  |  | 08/26/2023 | 09/25/2023 |  |  | 06/30/2025 |  | 09/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371219 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2023 |  |  |  |  | The user facility reported a 50-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the time flows were <1lpm and the motor current was flat. the physician noted signs of a clot ingestion were apparent. the physician made the decision to exchange the cp and reinsert a new cp. there was no patient harm. | [5 empty] | The investigation into the thrombus and the low pump flow issues has been completed. the device was not returned for benchtop investigation. data logs show low flow from the beginning of the case with suction conditions while running at p-level 9. the total case run time was 8 minutes. a motor current spike was reported while pump was running on a steady p-level. according to clinical reports, the doctor also noted an lv thrombus. the cause of the low flow issue was most likely biomaterial, based on given clinical details and data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34365 | 1220648-2023-03380 | 17813805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17813805&pc=OZD | 08/26/2023 | 11/14/2025 |  | 09/25/2023 |  |  | 08/26/2023 | 09/25/2023 |  |  |  |  | 09/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251615 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 08/26/2023 |  |  |  |  | The abiomed field service representative reported the user facility's automated impella controller had a ¿system self check failed¿ alarm when booting up. staff attempted to reboot, but was unsuccessful. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: console imc logs confirm that the system self check failure was issued. [26/08/23 12:09:08 post: syscall10 checkfirmware failed]. there preceded by numerous cpld and epm communication errors indicating that the there was communication between impellatronics and the epc. device analysis: console internal circuitry was inspected. there was burned components on the bottom layer of the pbm most probably due to a short circuit. replacing the pbm with a known good one didn't resolve issue. replacing the impellatronics board with a known good one resolved issue. conclusion: the root cause of the system self check failure was due to a defective impellatronics board. d4 model number updated.
34366 | 1220648-2023-03381 | 17813869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17813869&pc=OZD | 08/26/2023 | 01/25/2024 |  | 09/25/2023 |  |  | 08/26/2023 | 09/25/2023 |  |  | 06/30/2025 |  | 09/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356954 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2023 |  |  |  |  | The user facility reported a 55-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support while waiting for heart transplant. the patient was on support for several weeks when the device lost its placement signal and had suction/flow problems. the patient was to be transplanted with new heart soon so the physician decided to exchange the impella 5.5 for an impella cp in the interim. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the optical signal issue issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported. the cause of the low flow issue was most likely patient condition related with suction alarms and low flow at (b)(4) per log review and poor patient condition based on clinical review. the cause of the optical signal issue was most likely due to sensor silicone wear with placement signal drift observed impacting 5s parameters due to log review. the sensor wear occurred within the design life period of the 5.5 pump.
34367 | MW5146133 | 17817441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17817441&pc=OZD | 03/24/2023 |  |  | 09/25/2023 |  |  |  | 09/25/2023 |  |  |  |  | 09/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA HEART PUMP |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. | [8 empty] | No answer provided | [9 empty] | Voluntary report | Not Available | Patient Family Member or Friend | [6 empty] | 61 YR |  | Male |  |  | White | Death | Yes | No | [7 empty] | Cardiac Enzyme Elevation | Skin Discoloration | Ventilator Dependent | [10 empty] | Death | [5 empty] | VENTILATOR | [22 empty] | Initial submission | [9 empty] | On (b)(6) 2023 around 11:00 am, she took her father to (b)(6) (emergency room) in (b)(6); he was admitted. approximately 20 minutes later the nurse returned, informed her that her father had an elevated tropin indicating a heart problem. he was taken to their catheterization laboratory to have his heart examined. when he returned from the examination, there was an impella embedded in his body (heart) and he was on a ventilator. he was taken to icu. in icu she noticed his feet and arms had changed to the color purple, knees to his elbows. a nurse stayed over night, on noon (b)(6) 2023 the doctor removed the impella and implanted a heart stent. after the surgery, she requested her father be transferred to (b)(6). he was transported to (b)(6) and admitted in icu. he was placed on dialysis machines for one week. on (b)(6) 2023 he passed away.
34368 | 1220648-2023-03376 | 17805250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17805250&pc=OZD | 08/26/2023 | 02/07/2024 |  | 09/24/2023 |  |  | 08/26/2023 | 09/24/2023 |  |  | 07/31/2025 |  | 09/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371580 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 40-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support and the 5.5 support was ongoing for the patient who had the support escalated after eight days to the impella 5.5. the pump was removed on the second day of the impella 5.5 support due to flow saturation. | [5 empty] | The investigation into the reported saturated flow has been completed since the original report was filed. data logs confirmed the failure mode. about 26 hours into the case, the mc spiked to 1200 ma & the flow increased while p level unchanged. this event remained to the rest of the case. visual inspection found biomaterial inside the purge gap, if & of cages. no other issues were found on the rest of the pump. the biomaterial then was removed & pump ran without issue with flow in specification under bench test. the root cause of saturated flow was due to biomaterial ingestion. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34369 | 1220648-2023-03377 | 17805253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17805253&pc=OZD | 08/26/2023 | 07/31/2024 |  | 09/24/2023 |  |  | 08/26/2023 | 09/24/2023 |  |  | 05/31/2023 |  | 09/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365442 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the ct surgeon called regarding explant of device due to hemolysis. the impella was in adequate position. ecmo has had frequent suction for last 24 hours. multiple fluid blouses were given to help with ecmo flows. discussions on causality and troubleshooting took placed with surgeon the possible reasons for the hemolysis but the surgeon was persistent that it was the impella. the impella cp removed in cath lab with manual pressure. | [5 empty] | The investigation for hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34370 | 1220648-2023-03366 | 17804930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804930&pc=OZD | 09/13/2022 | 09/23/2023 |  | 09/23/2023 |  |  | 09/13/2022 | 09/23/2023 |  |  | 04/30/2024 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211194 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported 78yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during a purge cassette change, the purge sidearm broke off. a purge bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause, given previous investigations for the same failure. it was concluded that the cause of the purge leak was sidearm yellow luer damage was as a result of excessive force. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34371 | 1220648-2023-03367 | 17804933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804933&pc=OZD | 09/18/2022 | 09/23/2023 |  | 09/23/2023 |  |  | 09/18/2022 | 09/23/2023 |  |  | 06/30/2024 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234226 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported 65yo male with cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a "purge system open" alarm and leaking noted at the yellow connector. additionally, the sterile sheath was found to be damaged. a purge bypass was performed and the pump was exchanged the next day. there was no patient harm reported. | [5 empty] | The investigation for the reported purge leak and the sterile sleeve damage has been completed. the root cause of the purge leak was determined to be use related as the sidearm was most likely damaged while moving the patient. the exact location of the leak is unknown as no product was returned for analysis. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34372 | 1220648-2023-03368 | 17804934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804934&pc=OZD | 01/30/2023 | 09/23/2023 |  | 09/23/2023 |  |  | 01/30/2023 | 09/23/2023 |  |  | 03/31/2024 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208293 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a delivery issue with the impella 5.5 device due to patient's anatomy. the pump insertion was aborted. there was no harm to the patient. lt. multiple attempts made and the catheter proximal to the motor was kinked. eventually the staff aborted. | [5 empty] | The investigation into the reported delivery issue has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the impella would not advance past the innominate artery. the physician commented on the tortuosity in the vessel as making it more difficult. multiple attempts made and the catheter proximal to the motor was kinked. the surgeon did not think a impella cp would pass either. the root cause of the delivery issue was patient condition related and related to the tortuosity of the patient's vessels making it impossible to advance the pump past the innominate artery.
34373 | 1220648-2023-03369 | 17804935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804935&pc=OZD | 08/27/2023 | 05/02/2024 |  | 09/23/2023 |  |  | 08/27/2023 | 09/23/2023 |  |  |  |  | 09/23/2023 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male in st elevated myocardial infarction was implanted with an impella device for mechanical circulatory support. when changing the purge cassette (pc) the team observed the clip of the automated impella controller (aic) at the pc housing to be broken. the team followed ifu and replaced the aic for continued support. there was no harm or injury to the patient. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the console was tested on an engineering lab bench. the blue purge retainer and the flag were broken / missing, which confirms the reported failure mode. data logs were reviewed finding there were no changes in the purge pressure during ¿fast prime mode¿. this symptom aligns with the reported failure mode, and most likely at this point the purge retainer might have partially broken/dislocated. though the ¿bag & cassette change¿ procedure was completed, the console received a ¿purge system open¿ alarm. user again performed the ¿bag & cassette change¿ procedure. this time, the same cassette was removed and reinserted. then the console was not able to pass through the fast prime mode due to no change in the purge pressure, and so the aic was not able to complete the procedure. at this point of time, purge retainer might be broken/dislocated. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
34374 | 1220648-2023-03370 | 17804936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804936&pc=OZD | 02/27/2023 | 09/23/2023 |  | 09/23/2023 |  |  | 02/27/2023 | 09/23/2023 |  |  | 12/31/2024 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023284027 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/27/2023 |  |  |  |  | The user facility reported 55yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that a "placement signal not reliable" alarm suddenly occurred during impella support, but eventually resolved. several days later, while a heart transplant was performed on the patient, a high purge pressure alarm occurred. it was noted that the clinicians had clamped the pump across the catheter while the device was in surgical mode. however, the patient continued with support and was explanted 2 days later. | [5 empty] | The investigation for the reported placement signal and high purge pressures issues has been completed. data logs showed numerous motor current spikes with simultaneous drops in motor speed occurred in the 3 hours leading up to the high purge pressure event; these events were accompanied by erratic flows and suction alarms. after the motor current spikes increased in frequency, the purge pressure rose for 6 minutes until the "purge pressure high" alarm triggered; which sustained for the remainder of support. visual inspection of the returned product revealed the presence of a significant mass of biomaterial on the impeller and within the purge gap and a significant catheter kink was observed next to the kink protector. in conclusion, it was determined that the cause of the high purge pressure was ingested biomaterial occluding the purge gap. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34375 | 1220648-2023-03371 | 17804937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804937&pc=OZD | 08/26/2023 | 08/05/2024 |  | 09/23/2023 |  |  | 08/26/2023 | 09/23/2023 |  |  | 01/31/2025 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333984 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp easy guide lumen was defective. the wire was unable to feed through the pump; therefore, another device was successfully used. | [5 empty] | The investigation for the delivery issue has been completed. product was not returned for investigation. as per clinical, doctor were unable to load .018 wire through red cheater and red lumen reported to be defective. the cause of the delivery issue was most likely red lumen issue although the red lumen issue is unknown. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34376 | 1220648-2023-03372 | 17804940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804940&pc=OZD | 02/26/2023 | 09/23/2023 |  | 09/23/2023 |  |  | 02/26/2023 | 09/23/2023 |  |  | 10/31/2024 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023268101 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/26/2023 |  |  |  |  | The user facility reported 44yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were multiple high purge pressure alarms on the pump. the staff reported that there was a kink in the white impella cable and the alarms resolved. later the pump stopped three times and the patient became symptomatic. the pump was restarted but it was eventually replaced. | [5 empty] | The investigation for the reported pump stop issue has been completed. the field data log was reviewed. the returned product was visually inspected and a wrap was found around the reported catheter kink next to the red handle kink protector. the wrap was removed and the catheter kink was revealed. the motor was torn down and no blood, biomaterial, or denatured proteins were observed in the motor. the cause of the high purge pressure and resulting pump stop was a kinked purge line in the catheter next to the red handle kink protector. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34377 | 1220648-2023-03373 | 17804941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804941&pc=OZD | 08/26/2023 | 12/07/2023 |  | 09/23/2023 |  |  | 08/26/2023 | 09/23/2023 |  |  | 11/30/2024 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297384 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported while on impella support the motor current gradually increased and then just prior to failure it spike to >700. the pump stopped, the hospital staff attempted multiple times to restart, but were unsuccessful. hospital did not try changing consoles. attending cardiologist instructed intensivist to pull pump back into the descending aorta near ecmo arterial cannula and they will remove impella. patient was on ecmo and the doctor did not feel the need to re-implant impella. | [5 empty] | The investigation into the pump stop has been completed. pump was not returned for investigation. data logs were returned and reviewed. there was a motor current spike observed with the p-level change suggesting biomaterial ingestion. per the clinical information provided and data logs, the root cause of the pump stop issue was most likely biomaterial. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34378 | 1220648-2023-03374 | 17804942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804942&pc=OZD | 08/26/2023 | 01/25/2024 |  | 09/23/2023 |  |  | 08/26/2023 | 09/23/2023 |  |  | 04/30/2025 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347750 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported while the physician was repositioning in the unit, the impella was pulled across valve and unable to re-cross. the patient was brought back to the cath lab and a new impella inserted successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for benchtop investigation. according to clinical information, the impella was pulled out into the aorta during the repositioning of the pump, and the doctor was unable to reposition it afterward. the cause of the positioning issue was improper repositioning technique resulting in the pump pulling out of the ventricle, based on given clinical information.
34379 | 1220648-2023-03375 | 17804943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804943&pc=OZD | 12/23/2022 | 09/23/2023 |  | 09/23/2023 |  |  | 12/23/2022 | 09/23/2023 |  |  | 07/31/2024 |  | 09/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248927 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2022 |  |  |  |  | The user facility reported 63yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were alternating purge pressure high and purge system blocked alarms. no kinks observed in purge line and an x-ray did not show any observable kinks in the implanted catheter. tissue plasminogen activator (tpa) was hung and infused over 4-5 hours and the alarms resolved. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure issue has been completed. no product was returned for investigation. the data log shows that one hour into the case the purge pressure began steadily rising and continued to rise over the course of ten minutes following a motor current spike. the high purge pressure alarm was then triggered. this motor current spike is indicative of biomaterial ingestion. at this time of the spike there were also no changes in p-level. the data log shows that the tpa infusion then resolved the high purge pressure issue. therefore, the cause of the high purge pressure was most likely biomaterial since tpa resolved the issue. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34380 | 1220648-2023-03335 | 17797876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17797876&pc=OZD | 08/25/2023 | 01/23/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 04/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337377 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella 5.5 was attempted to be implanted multiple times, but the patient had tortuous anatomy and the implant was aborted. when the pump was pulled out into an artery, a dissection occurred. no blood products were required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the delivery issues- anatomy and the vascular damage - peripheral vessel has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient had tortuous anatomy and multiple attempts made to place. artery was dissected when pump was pulled out during removal due to tortuous vessels. the cause of the delivery issue was patient condition related with tortuous vessel condition per clinical. the cause of the vascular damage issue was patient condition related with dissection during removal due to tortuous vessels. correction: section h6: the medical device problem code and component code was inadvertently left out in the initial report which has been corrected to this report.
34381 | 1220648-2023-03336 | 17797950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17797950&pc=OZD | 08/25/2023 | 01/23/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 05/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024359389 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender received hemodynamic support with an impella cp smartassist heart pump and extracorporeal membrane oxygenation (ecmo) due to acute myocardial infarction/cardiogenic shock. the automated impella controller (aic) repeatedly alarmed due to impella heart pump priming. low pump flow has been reported that was not resolved by repositioning the impella cp. the impella cp was then removed and the patient remained on ecmo. | [5 empty] | The investigation into the low pump flow has been completed. data & product were not returned for review. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34382 | 1220648-2023-03337 | 17797959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17797959&pc=OZD | 08/23/2023 | 01/25/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 10/31/2023 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023164365 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 52-year-old male in st elevated myocardial infarction was implanted with an impella 2.5 device for mechanical circulatory support. the impella 2.5 was implanted and distal pulses were confirmed. patient was transferred and distal pulses were not found on the doppler. patient was hemodynamically stable, labs were stable, and patient was not on any pressors. the team made the decision to explant the 2.5. distal pulses came back. | [5 empty] | The investigation into the limb ischemia issue has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of limb ischemia was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34383 | 1220648-2023-03338 | 17797969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17797969&pc=OZD | 08/24/2023 | 01/16/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 03/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024333730 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 79-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump suddenly stopped on the fourth day from the start of support. the pump was being weaned at the time and the pump was not exchanged. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The medical center did not return the impella device so no benchtop analysis was possible. based on clinical description and data review, the subject was ingested into the pump could be biomaterial. however, without product return, the root cause of the pump stop was unable to be determined.
34384 | 1220648-2023-03339 | 17798006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17798006&pc=OZD | 08/23/2023 | 07/25/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 05/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363202 |  | 0048-0014 EU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | Retroperitoneal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 65-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had ischemia on the right femoral percutaneous insertion (pci) access site which was unrelated to the impella. the patient also had access site bleeding, retroperitoneal bleeding and vascular damage on the left femoral access site. the device was removed and no additional surgery was required. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-b5 mso review, h6 component code 925 added- d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34385 | 1220648-2023-03340 | 17798069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17798069&pc=OZD | 08/24/2023 | 01/19/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 04/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351041 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella was placed in the left femoral illiac space. attempts were made to advance the catheter and were not successful. pump removed, redo angio demonstrated illiac stenosis and distal aortic disease. abiomed rep advised to replace the device, but a decision was made to flush and reinsert after the long oscar was placed. long oscar placed and left ventricle (lv) was re-wired. the original pump was reinserted and delivered to the left ventricle (lv); flows initiated. suction occurred immediately. pump reposition occurred. no relief of suction, the first pump was removed, and a second pump was inserted for support. | [5 empty] | The investigation for the low pump flow has been completed. the impella device was returned for evaluation. visual inspection found no issues on the pump. the pump ran without issue under bench test. data logs confirmed the low flow when the pump started and remained to the rest of the case that triggered auto suction response and suction alarms. the root cause of low flow was most likely due to patient condition as no issues were found on the pump and the failure mode could not be reproduced. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34386 | 1220648-2023-03341 | 17798124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17798124&pc=OZD | 08/24/2023 | 07/17/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 06/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360213 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed hematoma after peel away sheath was removed and repo sheath advanced. hematoma was easily resolved with light manual pressure, but patient continued to develop hematoma after manual pressure. patient moving contributed to hematoma. the patient did not need any blood products. the impella was removed and hemostasis was successful with placement of angioseal x2 and manual pressure for approximately five minutes. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be patient movement because patient continued to develop hematoma due to patient movement. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34387 | 1220648-2023-03342 | 17798164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17798164&pc=OZD | 08/25/2022 | 09/22/2023 |  | 09/22/2023 |  |  | 08/25/2022 | 09/22/2023 |  |  | 05/31/2024 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223973 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported that a 85-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak at the yellow luer on day 3 of support. a purge bypass was performed. sodium bicarbonate was the purge solution. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was most likely a damaged yellow luer due to sodium bicarbonate in the purge as reported in the clinical account. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34388 | 1220648-2023-03343 | 17798185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17798185&pc=OZD | 08/25/2023 | 09/25/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 12/31/2024 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326233 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported the impella appeared to be in the aorta. the impella was repositioned. echo was performed revealing optimal placement. the patient was on impella support for 70.68 hours before care was withdrawn and the patient expired. no allegations were made against the impella for the cause of death. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03343 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34389 | 1220648-2023-03344 | 17799781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17799781&pc=OZD | 08/23/2023 | 10/27/2025 |  | 09/22/2023 |  |  | 08/23/2023 | 09/22/2023 |  |  | 06/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367327 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 08/23/2023 |  |  |  |  | The complainant also reported that the patient experienced erroneous "position in aorta" alarms. | The user facility reported a 55-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the support was ongoing when the pump alarmed for incorrect positioning. the team found that the pump additionally had a thrombus burden at the pump at explant. additionally, after the pump explant, the team observed the graft tore off and the team had physicians intervene by ligation and giving of 3 units of red blood cells. the team attempted to replace the impella 5.5 but were unsuccessful due to the patient's anatomy. |  |  |  |  | G3 ¿date received by manufacturer¿ corrected. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. b7 other relevant history, including preexisting medical conditions updated. h6 updated to include placement signal coding. d6a / d6b updated d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank. g1 reporting contact email updated. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | D6b "if explanted, give date" corrected. b1 and h1: corrected reportability type. | The investigation into the thrombus, the vascular damage - great vessel, and the positioning issues has been completed. the device was returned for benchtop evaluation and investigation. visual inspection found a kink on catheter right below the mh. no other issues were found on the rest of the pump. the cause of the false position in aorta alarm was patient condition related since position was confirmed via echo and software version was after v11.1. the cause of the thrombus on explant was patient condition related. the cause of the vascular damage- great vessel was patient condition related since the patient had a compromised right aortic arch (raa) and small left atrial appendage (laa).
34390 | 1220648-2023-03345 | 17800671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17800671&pc=OZD | 08/24/2023 | 11/03/2023 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 03/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324302 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 55-year-old male (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, that a hematoma was noted at the impella 5.5 axillary insertion site. the patient was brought to the or for a washout and a unit of packed red blood cells was administered. the patient is stable. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding was not determined since insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34391 | 1220648-2023-03362 | 17801589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17801589&pc=OZD | 08/24/2023 | 07/31/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 06/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360088 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2023 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was oozing from the impella insertion site and femostop was applied. it was further reported that dark red urine was noticed and impella was explanted. the rn stated the patient had nitroglycerine infusion and iv push hydralazine throughout the evening which may have contributed to both conditions. | [5 empty] | The investigation for access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03362: f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34392 | 1220648-2023-03360 | 17802212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802212&pc=OZD | 08/25/2023 | 07/25/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 05/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361575 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2023 |  |  |  |  | The user facility reported an 80-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician placed a 14fr short sheath for a delivery of the cp to the anatomy. the dilator was removed from the sheath and at that time one half of the orange ring fell off and there was some bleeding. the team remedied the issue by removal of the short 14fr and placement of the long 14fr. no lasting harm or injury that prompted further intervention, no blood products noted. | [5 empty] | The investigation for the introducer damage has been completed. the introducer was returned for evaluation. the orange valve was seen to be broken and adhesive residue was found below it suggesting insufficient adhesive was used during manufacturing. data log analysis was not necessary for this complaint. according to the clinical, after the dilator was removed, the 14fr short sheath was inserted. while inserting the 14fr short sheath half of the orange ring fell off and bleeding began to occur. the dilator was then pushed back into place as the sheath was exchanged for the 24fr long sheath. the cause of the mechanical introducer damage was a damaged valve. added-d6a/d6b f6/f8 should have been left blank in the initial report. h6 component code 4761 changed to 481. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34393 | 1220648-2023-03353 | 17802275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802275&pc=OZD | 08/24/2023 | 08/02/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 05/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The patient expired during the procedure. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a 42-year-old male in supraventricular tachycardia was implanted with an impella cp device for mechanical circulatory support. after implant, there was trouble achieving flow due to positioning in the mitral apparatus. the cardiologist attempted to reposition numerous times. patient was dependent on device to unload while on va ecmo. the implella was exchanged and a new impella cp was implanted successfully. |  |  |  |  | The investigation of positioning issues has been completed. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issue was use related. updated with additional information the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03353: d6a and d6b added; f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34394 | 1220648-2023-03363 | 17802440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802440&pc=PYX | 08/24/2023 | 03/04/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 01/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024329022 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 64-year-old female in st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. the rp device was inserted via the femoral and the team made choice to keep the 23fr peel-away in. the retention caused the limb to become ischemic. the team treated the thrombus with thrombectomy, balloon, and lytics. the team then proceeded to have surgery and a fasciotomy. | [5 empty] | The investigation into the reported limb ischemia - reversible has been completed since the original report was filed. per the clinical information provided, the best practices as stated per ifu were not followed and the peel away sheath was not removed after impella implant. the root cause of the limb ischemia issue was use related to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
34395 | 1220648-2023-03346 | 17802486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802486&pc=OZD | 08/24/2023 | 12/14/2023 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 01/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338559 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inadequate User Interface | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 61-year-old male noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. due to the patient's anatomy with left ventricle was more horizontal than normal, there was a concern for the impella catheter being long enough to cross the aortic valve. the physician cut the anti-contamination sleeve off the device which is "off label." the impella cp was weaned and removed successfully after percutaneous coronary intervention. | [5 empty] | The investigation into product damager (sterile sleeve) has been completed. no product was returned. as per clinical upon investigation, the doctor needed approximately 5 more centimeter of length access and had to cut sterile sleeve off the catheter. the cause of the product damage (sterile sleeve) issue was use related with physician cutting down the anti-contamination sleeve during support per clinical review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34396 | 1220648-2023-03358 | 17802533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802533&pc=OZD | 08/25/2023 | 02/14/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 06/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357329 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2023 |  |  |  |  | The user facility reported an 82-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during surgery. the cp was cross-clamped but after removal of the clamp and surgical coronary artery bypass grafting, the pump came across the valve and was unable to be re-positioned. the physician decided risk/benefit not worth putting impella back via other methods. patient was hemodynamically stable. pump removed and insertion site closed with perclose sutures successfully. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the positioning issue was use related due to improper technique when cross clamp removed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34397 | 1220648-2023-03352 | 17802546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802546&pc=OZD | 08/18/2023 | 09/22/2023 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  |  |  | 09/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1594271 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 08/24/2023 |  |  |  |  | The user facility's automated impella controller (aic) had a battery failure alarm during a procedure. the aic was replaced successfully with no patient harm. patient demographics are not available. | [5 empty] | The investigation into the aic battery failure has been completed. the device was returned and evaluated. in addition, data logs were evaluated. the root cause for the battery failure was determined to be single cell failure. the failure modes will be monitored and trended.
34398 | 1220648-2023-03350 | 17802573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802573&pc=OZD | 05/05/2022 | 09/22/2023 |  | 09/22/2023 |  |  | 05/05/2022 | 09/22/2023 |  |  | 12/31/2023 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023171574 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [5 empty] | Unknown |  |  |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/05/2022 |  |  |  |  | The user facility reported that when the impella 5.5 was first plugged into the automated impella controller (aic) during case setup, the "impella defective" alarm appeared. the device was disconnected and inspected, and no issues were found. however, the alarm reappeared when the pump was reconnected to the console. subsequently, the pump was replaced with a new device, which functioned with no issues with the same console. | [5 empty] | The investigation into the reported pump did not start issue was completed. the device and data logs were returned for evaluation. analysis of the data logs confirmed the occurrence of the "impella defective" alarm. visual inspection of the pump did not reveal any damage or abnormalities. functional testing of the pump did not duplicate the reported issue. the device functioned as intended. therefore, the root cause of the pump did not start issue was undetermined as the failure mode was unable to be reproduced. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34399 | 1220648-2023-03354 | 17802645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802645&pc=OZD | 08/07/2023 | 01/22/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 02/28/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024329793 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2023 |  |  |  |  | The user facility in japan reported a patient (unknown age and gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient was on impella 5.5 support for 258.17 hours and the impella was removed to wean the patient off support for heart recovery. the impella operated without issues. the patient expired after the impella 5.5 removal. the date of patient death is unknown, however the patient expired within 30 days of impella explant. the patient cause of death was non-occlusive mesenteric ischemia. the relationship between the impella and the cause of death is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not returned. the cause of the limb ischemia occurred 30 days after the device was removed. there was no relationship between the injury and the use of the device as the device functioned without any issues. the cause of the limb ischemia issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34400 | 1220648-2023-03351 | 17802849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17802849&pc=OZD | 12/28/2022 | 09/22/2023 |  | 09/22/2023 |  |  | 12/28/2022 | 09/22/2023 |  |  | 06/30/2024 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023232905 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/28/2022 |  |  |  |  | The user facility reported that a 69-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. there were no reports of alarms. due to the known malposition, the plan was to replace with a new device the following day and wean off ecmo. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. clinical data confirmed multiple comorbidities the patient had at the time of the reported event. additionally, analysis of data logs revealed positioning alarms. based on the available information provided and no device returned, the root cause cannot be determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34401 | 1220648-2023-03348 | 17803244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17803244&pc=OZD | 08/25/2023 | 11/14/2025 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  |  |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1231456 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the automated impella controller (aic) had a system self-check failure. the aic will be returned for investigation and the ifu followed by the use a backup aic for patient use. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation has been completed. data analysis: preliminary analysis of console logs obtained from the cloud indicated sau_ser6 (eep/epm) communication failure. console logs could not be initially obtained from the device due to file corruption (presumably due to frequent reboots), but after file system was repaired and logs downloaded, more recent entries confirmed initial observations, as the imc logs were flooded with repeating error messages related to epm communication failure: (b)(6) 2023 22:24:41 send_heartbeat_messages: name_open ncc_sau_simul did not succeed err=no such file or directory. (b)(6) 2023 22:24:41 shmemtools:calmain:: timeout wait for communication channel /dev/name/local/sau_sens_eep_1 10sec. (b)(6) 2023 22:24:41 shmemtools:scumain:: timeout wait for communication channel /dev/name/local/sau_sens_eep_1 10sec. (b)(6) 2023 22:24:42 send_heartbeat_messages: name_open ncc_sau_simul did not succeed err=no such file or directory. device analysis: console was tested in danvers and issue was reproduced. persistent inability to communicate with epm subcircuitry on impellatronic had been confirmed. analysis of the microprocessor responsible for epm/eep operation revealed corruption of its software, as evidenced by mismatched content of code extracted from the microprocessor and reference programming code. after re-flashing the microcontroller, issue was resolved. because the exact root cause of the epm software corruption could not be determined, the impellatronic board should be replaced as a precaution. root cause: console¿s self-check failure at boot-up was caused by epm/eep firmware corruption on impellatronic. d2 common device name updated. g1 reporting contact fax number updated. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. h6 type of investigation code added.
34402 | 1220648-2023-03349 | 17803362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17803362&pc=OZD | 08/25/2023 | 03/15/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 04/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336749 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella 5.5 had two aics swapped out due to controller alarms. the team was concerned about the pump itself and planned to swap 5.5 out in or. while in or, a "motor stopped" alarm occurred and the pump stopped. patient was emergently placed on va ecmo. both consoles (ic10546 and ic2927) with controller failure alarms sent in to in for investigation along with impella 5.5 that stopped. | [5 empty] | The investigation into the pump stop and the access site bleeding - minor issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. data logs were returned, and a motor current spike was recorded leading to pump stop. the catheter was kinked as well. per the investigation, the root cause of the pump stop issue was most likely due to open electrical lines from excessive force due to evidence of the catheter kink. the root cause of the access site bleeding ¿ minor was not determined. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34403 | 1220648-2023-03355 | 17803453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17803453&pc=OZD | 08/25/2023 | 08/09/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 06/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370579 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2023 |  |  |  |  | The outcome of the procedure was the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a 60-year-old male in st elevated myocardial infarction was implanted with an impella cp device and an impella rp device for mechanical circulatory support. there was concern for thrombus burden present. the cp was explanted and the rp remained on for hemodynamic support. |  |  |  |  | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank on manufacturer device report 1220648-2023-03355. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34404 | 1220648-2023-03356 | 17803684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17803684&pc=OZD | 08/24/2023 | 01/17/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 05/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369976 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | Aortic Valve Stenosis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 82-year-old male in st elevated myocardial infarction was to be implanted with an impella cp device for mechanical circulatory support. radial diagnostic revealed severe ostial left main (lm). drug eluding stent (des) placed but balloon could not be removed. shaft separation occurred. multiple attempts to snare device. the patient coded due to lack of flow. the impella was prepped, and sheath placed. the patient actively coding while attempting to place the impella. due to severe as (aortic stenosis), impella could not be placed. the patient expired in the procedural area. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the positioning issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely due to patient condition as the patient was coding and has severe as (aortic stenosis).
34405 | 1220648-2023-03359 | 17804095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804095&pc=OZD | 08/25/2023 | 09/19/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 05/31/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362350 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/25/2023 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 40-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella was inserted and impella started in auto, ramping up to 3.8l/min and then return of spontaneous circulation was achieved after a few minutes. right radial access was used for left heart catheterization and a clot was found in the mid left anterior descending artery, 100% occlusive. penumbra was used to aspirate the clot and then intravascular ultrasound (ivus) was used to assess the vessel. a 2.75x20 des was placed and ivus again to ensure stent apposition. eventually, the patient began to decompensated rapidly, requiring escalation of multiple pressors. urinary output ceased. central venous pressure increased from 16 to 25 and impella flows became reduced with constant suction. staff believes the catheter ingested vegetation or a clot. family was made aware and patient was made a dnr and expired after 14.27 hours of impella support. there are no allegations towards the impella for the cause of death. |  |  |  |  | The investigation into the into the low pump flow and the thrombus issues has been completed since the original report was filed. the pump was returned for investigation. a long string of biomaterial was found in the cannula. data logs show low flow while the pump was running on p4, with several suction alarms. the root cause of the low flow issue was biomaterial found in cannula. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34406 | 1220648-2023-03364 | 17804118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804118&pc=OZD | 08/24/2023 | 07/22/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 10/31/2024 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024284929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported an 82-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp repo sheath was continuing to ooze steadily requiring multiple dressing changes in cath lab. abiomed best practice recommendations followed for 14 x 25cm peel away to repositioning sheath exchange except first instance where repo sheath was briefly sutured by cath lab tech towards patient¿s back, oozing occurring. local clinical rep onsite recommended cutting sutures and re-suturing with cranial forward suture, angle match was maintained nicely. then oozing significantly continued after patient helicoptered arrived at upmc presbyterian. no systemic anticoagulation started. bicarb in impella cp purge, total of approximately six units lost within a six hour span. two units of packed red blood cells given and three litres of iv fluids. patient brought to cath lab for removal of cp in stable condition. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site oozing was lack of vessel recoil as the bleeding occurred after switch to gwru and all best practices were followed. physician specifically noted poor vessel recoil. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations. section: entering cardiac output. section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34407 | 1220648-2023-03347 | 17804186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804186&pc=OZD | 08/24/2023 | 03/06/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 02/28/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340951 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 48-year-old female/male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the team made observation of signs of hemolysis. the urine color had changed and the pump was under the papillary muscle. the team transferred the patient to a tertiary center and explanted the pump due to the hemolysis and decision to escalate care to the 5.5 pump. | [5 empty] | The investigation for the reported hemolysis has been completed. the pump was not returned for investigation. the data logs indicate low flow at the start of the case followed by saturated flow. also, several 'impella position unknown' alarms were reported during the case start. the pump was repositioned after initial delivery. the patient had a small left ventricle (lv). the hemolysis resolved after the pump explant. the cause of the hemolysis issue was not determined since product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.
34408 | 1220648-2023-03365 | 17804605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804605&pc=OZD | 08/24/2023 | 01/23/2024 |  | 09/22/2023 |  |  | 08/24/2023 | 09/22/2023 |  |  | 04/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352876 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2023 |  |  |  |  | The user facility reported a 64-year-old female noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the impella co encountered low flows and suction above p2. the physician elected to proceed with the pci. the impella was explanted after the pci procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. data logs confirmed the low flow when pump started & remained to the rest of the case that triggered auto suction response and suction alarms. logs also showed pump was in aortic area for 5 minutes then was resolved. product was not returned for review. based on clinical description & data review, the root cause of low flow was most likely due to patient condition.
34409 | 1220648-2023-03361 | 17804752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17804752&pc=OZD | 08/25/2023 | 08/02/2024 |  | 09/22/2023 |  |  | 08/25/2023 | 09/22/2023 |  |  | 06/30/2025 |  | 09/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364377 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2023 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had severe hemolysis while on impella support. the staff was unclear if it is impella related. the patient was on impella support for 222.53 hours before the patient expired. no allegations were made towards the impella for the cause of death. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. per medical safety officer review use of the device cannot conclusively be associated with the outcome. h6 component code added. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03361. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34410 | 1220648-2023-03311 | 17789241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17789241&pc=OZD | 08/24/2023 | 07/24/2024 |  | 09/21/2023 |  |  | 08/24/2023 | 09/21/2023 |  |  | 03/31/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341708 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 40-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that when explanting the impella cp, unknown tissue was removed (fem intima vs thrombus, sent to pathology). vascular surgery called in to repair. | [5 empty] | The investigation for vascular damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: section: warnings and cautions: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ h6 component code added 925 the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03311: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. h.6 code 4627 and 4625 were reported incorrectly. new code was added to health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34411 | 1220648-2023-03310 | 17789312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17789312&pc=OZD | 08/24/2023 | 09/22/2024 |  | 09/21/2023 |  |  | 08/24/2023 | 09/21/2023 |  |  | 04/30/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361677 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 11 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 11-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella support and experienced a pump stop. family decided to withdraw care, va-ecmo stopped shortly after along with any additional support. | It was reported the patient expired after care was withdrawn. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the pump stop issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the root cause of the pump stop was most likely due to ingested biomaterial. d.4 revised udi number as previously entered incorrectly. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03310 was provided in sections b2, b5, d6a, d6b, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34412 | 1220648-2023-03309 | 17789313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17789313&pc=OZD | 08/23/2023 | 09/19/2024 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  | 12/31/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024304756 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Hemorrhagic Stroke | [11 empty] | Other | Life Threatening | Death | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/23/2023 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. a head ct scan revealed subarachnoid hemorrhages so systemic heparin stopped. plan to restart heparin gtt today and repeat head ct. pt has blood infection with questionable vegetation grown on aortic valve. the patient was being treated with anti-biotics. plan to continue continuous renal replacement therapy and bipella support throughout icu course of care. | Additional information noted that the care was withdrawn, and the patient expired.medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported stroke and infection has been completed since the original report was filed. the device was not returned by the medical facility. data logs were reviewed, and no motor current spike was observed. the cause of the hemorrhagic stroke issue was most likely patient condition related and unknown relation to impella per clinical and data log review. the cause of the infection/sepsis issue was most likely patient condition related and unknown relation to impella based on clinical review. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34413 | 1220648-2023-03308 | 17789314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17789314&pc=OZD | 08/23/2023 | 11/17/2025 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  |  |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162556 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 08/23/2023 |  |  |  |  | A biomed team member reported that the user facility's automated impella controller (aic) had a battery failure. a loaner was sent to the facility and the aic will be returned for investigation. there was no noted patient use at the time of the observation. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data review: after reviewing the imc logs, the reported battery failure issue with (b)(6) was confirmed. there were numerous instances of battery failure alarms (transport connection blown/missing) (event set #360) observed on the reported complaint date. device analysis: console was tested after arriving in field service. the battery failure issue was able to be reproduced. results of running the batttest utility revealed that cell 2 in battery1 had a voltage that was significantly less than the rest of the cells in the battery. lcd screen of battery one was blinking indicating battery error. conclusion: the reported battery failure issue was determined to be caused by internal battery cell imbalance. d4 model number updated.
34414 | 1220648-2023-03312 | 17789345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17789345&pc=OZD | 08/24/2023 | 01/19/2024 |  | 09/21/2023 |  |  | 08/24/2023 | 09/21/2023 |  |  | 05/31/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368031 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that after the patient was taken to the intensive care unit (icu) the patient became hypertensive and oozing around the sheath started. the sheath was advanced and systemic heparin was stopped. manual pressure was held for 30 minutes. the bleeding did not slow down, and the patient was taken back to the catheterization lab where the impella was weaned and successfully explanted. the patient returned to icu in good and stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major has been completed since the original report was filed. based on the clinical account, oozing occurred around the gwru in the intensive care unit. all best practices were reported to be followed. based on this, the cause of the bleeding is likely lack of vessel recoil, as the patient's vessel likely failed to recoil from 14fr to 9fr when the gwru was inserted, despite best practices.
34415 | 1220648-2023-03313 | 17790205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17790205&pc=OZD | 08/24/2023 | 01/15/2024 |  | 09/21/2023 |  |  | 08/24/2023 | 09/21/2023 |  |  | 11/30/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023280820 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 31-year-old patient received support from an impella cp pump and va-ecmo during high-risk percutaneous coronary intervention (hrpci). after about 10 days, the va-ecmo could be removed. after about 12 days of impella support, ct scans of the brain showed multiple small infarcts with bleeding. further investigations revealed that the thrombi that caused the infarcts probably detached from a thrombosed patent foramen ovale. according to the treating physician's assessment, the patient's cerebral hemorrhages did not require any intervention. after approximately 17 days, the patient was weaned off the impella support. the patient was hemodynamically stable, sedated and ventilated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke has been completed. the impella device was not returned for evaluation. however clinical information provided was sufficient to determine the root cause. as per clinical patient had stroke during impella pump and va-ecmo support. the cause of the stroke issue was most likely patient condition related.
34416 | 1220648-2023-03314 | 17790228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17790228&pc=OZD | 09/27/2022 | 09/29/2025 |  | 09/21/2023 |  |  | 09/27/2022 | 09/21/2023 |  |  | 06/30/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247421 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical staff was having difficulty delivering an impella pump. after several attempts they were able to implant the pump, however, the pump was giving intermittent suction. the medical staff was unable to get better position of the pump at that time. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03314 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03314 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03314 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03314.
34417 | 1220648-2023-03315 | 17790785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17790785&pc=OZD | 08/24/2023 | 08/19/2025 |  | 09/21/2023 |  |  | 08/24/2023 | 09/21/2023 |  |  |  |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1150570 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella console had a "controller failure" red alarm. this was the same alarm that occurred on the previously used console for the same impella 5.5 pump. no change in flows and purge. console swapped out for new console. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | Corrections to the initial mfr report: g.1. mdr reporting contact email. d.2. device name has been updated.
34418 | 1220648-2023-03316 | 17791048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17791048&pc=OZD | 11/10/2022 | 03/11/2025 |  | 09/21/2023 |  |  | 11/10/2022 | 09/21/2023 |  |  | 07/31/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023249701 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the md accidentally pulled back the impella in the aorta. the decision was made to attach an additional graft, rewire the impella, and use a new axillary kit to place the impella back into the ventricle. after the impella was restarted, they could not get above p4 without causing suction alarms. the impella position was not optimal, but position was not changed due to the prior surgical intervention to the left ventricle. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34419 | 1220648-2023-03317 | 17794429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17794429&pc=OZD | 08/22/2023 | 01/11/2024 |  | 09/21/2023 |  |  | 08/22/2023 | 09/21/2023 |  |  | 11/30/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274879 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 19-year-old male noted in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had a pulseless, cold leg likely due to clot at access site. the team was discussing options for treatment versus removal. decision was made to attempt distal antegrade reperfusion sheath in right superficial femoral artery (sfa). patient was taken to cardiac catheterization lab (ccl), but the interventional cardiologist(ic) was unable to access vessel successfully. the pump was removed over wire and substituted for 6fr sheath with previously placed perclose devices cinched down. multiple angiograms found occluded flow at the level of the popliteal. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia - reversible has been completed. the product was not returned. per the clinical information provided, the cause of the limb ischemia - reversible issue most likely was patient condition related.
34420 | 1220648-2023-03319 | 17794625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17794625&pc=OZD | 03/21/2023 | 10/19/2023 |  | 09/21/2023 |  |  | 03/21/2023 | 09/21/2023 |  |  | 12/31/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291469 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2023 |  |  |  |  | The user facility in the eu reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. purge sidearm may have been damaged from alcoholic disinfectant. purge sidearm bypass successfully performed and pump remained on for support without harm. | [5 empty] | The investigation into the purge leak issue has been completed since the original report was filed. the medical center in the eu did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the purge leak was sidearm filter damage due to interaction with ipa. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
34421 | 1220648-2023-03321 | 17794699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17794699&pc=OZD | 08/25/2023 | 11/30/2023 |  | 09/21/2023 |  |  | 08/25/2023 | 09/21/2023 |  |  | 06/30/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359704 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient had a hematoma and the staff turned the systematic heparin off. the hematoma became stable. | [5 empty] | The investigation into the patient's access site bleeding (hematoma) has been completed. no product was returned. as per clinical patient had a hematoma at axillary impella site, and all other information was unknown. pressure dressing done to achieve hemostasis and device was not explanted. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34422 | 1220648-2023-03322 | 17794918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17794918&pc=OZD | 08/25/2023 | 01/18/2024 |  | 09/21/2023 |  |  |  | 09/21/2023 |  |  | 07/31/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371828 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported had a hematoma/ bleeding from the access site and the patient received one unit of rbcs. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34423 | 1220648-2023-03323 | 17794924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17794924&pc=OZD | 09/07/2021 | 09/21/2024 |  | 09/21/2023 |  |  |  | 09/21/2023 |  |  | 04/30/2023 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022074266 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Activation, Positioning or Separation Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/07/2021 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical team was having difficulty delivering an impella pump to a patient in need of impella support. once they gained access into the left ventricle (lv) and tried to advance, impella continued to get stuck and would not advance any further. when impella was removed for the 3rd time they saw that the pigtail was almost straight and did not want to attempt again. there was no harm reported to the patient as a result of this incident. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation for the reported delivery issues and introducer damage has been completed. the device was not returned. however, clinical information provided was sufficient to determine the root cause of the delivery issues and introducer damage. according to the clinical data, it was reported that there was a loss of control of the wire and the wire came out of the ascending aorta and had to be removed. during removal, it was not observed that the sheath had folded in on itself and had severely kinked. therefore, the most likely root cause of the delivery issues was user-related. the most likely root cause of the introducer sheath damage was due to sheath kink. added: medical device problem code 2889. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. note: corrected information provided in h6 is an addition to previous coding.
34424 | 1220648-2023-03325 | 17795010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795010&pc=OZD | 10/25/2022 | 09/21/2023 |  | 09/21/2023 |  |  |  | 09/21/2023 |  |  | 05/31/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023219708 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported that a 73-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge side arm leak after a cassette change. sodium bicarbonate was in the purge solution with no extension tubing. the patient subsequently had a purge bypass completed. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the root cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34425 | 1220648-2023-03326 | 17795012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795012&pc=OZD | 11/17/2022 | 09/21/2023 |  | 09/21/2023 |  |  |  | 09/21/2023 |  |  | 07/31/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023248407 |  | 0550-0008 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported that a 75-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. when the patient coughed, the device became displaced into the aorta. the physician attempted to re-access the ventricle without a guidewire, but the pump was unable to re-cross the aortic valve. the decision was then made to explant the pump and replace it with an impella cp. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. after reviewing the clinical information, it was determined that the root cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34426 | 1220648-2023-03318 | 17795051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795051&pc=OZD | 08/22/2023 | 01/24/2024 |  | 09/21/2023 |  |  | 08/22/2023 | 09/21/2023 |  |  | 02/28/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335143 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHARMACIST | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2023 |  |  |  |  | The user facility reported a 70-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. as the md was inserting the peel-away sheath, fluoroscopy was not used and the md met a lot of resistance. once the md visualized the sheath on fluoroscopy imaging, the physician saw the sheath tip doubled back on itself. the sheath was pulled back, advanced the stiff wire further into the artery, and then advanced the peel away into the artery with ease. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed. the device was not returned for benchtop investigation. per the clinical information provided, the root cause of the introducer damage issue was an introducer sheath kink due to patient¿s difficult vasculature (femoral artery).
34427 | 1220648-2023-03327 | 17795258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795258&pc=OZD | 08/23/2023 | 07/16/2024 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  | 05/31/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354136 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2023 |  |  |  |  | The user facility reported a 75-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during transit to the icu, the patient became agitated and, upon arrival, a hematoma was noted in the right femoral artery (rfa) at site of the impella insertion. manual pressure applied. the patient became extremely hypotensive. loss of pulsatility and reduced flows noted. albumin and 2 units of packed red blood cells and normal saline given. after hemostasis was obtained and blood pressure (bp) stabilized, femostop was applied. the site is soft and ecchymosis. the patient remained on impella cp device. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the major access site bleeding was anticoagulation management as acts were measured in the 400s. bleeding was resolved with manual pressure and femstop with blood products as well. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34428 | 1220648-2023-03330 | 17795343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795343&pc=OZD | 08/23/2023 | 07/31/2024 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  |  |  | 09/21/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1404002 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | During a test, the battery error alarm message occurred on the automated impella controller (aic). changing the socket and pressing the transport switch on the bottom of the device did not resolve the error. the defective controller was returned for examination and repair. | [5 empty] | The investigation for the console (aic) battery failure red alarm has been completed. the console (aic) was returned for evaluation. upon functional testing, the console displayed ¿unexpected controller shutdown ¿ alarm¿ event #360 was reproduced by having the battery switch disengaged and turning on aic only with ac power. upon turning on the battery switch, the alarm resolved itself he root cause for the battery failure alarm was due to the use issue. the data logs were reviewed which showed upon initial boot up, the console received the battery failure (transport connection blown/missing) ¿ alarm, event set #360, which aligns with the reported failure mode. after the initial bootup, the aic was most likely shut down by disconnecting the ac power (indicating improper shutdown), as no sequence of shutdown logs was recorded. after the second bootup, in addition to event #360, the console displayed ¿unexpected controller shutdown ¿ alarm¿ event #360 due to the previous improper shutdown. then, the console was booted two times, most likely with the battery switch disengaged, and the aic received the battery failure alarm every time. the root cause for the battery failure alarm was due to the use issue. d2-corrected common device name. h6 med dev prob code 3270 changed to 2885. h8 changed to reuse. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
34429 | 1220648-2023-03306 | 17795547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795547&pc=OZD | 09/12/2022 | 09/21/2023 |  | 09/21/2023 |  |  | 09/12/2022 | 09/21/2023 |  |  | 06/30/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229193 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported that a 77-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a kinked cannula. the physician noticed the catheter portion was kinked upon attempting to implant the pump. the patient was reported to have calcified vessels. another pump was implanted without issue. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported cannula kink was completed. the pump, purge cassette, guidewire, and data stick were returned for evaluation. based on the available information, the root cause of the reported event was patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34430 | 1220648-2023-03328 | 17795592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795592&pc=OZD | 08/23/2023 | 01/15/2024 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  | 04/30/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient started having red urine and the pump migrated back. the surgeon slightly repositioned under formal echo; however, the pump left slightly deep at about 5 cm (pump jumping in and out at 3.5 cm). there were placement signal low alarms and hemolysis. the pump was removed and replaced with intra-aortic balloon pump (iabp). | [5 empty] | The investigation into the positioning issues and the hemolysis has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34431 | 1220648-2023-03333 | 17795604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795604&pc=OZD | 08/24/2023 | 10/07/2025 |  | 09/21/2023 |  |  | 08/24/2023 | 09/21/2023 |  |  | 12/31/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307422 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 65-year-old male noted as high risk coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that during the procedure there was a question as to whether the patient was oozing from the incision site. 1 unit of blood was given, and the patient was transferred to the intensive care unit in stable condition. | The complainant also reported a 'placement signal not reliable' alarm. the placement signal spontaneously increased before disappearing. the waveforms reappeared, temporarily resolving the alarm, but disappeared again shortly after. |  |  |  |  | The investigation into the access site bleeding ¿ major and the optical signal issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the optical signal issue cannot be determined due to lack of product returned. the root cause of the access site bleeding cannot be determined due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-03333 d6a / d6b updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
34432 | 1220648-2023-03329 | 17795630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795630&pc=OZD | 08/23/2023 | 01/31/2024 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  | 05/31/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362395 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2023 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was emergently implanted with an impella cp device for mechanical circulatory support. once the pump was out of the peel-away sheath, it kinked along with 0.18 wire and the physician was unable to advance it further. the patient was hemodynamically unstable, so the pump was completely pulled out of the body without harm to the patient or device. a new impella cp device was implanted through the same access site without any issues. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues- use has been completed. the introducer was not returned for review. the guidewire was returned in multiple bent condition (most likely occurred during packaging to return to abiomed). visual inspection found no issues on the pump. the pump passed od ring gauge test indicating that the pump od was in specification. the root cause of the difficulty inserting pump through introducer was unable to be determined as limited clinical details were provided and the introducer was not returned for review.
34433 | 1220648-2023-03332 | 17795632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795632&pc=OZD | 08/23/2023 | 01/17/2024 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  | 04/30/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352657 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2023 |  |  |  |  | The user facility reported a 56-year-old female/male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp support was ongoing when the team noted an access site hematoma. the patient anatomy of a large pannus contributed to the injury. the team had surgery perform a cutdown and placed a wound vac to the site. the patient remains impella cp support for over 1 week to date. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, per clinical details, the patient had a hematoma due to a large pannus and anatomy complication. the pump was explanted in the operating room and surgery was done to place wound vac to resolve the bleeding. the cause of the access site bleeding issue was patient condition related due to the patient having a large pannus and obesity anatomy leading to bleeding issue per clinical review.
34434 | 1220648-2023-03307 | 17795640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795640&pc=OZD | 09/26/2022 | 09/21/2023 |  | 09/21/2023 |  |  | 09/26/2022 | 09/21/2023 |  |  | 12/31/2023 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170242 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported that a 15-year-old male patient implanted with the impella 5.5 for mechanical circulatory support had a significant amount of clot visualized and removed from the graft during explant of the device. no clot had adhered to the device itself. additionally, during the removal process, the right radial arterial line was lost, causing concern that the clot had traveled down the patient's arm. post-procedure, arterial and venous imaging results were negative for clot and the patient continued to have a pulse. no other concerns were noted. there was no harm reported to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation. based on the available information, the root cause of the positioning issue was determined to be thrombus observed on the pump after device removal. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34435 | 1220648-2023-03331 | 17795690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795690&pc=OZD | 08/23/2023 | 09/21/2023 |  | 09/21/2023 |  |  | 08/23/2023 | 09/21/2023 |  |  |  |  | 09/21/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1180869 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 08/23/2023 |  |  |  |  | Abiomed field service representative reported the user facility's automated impella controller (aic) was cracked and needed repair. the user facility sent in the affected aic for investigation and repair. | [5 empty] | The investigation into the cracked automated impella controller has been completed. the user facility returned the device. the root cause of the crack was use related. the failure modes will be monitored and trended.
34436 | 1220648-2023-03334 | 17795691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795691&pc=OZD | 08/24/2023 | 07/25/2024 |  | 09/21/2023 |  |  | 08/24/2023 | 09/21/2023 |  |  | 04/30/2025 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348283 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was taken back to the catheterization lab for repositioning. the impella was very difficult to reposition due to large papillaries and mitral valve structures. the impella came out of the ventricle all the way across the aortic valve and was unable to be reinserted. a new impella was inserted after removal. | [5 empty] | The investigation for positioning issues has been completed. the impella device was not returned for evaluation. data logs were reviewed, and it was found out that there were "impella in aorta" alarms followed by "impella disconnected while running" alarm. the patient was big in size with big papillary muscles. per the clinical information and data logs, the root cause of the positioning issue was most likely patient movement. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03334: g1 reporting contact fax number missing h.6 code 4629, 2906, and 3038 were reported incorrectly. new code was added to medical device problem code and component code. d6a/d6b added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34437 | 1220648-2023-03320 | 17795884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17795884&pc=OZD | 01/09/2023 | 09/21/2023 |  | 09/21/2023 |  |  | 01/09/2023 | 09/21/2023 |  |  | 11/30/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278443 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/09/2023 |  |  |  |  | The user facility reported 64yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge pressure rose until purge alarm for purge block occurred. changing the purge cassette did not resolve the issue, neither did adding heparin to the purge solution. the physician decided to remove the impella 5.5 to avoid a possible pump stop. patient successfully weaned off impella and stable after surgical wound closure. | [5 empty] | The investigation for the high purge pressure has been completed. the data logs confirm a rapid increase in purge pressure within 5 minutes. during this time there were spikes in motor current. inconsistent placement signal and purge system blocked alarms were triggered. no product was returned so the location of the purge blockage could not be confirmed. the root cause of the high purge pressure was unable to be determined as no product was returned for evaluation. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34438 | 1220648-2023-03324 | 17796016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17796016&pc=OZD | 11/10/2022 | 01/12/2024 |  | 09/21/2023 |  |  | 11/10/2022 | 09/21/2023 |  |  | 08/31/2024 |  | 09/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support, the patient who was being escalated from impella cp to impella 5.5. the patient experienced bleeding at the impella cp access site when the pump was removed. the medical team had to apply manual pressure and give multiple blood products to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that the bleeding occurred after insertion. therefore, the root cause of the access site bleed was due to patient condition.
34439 | 1220648-2023-03273 | 17779685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17779685&pc=OZD | 08/21/2023 | 11/13/2023 |  | 09/20/2023 |  |  | 08/21/2023 | 09/20/2023 |  |  | 05/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024348704 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had lower limb ischemia on the impella side. the doctor reported there was stenosis in the common iliac artery and balloon dilation was performed, but after pci, lower limb ischemia on the impella side resulted in the removal of the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the cause of the limb ischemia issue was patient condition related small vassal since doctor reported narrow vessel condition and limb ischemia resolved after pump removal.
34440 | 1220648-2023-03274 | 17779699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17779699&pc=OZD | 08/22/2023 | 01/23/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 06/30/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371350 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that initial imaging of the groin site was appropriate for impella placement. impella placed in the usual manner without complication. once placement was verified, physician then noted reduced flow past the point of insertion when performing distal perfusion imaging. physician asked for options to increased perfusion. external ipsilateral femoral bypass was discussed and a 5fr retrograde sheath placed. nitro was introduced for patency bypass. physician was pleased with the results noting immediate improvement of color and temperature to distal extremity. | [5 empty] | The investigation into the patient's limb ischemia has been completed. product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition being diseased/small vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34441 | 1220648-2023-03275 | 17779711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17779711&pc=OZD | 08/22/2023 | 04/02/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 03/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347269 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that impella cp positioned in mitral apparatus and the patient was found to be hemolyzing. attempted to wean to p-7 but increased back to p-9. the following day the patient escalated to impella 5.5. the patient is stable. | [5 empty] | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. based on the clinical description and data analysis, the root cause of hemolysis was most likely due to improper positioning of the pump. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34442 | 1220648-2023-03276 | 17779715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17779715&pc=OZD | 08/22/2023 | 01/23/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 05/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024359240 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 69-year-old female patient was treated for acute myocardial infarction/cardiogenic shock and received hemodynamic support with an impella cp smartassist heart pump. access site bleeding occurred after removal of the 14 fr peel-away introducer sheath. the bleeding caused a hematoma to form in the thigh, which pressed distally on the femoral artery and resulted in slightly reduced blood flow to the leg. the impella cp had to be surgically removed. a new impella heart pump was not used. after the hematoma was removed during the surgical procedure, the leg had good blood circulation again. | [5 empty] | The investigation has been completed. the device was not returned for benchtop evaluation and investigation. clinical review stated that a complication hemorrhage was noticed after removal of the pell away sheath. no pre-dilation was done prior to implant. hemorrhage into the thigh with hematoma formation due to which compressions of the femoral artery was observed resulting in reduced perfusion of the leg. the pump had to be surgically removed. after removal/evacuation of the hematoma in the course of the pump explant, the leg was again well perfused and warm (no ischemia). no new pump was implanted. the root cause of the access site bleeding issue was use related to improper technique. the root cause of the ischemia issue was use related to compression of the femoral artery caused by hemorrhage formation during peeling away sheath. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34443 | 1220648-2023-03277 | 17779749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17779749&pc=OZD | 08/21/2023 | 11/02/2023 |  | 09/20/2023 |  |  | 08/21/2023 | 09/20/2023 |  |  |  |  | 09/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1205399 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Operating System Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 56-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the automated impella controller (aic) alarmed "controller failure". the aic was switched out and was running properly. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. product was returned for evaluation. the cause of the aic issue was a defective impellatronics board drawing too much current causing a lockout on the pbm.
34444 | 1220648-2023-03278 | 17780084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17780084&pc=OZD | 01/12/2023 | 09/20/2023 |  | 09/20/2023 |  |  | 01/12/2023 | 09/20/2023 |  |  | 05/31/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023224593 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Dislodged or Dislocated | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2023 |  |  |  |  | The user facility reported that a 41-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the device had gotten pulled back into the aorta when patient went to the restroom. the pump was turned down to p-2 and attempts were made to reposition at bedside without success. the patient was brought to the operating room (or) to attempt to reposition again; however, repositioning was once again unsuccessful. the pump was subsequently explanted. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. analysis of the field logs confirmed the position in the aorta at the end of the run. based on the available information, the root cause of the reported event was patient movement. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34445 | 1220648-2023-03279 | 17780105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17780105&pc=OZD | 01/13/2023 | 09/20/2023 |  | 09/20/2023 |  |  | 01/13/2023 | 09/20/2023 |  |  | 09/30/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023257464 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/13/2023 |  |  |  |  | The user facility reported that a 66-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. after insertion, the device flows decreased. there were no issues with left ventricular placement. suddenly, the left ventricular wave form became significantly higher than ao wave form, then it equalized. per clinical staff, the device was going to force stop/saturated flows. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. visual inspection of the pump visually revealed a big chunk of biomaterial around impella. based on the available information, the root cause of the low pump flows was biomaterial ingestion. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34446 | 1220648-2023-03280 | 17780130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17780130&pc=OZD | 12/07/2022 | 09/20/2023 |  | 09/20/2023 |  |  | 12/07/2022 | 09/20/2023 |  |  | 07/31/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252300 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported that a 67-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced constant suction with the device prompting removal/replacement. there were no issues reported with the replacement pump. no further information or clinical details were provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. based on the available information, the root cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34447 | 1220648-2023-03281 | 17781288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17781288&pc=OZD | 08/22/2023 | 09/22/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 03/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343773 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/22/2023 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp began to have high purge pressure and purge pressure alarms. pump was replaced with the impella 5.5. there was no patient harm. | The procedure outcome was that patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the high purge pressure has been completed since the original report was filed. the product was returned and investigated. the logs were evaluated. data logs shows there was a gradual rise in the purge pressure and while product investigation, biomaterial was found in the outflow cage and purge gap. the root cause of the issue was biomaterial found during return product investigation and data logs. the instructions for use states the following for high purge pressure: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ h6 impact code: change from 2199 to 4629. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03281 was provided in sections b1, b5, d6, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34448 | 1220648-2023-03282 | 17781368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17781368&pc=OZD | 01/01/2023 | 09/20/2023 |  | 09/20/2023 |  |  | 01/01/2023 | 09/20/2023 |  |  | 09/30/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023261069 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 64-year-old female was placed on impella 5.5 and ecmo support. a day post impella placement, there was an access site bleed that prompted 1 unit of rbcs to be given for the postop surgical patient. | [5 empty] | The investigation into the reported access site bleeding has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that root cause of the issue was patient condition related. instructions for use are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34449 | 1220648-2023-03283 | 17781905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17781905&pc=OZD | 08/22/2023 | 01/18/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 03/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP - JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024333882 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 36-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp and ecmo support ongoing when there was cardiac tamponade. the team was unsure of the cause/etiology of the bleeding effusion. the bleed was due to either the cpr, intervention, and/or anticoagulation. the impella was successfully explanted eight days after original implant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. impella was removed to reduce heparin dose as the patient was prone to bleeding. the cause of thrombocytopenia was due to patient condition related.
34450 | 1220648-2023-03284 | 17782092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17782092&pc=OZD | 08/22/2023 | 08/09/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 05/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the staff was unable to remove the guidewire from the impella cp pump after placing the pump in the left ventricle. a new pump was inserted successfully, without any harm or injury. | Following the pump exchange, the patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of product damage (guidewire damage - peripheral vessels) has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was not determined since product was not returned and sufficient clinical information was not provided. b5 updated information provided. d6a and d6b added. g1 reporting contact fax number missing on manufacturer device report 1220648-2023-03284.
34451 | 1220648-2023-03285 | 17782148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17782148&pc=OZD | 05/08/2021 | 01/26/2024 |  | 09/20/2023 |  |  | 05/08/2021 | 09/20/2023 |  |  | 02/28/2023 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022047529 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2021 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella 5.5 support experienced access site bleeding. the patient had to receive two units of blood. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was not returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34452 | 1220648-2023-03286 | 17782229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17782229&pc=OZD | 05/04/2021 | 09/16/2024 |  | 09/20/2023 |  |  | 05/04/2021 | 09/20/2023 |  |  | 02/28/2023 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054981 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/04/2021 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had a computed tomography¿¿(ct) scan which showed that the patient had experienced a brain bleed. the pump remained on for support of over 51 hours til care was withdrawn. causality of the brain bleed was not known. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the stroke/cva issue was patient condition related, there was no spike in the motor current in the data logs and the relation to the impella is unknown. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03286 was provided.
34453 | 1220648-2023-03287 | 17784987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17784987&pc=OZD | 08/21/2023 | 01/12/2024 |  | 09/20/2023 |  |  | 08/21/2023 | 09/20/2023 |  |  | 11/30/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295460 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the "purge lumen blocked" alarm and "purge pressure high" alarm triggered. changing purge cassette did not resolve the issue. the patient was escalated to an impella 5.5 device. | [5 empty] | The investigation into the access site bleeding ¿ minor and the high purge pressure has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the access site bleeding issue was use related with introducer sheath not peeled away completely outside of the body per clinical review. the root cause of the high purge pressure issue was most likely a kinked purge line, per data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34454 | 1220648-2023-03289 | 17785194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785194&pc=OZD | 08/22/2023 | 01/24/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 04/30/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352031 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2023 |  |  |  |  | The user facility reported a 75-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the device was providing suboptimal flows and a kink on the upper end of the cannula was reported. the impella cp was immediately removed and replaced with another impella cp. there was no patient harm. | [5 empty] | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. data logs were not either for review. based on clinical description, the root cause of low flow was due to cannula kink. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34455 | 1220648-2023-03288 | 17785207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785207&pc=OZD | 08/21/2023 | 02/07/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 05/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363007 |  | 0048-0014 EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2023 |  |  |  |  | The user facility reported the impella cp pump would not show any waveforms after implant. the team used the same console, but a new impella cp pump and there were no further issues. no patient details have been provided and there is no mention of patient injury. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported optical signal issue has been completed since the original report was filed. the root cause for the optical signal issue is most likely due to biomaterial on the sensor causing the waveforms to not accurately display. the product was returned and all the parameters were within specification. the pump was soaked and the placement signal waveform was able to be displayed on the console in the lab.
34456 | 1220648-2023-03290 | 17785400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785400&pc=OZD | 08/24/2023 | 01/18/2024 |  | 09/20/2023 |  |  | 08/24/2023 | 09/20/2023 |  |  | 03/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342642 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that at the access site, a covered stent was deployed in the subclavian artery. the irregualr access site was necessary as the femoral artery would not accomodate the impella. the femoral artery was calcified. hemostasis was achieved with the covered stenting. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported vascular damage has been completed. the root cause of vascular damage was unable to be determined as limited clinical details were provided and no product was returned for review.
34457 | 1220648-2023-03291 | 17785419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785419&pc=OZD | 08/23/2023 | 08/15/2025 |  | 09/20/2023 |  |  | 08/23/2023 | 09/20/2023 |  |  |  |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1729761 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a controller failure, red alarm occurred. the nurse swapped out the automated impella controller (aic) without issue. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | D.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03291 was submitted. d.4 model number and serial number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03291 was submitted. e.1 revised address line1, us city and us zip code as they were entered incorrectly on the initial manufacturer device report number 1220648-2023-03291. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-03291 in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03291 was submitted. h.6 codes 2199 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2023-03291 in accordance with updated procedures. h.8 revised usage of device as it was reported incorrectly on initial manufacturer device report number 1220648-2023-03291.
34458 | 1220648-2023-03292 | 17785478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785478&pc=OZD | 08/23/2023 | 01/23/2024 |  | 09/20/2023 |  |  | 08/23/2023 | 09/20/2023 |  |  | 07/31/2023 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022129469 |  | 005062-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male noted as high risk percutaneous cardiac intervention was implanted with an impella 5.0 device for mechanical circulatory support. the 5.0 support was ongoing when on day 34 of the support there was a purge leak noted a the extension line. the pump leak prompted the team to replace the pump. the team replaced the 5.0 with a 5.5 pump. care was escalated. | [5 empty] | The investigation has been completed since the original report was filed. the device was returned for investigation and was tested upon return. data logs confirmed the failure mode. visual inspection found a crack in the yellow luer of pss. no leak was found from yellow luer under leak test. pump ran without issue under bench test with pp/pf in normal range. the root cause of the purge leak was most likely due to a crack in the yellow luer. this occurred after 34 days of run time, which was well beyond the design life of 28 days. | The impella device was received from the customer and the investigation of it is ongoing a this time. upon investigation closure, a supplemental mdr will be filed.
34459 | 1220648-2023-03293 | 17785486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785486&pc=OZD | 05/25/2023 | 10/14/2024 |  | 09/20/2023 |  |  | 05/25/2023 | 09/20/2023 |  |  | 10/31/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290641 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2023 |  |  |  |  | The user facility reported a 54-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during attempted repositioning of the impella cp pump that was positioned too deep, the device kinked just distal to the outlet. additional troubleshooting was done in an attempt to relieve the kink however the issue remained. the pump was subsequently removed and replaced as a result of the noted kink. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was returned for evaluation. returned complaint pump sn(b)(6) visually inspected. cannula kink distal to the outlet observed. no other abnormality found. root cause of the kink and positioning issue was use issue related due to improper positioning. the failure modes will be monitored and trended. | The investigation has been completed for the reported issue of positioning issue. the device was received for evaluation. cannula kink distal to the outlet was observed. no other abnormality found. root cause of the kink and positioning issue was use related due to improper technique. additional information for the initial mfr (b)(4) was provided in section d6a, d6b, h1, and h6. this report is being filed as part of a retrospective review of historical records
34460 | 1220648-2023-03294 | 17785497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785497&pc=OZD | 08/23/2023 | 12/01/2025 |  | 09/20/2023 |  |  | 08/23/2023 | 09/20/2023 |  |  |  |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1699285 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 08/23/2023 |  |  |  |  | The user facility's automated impella controller (aic) failed at case start-up. the aic failed to recognize the the pump and was not priming the impella device. the team followed ifu recommendations and utilized a backup controller for use. no harm or injury. | [5 empty] | The impella console has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email . d2 device name has been updated . d4 model number has been updated. f6 and f8 should have been left blank. h6 type of investigation code added.
34461 | 1220648-2023-03295 | 17785521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785521&pc=OZD | 03/17/2023 | 09/20/2023 |  | 09/20/2023 |  |  | 03/17/2023 | 09/20/2023 |  |  | 12/31/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292143 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Increase in Suction | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/17/2023 |  |  |  |  | The user facility reported a 77yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was continuous suction and a low pump flow during impella support. the pump was explanted as a result of the noted issue and thrombi were noted on the inflow and outflow areas. the patient was placed on a venoarterial extracorporeal membrane oxygenation (va-ecmo) for ongoing support. there was no patient harm reported. | [5 empty] | The investigation for the reported low flow and thrombus formation has been completed. the device was not returned for evaluation. however, review of the data confirmed the failure mode. about 48 hours into support, flow decreased to 1.8 l/min that triggered suction response & suction alarms. low flow remained to the rest of the case. no other issues were found on the logs. based on clinical description, the root cause of low flow was most likely due to biomaterial ingestion and the root cause of thrombus was most likely due to patient condition. the instructions for use provides details on how prevent thrombus formation in the pump. it states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34462 | 1220648-2023-03297 | 17785524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785524&pc=OZD | 05/31/2023 | 09/20/2023 |  | 09/20/2023 |  |  | 05/31/2023 | 09/20/2023 |  |  | 08/31/2024 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023260988 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 23 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/31/2023 |  |  |  |  | The user facility reported a 23-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was accidentally pulled back into the aorta during a procedure for an unrelated issue. the pump was subsequently removed and replaced with another impella cp pump for ongoing support. there was no patient harm. | [5 empty] | The investigation into the positioning issue has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. based on clinical description, the root cause of positioning issue was most likely due to use issue because the heart was elevated during the procedure. the failure modes will be monitored and trended.
34463 | 1220648-2023-03296 | 17785525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785525&pc=OZD | 07/01/2023 | 09/23/2023 |  | 09/20/2023 |  |  | 07/01/2023 | 09/20/2023 |  |  | 05/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345522 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed a hematoma around the impella access site and coincidingly experienced a drop in hemoglobin levels. the patient was administered one unit prbc to counteract the condition. there was no lasting patient harm. | [5 empty] | The investigation into the reported "access site bleeding" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the patient developed a hematoma around the impella access site and coincidingly experienced a drop in hemoglobin levels. the patient was administered one unit prbc to counteract the condition. there was no lasting patient harm. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34464 | 1220648-2023-03298 | 17785547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785547&pc=OZD | 08/22/2023 | 01/25/2024 |  | 09/20/2023 |  |  | 08/23/2023 | 09/20/2023 |  |  |  |  | 09/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/23/2023 |  |  |  |  | The user facility reported a 58-year-old male in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the team noticed the purge cassette was leaking. the nurse changed the purge cassette and the leak persisted. the nurse changed the automated impella controller and the leaking was resolved. the user facility will return the aic for investigation and maintenance. there was no patient harm related to this event. | [5 empty] | The investigation into the aic - purge issue ¿ other has been completed since the original report was filed. the console was returned and inspected for visible damage along the purge pressure sensor area with no sign of damage. a known good pump and purge system was ran on the console without any noticeable malfunction. the root cause for the issue was unable to be determined due to not being able to reproduce the failure. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34465 | 1220648-2023-03299 | 17785564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785564&pc=OZD | 09/15/2021 | 09/20/2023 |  | 09/20/2023 |  |  | 09/15/2021 | 09/20/2023 |  |  | 06/30/2023 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022106264 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/15/2021 |  |  |  |  | The user facility reported a 71yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was breakage of the sterile sleeve that occurred during use of the device. there was no patient harm reported. patient remained on support until weaned off. | [5 empty] | The investigation for the reported damaged sterile sleeve has been completed. the device was not returned for evaluation. therefore, the exact root cause of the reported issue could not be determined. the instructions for use has provided information on how to handle the device. it states: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34466 | 1220648-2023-03300 | 17785571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785571&pc=OZD | 06/05/2023 | 09/20/2023 |  | 09/20/2023 |  |  | 06/05/2023 | 09/20/2023 |  |  | 01/31/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024307793 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Advance | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported a 51-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device had ongoing suction issues during patient's impella support. repositioning was attempted to troubleshoot the issue, however the pump was pulled into the aorta and would not re-cross the aortic valve. the pump was subsequently removed and replaced for continued support. there was no patient harm. | [5 empty] | The investigation for malfunctions "low pump flow" and "positioning issues" has been completed. the user facility returned the impella products for evaluation. in addition, data logs were provided to aid in the investigation. low pump flow malfunction: visual inspection found biomaterial in ph & wrapped around the impeller. no other defects were found on the rest of the pump. positioning issues: the root cause of the positioning issue was determined to be use related as the pump was pulled back across the aortic valve during repositioning. the failure modes will be monitored and trended.
34467 | 1220648-2023-03301 | 17785579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785579&pc=OZD | 08/23/2023 | 01/16/2024 |  | 09/20/2023 |  |  | 08/23/2023 | 09/20/2023 |  |  | 04/30/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353520 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 31-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced a retroperitoneal (rp) bleed. one unit of blood was given, and systemic heparin was on hold. unknown whether rp bleed was related to the impella use. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34468 | 1220648-2023-03302 | 17785621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785621&pc=OZD | 07/07/2023 | 01/22/2024 |  | 09/20/2023 |  |  | 07/07/2023 | 09/20/2023 |  |  | 04/30/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported an 83-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the physician noted a kink in 14/25cm sheath on fluro prior to implant. they attempted to implant and impella motor would not cross. the physician exchanged sheath for 14/13cm and impella crossed and implanted without incidence. | [5 empty] | The investigation into the product damage (sterile sleeve damage) issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, sheath kink was noted on 14x25 introducer, and it was replaced with short 14x13 introducer. the patient had known tortuous vessel condition. the root cause of the introducer issue was patient condition related to torturous vessel. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34469 | 1220648-2023-03303 | 17785624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785624&pc=OZD | 12/02/2021 | 09/16/2024 |  | 09/20/2023 |  |  | 12/02/2021 | 09/20/2023 |  |  | 07/31/2023 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130633 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/02/2021 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient who had transitioned from impella cp to impella 5.5 was bleeding from the impella 5.5 access site and the systemic anticoagulation had to be stopped. two days after, the patient's condition did not improve and decision was made to withdraw care. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleed was unable to be determined as insufficient clinical information and product was provided for the investigation. | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemen al mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03303 was provided.
34470 | 1220648-2023-03304 | 17785647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785647&pc=OZD | 08/22/2023 | 03/19/2024 |  | 09/20/2023 |  |  | 08/22/2023 | 09/20/2023 |  |  | 06/30/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024360362 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | Activation, Positioning or Separation Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the impella 5.5 was placed direct post coronary artery bypass graft (cabg). pump was inserted successfully and there were no issues weaning off bypass and on to impella support. upon chest cavity closure, patients flows dropped. volume was given and pump was repositioned under echo. surgeon found that most adequate flows were achieved when pump was deep and opted to leave pump. upon repositioning, flows immediately saturated at 4.3/4.3l on p4 and purge pressure dropped to 2cc/hr. patient became hemodynamically unstable and position alarms became constant despite pump being in good position. impella proceed to briefly stop twice. decision was made to emergently take patient to or for explant once in or, pump stopped two additional times but restarted immediately. after visualizing pump post explant, a large thrombus was noted in the inflow cage as well as around the motor housing. unclear whether pump ingested this clot in or in ccu upon repositioning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported positioning issue, pump stop, and thrombus has been completed. the pump was returned for investigation. significant biomaterial was observed on inlet cage, cannula and impellor. no other physical abnormalities were observed. during testing, the pump ran successfully after removal of biomaterial. data logs shows the pump stop at the end of case with motor current high alarm but then it is able to restart and support resumed until it stops again and is not able to be restarted. the cause of the temporary pump stop issue was most likely due to biomaterial. the cause of the positioning issue was most likely use related based on provided clinical information.
34471 | 1220648-2023-03305 | 17785648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17785648&pc=OZD | 07/26/2023 | 01/23/2024 |  | 09/20/2023 |  |  | 07/26/2023 | 09/20/2023 |  |  | 06/30/2025 |  | 09/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364871 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2023 |  |  |  |  | The user facility reported a 52-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, the impella 5.5 pump came out of position. the locking mechanism was malfunctioning. the pump was replaced and the second impella 5.5 pump successfully supported the patient for two weeks until transplant decisions were made. | [5 empty] | The impella device was received from the customer and an evaluation of the device is pending. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus and low pump flow issues has been completed since the original report was filed. the device was returned for investigation. the returned complaint 5.5 s2 pump was visually inspected. catheter kink observed at line 45. no issue with catheter lock mechanism. no other abnormality observed. the cause of the positioning issue was use related that led to catheter kink and the catheter lock was unable to pass the kinked area.
34472 | 1220648-2023-03241 | 17773507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17773507&pc=OZD | 08/20/2023 | 11/06/2025 |  | 09/19/2023 |  |  | 08/20/2023 | 09/19/2023 |  |  | 06/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367307 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/20/2023 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient had been on impella 5.5 support doing well and showing signs of recovery. the team noted during the support, the bicarbonate purge was removed to do a tissue plasminogen activator (tpa) due to high purge pressure and low purge flow. the patient was starting to be weaned from 5.5 support a couple hours later pump stop occurred. controller was exchanged but the pump was unable to restart. the physician was called and removed the impella 5.5 at bedside. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. per the ifu: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | As noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.6a and d.6b added since the initial submission of manufacturer device report number 1220648-2023-03241 in accordance with updated procedures. e.1 fax number was inadvertently left off the previously submitted manufacturer device report number 1220648-2023-03241 and was added. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-03241 and should not have been. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-03241 in accordance with updated procedures. h.6 codes 2199, 10 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2023-03241 in accordance with updated procedures. h.10 updated instructions for use since the initial submission of manufacturer device report number 1220648-2023-03241 in accordance with updated procedures. | The investigation of pump stop and high purge pressure has been completed. the pump was returned for investigation. data logs were returned and analyzed. high purge pressure: the data logs show that the purge flow and pressure were stable for about 88 hours until there was a gradual rise in purge pressure. this gradual rise was about 12 minutes log to trigger high purge pressure alarms. during this timeframe there were no changes in motor current or pump flow. high purge pressure lasted till the remainder of the case which was about 48 hours after. the returned product was evaluated and it was seen that the high purge pressure alarm was reproduced when the pump was purged in the lab. there was no blood in the purge line, no biomaterial around the impeller blades or in the cannula, with no catheter or cannula kinks. biomaterial was found in the purge gap. the root cause of the high purge pressure was determined to be biomaterial buildup found in the purge gap based on the product evaluation and data log analysis. pump stop: the data logs show that there was high purge pressure and low purge flow with no changes in motor current or pump flow up until the pump stopped. there were #115 and #13 pump stop alarms triggered on the 20th. the returned product was evaluated and there were open electrical lines. the red handle was opened and it revealed all of the motor cable lines to be snapped on the catheter side, and the pump would not start. based on the data log analysis and returned product, the root cause of the pump stop is most likely due to open electrical lines likely due to excessive force.
34473 | 1220648-2023-03242 | 17773577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17773577&pc=OZD | 12/22/2022 | 03/01/2024 |  | 09/19/2023 |  |  | 12/22/2022 | 09/19/2023 |  |  | 08/31/2024 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023276913 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2022 |  |  |  |  | The user facility reported a 51-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a significant large post infarct left ventricle (lv) pseudo aneurysm, while supported on impella 5.5 as bridge to ventricular assist device (vad). the following day, the patient was brought to or for impella 5.5 removal. | [5 empty] | The investigation into the reported stroke/cva has been completed since the original report was filed. since enough clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34474 | 1220648-2023-03243 | 17773600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17773600&pc=OZD | 08/20/2023 | 07/24/2024 |  | 09/19/2023 |  |  | 08/20/2023 | 09/19/2023 |  |  | 05/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356694 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2023 |  |  |  |  | The user facility reported an 81-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support prior to cardiac surgery. the physician incorrectly removed peel-away and created a significant bleed around repo sheath. the bleeding was managed when impella was temporarily removed and a 16fr sheath was placed. the decision was made by the physician to re-insert the same impella against advice to open new pump. | It was reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | B.5 additional information was received, and abiomed's medical safety officer review was completed. the investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be use issue because the physician incorrectly removed the peel-away causing the bleed. hemostasis was achieved by temporarily removing the pump and 16 french sheath inserted and pump reinserted. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 added information as it was previously entered incorrectly on initial submission of manufacturer device report number 1220648-2023-03243. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-03243 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-03243 and should not have been. h.6 added code 925 since the initial submission of manufacturer device report number 1220648-2023-03243 in accordance with updated procedures. h.6 code 4641 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2023-03243. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2023-03243 . a new code has been added to type of investigation codes h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-03243 in accordance with updated procedures. additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2023-03243 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
34475 | 1220648-2023-03244 | 17773646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17773646&pc=OZD | 10/14/2022 | 10/08/2025 |  | 09/19/2023 |  |  | 09/15/2023 | 09/19/2023 |  |  | 03/31/2023 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022057810 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Air Embolism | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 09/15/2023 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella cp was being transitioned to an impella 5.5. the medical staff had difficulty with delivering the impella 5.5 due to the patient's tortuous anatomy and the vascular surgeon was called for help. the artery finally exposed and it was of smaller caliber of roughly 6 mm but resulted in bleeding requiring multiple blood products. | There was note of air embolism that occurred. while inserting the device, air was noted in both the right and left ventricles. no health impact was noted as a result of this and no intervention was performed. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | B(5): included note of air embolism that occurred. while inserting the device, air was noted in both the right and left ventricles. no health impact was noted as a result of this and no intervention was performed. h(6) added appropriate health effect coding. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation of delivery issues, access site bleeding, and embolism has been completed. the impella device was not returned for evaluation. delivery issue: the physician experienced difficulty exposing right axillary artery since the pocket is very deep. artery finally exposed and of smaller caliber of roughly 6 mm but resulted in bleeding requiring multiple blood products. no vessel disease and artery soft, but vessels were noted to be tortuous. the physician advised of possible complications and difficulties of smaller caliber. the root cause of the delivery issue was most likely due to the patient condition i.e., small/tortuous vessels and deep pocket. access site bleeding: right axillary artery finally exposed and of smaller caliber of roughly 6 mm but resulted in bleeding requiring multiple blood products. graft anastomosis made to artery with great difficulty and took over an hour to obtain hemostasis the cause of the access site bleeding was most likely due to the patient condition i.e., small vessels/ deep pocket. embolism: air was noted in the ventricle; however, no intervention was performed for this. the cause of the embolism could not be determined since no product was returned and insufficient clinical details were provided. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella &section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ b3 date of event was erroneously entered on the initial manufacturer device report number. d6a and d6b revised from the initial manufacturer device report in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 health effect - clinical code 4438 and medical device problem code 2682 was inadvertently omitted on the initial manufacturer device report number. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number.
34476 | 1220648-2023-03245 | 17774485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17774485&pc=OZD | 08/20/2023 | 05/13/2024 |  | 09/19/2023 |  |  | 08/20/2023 | 09/19/2023 |  |  | 06/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367332 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2023 |  |  |  |  | The user facility reported a 62-year-old female with heart failure was implanted with an impella 5.5 for mechanical circulatory support as escalation of care from an intra-aortic balloon pump. the patient had impella 5.5 implanted around 24 hours when the purge system started leaking. dripping fluid was observed coming from the plastic casing around the purge sidearm. specifically, the dripping is coming from the connection between the clear plastic case and the white plastic just distal. inside of the clear plastic case, it is difficult to determine where exactly the leak is. drops were noted on the blue diaphragm and on the connection between the diaphragm and the filter. the drip is continuous, but slow. no changes have been noted with the purge pressure, flows or motor current. there has been no obvious cause of the leak (bumping into anything, etc). the team is aware of the leak. care for the patient continued as normal, with an emphasis on watching purge flow, purge pressure and motor current trends. | [5 empty] | The investigation into the purge leak ¿ pump issues has been completed since the original report was filed. the pump was returned and evaluated. the pump was purged with a purge cassette from the lab for 15 minutes and it maintained steady purge pressure and no leaks from the sidearm were noted. the sidearm was opened and cut and the reservoir was purged with a backpressure to be able to identify any leaks. no leaks were noted or any damaged. the leak could not be reproduced. the root cause of the purge leak - pump is most likely due to use. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34477 | 1220648-2023-03246 | 17774522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17774522&pc=OZD | 08/20/2023 | 01/23/2024 |  | 09/19/2023 |  |  | 08/20/2023 | 09/19/2023 |  |  | 05/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365487 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2023 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. when converting to repositioning sheath, groin began to bleed without control. decision made to remove the impella cp because the patient was hemodynamically stable, had complete revascularization and was showing evidence of anoxic brain injury. the patient was on angiomax and not heparin. the physician elected to remove the pump and use a strategy of manual pressure with a femostop. | [5 empty] | The investigation into the access site bleeding issue has been completed. the device was not returned for investigation. product was not returned for review. based on clinical description, the patient began to bleed from repo unit/ introducer hub valve when converted to repositioning sheath. the pump was explanted, and manual pressure held with fem-stop in place. the cause of the access site bleeding issue was not determined since product was not returned with location of bleeding reported from repo unit/ introducer hub valve per clinical. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34478 | 1220648-2023-03247 | 17774820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17774820&pc=OZD | 08/20/2023 | 01/25/2024 |  | 09/19/2023 |  |  | 08/21/2023 | 09/19/2023 |  |  |  |  | 09/19/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1257649 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported they were setting up a procedure for impella cp insertion when the automated impella controller (aic) monitor froze up for some moments. there was no problem with impella durng this time. there was no patient harm. after turning the monitor back on, there were no further issues. | [5 empty] | The investigation into the aic - shutdown or restart issue has been completed since the original report was filed. the console was returned and tested after arriving in fs. no issue was found in the console hardware, unable to reproduce the reported failure. the root cause for the frozen screen was not determined due to the inability to reproduce. added d9-device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. follow up with the field service engineer has not been completed. upon investigation closure, a supplemental mdr will be filed.
34479 | 1220648-2023-03248 | 17774900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17774900&pc=PYX | 08/19/2023 | 10/08/2024 |  | 09/19/2023 |  |  | 08/21/2023 | 09/19/2023 |  |  | 01/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024310608 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 50-year-old female with heart failure was implanted with an impella rp for mechanical circulatory support. during support, the impella had a "self resolving pump stop" alarm and the impella stopped when rolling the patient. alarm resolved itself without having to manually be restarted. there was no patient harm. | [5 empty] | The impella device has been received from the customer, however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. per the ifu: "unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen." | The investigation for the pump stop has been completed. according to the clinical, a high motor current pump stop occurred and resolved itself shortly afterward. this alarm was recorded when the patient was rolled to remove the bed pan. no other issues occurred and the patient was weaned off of support two days later. product evaluation confirmed there was no issue with the device. after running the pump in the lab no pump stop alarm was reproduced. data logs confirm a high motor current pump stop resolves itself seconds after it is recorded. the cause of the temporary pump stop was not determined because not enough clinical information was given. corrections to the initial mfr report:
34480 | 1220648-2023-03249 | 17775005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17775005&pc=OZD | 10/19/2022 | 03/11/2025 |  | 09/19/2023 |  |  | 10/19/2022 | 09/19/2023 |  |  | 07/31/2024 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023254819 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2022 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had the pump be caught up and pulled into the aorta. attempts to reposition with echocardiogram (echo) guidance failed. decision was made to return to or to replace 5.5. during induction, patient went into vf arrest. extracorporeal cardiopulmonary resuscitation (ecpr) performed and placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | No product was returned. the root cause of the positioning issue was most likely use related, attempts to reposition with echo guidance as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
34481 | 1220648-2023-03250 | 17775123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17775123&pc=OZD | 11/30/2022 | 09/19/2023 |  | 09/19/2023 |  |  | 11/30/2022 | 09/19/2023 |  |  | 07/31/2024 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023250373 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2022 |  |  |  |  | The user facility reported a 41yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient experienced hemolysis despite adequate positioning and giving additional preload. a left ventricle thrombus was observed. the patient remained on support. | [5 empty] | The investigation for the hemolysis has been completed. the pump was returned for evaluation. a visual inspection of the pump revealed biomaterial on the impeller. after soaking the pump in warm water the biomaterial remained on the impeller. no other abnormalities were observed. the cause of the hemolysis was biomaterial. the instructions for use for the event states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ it also states ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ it also states the following as it relates to thrombus: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34482 | 1220648-2023-03251 | 17775157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17775157&pc=OZD | 11/19/2021 | 09/19/2023 |  | 09/19/2023 |  |  | 11/19/2021 | 09/19/2023 |  |  | 07/31/2023 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130559 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/19/2021 |  |  |  |  | The user facility reported that a 50-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. during use, the tuohy valve was reportedly unable to hold the catheter in place. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the pump and purge cassette were returned for evaluation. analysis of the device revealed that the tuohy locked and gripped the catheter without any issue when the catheter was cleaned and dried and when it was slick, the catheter was able to move with significant force when the tuohy was fully engaged. given this evaluation, it is likely that the catheter was able to move as a result of blood between the catheter and the tuohy reducing the friction. the product worked as intended so this was determined to be a use-related issue. the root cause of the positioning issue was use related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34483 | 1220648-2023-03252 | 17775177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17775177&pc=OZD | 02/04/2022 | 09/19/2023 |  | 09/19/2023 |  |  | 02/04/2022 | 09/19/2023 |  |  | 10/31/2023 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154271 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Image Display Error/Artifact | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/04/2022 |  |  |  |  | The user facility reported that a 74-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge flow stopped due to the program being ¿stuck¿ during the purge cassette change. urgent attempts to separate the luers led to breakage of the purge sidearm, requiring a purge bypass and eventually, removal of the device. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge flow stopped was completed. the device was not returned for evaluation. based on the available information, the root cause of the stuck purge cassette change screen issue was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34484 | 1220648-2023-03253 | 17775299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17775299&pc=OZD | 03/08/2022 | 09/19/2023 |  | 09/19/2023 |  |  | 03/08/2022 | 09/19/2023 |  |  | 07/31/2023 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130619 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2022 |  |  |  |  | The user facility reported that a 43-year-old male patient implanted with the impella 5.5 for mechanical circulatory support had an aortic thrombus/tissue that was visualized on transesophageal echocardiogram (tee) during device explant. the physicians believed the thrombus to be related to the device. the physicians also believed that the thrombus was not life threatening, thus no intervention additional was performed. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation. based on the available information, the root cause of the reported event was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34485 | 1220648-2023-03254 | 17775346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17775346&pc=OZD | 10/25/2021 | 09/19/2023 |  | 09/19/2023 |  |  | 10/25/2021 | 09/19/2023 |  |  | 07/31/2023 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022117924 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/25/2021 |  |  |  |  | The user facility reported a 50 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a pump stop occurred after the white cord got caught on the wheel while patient was standing. the white cable was reinserted but the pump failed to restart. the team attempted to push the white cable in deeper with no success. the console was swapped but it still did not restart. the pump was eventually removed. there was no patient harm reported. | [5 empty] | The investigation for the pump stop has been completed. the returned pump was visually inspected, and a severe catheter kink was observed adjacent to the pump handle. the pump was plugged into a console and would not start, even at various p-levels. electrical resistance testing revealed multiple open lines. further electrical continuity testing confirmed the kink as the location of the open lines. therefore, the root cause of the reported pump stop was determined to be multiple open lines due to excessive force. the instructions for use states the following for "impella stopped": "if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34486 | 1220648-2023-03255 | 17776793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17776793&pc=OZD | 08/21/2023 | 01/23/2024 |  | 09/19/2023 |  |  | 08/21/2023 | 09/19/2023 |  |  | 04/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 69-year-old male in worsening cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. an echo was obtained on arrival at the icu to verify impella position. the patient was awake at this time in no distress. on echo intermittently (every 5-10 sec, the left ventricle would fill up with micro bubbles and then they would disappear and come back. purge tubing observed and no bubbles or anything seen in purge tubing. unsure of the source. while impella cp was prepping in catheter lab it had time to prime for 5-10 minutes before going in the body. it was observed that the patient was having no adverse events from this event, but unsure where the source was. the physician was attempting repositioning and the micro bubbles continued with the intermittent pattern during manipulation as well as with impella stationery. the pump support remained for just under 2 days til care withdrawn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the embolism has been completed. the device was not returned for benchtop investigation. based on the clinical data, no issues during impella insertion and priming. the cause of the embolism issue was most likely patient condition related. unknown relation to impella.
34487 | 1220648-2023-03256 | 17776848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17776848&pc=OZD | 08/21/2023 | 11/03/2023 |  | 09/19/2023 |  |  | 08/21/2023 | 09/19/2023 |  |  | 03/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346266 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. prior to surgery ccu bedside nurse reported left leg was pulseless. the physician evaluated limb and in the or and was uncertain if limb was totally occluded from impella placement. after transfer patient continued to decompensate and with a questionable limb, the decision was made to place ecmo and replace the femoral cp with an impella 5.5 via axillary insertion. | [5 empty] | The investigation into the limb ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the cause of the limb ischemia issue was patient condition related to classified vessel condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
34488 | 1220648-2023-03257 | 17776870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17776870&pc=OZD | 08/21/2023 | 02/07/2024 |  | 09/19/2023 |  |  | 08/21/2023 | 09/19/2023 |  |  | 02/28/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024341065 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was experiencing hemolysis. the impella cp was upgraded to a 5.5 and the patient is doing well | [5 empty] | The investigation into the hemolysis has been completed. the device was not returned for investigation. data logs showed pump ran without issue for entire case. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34489 | 1220648-2023-03258 | 17776910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17776910&pc=OZD | 06/13/2023 | 09/19/2023 |  | 09/19/2023 |  |  | 06/13/2023 | 09/19/2023 |  |  | 01/31/2024 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023190677 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 69-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella cp was removed due to left ventricular thrombus and poor positioning due to thrombus. the medical staff reported "impella in aorta" alarms additionally, even with pulsatile placement signal and pulsatile motor current. echo confirmed proper placement across valve. the patient continued on ecmo support. there was no known harm or injury. | [5 empty] | (h3) the investigation into the reported "positioning issues" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. as per clinical, the impella position became shallow due to left ventricular thrombus and the impella cp was explanted due to complication. the cause of the positioning issue was due to patient's condition related with patient having left ventricle (lv) thrombus complicating pump repositioning.
34490 | 1220648-2023-03259 | 17777020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17777020&pc=OZD | 07/05/2023 | 09/19/2023 |  | 09/19/2023 |  |  | 07/05/2023 | 09/19/2023 |  |  | 12/31/2024 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024301613 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/05/2023 |  |  |  |  | The user facility reported a 52-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the automated impella controller (aic) displayed an alarm "impella defective". it was reconnected to the original aic and same alarm appeared. patient was hemodynamically stable and was taken to operating room (or) for 5.5 replacement so they could be transferred. | [5 empty] | (h3) the investigation into the reported pump stop has been completed. the device was visually inspected and found upon closer examination, blood was discovered in the red handle, and a catheter hole was observed near the 30cm mark. data logs shows the pump stop (alarm 115) event during aic transfer. additionally, during the test run, the impella deflective alarm was triggered right from the beginning. the device successfully passed the electrical testing. as per clinical investigation, it was reported the automated impella controller (aic) displayed an alarm "impella defective". it was reconnected to the original aic and same alarm appeared. patient was hemodynamically stable and was taken to or for 5.5 replacement so they could be transferred. the root cause of the issue was an short electrical line due to blood in red handle from the hole found in catheter.
34491 | 1220648-2023-03260 | 17777828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17777828&pc=OZD | 08/22/2023 | 01/27/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 01/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333944 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was having hemolysis with pfhb >300 and requiring three pressors. patient brought to the or for escalation to impella 5.5. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the product was not returned for investigation. data logs show no signs related to improper positioning or biomaterial ingestion. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34492 | 1220648-2023-03261 | 17777838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17777838&pc=OZD | 08/21/2023 | 01/12/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 12/31/2024 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320266 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old female noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the access site had a hematoma. the team treated by blood products and re-adjusting the repo sheath and sutures. the hematoma was expressed manually by the technologist in the lab/at bedside. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. per the clinical information provided, the root cause of the access site bleeding issue was most likely use related to improper technique. this was resolved by suturing and angle matching as per the best practices to achieve hemostasis which suggests best practices were not used before.
34493 | 1220648-2023-03262 | 17778009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778009&pc=OZD | 08/21/2023 | 01/30/2025 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  |  |  | 09/19/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1124584 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the staff performed a console check, and ¿battery failure¿ alarm was present. battery cycled on and off under console with no success. the console was sent to the hospital biomed department. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) battery failure -red alarm has been completed. after reviewing the imc logs, the reported battery failure issue was confirmed. there were numerous instances of battery failure alarms (transport connection blown/missing) (event set #360) observed on the reported complaint date. the aic was returned for evaluation. the battery failure issue was able to be reproduced. results of running the batttest utility revealed that cell 4 in battery 2 had a voltage that was significantly less than the rest of the cells in the battery. the reported battery failure issue was determined to be caused by internal battery cell imbalance.
34494 | 1220648-2023-03263 | 17778012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778012&pc=OZD | 08/21/2023 | 01/25/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 01/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306229 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient has history of bilateral peripheral artery disease. angiography showed the left side was better than the right and after impella insertion there was flow distal to the sheath. a dopplerable left posterior tibialis pulse was found in the cath lab, but lost on arrival to main campus icu. interventional cardiology already contacted and planning to place a reperfusion catheter down the left leg. | [5 empty] | The medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the cause of the limb ischemia - reversible was patient condition related because patient had history of bilateral pad. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34495 | 1220648-2023-03264 | 17778017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778017&pc=OZD | 08/20/2023 | 01/22/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 06/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363822 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 71-year-old patient with known coronary artery disease was treated with high-risk percutaneous coronary intervention (hrpci) and received hemodynamic support from an impella cp smartassist heart pump during the pci procedure. the same femoral access site was used as the access site for the pci catheter as for the impella catheter. after the complicated pci procedure in which several stents were placed, the impella cp remained in the patient to allow the heart to recover. the leg below the femoral access site was reported to show pallor/color change. the impella cp heart pump was then removed and manual compression was applied to achieve hemostasis. after manual compression at the access site, acute ischemia of the leg was observed, and femoral artery surgery was performed to resolve the ischemia. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's limb ischemia has been completed. product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition--diseased/small vessels. d4, corrected unique identified (udi) number.
34496 | 1220648-2023-03265 | 17778020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778020&pc=OZD | 08/21/2023 | 01/11/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 05/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362060 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 76-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the device was inserted via the femoral artery. both iliofemoral anatomy/limbs were evaluated for the access as there was peripheral artery disease and calcification. the insertion did cause/contribute to a vessel perforation that was treated by covered stenting. after stenting the pump would not deliver and, as such, the case was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported vessel damage has been completed. the device was not returned for evaluation. clinical data revealed the use of multiple dilation and shockwave devices to access arteries. however, initial delivery attempts failed on both arteries due to extensive calcification, and vascular damage was identified in the right femoral artery. the cause of the vascular damage issue was patient condition related to calcified vessel condition.
34497 | 1220648-2023-03266 | 17778027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778027&pc=OZD | 08/21/2023 | 09/21/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 06/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356952 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 52-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had a 5.5 support that was ongoing when there was persistent suction alarms despite adequate volume and positioning. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the low pump flow and the access site bleeding ¿ minor has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the low pump flow and the access site bleeding ¿ minor issues was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03266 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34498 | 1220648-2023-03267 | 17778032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778032&pc=OZD | 08/21/2023 | 07/01/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 06/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359739 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 76-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. when advancing catheter and wire, the physician experienced difficulty delivering to aorta. cardiologist called to assist. contrast injection revealed dissection in axillary artery. cardiologist re-wired vessel in true lumen and then successfully delivered impella to left ventricle. | [5 empty] | The investigation into the vascular damage - peripheral vessel has been completed since the original report was submitted. the device was not returned for investigation. it was reported that before beginning impella support, the team had trouble inserting the pump into the aorta. the most likely cause of the peripheral vessel vascular damage was use related. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. removed f6 and f8. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
34499 | 1220648-2023-03268 | 17778034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778034&pc=OZD | 08/21/2023 | 01/24/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 07/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370777 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a71-year-old male undergoing emergent cardiopulmonary bypass surgery was implanted with an impella 5.5 device for mechanical circulatory support. the surgeon attempted to implant impella 5.5 without the guidewire. it was difficult delivery through the anastomosis site causing excess bleeding from suture line which required repair. | [5 empty] | The investigation into the access site bleeding issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to use issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34500 | 1220648-2023-03269 | 17778042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778042&pc=OZD | 08/19/2023 | 01/22/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  |  |  | 09/19/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1302874 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  | White | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | An 86-year-old patient was treated for acute myocardial infarction/cardiogenic shock. an impella cp smartassist heart pump was used for hemodynamic support. during purge system preparation, the automated impella controller (aic) did not recognize the purge cassette. visual inspection revealed damage to the aic. the issue was resolved by resorting to the backup controller. the defective controller was sent in for examination and repair. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the device was returned for investigation. it was determined that the root cause of this issue was confirmed to be a cracked purge retainer. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34501 | 1220648-2023-03270 | 17778046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778046&pc=OZD | 08/22/2023 | 01/22/2024 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 04/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024349664 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 44-year-old female with cardiogenic shock without known cause received support from an impella cp smartassist cardiac pump and va-ecmo to rest the heart. after more than 10 days of support, the patient had successfully been weaned from support. the day after impella removal, the condition of the patient worsened. trans-esophageal echo revealed massive mitral valve insufficiency with loose chordae. the patient underwent emergency surgery to repair the damaged mitral valve. the patient has been stable afterwards. (of note: impella cp smartassist pump is indicated for 5 days of use.) | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the heart valve damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the cause of the heart valve damage issue was not determined since product was not returned and insufficient clinical information were provided.
34502 | 1220648-2023-03271 | 17778063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778063&pc=OZD | 08/22/2023 | 03/27/2025 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 04/30/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Ventricular Fibrillation | [10 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient had a large hematoma around the left femoral artery (lfa) access site. patient began to decompensate requiring increased pressor support, while also having noticeable decrease in pulsatility on impella and arterial line. hgb noted to have decreased from 13 to 10.5, and multiple units of blood were transfused. patient ventricular fibrillation (vf) arrested multiple times, requiring shock. unable to obtain return of spontaneous circulation (rosc), so decision made to stop resuscitation efforts. patient was on impella support for 4.67 hours before expiring. rep. confirmed that there were no allegations made towards the impella for cause of death. | The patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleed, limb ischemia, and ventricular fibrillation (vf) has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as all the necessary information was not provided, the root cause of the limb ischemia and ventricular fibrillation was not determined. as noted in the impella instructions for use: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.5 abiomed's medical safety officer review was added. h.6 code 1884 was added due to investigation results. h.10 additional instructions for use were added due to investigation results. b.2 revised selections due to medical safety review since initial manufacturer device report number 1220648-2023-03271 was submitted. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-03271. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-03271. d.6a and d.6b implant and explant dates were added since the initial manufacturer device report number 1220648-2023-03271 was submitted in accordance with updated procedures. e.1 fax number was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-03271. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2023-03271 was submitted in accordance with updated procedures. h.1 revised selection due to medical safety review since initial manufacturer device report number 1220648-2023-03271 was submitted. h.6 code 925 was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-03271. code 1802 was added due to medical safety review since initial manufacturer device report number 1220648-2023-03271 was submitted.
34503 | 1220648-2023-03272 | 17778065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17778065&pc=PYX | 08/22/2023 | 11/13/2023 |  | 09/19/2023 |  |  |  | 09/19/2023 |  |  | 01/31/2025 |  | 09/19/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024315026 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 24-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella rp device for mechanical circulatory support. the complainant reported there was a thrombus on the impella rp. the pump was explanted and staff planned on escalating to impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the low pump flow issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the low pump flow is most likely due to ingested biomaterial.
34504 | 1220648-2023-03222 | 17764346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17764346&pc=OZD | 06/11/2022 | 01/10/2024 |  | 09/18/2023 |  |  | 06/11/2022 | 09/18/2023 |  |  | 09/30/2023 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143119 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2022 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced positioning issues and the pump had to be repositioned with echo guidance. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issue has been completed since the original report was filed. the root cause of the positioning issue was determined to be use issue. it was confirmed using echo that impella turned toward mitral and impella was repositioned with echo guidance. pump now midventricular at 4.7cm and systemic heparin was discontinued. there was no additional clinical details provided and neither data logs nor the product were returned.
34505 | 1220648-2023-03223 | 17764487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17764487&pc=OZD | 08/23/2023 | 09/20/2024 |  | 09/18/2023 |  |  | 08/23/2023 | 09/18/2023 |  |  | 12/31/2024 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024304047 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 55-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the medical center had a patient made a dnr and expired on impella cp support after the comorbidities of cerebral infarct and deep vein thrombosis (dvt) caused the team to make this decision. the cerebral infarct relationship to impella is not known. there was no harm, injury or malfunction attributed to the use of impella. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/cva has been completed since the original report was filed. as enough clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03223 was provided in sections b2, b5,d6, h1, and h6.
34506 | 1220648-2023-03224 | 17764712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17764712&pc=OZD | 06/19/2022 | 01/26/2024 |  | 09/18/2023 |  |  | 06/19/2022 | 09/18/2023 |  |  | 03/31/2024 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208290 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding and hematoma at the impella access site. a bedside washout had to be performed and new dressing placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding - major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34507 | 1220648-2023-03225 | 17765682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17765682&pc=OZD | 12/31/2022 | 05/04/2024 |  | 09/18/2023 |  |  | 01/14/2023 | 09/18/2023 |  |  | 07/31/2024 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248411 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/14/2023 |  |  |  |  | Prior to the reported pump stop, it was noted that mild oozing was seen at the abdomen where the pump was fixated. additional sutures were placed at the site to assist with the bleeding. support continued with the implanted pump following the intervention. | The user facility reported a 49yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a pump stop but the staff could not determine the cause. upon troubleshooting, the staff noted a brief motor current spike and a subsequent pump stop. the nurse changed the settings and all waveforms were as they were prior to the pump stop. there were no further issues. the pump remained on for support. there was no patient harm reported. |  |  |  |  | Additional information was reviewed pertaining to the noted case. sections b, g4, h1, and h6 have been updated to coincide with the additional report of oozing/bleeding related to the implanted pump. upon completion of the investigation, a supplemental report will be filed. | The investigation for the reported access site bleed has been completed. without additional clinical details, the root cause of the access site bleed could not be determined. the impella 5.5® with smartassist® for use during cardiogenic shock instructions for use states in section: warnings and cautions to ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. it also states ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ to conform to updated procedures, sections d6 & h10 were revised with updated information. | The investigation into the pump stop has been completed. the returned pump was visually inspected and biomaterial was observed on the impeller. the cause of the issue was biomaterial. the instructions for use for the pump stop is as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34508 | 1220648-2023-03226 | 17765946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17765946&pc=OZD | 09/14/2023 | 12/12/2023 |  | 09/18/2023 |  |  | 09/14/2023 | 09/18/2023 |  |  | 05/31/2024 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231104 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during preparation for the impella implant, staff went to put a slight curve on the impella guidewire, and the wire curled excessively. local support confirmed the staff did not put excessive force on the wire in any manner. the curled wire was discarded and a new impella wire was used successfully. additionally, multiple attempts were made to advance impella and position, but the staff were unable to do so due to aortic arch anomaly. decision made to abort procedure and consult the doctor for impella 5.5 placement for complex pci procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into delivery issues/product damage (guidewire) has been completed. the product was not returned for evaluation. the root cause of the delivery issues/product damage is most likely due to use since the product was visualized prior to preparation for use and the team did not identify any issues.
34509 | 1220648-2023-03227 | 17767358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17767358&pc=OZD | 08/19/2023 | 04/19/2024 |  | 09/18/2023 |  |  | 08/19/2023 | 09/18/2023 |  |  |  |  | 09/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1320578 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/19/2023 |  |  |  |  | The user facility reported an 80-year-old male with cardiomyopathy was to be implanted with an impella cp for mechanical circulatory support. during set-up for the procedure, the automated impella controller (aic) was not making audible beeps with any alarms. the team switched to a new aic and the issue resolved. no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic ¿ other issue has been completed since the initial report was submitted. the console was returned and tested. the failure was reproduced on this console. alarms did not sound, and the alarm cable was found to not be connected. the root cause of the lack of alarm was due to the alarm being unplugged. it is unclear what caused the cable to become unplugged.
34510 | 1220648-2023-03228 | 17767480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17767480&pc=PYX | 09/21/2022 | 09/18/2023 |  | 09/18/2023 |  |  | 09/21/2022 | 09/18/2023 |  |  | 02/28/2024 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023194542 |  | 004334 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Suction Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported that a 78-year-old male patient implanted with the impella rp for mechanical circulatory support experienced low pump flows and suction alarms. the patient received added "blood volume" and the pump was repositioned multiple times without success. due to the unresolved suction, the decision was made to explant the device. of note, the patient was also on concurrent venous-arterial extracorporeal mechanical circulatory (va ecmo) support. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows and suction alarms was completed. the device was not returned for evaluation. based on the available information, the root cause of the low pump flows and suction alarms was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34511 | 1220648-2023-03229 | 17767625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17767625&pc=OZD | 07/14/2021 | 09/18/2023 |  | 09/18/2023 |  |  | 07/14/2021 | 09/18/2023 |  |  | 04/30/2023 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022070055 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/14/2021 |  |  |  |  | The user facility reported that a 32-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a pump stop when the rn raised the bed and the white cable got caught in the bed rail, also pulling out the red handle. the device immediately stopped and, as a result, it was explanted at bedside. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported pump stop was completed. the pump and data logs were returned for evaluation. analysis of the data logs confirmed the sudden pump stop. visual inspection of the pump revealed that the catheter was cut at the 44 cm mark and the catheter was confirmed to have been pulled out of the red handle with all the motor cables fractured. the root cause of the reported event was use-related excessive force. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34512 | 1220648-2023-03230 | 17769027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769027&pc=OZD | 09/18/2022 | 09/18/2023 |  | 09/18/2023 |  |  | 09/18/2022 | 09/18/2023 |  |  | 06/30/2024 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229045 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 66yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a pump stop due to a frayed white cable. the staff further noted that the catheter got caught on the bed while turning the patient, and the catheter came out of the kink protector. the pump was removed. there was no patient harm reported. | [5 empty] | The investigation into the pump stop has been completed. the returned product was visually inspected and catheter to kink protector detachment was observed with damaged motor cable lines. the cause of the pump stop was an open line due to a catheter to kink protector detachment. the cause of the pump stop was an electrical open during to patient movement. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34513 | 1220648-2023-03231 | 17769135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769135&pc=OZD | 08/21/2023 | 01/11/2024 |  | 09/18/2023 |  |  | 08/21/2023 | 09/18/2023 |  |  | 05/31/2025 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345515 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male had a 5.5 support ongoing for a st elevated myocardial infarction patient. the impella 5.5 was placed for support of just under five days when weaned and explanted. during the support there was a red plug leak that was not treated in any way. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the reported product damage (hole in the catheter) has been completed since the original report was filed. no external damage or abnormalities were found on the returned pump. the red handle was opened, and purge fluid was observed to be leaking into the handle from the catheter. a small hole was observed in the purge lumen proximal to the motor housing after stripping the catheter; there was no evidence of use-related issues. therefore, it was concluded that the cause of the purge leak was damaged purge tubing stemming from manufacturing.
34514 | 1220648-2023-03232 | 17769535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769535&pc=OZD | 08/21/2023 | 01/11/2024 |  | 09/18/2023 |  |  | 08/21/2023 | 09/18/2023 |  |  | 11/30/2024 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had a long 25cm 14fr sheath that was placed in a 80-year-old female patient for impella cp delivery. the anatomy was tortuous. the team replaced the guidewire and removed the dilator. the case completed. at the time of explant the tip of the sheath was seen to be cracked/split. the sheath split caused no harm or patient injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the introducer damage ¿ mechanical has been completed. the product was not returned. as per the clinical information provided, the root cause of introducer damage was most likely tortuous vasculature patient condition related.
34515 | 1220648-2023-03233 | 17769826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769826&pc=OZD | 08/19/2023 | 04/23/2024 |  | 09/18/2023 |  |  | 08/19/2023 | 09/18/2023 |  |  | 04/30/2025 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353513 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2023 |  |  |  |  | The user facility reported a 74-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella alarmed for being in the ventricle during an icu bath. the p-level for the device was adjusted, suction alarms ensued and diastolic impella flows went and remained at zero. impella was reduced down to p2 and remained in suction. echo was called and showed impella was pulsating in and out of the left ventricle. the physician repositioned the impella to optimize stability and distance to the valve, however, suction alarms continued, and the left ventricle waveform remained decoupled from aorta wave form. impella flows were below 1 l and milrinone was added to support patient. clinically, the patient states feeling good. the physician repositioned impella under echo as pumped was angling to posterior wall. the impella was optimized in left ventricle apex in free space, however, impella flows remained at a low state. the decision from the cardiologist is to manage the patient on p2 with aggressive diuresis with plans for possible explant in the next day or two. if the patient decompensates, the physician will replace cp with another impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. data logs confirmed that the low flow occurred while the pump was in the ventricle, where it remained to the rest of the case, and triggered auto suction response and suction alarms. based on clinical description and data analysis, the root cause of positioning issue was due to use related.
34516 | 1220648-2023-03234 | 17769835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769835&pc=OZD | 08/20/2023 | 01/22/2024 |  | 09/18/2023 |  |  | 08/20/2023 | 09/18/2023 |  |  | 02/28/2025 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024323288 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 70-year-old patient received hemodynamic support from an impella cp smartassist heart pump after high-risk surgery due to postcardiotomy cardigenic shock and low cardiac output syndrome (lcos). after 15 days of running, the pump stopped with an alarm due to high motor current. the pump could not be restarted. a new impella cp heart pump was used. the patient survived the event but continues to require hemodynamic support to bridge the gap to permanent lvad placement. (note: intended duration of use of the impella cp heart pump according to the instructions for use: up to 5 days). | [5 empty] | The investigation into the pump stop has been completed. the impella cp was returned with the catheter cut open near to the red plug. visually inspected and there was a large purge gap present near the motor. no other abnormalities were found. data logs were not returned for review. as per clinical, the impella cp device was run for more than 15 days before pump stop failure with high motor current alarm observed. the cause of the pump stop issue was due to bearing wear failure which occurred after 15 days in support which is beyond design life per field investigation and clinical review. the impella cp optical device is not indicated for use more than 8 days of design life period according to design trace matrix. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34517 | 1220648-2023-03235 | 17769843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769843&pc=OZD | 08/20/2023 | 04/02/2024 |  | 09/18/2023 |  |  | 08/20/2023 | 09/18/2023 |  |  | 06/30/2025 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371138 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported during the removal of the abiomed introducer, the blue suture hub was advanced into skin track leading to site bleeding that was controlled with very positional repo sheath to maintain hemostasis. vascular surgery was called in to assess groin bleeding despite many efforts to allow the staff to redress the groin first. the vascular team applied 3-4 figure 8 sutures and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. as per clinical details, during removal of introducer blue suture hub was advanced into skin track leading to some site bleeding. the root cause of the access site bleeding issue was use related since blue suture hub was found into skin during introducer removal.
34518 | 1220648-2023-03236 | 17769889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769889&pc=OZD | 08/19/2023 | 09/19/2024 |  | 09/18/2023 |  |  | 08/19/2023 | 09/18/2023 |  |  | 05/31/2025 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360903 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Power Up | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/19/2023 |  |  |  |  | The user facility reported a 60-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. p\prior to impella insertion aggrastat bolus was given. the sheath aspirated and flushed, and impella was advanced into left ventricle. upon removal of the guidewire and initiation of support, a high motor current alarm sounded and the pump did not start. support levels were increased to p-7 and saturated flows were noted. the decision was made to replace the pump for continued support. there was no patient harm. | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported saturated flows and pump stop has been completed. the pump was returned for evaluation and visually inspected. biomaterial was observed around the impella and purge gap. pump was connected to the console and attempted to run three times. pump stopped and did not start and a high motor current alarm occurred. pump was soaked in water and an enzyme detergent to dissolve the biomaterial. biomaterial was removed and the pump ran in investigation lab at different p-levels. no issues observed and the pump ran normally. therefore, the root cause of the reported issue was due to biomaterial. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03236 was provided. note: corrected information provided in h6 is an addition to previous coding.
34519 | 1220648-2023-03237 | 17769906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769906&pc=OZD | 08/21/2023 | 01/24/2024 |  | 09/18/2023 |  |  | 08/21/2023 | 09/18/2023 |  |  | 06/30/2025 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367180 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the patient was without major residual status post right middle cerebral artery (rmca) syndrome/stroke in setting of peripheral heparin. low-dose was kept for subtherapeutic partial thromboplastin time (ptt). patient had slight right-sided facial droop which was reported to have resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ minor and the stroke/cva issues has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. the root cause of the stroke issue was most likely patient condition related but the relation to impella is unknown.
34520 | 1220648-2023-03238 | 17769925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769925&pc=OZD | 08/21/2023 | 02/14/2024 |  | 09/18/2023 |  |  | 08/21/2023 | 09/18/2023 |  |  | 05/31/2025 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during the impella insertion the patient developed a large hematoma requiring femostop. hematoma ¿milked¿, and fem stop placed 3 cm above site on 60mmhg of pressure. the decision was made to keep impella in place as 250 of contrast was used, but after 3 hours in icu impella weaned and explanted due to hematoma concerns. hemostasis achieved via 2 hours of femostop. patient was stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the bleeding was a lack of vessel recoil onto the repo sheath because the patient was large, and the repo unit is not as long as the14f x 25 cm abiomed sheath that was peeled away.
34521 | 1220648-2023-03239 | 17769933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17769933&pc=OZD | 08/21/2023 | 12/01/2025 |  | 09/18/2023 |  |  | 08/21/2023 | 09/18/2023 |  |  |  |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the controller was turned on and a yellow banner appeared reading controller error. the controller was exchanged as a result of the noted issue. there was no patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | A.2 age should have been left blank in accordance with updated procedures on the initial manufacturer device report number 1220648-2023-03239 as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report number 1220648-2023-03239 as it is unknown. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03239 was submitted. d.4 model number, catalog number and serial number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-03239 was submitted. e.1 name prefix/title was revised as it was previously entered incorrectly on initial manufacturer device report 1220648-2023-03239. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-03239 in accordance with updated procedures. h.6 codes 2199 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2023-03239 in accordance with updated procedures.
34522 | 1220648-2023-03240 | 17770238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17770238&pc=OZD | 11/02/2021 | 09/18/2023 |  | 09/18/2023 |  |  |  | 09/18/2023 |  |  | 08/31/2023 |  | 09/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130667 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2021 |  |  |  |  | The user facility reported that a 71-year-old male patient implanted with the impella 5.5 for mechanical circulatory support had a crack in the purge sidearm filter. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge sidearm crack/ purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge sidearm crack/ purge leak was sidearm filter damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34523 | 1220648-2023-03219 | 17762258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17762258&pc=OZD | 11/17/2022 | 09/17/2023 |  | 09/17/2023 |  |  | 11/17/2022 | 09/17/2023 |  |  | 07/31/2024 |  | 09/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023249663 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported that a 74-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced suboptimal flows despite multiple echos and repositioning of the pump. echo results revealed that the pump was in the left ventricle (lv) free space measuring 4.8-5.1 cm from aortic valve (av) annulus. the physician reported the insertion to be challenging. there was significant manipulation of the catheter. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. visual inspection of the pump revealed a large clot in the outflow cage. based on the available information, the root cause of the low pump flow is most likely due to biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34524 | 1220648-2023-03220 | 17762259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17762259&pc=OZD | 10/18/2022 | 09/17/2023 |  | 09/17/2023 |  |  | 10/18/2022 | 09/17/2023 |  |  | 04/30/2024 |  | 09/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023215053 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported that a 72-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak after 49 days of support due to a break in the purge sidearm. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the issue could not be determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34525 | 1220648-2023-03221 | 17762260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17762260&pc=OZD | 10/16/2022 | 09/17/2023 |  | 09/17/2023 |  |  | 10/16/2022 | 09/17/2023 |  |  | 12/31/2022 |  | 09/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022026879 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/16/2022 |  |  |  |  | The user facility reported that a 66-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the device was placed, started, and position verified via echo, the physician did not see the improvement he wanted/needed for the patient and the device was explanted to put the patient on extra corporeal membrane oxygenation (ecmo) to stabilize hemodynamic support. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. based on the available information, the root cause of the issue was not determined as the device was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34526 | 1220648-2023-03188 | 17754685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17754685&pc=OZD | 09/27/2022 | 09/15/2023 |  | 09/15/2023 |  |  | 09/27/2022 | 09/15/2023 |  |  | 06/30/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023240253 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported that a 34-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the device was pulled back into the aorta while the patient attempted to place himself on the bedpan. the pump was subsequently removed at beside and the patient, remained on va ecmo overnight and it was successfully removed the following day. follow-up with staff, revealed that the was told multiple times not to move on his own and he failed to comply. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the available information, the root cause of the positioning issue was most likely due to patient movement. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34527 | 1220648-2023-03189 | 17754720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17754720&pc=OZD | 09/20/2022 | 09/15/2023 |  | 09/15/2023 |  |  | 09/20/2022 | 09/15/2023 |  |  | 06/30/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237520 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/20/2022 |  |  |  |  | The user facility reported that a 65-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a high purge pressure alarm about 16 hours into support. the high purge pressure persisted even though no catheter kinks or purge system obstructions were observed. the purge cassette was replaced, but the issue remained. the p level was reduced without resolution and the decision was made to explant the pump. the patient remained stable on extracorporeal membrane oxygenation (ecmo) following explant. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure was completed. the device was not returned for evaluation, but the data logs were received. analysis of the data logs showed the "purge pressure high" alarm triggered following a gradual rise in purge pressure over 4 hours; "purge system blocked" alarms were also present at this time. heparin levels were at zero through the entirety of the case. however, since the device was not returned, the root cause of the high purge pressure was undetermined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34528 | 1220648-2023-03190 | 17754743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17754743&pc=OZD | 01/31/2022 | 09/15/2023 |  | 09/15/2023 |  |  | 01/31/2022 | 09/15/2023 |  |  | 10/31/2023 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  |  |  | 2022157667 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2022 |  |  |  |  | The user facility reported that a 65-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the clinical staff was unable to get the device into a stable position with adequate flows. the device was removed, a new one was prepped and inserted without issue. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. the root cause could not be determined as neither product nor data was returned. no further information was provided. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34529 | 1220648-2023-03191 | 17754841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17754841&pc=OZD | 08/17/2023 | 11/30/2023 |  | 09/15/2023 |  |  | 08/17/2023 | 09/15/2023 |  |  | 04/30/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348674 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. a pump exchange occurred due to care team concern for infection. extracorporeal membrane oxygenation (ecmo) was being decannulated after pump exchange so no heparin was given prior to device insertion. sheath was aspirated and flushed just prior to pump insertion. throughout the night, pump developed flow saturation. cath lab team replacing pump again to prevent failure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into saturated flow and infection/sepsis has been completed. in the logs it was seen that there were suction conditions present and then a spike in purge pressure as well as motor current. after these events, the flow became saturated. additionally, the product was returned to the lab and biomaterial was wrapped around the impeller blades. the patient had an unknown active infection before this cp was placed. the root cause for the saturated flow is most likely due to biomaterial ingested during this period and the root cause of the infection is most likely due to the patients condition.
34530 | 1220648-2023-03192 | 17754853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17754853&pc=OZD | 08/17/2023 | 09/22/2024 |  | 09/15/2023 |  |  | 08/17/2023 | 09/15/2023 |  |  | 02/28/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319689 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 51-year-old male patient was scheduled for and av node ablation with crt-d placement. patient started to decompensate, and physician was called in to place impella cp to complete the placement of the crt-d and av node ablation successfully. the patient went back to icu room and impella position was confirmed via bedside. however, the patient¿s urine went from yellow to black throughout the night. therefore, the impella was explanted a day later due to concerns of hemolysis and decision was made to withdraw care. the patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis has been completed. the impella device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03192 was provided in sections b2, b5, d6a, d6b, and h6. note: corrected information provided in h6 is an addition to previous coding.
34531 | 1220648-2023-03193 | 17754880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17754880&pc=OZD | 10/26/2022 | 01/26/2024 |  | 09/15/2023 |  |  | 10/26/2022 | 09/15/2023 |  |  | 05/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222340 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced pain and a hematoma at the access site. the patient was taken to the or to have the hematoma evacuated. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34532 | 1220648-2023-03194 | 17756183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17756183&pc=OZD | 10/27/2022 | 03/11/2025 |  | 09/15/2023 |  |  | 10/27/2022 | 09/15/2023 |  |  | 03/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023199537 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Fluid/Blood Leak | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a purge fluid leak coming out of the reservoir chamber. a day after this event, the patient experienced ventricular tachycardia. the pump position was checked and it was in the correct position. in the following days the patient condition deteriorated and decision was made to withdraw care. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | Purge leak pump- the root cause of the purge leak was unable to be determined as the pump was not returned for investigation. the root cause of the vf/vt/af/at could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
34533 | 1220648-2023-03195 | 17756199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17756199&pc=OZD | 08/18/2023 | 12/07/2023 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 05/31/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363500 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention. hemolysis has been reported. abiomed customer support provided advice regarding possible impella-related causes and treatment options for hemolysis. the attending physician had the correct position of the impella heart pump in the left ventricle confirmed using an echo to rule out hemolysis caused by suction. thereafter, no hemolysis or other problems related to the impella cp were reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into hemolysis has been completed. the impella products and data logs were not returned for review. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
34534 | 1220648-2023-03196 | 17756234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17756234&pc=OZD | 08/18/2023 | 11/06/2025 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 05/31/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359796 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 75-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the physician was concerned due to motor current and placement signal irregularities. patient was on bivalrudin and bicarb in purge. physician was worried this could be thrombus related. pump was removed. thrombus was noted on pigtail after pump was removed. | Additional information was received via abiomed's long-term outcome and quality indicator impella registry (loqi) indicating this patient to have endured hemolysis with a probably relationship to the impella device: "impella was complicated with hemolysis (ldh > 1189>2200>2349, cpk 590>2203, and low haptoglobin)" resulting in device removal. additionally, the day prior (on august 17th) thrombocytopenia was noted; also with a probable relationship to the device: ""high suspicion for hit syndrome given significant bleeding and thrombosis despite therapeutic act values on heparin." |  |  |  |  | The investigation of thrombus, hemolysis, and thrombocytopenia has been completed. the pump was returned for investigation. data logs were returned for analysis. data log analysis confirmed suction alarms during the case, which occurred in conjunction with placement signal low alarms, around 9am on (b)(6) 2023. p-level was reduced to p-6 around this time. there were no motor current trends or spikes outside of p-level changes. there were no purge pressure spikes outside of purge bag changes. there were no placement signal spikes. product evaluation confirmed there was no biomaterial in the inflow and outflow cage or on impeller. however, scratches were observed on the impeller wings. the fracture was characteristic of an interaction with another device, but this was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. additionally, no thrombus was observed on the pigtail of the returned pump. hemolysis testing was also performed on the returned pump and the resulting mih values were greater than 20 during the first and second rounds of testing. finally, microscopic imaging confirmed blood between the pump delo joint and delaminated epo-tek. hemolysis: the root cause of the hemolysis was most likely damaged impeller based on product evaluation and clinical details. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. thrombocytopenia: the root cause of the thrombocytopenia was not determined due to insufficient clinical details. instructions for use: thrombus: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." hemolysis: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombocytopenia: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d6a and d6b revised from the initial manufacturer device report 1220648-2023-03196 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03196 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03196. g2 report source; study was missing from manufacturer device report 1220648-2023-03196 h6 component code revised from the initial manufacturer device report number 1220648-2023-03196 in accordance with updated procedures. h10 instructions for use inadvertently omitted on the manufacturer device report number 1220648-2023-03196 and 1220648-2023-03196-1. | The impella device was received from the customer and, an evaluation of the device is ongoing at this time.. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | B(5): included new information of thrombocytopenia and hemolysis. investigation is still in process and a supplemental will be filed upon conclusion.
34535 | 1220648-2023-03197 | 17757156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17757156&pc=OZD | 12/03/2022 | 09/15/2023 |  | 09/15/2023 |  |  | 12/03/2022 | 09/15/2023 |  |  | 06/30/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234672 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported a 63 year old male noted as high risk for cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was flow saturation issues. there was no patient harm reported. no additional information was provided. | [5 empty] | The investigation into the flow saturation issue has been completed. the pump and the data logs were returned for evaluation. the data logs were reviewed and show a sudden purge pressure increase after a purge cassette change. the purge cassette change was longer than recommended taking over 3 minutes. the root cause of the high purge pressure/ saturated flow was determined to be blood ingress that occurred during a long purge cassette change, this is most likely a use related issue. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34536 | 1220648-2023-03198 | 17757233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17757233&pc=OZD | 08/19/2023 | 01/24/2024 |  | 09/15/2023 |  |  | 08/19/2023 | 09/15/2023 |  |  | 12/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310668 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 44-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. a right groin hematoma was noted at the end of the procedure, and so a femstop was applied. upon explant, surgeon was unable to achieve hemostasis with perclose. the patient was brought to or, surgical repair with 2prbc and 2 platelets transfusions as well. patient remained on impella support for the next two days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. per the clinical information provided, the root cause of the access site bleeding issue was most likely use related due to improper technique. high stick is not following best impella practices.
34537 | 1220648-2023-03199 | 17757426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17757426&pc=OZD | 08/17/2023 | 03/13/2024 |  | 09/15/2023 |  |  | 08/17/2023 | 09/15/2023 |  |  | 08/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272562 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the pump was inserted but was not achieving optimal flows not even after repositioning. the pump and sheath were removed, and the sheath was clotted off. a new sheath and pump were inserted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella device was not returned for evaluation. data logs show a spike in motor current and active suction alarm, followed by low pump flow during the start of the case. according to clinical records, the pump was replaced after observing the motor current spike and decrease in pump flow. the cause of the low pump flow issue was most likely biomaterial, based on data log review. added-h6 impact code 4628.
34538 | 1220648-2023-03200 | 17757509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17757509&pc=OZD | 08/18/2023 | 11/26/2025 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  |  |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1128492 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 08/18/2023 |  |  |  |  | The abiomed field service representative reported that the user facility's automated impella controller (aic) for cardiovascular operation room (cvor) in-service, got an alarm ¿controller error: switch to back-up controller¿, when connecting the white plug. the white plug was removed and reconnected in an attempt to troubleshoot the issue, but an optical sensor system error alarm appeared. the aic was swapped out with no further issues. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-03200 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-03200. e2 was erroneously selected on the initial manufacturer device report number 1220648-2023-03200. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03200. g1 reporting contact email revised on manufacturer device report 1220648-2023-03200 in accordance with updated procedures. | The investigation has been completed. data analysis: on august 18, 2023, the logs recorded abnormal optical signal parameters: the light was and the snr was. the console then displayed ¿controller error (defective optical sensor lamp) ¿ alarm¿ (event #1039). even after manually rebooting the console, the optical parameters remained abnormal and event #1039 persisted, this confirms the reported failure mode. device analysis: the console was evaluated on an engineering lab bench, but the reported failure mode ¿controller error¿ could not be reproduced. the 5s parameter measured with the test cavity was within specification. root cause: the root cause of the intermittent controller error could not be determined due to the inability to reproduce it. h3 updated.
34539 | 1220648-2023-03201 | 17757521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17757521&pc=OZD | 08/17/2023 | 09/20/2024 |  | 09/15/2023 |  |  | 08/17/2023 | 09/15/2023 |  |  | 03/31/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346150 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 11 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 11-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was taken to the operation room because impella pump stopped while intubated on va ecmo. the pump was successfully exchanged. | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was not returned for investigation. according to clinical details, the pump was exchanged with a new one after the pump stop event. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03201 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
34540 | 1220648-2023-03203 | 17757921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17757921&pc=OZD | 08/18/2023 | 11/29/2023 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 02/28/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024319714 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on ecmo and impella support for 237 hours and the patient had a cerebral hemorrhage. after consulting with the patients family, the patient expired. the cause of the cerebral hemorrhage is unknown. patient was on mcs, echmo and impella support. in addition, the patient had multiple organ failure and coagulopathy. the cause of death is unknown, but was determined to be hypocardiac function and cerebral hemorrhage. cause of death is potentially cerebral hemorrhage which was possibly caused by impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center in japan did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the root cause of the cva event was most likely patient condition related, the relationship to the impella cp is unknown.
34541 | 1220648-2023-03202 | 17757922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17757922&pc=OZD | 12/08/2022 | 03/11/2025 |  | 09/15/2023 |  |  | 12/08/2022 | 09/15/2023 |  |  | 10/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023268067 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient mistakenly got a leg get caught in the purge/driveline and tried kicking it loose and purge sidearm broke between filter and red plug. the medical staff used a needle and connected the yellow lure to it to bypass, and then took patient to or for change out to new device. there was no harm to the patient reported as a result of this incident. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." . ¿clean the connector cable with 70% isopropyl alcohol.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34542 | 1220648-2023-03204 | 17758038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17758038&pc=OZD | 08/17/2023 | 01/24/2024 |  | 09/15/2023 |  |  | 08/17/2023 | 09/15/2023 |  |  |  |  | 09/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1322913 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The abiomed field service representative reported the user facility's automated impella controller (aic) had a controller yellow alarm. a loaner aic was sent to the user facility. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) has been completed since the original report was filed. impella console was returned. the error was found to be caused by disruptions in communication between the 4-in-1 and the keyboard. the error was not able to be reproduced despite extensive testing, however preliminary evidence suggests that the issue is specifically caused by a mismatch between the power-on sequences of the 4-in-1 and keyboard controller. the cause of the aic issue was a defective kbd/display.
34543 | 1220648-2023-03205 | 17759936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17759936&pc=OZD | 03/14/2023 | 09/15/2023 |  | 09/15/2023 |  |  | 03/14/2023 | 09/15/2023 |  |  | 10/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266035 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/14/2023 |  |  |  |  | The user facility reported a 66 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while trying to come off bypass, the impella stopped and a "impella stopped/high motor current" alarm occurred. an attempt was made to restart by restarting the console, but it did not restart. the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation for the pump stop issue has been completed. the field data logs were reviewed and no high purge pressure alarms were observed. low pump flow leading to high motor current pump stops were noted. the returned pump was visually inspected and biomaterial was observed in the cannula. blood was observed in the purge gap. the pump was attempted to be activated but an immediate high motor current pump stop reproduced. the purge gap was again inspected and remnants of biomaterial were observed in the purge gap. the cause of the pump stop was biomaterial. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34544 | 1220648-2023-03206 | 17760677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17760677&pc=OZD | 08/18/2023 | 07/08/2024 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 04/30/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352835 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2023 |  |  |  |  | The user facility reported an 80-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. an iliac artery angiogram revealed small perforation which was successfully tamponaded with a balloon. | It was further reported that there was difficulty achieving optimal flows likely attributed to very small left ventricle (lv) cavity and acute bend of cannula entering av. the pump was then weaned and explanted. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the vascular damage has been completed. pump was not returned for investigation. data logs were not provided. the patient was reported to have a small ventricle. the cause of the low pump flow issue was most likely patient condition related, based on given clinical information. the cause of the vascular damage issue cannot be determined due to insufficient clinical information provided. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- b1-selected product problem, b5-additional narrative, h6 impact code 4627, h6 med dev prob code 1248, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
34545 | 1220648-2023-03207 | 17760680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17760680&pc=OZD | 08/18/2023 | 04/16/2024 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 06/30/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367259 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient was moved from the bed to the chair and the automated impella controller (aic) alarmed, pump in the aorta. it was confirmed by echo that the pump was in the aorta. the p level was reduced to p2 and the patient was taken to the operation room to have the pump removed. the pump was explanted without any complications. the patient is stable off the pump at this time. | [5 empty] | The impella device was received and an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. a catheter twist was observed around the locked area due use force. the cause of the positioning issue was patient movement related.
34546 | 1220648-2023-03208 | 17760685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17760685&pc=OZD | 08/19/2023 | 07/17/2024 |  | 09/15/2023 |  |  | 08/19/2023 | 09/15/2023 |  |  | 06/30/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357367 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 47-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient was on impella support and the ct showed a small infarct, however they had some slight movement in his rue. concern for gastrointestinal (gi) bleed due to significant drop in hgb, patient was transfused with 2 units prbc on 1 plt. despite possible bleed, bival infusion will continue due to new left ventricle (lv) thrombus seen on echo and cerebral infarct. family withdrew care after 192.5 hours and the patient expired. death is unrelated to the impella. | It was further reported that the patient expired after the procedure. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported vascular damage /gi bleed, thrombus, and stroke has been completed. the product and data logs were not returned for review. the cause of the vascular damage issue was most likely patient condition related and unknown relation with impella due to gastrointestinal (gi) bleeding observed per clinical review. without additional clinical details, the root cause of the thrombus and the stroke could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock in the section: potential adverse events (united states), which now includes the following statement "cardiac or vascular injury (including ventricular perforation)¿. added- b5 mso review d4-corrected udi number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
34547 | 1220648-2023-03209 | 17760691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17760691&pc=OZD | 12/21/2022 | 09/29/2025 |  | 09/15/2023 |  |  | 12/21/2022 | 09/15/2023 |  |  | 10/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266173 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2022 |  |  |  |  | The user facility reported a 40-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was trending haptoglobin for acute hemolysis. hemolysis was also questioned due to slight hematuria. plan remains to explant device once renal recovery has been achieved. pump was weaned and explanted 9 days after hemolysis concerns. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03209 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03209 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03209 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03209. h6 component code revised from the initial submission in accordance with updated procedures.
34548 | 1220648-2023-03210 | 17761418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761418&pc=OZD | 08/18/2023 | 07/18/2024 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 04/30/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349861 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2023 |  |  |  |  | The user facility reported a 59-year-old male post open heart surgery and on va ecmo was implanted with an impella cp for mechanical circulatory support and to vent. the patient started bleeding significantly from impella site, so femostop was placed. several hours later femostop removed, and the site was redressed but still oozing. blood products were given. the patient remained on support for ten days. | It was further reported that care was withdrawn and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleed has been completed. the device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the major access site bleed was patient condition as the patient had low platelets. it is unknown if the bleed resolved, but the patient required blood products, femstop, and site redressing. the instructions for use for the reported impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
34549 | 1220648-2023-03211 | 17761421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761421&pc=OZD | 08/18/2023 | 01/22/2024 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 04/30/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024349052 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2023 |  |  |  |  | The user facility reported a 66-year-old male (race unknown) with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. two days after implant, medical team discovered a left ventricle (lv) thrombus, located near the inlet area. it was not visible under transesophageal echocardiography the days before. to prevent any embolism, they decided to explant the pump and performed an atrioseptostomy to decrease left aortic/left ventricle pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus has been completed. the device was not returned for investigation. based on clinical description, the cause of the thrombus issue was most likely patient condition related since lv thrombus observed in patient during imaging per clinical review.
34550 | 1220648-2023-03212 | 17761427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761427&pc=OZD | 03/13/2023 | 09/15/2023 |  | 09/15/2023 |  |  | 03/13/2023 | 09/15/2023 |  |  | 11/30/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024294407 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/13/2023 |  |  |  |  | The user facility reported a male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a pump was started before the guidewire was fully removed from the device. immediately a pump stopped occurred. the pump was removed with no issue or visible damaged and the case was successfully completed with new pump. there was no patient harm reported from this user error. | [5 empty] | The investigation for the pump stop has been completed. the device and data logs were returned for evaluation. no damage or abnormalities were found on the returned pump. data logs show an immediate high motor current pump stop occurred when attempting to start the device. after reviewing the clinical information, it was determined that the cause of the pump not starting was use related, specifically failure to remove the guidewire prior to activating the device. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34551 | 1220648-2023-03213 | 17761450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761450&pc=OZD | 08/29/2022 | 09/15/2023 |  | 09/15/2023 |  |  | 08/29/2022 | 09/15/2023 |  |  | 05/31/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222067 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported that a 29-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. on the third day of support the sidearm was inadvertently broken by staff during therapy in an attempt to stop the patient from sliding off of the bed. the location of the break was noted to be near the yellow luer. the retainer was in place at the time. a purge bypass was performed, until the pump could be exchanged. of note, it was also reported that excessive force was reportedly applied. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was most likely use related as the purge sidearm broke during patient transfer when the staff applied excessive force. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34552 | 1220648-2023-03214 | 17761457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761457&pc=OZD | 05/21/2022 | 09/15/2023 |  | 09/15/2023 |  |  | 05/21/2022 | 09/15/2023 |  |  | 02/29/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023194866 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported that a 60-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the clinical team reported that they observed numerous placement signal spikes and decreased flows, even though the motor current remained stable. the placement signal metrics did not match the patient's measured blood pressure. the following day, the pump was repositioned under echo imaging. after repositioning, all placement signal issues resolved. the patient remained hemodynamically stable. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows and placement signal issue was completed. the device was not returned for evaluation, but data logs were received. analysis of the data logs showed the placement signal waveform was erratic with frequent spikes and concurrent, erratic, sudden drops in flow. after reviewing the clinical information, the root cause was most likely use-related since repositioning resolved the issue. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34553 | 1220648-2023-03215 | 17761459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761459&pc=OZD | 04/13/2022 | 09/15/2023 |  | 09/15/2023 |  |  | 04/13/2022 | 09/15/2023 |  |  | 01/01/2024 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186158 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported that a 58-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues a few days post implant. the clinical team is unsure what happened but suspects an issue with the was a tuohy borst. no further information was provided. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the available information, the root cause of the issue was not determined as the device was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34554 | 1220648-2023-03216 | 17761465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761465&pc=OZD | 06/08/2023 | 11/13/2023 |  | 09/15/2023 |  |  | 06/08/2023 | 09/15/2023 |  |  | 01/31/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328991 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 82-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant reported that after replacing the purge cassette, the purge flow never returned to the previous value. multiple checks of the line connections, cassette and transducer showed no active leaks. yet, the purge pressure remained within normal limits. the bedside nurse called approximately seven hours later stating they noticed fluid coming from the purge cassette housing. the nurse switched to a different purge cassette and the flows return to previous value. additionally, a hematoma was noted and it had appeared to resolve before going to the icu. it started again and a femostop was applied distal to cp access site and two units of blood were given due to a hemoglobin drop. | [5 empty] | The impella pump device was not received from the customer, only the purge cassette was returned and therefore, an evaluation to root cause of the bleed was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding has been completed since the original report was filed. only the purge cassette was returned for evaluation.the introducer was not returned for evaluation. as per clinical, during access through left femoral artery using single access technique doctor noted hitting something hard with the needle leading to bleeding. the cause of the access site bleeding issue was use related since during the left femoral access doctor had some resistance hitting with the needle during single access technique.
34555 | 1220648-2023-03217 | 17761505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761505&pc=OZD | 06/09/2023 | 09/22/2024 |  | 09/15/2023 |  |  | 06/09/2023 | 09/15/2023 |  |  |  |  | 09/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC |  |  | US |  | 1401317 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant called requesting assistance with changing a cassette due to ¿purge disc not detected¿ alarm. after the abiomed clinical support consultant troubleshot the issue, it was ¿failed to prime¿ alarm. a second cassette attempt was unsuccessful. the nurse obtained a backup console and transferred it over. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported "automated impella controller (aic) - purge disc/flag issues" has been completed. product was returned for investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. during data analysis, the mc logs from the complaint date revealed that the console issued the purge disc not detected alarm several times and was unable to complete case wizard. the root cause of the issue was a broken/missing purge flag. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03217 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34556 | 1220648-2023-03218 | 17761575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17761575&pc=OZD | 08/18/2023 | 03/12/2024 |  | 09/15/2023 |  |  | 08/18/2023 | 09/15/2023 |  |  | 05/31/2025 |  | 09/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363587 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2023 |  |  |  |  | The user facility reported a 64-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the doctor obtained radial access for repeat angiography and found dissection at the bifurcation. doctor emergently placed the impella and called the doctor to take patient for emergent coronary artery bypass grafting (cabg). once peel-away sheath removed, physician noted major bleeding. they stated that tech pushed part of sheath back into artery while they were peeling away which caused tear that required surgical cutdown and repair. patient transported to cvor for impella removal, cabg x 2 and direct placement of impella 5.5 for escalation. physician reported the left femoral artery was repaired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the vascular damage- great vessel issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the vascular damage issue was determined to be related to usage, as the introducer was pushed into the artery during the peeling process, as indicated in the provided clinical information.
34557 | 1220648-2023-03152 | 17743465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17743465&pc=OZD | 10/20/2021 | 09/14/2023 |  | 09/14/2023 |  |  | 10/20/2021 | 09/14/2023 |  |  | 07/31/2023 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022119320 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/20/2021 |  |  |  |  | The user facility reported that a 48-year-old male patient implanted with the impella 5.5 for mechanical circulatory support had the pump ingest a clot into motor resulting in a pump malfunction. the pump was subsequently replaced. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus and pump stop was completed. the device was not returned for evaluation. based on the available information and no device return, the root cause of the reported event is undetermined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34558 | 1220648-2023-03153 | 17743480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17743480&pc=OZD | 07/29/2022 | 09/14/2023 |  | 09/14/2023 |  |  | 07/29/2022 | 09/14/2023 |  |  | 05/31/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222019 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported that a 44-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak near the red handle. a purge sidearm bypass was performed as a result and the device was later exchanged for another 5.5. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the root cause of the purge leak was unable to be determined as no product was returned for investigation and insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34559 | 1220648-2023-03154 | 17743489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17743489&pc=OZD | 09/22/2022 | 09/15/2024 |  | 09/14/2023 |  |  | 09/22/2022 | 09/14/2023 |  |  | 06/30/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023230452 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/22/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site. the surgeon had to intervene by placing a jackson-pratt drain (jp drain) and was planning to stop suction and let drain clot off in chest. the patient condition did not improve and 5 days after this event decision was made to withdraw care. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. - medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the access site bleeding issue was most likely low/high patient act values. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03154 was provided in sections b2, b5, h1, and h6.
34560 | 1220648-2023-03155 | 17744363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17744363&pc=OZD | 08/17/2023 | 01/18/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 04/30/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350366 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 | [10 empty] | The investigation into the reported acess site bleeding has been completed since the original report was filed. data logs showed sufficient flow for the respective p-level and not evidence of a low pump flow event. product was not returned for review. based on clinical description, the root cause of access site bleeding was high acts since the oozing from the site resolved when ptt returned to a normal level. correction : the clinical code in section h6, was submitted incorrect inadvertently in the initial report which has been corrected to this report. correction: the b1 and b2 section was submitted incorrect inadvertently in the initial report which has been corrected to this report. correction : the type of reportable event in h1 section was submitted incorrect inadvertently in the initial report which has been corrected to this report.
34561 | 1220648-2023-03156 | 17744664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17744664&pc=OZD | 10/20/2022 | 03/11/2025 |  | 09/14/2023 |  |  | 10/20/2022 | 09/14/2023 |  |  | 07/31/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023251455 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 51-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during the impella implant they perforated the artery. additional access was obtained distal to perforation and two covered stents were placed. the placement of the impella was aborted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34562 | 1220648-2023-03157 | 17744693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17744693&pc=OZD | 01/26/2023 | 09/14/2023 |  | 09/14/2023 |  |  | 01/26/2023 | 09/14/2023 |  |  | 09/30/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023261060 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/26/2023 |  |  |  |  | The user facility reported a 47 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a leak from the sidearm reservoir. a purge bypass was performed. there was no product returned or other details provided. there was no patient harm reported. this occurred after 1 month of support, and after the bypass was performed the impella support continued for 2 more months. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34563 | 1220648-2023-03158 | 17744720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17744720&pc=OZD | 08/17/2023 | 11/13/2023 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 06/30/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356774 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 12-year-old female with heart failure was implanted with an impella cp device for mechanical circulatory support. while performing echo imaging on day 3 of the support, the pump was found to be in the aorta. the decision was made to remove the pump. the pump and removed and manual pressure was held. the patient is currently only medically managed, as the impella was not replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, pump position was found to be in the aorta and later could not reposition with patient having diseased heart valve condition. the cause of the positioning issue was patient condition related based on clinical review.
34564 | 1220648-2023-03159 | 17744753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17744753&pc=OZD | 01/25/2023 | 09/14/2023 |  | 09/14/2023 |  |  | 01/25/2023 | 09/14/2023 |  |  | 02/29/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023193671 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/25/2023 |  |  |  |  | The user facility reported a 70 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the yellow luer broke after replacing the purge cassette. a purge bypass was completed. there was no patient harm reported. the patient remained on support. the leak occurred on day 7 and the pump remained on for 22 more days. | [5 empty] | The investigation into the broken yellow luer. the sidearm only was returned for evaluation. upon visual inspection, the sidearm yellow luer broken was observed. leaking was observed through the connection of the yellow luer and the tubing during the purge testing. the clinical report found that sodium bicarbonate was running through the purge. the root cause of the leaking was a damaged sidearm yellow luer due to use of sodium bicarbonate in the purge solution. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34565 | 1220648-2023-03160 | 17745214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17745214&pc=OZD | 01/28/2023 | 09/14/2023 |  | 09/14/2023 |  |  | 01/28/2023 | 09/14/2023 |  |  | 09/30/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023263126 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/28/2023 |  |  |  |  | The user facility reported a male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge leak at the sidearm filet to tubing connecting joint. purge flow and pressure stabilized within acceptable ranges. there was no patient harm reported. | [5 empty] | The investigation for the sidearm leak has been completed. the pump was returned for evaluation. upon evaluation, the leak at the sidearm filter to tubing joint was reproduced. review of the clinical report found that there was excessive patient movement prior to the time of the failure. the cause of this leak is most likely due to excessive force, as patient movement was reported during support. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34566 | 1220648-2023-03161 | 17745282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17745282&pc=OZD | 08/17/2023 | 07/08/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 03/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338618 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 70-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had a large hematoma at the groin site. manual pressure, compression at groin site, and two units packed red blood cells were given. repo sheath advanced and forward tension sutures placed. the site was dressed and after 45 minutes groin site stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed/hematoma has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since additional sutures were used to attain hemostasis. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
34567 | 1220648-2023-03163 | 17746011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746011&pc=OZD | 05/18/2022 | 03/11/2025 |  | 09/14/2023 |  |  | 05/18/2022 | 09/14/2023 |  |  | 02/29/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | 2023187586 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported she was taken to the or to have the impella removed, which was found to have a clot. the medical staff had it removed and allowed it to back bleed. there was no lasting harm from the thrombosis. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Analysis summary the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
34568 | 1220648-2023-03162 | 17746020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746020&pc=OZD | 03/02/2023 | 09/14/2023 |  | 09/14/2023 |  |  | 03/02/2023 | 09/14/2023 |  |  | 08/31/2023 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139019 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/02/2023 |  |  |  |  | The user facility reported a 22 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a leak at the yellow causing low purge pressure and purge system open alarms. a purge bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. no product was returned for investigation. the data logs were reviewed and show that there were purge pressure low alarms on the day of the reported purge leak. the clinical report was reviewed finding that sodium bicarbonate was running in the purge. therefore, the cause of the purge leak was most likely sidearm yellow luer damage due to bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34569 | 1220648-2023-03164 | 17746040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746040&pc=OZD | 03/02/2023 | 09/14/2023 |  | 09/14/2023 |  |  | 03/02/2023 | 09/14/2023 |  |  | 02/29/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023189537 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/02/2023 |  |  |  |  | The user facility reported a 78 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge side arm started to leak at the yellow luer after extension tubing was added to purge. a purge reservoir and filter bypass was performed. sodium bicarbonate was used in the purge. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. neither the product nor the data logs were returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34570 | 1220648-2023-03165 | 17746132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746132&pc=OZD | 08/17/2023 | 01/22/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  |  |  | 09/14/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1474181 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Erratic or Intermittent Display | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The abiomed field service representative reported that during preventive maintenance, the user facility's automated impella controller started two times with a black screen. only the green led light was on. all softkeys and the screen were off. the console could only be restarted after an emergency shutdown. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console issue has been completed. the console was returned and tested on an engineering lab bench. the issue was unable to reproduced upon testing.the root cause of the black screen and the disabled softkeys were not determined due to the inability to reproduce.
34571 | 1220648-2023-03166 | 17746171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746171&pc=OZD | 08/16/2023 | 01/18/2024 |  | 09/14/2023 |  |  | 08/16/2023 | 09/14/2023 |  |  | 02/28/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335105 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was brought to the operation room to change the impella cp for an impella 5.5 due to hemolysis. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported hemolysis has been completed. the impella was not returned for evaluation. however, the data logs were reviewed and revealed that the pump was started and ran without issue for about 30.5 hours, and then the pump was in improper position the rest of the case. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position.
34572 | 1220648-2023-03168 | 17746467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746467&pc=OZD | 08/17/2023 | 09/17/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 06/30/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361174 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 76-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the cp support that was initiated for approximately five minutes, when the pump came out of position. the pump could not be re-delivered across the valve and so the team replaced the pump. the second pump was placed but the patient coded and expired. the death occurred after over one hour of advanced cardiovascular life support (acls) measures and shock x15. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for investigation. per the clinical information provided, the inserted pump jumped back across the aortic valve and the pump was unable to re-cross the valve. the root cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
34573 | 1220648-2023-03167 | 17746487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746487&pc=OZD | 08/16/2023 | 01/18/2024 |  | 09/14/2023 |  |  | 08/16/2023 | 09/14/2023 |  |  | 05/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352136 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 83-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was taken back to the operation room for impella cp removal because no pulse or blood flow to the lower leg. a new impella cp was advanced axillary into the left ventricle alongside existing impella cp. the wire was removed and new impella was ramped up to p5. existing cp was decreased to p0 and physician explanted the original cp without any difficulty. the physician proceeded with thrombectomy to left femoral artery (lfa) which established flow and pulses were regained per doppler. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. the cause of limb ischemia- reversible was patient condition related due to patient diseases as thrombectomy occurred and flow restored.
34574 | 1220648-2023-03169 | 17746547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746547&pc=OZD | 08/17/2023 | 07/30/2025 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  |  |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1116353 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) had an error message of unexpected controller shutdown. switched to the backup controller. there was no patient harm. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-03169 in accordance with updated procedures. e2 and e3 was erroneously selected on the initial manufacturer device report number 1220648-2023-03169. g1 reporting contact email revised on manufacturer device report 1220648-2023-03169 in accordance with updated procedures.
34575 | 1220648-2023-03170 | 17746838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17746838&pc=OZD | 08/17/2023 | 09/21/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 03/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331057 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/17/2023 |  |  |  |  | The user facility reported that a 49-year-old female with heart failure had an impella cp replaced for a new impella cp for continued support. the second impella cp was implanted for three days when a ct scan showed multiple brain infarcts. awaiting on family to withdraw care. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/ cva issue has been completed since the original report was filed. product & data were not returned for review. the cause of the stroke was most likely due to patient condition and unknown relation to impella. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03170 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34576 | 1220648-2023-03171 | 17747129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17747129&pc=OZD | 08/17/2023 | 12/06/2023 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 03/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339534 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 81-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported multiple placement signal low alarms throughout procedure due to catheter manipulation during percutaneous coronary intervention (pci). repositioned each time by physician, but continued to migrate due to lack of proper handling. furthermore, hemolysis was noted in urine after insertion and labs hemolyzing. additionally, the peel away sheath was removed at the end of the procedure and repositioning sheath advanced. notable bleeding at insertion site requiring manual pressure and fem stop. the field rep. discussed with staff potential for additional bleeding. the large hematoma noted in right groin, but was stabilizing with femostop. no blood products were given. after 56.67 of impella support, care was withdrawn and the patient expired. no allegations were made towards the impella for the cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into positioning issues and hemolysis has been completed. the impella pump was not returned for investigation. data logs were not provided. the doctor reported multiple pump repositioning due to poor handling. the cause of the positioning issue was most likely use related. the cause of the hemolysis issue was most likely improper positioning of the device, based on given clinical details.
34577 | 1220648-2023-03173 | 17747176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17747176&pc=OZD | 08/17/2023 | 09/14/2023 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 01/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331502 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp for circulatory support. the device was experiencing low flow on the pump. the device was explanted and replaced with a new impella cp. there was no patient harm. when the device was returned for investigation, broken blades on the pump were found. | [5 empty] | The device was received and an evaluation was completed. data review: data confirmed the low flow. data showed flow was in normal range when pump started. about 5 minutes into the case, the mc spiked to 1600 ma & flow suddenly dropped to 0.6 l/min, then remained for the rest of the case and triggered auto suction response & suction alarms. product analysis: visual inspection found the blades of impeller were broken. no other issues were found on the rest of the pump. conclusion: the root cause of low flow was due to fractured impeller blades. because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. the failure modes will be monitored and trended.
34578 | 1220648-2023-03172 | 17747315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17747315&pc=OZD | 05/19/2022 | 01/25/2024 |  | 09/14/2023 |  |  | 05/19/2022 | 09/14/2023 |  |  | 03/31/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023201432 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a 42-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient who had just gone impella implant was found to have a small hematoma post procedure. the medical team stopped the heparin and had the patient go to the or to have it removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
34579 | 1220648-2023-03174 | 17747958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17747958&pc=OZD | 08/18/2023 | 11/06/2023 |  | 09/14/2023 |  |  | 08/18/2023 | 09/14/2023 |  |  | 09/30/2024 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023264480 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Frayed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2023 |  |  |  |  | The user facility in japan reported an 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the team found the guidewire to be interacting and not allowing for appropriate pump delivery. the team replaced the pump and wire and, and at the explant, observed the wire to be frayed. no harm or injury from delay in support. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage of the .018 guidewire has been completed since the original report was filed. the device was returned for investigation. it was determined that the cause of the guidewire damage issue was use related since excessive force applied during the removal of guidewire.
34580 | 1220648-2023-03175 | 17748838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17748838&pc=OZD | 08/17/2023 | 01/18/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 05/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362104 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp was removed after 5 days of support due to possible infection. a new impella cp was inserted. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported infection /sepsis has been completed since the original report was filed. no product was returned. the cause of the infection was most likely patient condition related since the device is sterilized under validated conditions and there was no evidence to indicate a break in the sterile barrier or pump contamination, unknown relation to impella.
34581 | 1220648-2023-03176 | 17748891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17748891&pc=OZD | 08/15/2023 | 01/16/2025 |  | 09/14/2023 |  |  | 08/15/2023 | 09/14/2023 |  |  |  |  | 09/14/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1078425 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Charging Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that despite plugging the console in and charging the battery, the battery would continually alarm as ¿battery not charging". the console was exchanged. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) battery charging issue has been completed. aic logs confirmed the reported complaint. data logs confirm battery failure battery failure (transport connection blown/missing) - alarm issued. there are numerous instances of the alarm listed above in the logs. the aic was returned for evaluation. the failure mode was reproduced in the lab. the console reported not being able to charge past 50% when connected to a.c. the lcd of battery one was blank, indicative of battery lockout. this led to not being able to recharge locked out battery. it was confirmed by batttest that cell2 of battery1 was at significantly lower voltage than the rest of the cells within the battery. additionally, presence of residue build up around purge pressure sensor area was observed. the reported complaint of the console not recharging was due to battery lockout. this form of battery failure was due to a defective battery. the root cause of the defective battery was single cell failure, a known pattern failure. corrections to the initial mfr report: d2 updated common device name g1 mdr reporting contact email updated h8 usage of device updated to reuse.
34582 | 1220648-2023-03179 | 17749032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749032&pc=OZD | 08/17/2023 | 11/03/2023 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 05/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365431 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on extracorporeal membrane oxygenation (ecmo) and impella and during guide catheter insertion, the impella was accidentally pulled out of ventricle and the doctor was unable to re-advance. the staff then proceeded to do complex percutaneous coronary intervention (pci), had some pci complications that caused the case to be extremely long. after pci, the staff then tried again and again to re-advance impella, but was unsuccessful. doctor was able to implant a new impella with success when re-insertion failed with the prior pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the root cause of the positioning issues is most likely due to the wireless re-access as the pump was accidently pulled back into the aorta and was unable to wirelessly re-cross.
34583 | 1220648-2023-03180 | 17749048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749048&pc=OZD | 08/16/2023 | 02/02/2024 |  | 09/14/2023 |  |  | 08/16/2023 | 09/14/2023 |  |  | 04/30/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024341826 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 56-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the nurse called abiomed rep and stated that the impella cp was upgraded to impella 5.5 and was running "fine" but suddenly had purge blocked alarm. there were no kinks or issues found. the patient has bicarb in d5w for purge fluid within the device. the medical staff was advised regarding tpa protocol. no other info available form nurse at this time. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported high purge pressure has been completed since the original report was filed. the pump was returned for investigation. a kink was observed near the 45cm mark. the pump ran successfully during the test without reporting any purge-related alarms. data logs indicated sudden spike in purge pressure with purge system block alarm, which occurred 5 hours after implantation and persisted until the end of the case, with the purge flow close to zero. the cause of the high purge pressure issue was a kinked purge line found on catheter near the 45cm mark. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | Corrected g3.
34584 | 1220648-2023-03177 | 17749081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749081&pc=OZD | 08/15/2023 | 01/19/2024 |  | 09/14/2023 |  |  | 08/15/2023 | 09/14/2023 |  |  | 04/30/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351105 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. a fem-fem bypass was done the night of (b)(6) 2023 with good result. there was signs of hemolysis noted. the impella was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported hemolysis and limb ischemia has been completed. product nor data logs were returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of limb ischemia was most likely due to patient's condition of diseased/small vessels.
34585 | 1220648-2023-03181 | 17749155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749155&pc=OZD | 08/17/2023 | 07/30/2025 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  |  |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090368 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The abiomed field service representative reported the user facility's automated impella controller (aic) had a controller failure alarm. this was discovered during a preventive maintenance. the console was diagnosed, and the power and battery management (pbm) module had a failure to communicate. after carefully inspecting the pbm the caps were leaking, and some corrosion had occurred. the pbm will be replaced. | [5 empty] | The abiomed field service engineer (fse) determined the power and battery management (pbm) caps were leaking and some corrosion had occurred. follow up with the fse has not been completed to confirm part was replaced and aic is back in service. a post market investigation is required. upon completion, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-03181 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-03181. e2 was erroneously selected on the initial manufacturer device report number 1220648-2023-03181. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03181. g1 reporting contact email revised on manufacturer device report 1220648-2023-03181 in accordance with updated procedures.
34586 | 1220648-2023-03178 | 17749172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749172&pc=OZD | 08/16/2023 | 01/11/2024 |  | 09/14/2023 |  |  | 08/16/2023 | 09/14/2023 |  |  | 01/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310711 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 82-year-old female noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp was implanted and turned on, soon after the impella decreased flows and immediately suction. fluoroscopy of the cannula showed noticeable kink. explant decided when impella flows had not improved. new impella cp implanted and used during procedure without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed since the original report was filed. the device was returned for investigation. data logs confirmed the low flow when pump started & remained to the rest of the case. visual inspection found a kink on can about 13 mm from of cage & can bonding site. no other issues were found on the rest of the pump. based on clinical description, product analysis & data review, the root cause of low flow was due to canula kink. the cause of cannula kink was due to patient condition. the failure mode will be monitored and trended.
34587 | 1220648-2023-03182 | 17749191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749191&pc=OZD | 08/16/2023 | 03/01/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 01/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024309520 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support on (b)(6) 2023. the patient suffered a cerebral hemorrhage on (b)(6) 2023. the relationship to the impella is unknown. patient received transfusions during impella support. patient received ecmo treatment as well, and the cerebral hemorrhage occurred post ecmo while on impella support. physician could not conclude if a thrombus occurred from iabp, impella, or ecmo. on (b)(6) 2023 the user facility reported that the patient was on impella support for 202.5 hours before the patient expired. the cause of death was the patient was discontinued from treatment and switched to palliative treatment because the neurological prognosis could not be expected to improve due to cerebral hemorrhage. the automated impella controller and the impella cp pump was functioning without problems until death. the relationship between impella and death is unknown because the relationship between cerebral hemorrhage and impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
34588 | 1220648-2023-03184 | 17749245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749245&pc=OZD | 08/17/2023 | 11/17/2025 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  |  |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087202 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 08/17/2023 |  |  |  |  | The user facility reported they were setting up for a procedure when the automated impella controller (aic) displayed a failure alarm. it would resolve on its own and then start again. the aic was taken out and new console used for the procedure. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-03184 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-03184. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03184. g1 reporting contact email revised on manufacturer device report 1220648-2023-03184 in accordance with updated procedures.
34589 | 1220648-2023-03183 | 17749271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749271&pc=OZD | 08/17/2023 | 01/22/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 05/22/2023 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338697 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 15-year-old female (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. the team experienced complications during insertion due to the patient's anatomy. the catheter cannula kinked after multiple attempts and the team was unable to advance. the decision was made to abort and go to ecmo. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for investigation. product was not returned for review. based on clinical description, the cause of the delivery issues- use was patient condition related due to vessels stenosis that obstructed the insertion process and cannula kink.
34590 | 1220648-2023-03185 | 17749274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749274&pc=OZD | 08/17/2023 | 09/22/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 04/30/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient was high risk and underwent a percutaneous coronary intervention (pci). patient decompensated and cardiopulmonary resuscitation (cpr) was initiated. the impella was pushed out of the left ventricle during cpr and was unable to cross back. the impella did cross and was on. the impella was removed due to positioning issues. the patient expired. | Medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the positioning issue was use related due to improper technique while impella pushed out of lv during cpr and was unable to cross back. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03185 provided in sections b1, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34591 | 1220648-2023-03186 | 17749332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749332&pc=OZD | 08/17/2023 | 01/23/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 04/30/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348518 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 75-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was taken to cath lab for explant. anticoagulation had been held for an extended amount of time due to gi bleeding. the impella was explanted and closed using perclose x 2. the team said they had pulses at time of explant. several hours later in ccu patient lost left leg pulses, patient was brought back down to cath lab around 8pm, found to had large clot burden, angiojet was able to aspirate down to knee. ic was unable to restore flow below knee area. vascular was consulted and patient was taken to or for fasciotomy and open thrombectomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as per clinical, after the pump explant, the patient lost pulses in the impella leg, and the doctor discovered a large clot extending down to the knee. the root cause of the limb ischemia issue was patient condition related to clot interaction.
34592 | 1220648-2023-03187 | 17749372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17749372&pc=OZD | 08/17/2023 | 03/13/2024 |  | 09/14/2023 |  |  | 08/17/2023 | 09/14/2023 |  |  | 05/31/2025 |  | 09/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369479 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while repositioning the impella cp, the doctor pulled the impella back across the valve into ascending aorta and was unable to successfully re-cross. patient brought back down to ccl to replace the impella cp. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for reported positioning issue has been completed. the pump was not returned for investigation. according to the provided clinical information, the impella was pulled out into the aorta during the repositioning of the pump. the cause of the positioning issue was most likely use related based on given clinical details.
34593 | 1220648-2023-03133 | 17735547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17735547&pc=OZD | 10/08/2021 | 09/13/2023 |  | 09/13/2023 |  |  | 10/08/2021 | 09/13/2023 |  |  | 07/31/2023 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118242 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/08/2021 |  |  |  |  | The user facility reported that a 43-year-old patient implanted with the impella 5.5 for mechanical circulatory support had a clot noted at the entry of the inlet cage that extended into the inner lumen of the cannula during device explant. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. visual inspection confirmed biomaterial in the outflow cage and on the catheter. kinks in the catheter were also noted. the root cause of the reported event was most likely biomaterial that was observed on the pump at the time of pump removal. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34594 | 1220648-2023-03134 | 17735590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17735590&pc=OZD | 07/11/2022 | 09/13/2023 |  | 09/13/2023 |  |  | 07/11/2022 | 09/13/2023 |  |  | 04/30/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217105 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported that a 54-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low purge flow. the physician suspected a clot ingestion, and the patient was subsequently taken for pump exchange. biomaterial was reported to be observed in the outflow. the patient's act values were unknown. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported flow alarm was completed. the device was not returned for evaluation. based on the provided clinical information, the root cause of the reported event was most likely due to biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34595 | 1220648-2023-03135 | 17735629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17735629&pc=OZD | 06/26/2022 | 09/13/2023 |  | 09/13/2023 |  |  | 06/26/2022 | 09/13/2023 |  |  | 03/31/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203088 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2022 |  |  |  |  | The user facility reported that a 66-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak due to a crack at the yellow luer on the 18th day of support. a purge bypass was performed without issue. the patient was reported to have sodium bicarbonate in the purge solution. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was determined to be due to a crack in the yellow luer. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34596 | 1220648-2023-03136 | 17735699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17735699&pc=OZD | 09/03/2022 | 09/21/2024 |  | 09/13/2023 |  |  | 09/03/2022 | 09/13/2023 |  |  | 05/31/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023219366 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/03/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient in a serious medical condition was taken to the or to have impella implanted. four days later when the medical team was weaning off sedation the computed tomography (ct) scan showed that the patient had experienced a stroke and the medical team questioned the medical status. the patient status did not improve which lead the medical team to withdraw care and the patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03136 was provided in sections b2, b5, h1, and h6.
34597 | 1220648-2023-03137 | 17735752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17735752&pc=OZD | 08/14/2023 | 01/18/2024 |  | 09/13/2023 |  |  | 08/14/2023 | 09/13/2023 |  |  | 04/30/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339354 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2023 |  |  |  |  | The user facility reported a 72-year-old patient (unknown gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp support was ongoing when the pump came out of position. the pump's fixing ring had not been tightened. the cp was removed and replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, touhy-borst was not tighten during time of positioning issue. the root cause of the positioning issue was loose touhy-borst related, based on given clinical details.
34598 | 1220648-2023-03138 | 17735817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17735817&pc=OZD | 08/25/2022 | 01/26/2024 |  | 09/13/2023 |  |  | 08/25/2022 | 09/13/2023 |  |  | 05/31/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023228246 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a 30-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was found to have continued bleeding and a hematoma at the access site. the medical team had to reopen pocket and evacuation of hematoma. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34599 | 1220648-2023-03139 | 17735832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17735832&pc=OZD | 08/15/2023 | 09/21/2024 |  | 09/13/2023 |  |  | 08/15/2023 | 09/13/2023 |  |  | 05/31/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352394 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 33-year-old female in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had a impella 5.5 pump support ongoing when on day three there was an observation made of neuro change/cva/stroke. the team made decision along with the family to withdraw care. the patient expired. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03139 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the reported stroke/cerebral vascular accident (cva) has been completed since the original report was filed. data was reviewed and no issues were found on data logs. the product was not returned for review. the cause of the stroke was most likely due to patient condition and unknown relation to the impella.
34600 | 1220648-2023-03140 | 17737552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17737552&pc=OZD | 08/15/2023 | 02/06/2024 |  | 09/13/2023 |  |  | 08/15/2023 | 09/13/2023 |  |  | 05/31/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363589 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was experiencing hemolysis. the urine output had decreased, and labs have been hemolyzed. the patient was successfully explanted the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation has been completed since the original report was filed. no device was returned for investigation. data logs showed about 35 hours into support, mc suddenly spiked to 1000 ma and flow dropped to zero that triggered auto suction response and suction alarms. the pump continued running in this situation for 3 hours 38 minutes, and then was stopped manually and disconnected from the console. no positioning issue was found on the data logs. based on clinical description and data analysis, most likely, biomaterial had been ingested into the pump and caused hemolysis. hemolysis was worse as the pump continued running with biomaterial inside the pump housing. the root cause of hemolysis was most likely was due to biomaterial ingestion.
34601 | 1220648-2023-03141 | 17737586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17737586&pc=OZD | 08/14/2023 | 07/31/2024 |  | 09/13/2023 |  |  | 08/14/2023 | 09/13/2023 |  |  | 04/30/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337520 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old female noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. the complaint reported that a pump exchange due to hemolysis and suction was performed. the patient was suspected to have a form of hit, with elevated fibrinogen levels causing near occlusive clot in the inlet. additionally, patient's ef is <5%, causing possible stasis in the left ventricle (lv). there was no thrombus noted. the pump was explanted and successfully exchanged. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for hemolysis has been completed. the impella device was not returned for evaluation. data logs were returned and investigated. there were numerous "impella position unknown" and "impella in aorta" alarms noticed throughout the case. there were also suction events noted throughout the case. no motor current spikes were observed. per the clinical information and data log review, the root cause of the hemolysis issue was likely improper positioning of the device due to improper technique. instructions for use : impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03141: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. g1 reporting contact fax number missing.
34602 | 1220648-2023-03142 | 17738551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17738551&pc=OZD | 08/16/2023 | 11/12/2023 |  | 09/13/2023 |  |  | 08/16/2023 | 09/13/2023 |  |  | 05/31/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023211906 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility in the eu reported patient had high purge pressure during impella 5.0 support. there was probably thrombus according to the medical team; however, no effect on impella flow. the pump remained on for a support of 6 days til explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported high purge pressure has been completed since the original report was filed. no product was returned for evaluation. data logs were not available for review of high purge pressure issue. as per clinical, 5.0 pump was used for 6 days, and high purge pressure alarm was found. the cause of the high purge pressure issue was not determined since product and data logs were not returned for review.
34603 | 1220648-2023-03143 | 17740508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740508&pc=OZD | 08/16/2023 | 07/16/2024 |  | 09/13/2023 |  |  | 08/16/2023 | 09/13/2023 |  |  | 04/30/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338704 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2023 |  |  |  |  | The user facility reported a 59-year-old female was implanted with an impella 5.5 as a bridge to transplant. the patient experienced a small hematoma at the access site and was scheduled for a washout. post hematoma washout, found to be seroma. a jackson pratt drain was put in place with minimal drainage. patient successfully weaned to heart transplant surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock in the following sections: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
34604 | 1220648-2023-03144 | 17740739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740739&pc=OZD | 08/14/2023 | 08/09/2024 |  | 09/13/2023 |  |  | 08/14/2023 | 09/13/2023 |  |  | 04/30/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355032 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2023 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the impella was primed and while removing the red cheater, the physician noticed part of "cheater" remained in the pump. a decision was made to replace the device before entering the pump through the sheath. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues- use has been completed. the impella device was not returned for evaluation. the cause of the issue was a red lumen removal issue. d6a/d6b added. h6 component code added f6/f8-should have been left blank on manufacturer device report 1220648-2023-03144.
34605 | 1220648-2023-03145 | 17740756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740756&pc=OZD | 08/17/2023 | 01/23/2024 |  | 09/13/2023 |  |  | 08/17/2023 | 09/13/2023 |  |  | 05/31/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024340913 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported a 60-year-old male was implanted with an impella 5.5 as a bridge to transplant. a hematoma started to form and has slowly gotten worse even with pressure. the patient had a hematoma evacuation. systemic heparin was stopped, and purge was switched back to sodium bicarb. patient received his heart transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding issue has been completed. the device was not returned for investigation. as per clinical, patient had hematoma around impella with calcified vessel conditions in axillary. hematoma evacuation surgical intervention was done. the cause of the access site bleeding issue was most likely patient condition related with patient having calcified vessel leading to bleeding around sheath per clinical review.
34606 | 1220648-2023-03146 | 17740844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740844&pc=OZD | 09/18/2022 | 01/15/2024 |  | 09/13/2023 |  |  | 09/18/2022 | 09/13/2023 |  |  | 06/30/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234266 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding from the access site and required blood products. additionally the site was re-dressed. the pump remained on for 3 more weeks. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34607 | 1220648-2023-03147 | 17740877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740877&pc=PYX | 08/17/2023 | 09/21/2024 |  | 09/13/2023 |  |  | 08/17/2023 | 09/13/2023 |  |  | 01/31/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024322361 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Deformation | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/17/2023 |  |  |  |  | The user facility reported an 83-year-old female with right heart failure was implanted with an impella cp and an impella rp for mechanical circulatory support. the rp was found to have a kink during insertion. in addition, the rp stopped after a day of support when moving the pump from femoral to axillary insertion. this move was due to the leg flow being compromised at the cp insertion. the pump stop was due to the team stopping the pump per conversation with field service representative. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03147 was provided. note: corrected information provided in h6 is an addition to previous coding. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation the root cause of the pump unable to start was a kinked cannula due the pump being inserted too deep. the pump was pulled back and successfully restarted.
34608 | 1220648-2023-03148 | 17740886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740886&pc=OZD | 09/13/2022 | 09/21/2024 |  | 09/13/2023 |  |  | 09/13/2022 | 09/13/2023 |  |  | 06/30/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023230381 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced thrombus in several locations and additionally the patient had poor flow in the leg. the medical team also said that the patient had bowel ischemia, however, after failed attempts they were unable to open up the superior mesenteric artery. the patient condition did not improve in the following days and decision was made to withdraw care leading to the patient expiry. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03148 was provided in sections b2, b5,d6, h1, and h6.
34609 | 1220648-2023-03149 | 17740906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740906&pc=OZD | 08/16/2023 | 01/19/2024 |  | 09/13/2023 |  |  | 08/17/2023 | 09/13/2023 |  |  | 05/31/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024351423 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2023 |  |  |  |  | The user facility in the eu reported a 74-year-old male (race unknown) with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the pump experienced reduced impella flow rates. the patient experienced bleeding form impella insertion site, which was questioned to be retroperitoneal bleeding. surgical intervention to repair vascular damage occurred and then the patient was escalated to an impella 5.5 and va ecmo. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the low pump flow and the access site bleeding ¿ major has been completed since the original report was filed. the pump was returned with the catheter cut off and missing the red pump plug. visual inspection found kink on cannula about 14 mm from of cage and cannula bonding site. no other issues were found on the rest of the pump. it was determined that the root cause of low flow was most likely due to a cannula kink and the root cause of access site bleeding was most likely due to lack of vessel recoil.
34610 | 1220648-2023-03150 | 17740922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17740922&pc=OZD | 09/19/2022 | 03/11/2025 |  | 09/13/2023 |  |  | 09/19/2022 | 09/13/2023 |  |  | 06/30/2024 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | 2023234459 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2022 |  |  |  |  | The user facility reported a 46-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team was having difficulty delivering the impella pump due to the patient's anatomy. during this attempt, as the long sheath passed into the graft-native vessel anastomosis the graft tore off the artery. case was aborted and vascular repair done. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34611 | 1220648-2023-03151 | 17741065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17741065&pc=OZD | 08/16/2023 | 01/18/2024 |  | 09/13/2023 |  |  | 08/16/2023 | 09/13/2023 |  |  | 05/31/2025 |  | 09/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366822 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 77-year-old female patient had a impella cp placed for hemodynamic support while awaiting bypass (cabg) surgery. when the patient was in the icu for continued monitoring there was a bleed around the reposheath. the team explanted and replaced the pump with a 14fr cook sheath and infused (b)(4) units of blood products. the cp pump had remained on for support of just 6.75 hours. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of access site bleeding was most likely due to a lack of vessel recoil.
34612 | 1220648-2023-03110 | 17729151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17729151&pc=OZD | 07/14/2022 | 09/22/2024 |  | 09/12/2023 |  |  | 07/14/2022 | 09/12/2023 |  |  | 09/30/2023 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022140818 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/14/2022 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a high purge pressure and low purge flow during impella support. a purge bypass was completed to resolve the noted issue and support continued. there was no patient harm. |  |  |  |  | The investigation into the reported high purge pressures has been completed since the original report was filed. no product was returned. the root cause of the high purge pressure issue was not determined since product was not returned and not enough clinical information is provided. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03110 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
34613 | 1220648-2023-03111 | 17729185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17729185&pc=OZD | 08/15/2023 | 07/08/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  | 03/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327934 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding/oozing at the introducer sheath. the team did not need any extensive intervention, or any vessel repair. the team continued with the 5.5 support. the team infused one unit of red blood cells. support is ongoing at this time without lasting harm or injury. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 impact code 4642, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34614 | 1220648-2023-03112 | 17729218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17729218&pc=OZD | 06/17/2022 | 09/12/2023 |  | 09/12/2023 |  |  | 06/17/2022 | 09/12/2023 |  |  | 02/29/2024 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197320 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2022 |  |  |  |  | The user facility reported that a 46-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the patient also had low purge pressure and high purge flows. a purge bypass was performed. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was returned for evaluation. visual inspection revealed that the purge tubing was cut for the bypass. the sidearm was not returned for evaluation. the root cause of the issue was a leak from the sidearm; however, the exact location of the leak was unknown. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34615 | 1220648-2023-03113 | 17729246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17729246&pc=OZD | 08/15/2023 | 03/25/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  | 07/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371809 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/15/2023 |  |  |  |  | Additional details have been added to the narrative for this event. the complainant reported that a patient was implanted with an impella 5.5 pump for bailout. the team had access site and delivery issues leading to multiple attempts before successful implant. post placement, blood was seen in the patient¿s chest and a perforation was discovered in the back of the left ventricle. the perforation proceeded to grow rapidly and the pump was removed. the patient was placed back on bypass and the lv was surgically repaired. a second pump was prepped for implant, but the team was unable to successfully deliver the second pump. the patient was medically managed and expired within an hour of arriving in the icu. | The user facility reported a 65-year-old female noted as post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 device for a bailout situation. the team had access site and delivery issues. post placement, the team observed the left ventricle (lv) to be perforated. the 5.5 was removed for surgical repair of the lv. afterwards the team was unable to deliver a new 5.5 pump and the patient was medically managed and expired. |  |  |  |  | The investigation for the reported ventricular perforation/vessel damage has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the vessel damage was use related. | Additional information added to b2, b5, h1, and h6 to coincide with the report of the patient's passing. should additional relevant information become available, a supplemental report will be submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
34616 | 1220648-2023-03114 | 17729311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17729311&pc=OZD | 08/09/2023 | 01/18/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  |  |  | 09/12/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1716739 |  | 0048-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/15/2023 |  |  |  |  | The user facility reported a 45-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp support was ongoing when the automated impella controller (aic) failed to document on the console display the purge. there was no harm or injury and the support proceeded. the team later noted the integrated amount to be displayed. | [5 empty] | The investigation into the aic - display issue has been completed since the original report was filed. the console was returned for evaluation and was tested after arriving in fs. the investigation determined that the console had performed within specification. there was no issue found during the case. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34617 | 1220648-2023-03115 | 17729374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17729374&pc=OZD | 08/13/2023 | 06/24/2024 |  | 09/12/2023 |  |  | 08/13/2023 | 09/12/2023 |  |  | 03/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346505 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the femoral access site developed a hematoma. the hematoma prompted the team to hold pressure, femostop, and suture, as well as deliver 1 of rbcs. the pump remained on for support of 3 days till wean and explant. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the product was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34618 | 1220648-2023-03116 | 17729428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17729428&pc=OZD | 08/15/2023 | 07/08/2024 |  | 09/12/2023 |  |  | 08/16/2023 | 09/12/2023 |  |  | 05/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024356114 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2023 |  |  |  |  | The user facility reported a 71-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician felt resistance when inserting and pulled the device out and found a kink in the shaft part proximal to the motor. the impella cp was replaced with another. there was no patient harm or injury. | [5 empty] | The investigation for the delivery issues has been completed. product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to use related. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. added-h6 impact code 4629; h6 med dev pro code 2682 added- d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34619 | 1220648-2023-03117 | 17730027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17730027&pc=OZD | 04/26/2022 | 11/12/2023 |  | 09/12/2023 |  |  | 04/26/2022 | 09/12/2023 |  |  | 07/21/2023 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022115149 |  | 0550-0008 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site which required the patient to receive several blood products. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the device was not returned by the medical facility. no product was returned. the root cause of the access site bleeding issue was most likely patient condition related, there were multiple site bleeding after the impella was successfully implanted and the relation to impella is unknown as per the clinical description provided. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.¿ | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ , ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ , ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ , this report is being filed as part of a retrospective review of historical records.
34620 | 1220648-2023-03118 | 17730618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17730618&pc=OZD | 08/15/2023 | 03/01/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  | 03/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343456 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant had been a protect iv study patient in (b)(6) 2023. in (b)(6) 2023 the crf form documented a harm that was unknown to the abiomed rep. the crf stated post procedure, the patient had groin complication requiring patient to go to or for femoral cut down and vascular repair due to limb ischemia. the patient was managed and discharged, then readmitted for further groin complication requiring wound vac treatment. | [5 empty] | The investigation into the limb ischemia ¿ irreversible event has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34621 | 1220648-2023-03119 | 17731611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17731611&pc=OZD | 08/14/2023 | 04/29/2024 |  | 09/12/2023 |  |  | 08/14/2023 | 09/12/2023 |  |  | 01/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328512 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Frayed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2023 |  |  |  |  | The user facility reported that the 0.018 guidewire for an impella cp kit was frayed upon opening the package. the wire was not used. another impella cp box was opened and a new 0.018 wire was utilized. no patient harm. | [5 empty] | The investigation for the product damage has been completed. the impella pump was not returned. clinical details were not sufficient to determine the root cause. the cause of the product damage (guidewire damage) issue was not determined since product/images were not returned for investigation. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states in the caution section: inspect the impella set packaging while opening. in the event that any key components, including its end seal labels, are damaged excessively during shipment, the use of a back-up impella set should be considered. to conform to updated procedures, sections d6 and h10 have been revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34622 | 1220648-2023-03120 | 17731940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17731940&pc=OZD | 08/14/2023 | 01/19/2024 |  | 09/12/2023 |  |  | 08/14/2023 | 09/12/2023 |  |  | 03/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024331602 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 40-year-old patient received support with extracorporeal membrane oxygenation (va-ecmo) and an impella 5.5 smartassist heart pump due to heart failure and cardiogenic shock. it was reported that the patient suffered a cerebral hemorrhage and died while on impella support. no further clinical details are currently available. | [5 empty] | The investigation into a patient's stroke/cerebral vascular accident has been completed. product was not returned. data logs were not provided. patient was given 25 (u/ml) of heparin during the case. the cause of the stroke issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34623 | 1220648-2023-03121 | 17732132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17732132&pc=OZD | 08/14/2023 | 01/17/2024 |  | 09/12/2023 |  |  | 08/14/2023 | 09/12/2023 |  |  | 03/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343836 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp triggered an impella stopped retrograde flow red alarm during support. troubleshooting was attempted, however a high motor current alarm then appeared. the power to the pump was cycled, but attempts were unsuccessful at restarting the pump. as a result of the persistent issue, the decision was made to explant the device. there was no patient harm. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. the pump was returned for investigation. a large impeller purge gap was observed. pump connected to aic and pump stop alarm reproduced. per design trace matrix 0048-9007, pump use exceeded the design life as the the pump stopped on day 11th. the root cause of the pump stop was bearing wear due to misuse of the product. | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿, ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿, ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34624 | 1220648-2023-03122 | 17732320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17732320&pc=OZD | 08/15/2023 | 01/24/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  |  |  | 09/12/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1441400 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/15/2023 |  |  |  |  | The user facility reported the automated impella controller (aic) displayed an error alarm. the team exchanged the aic for another. there are no patient details. no harm was reported. | [5 empty] | The investigation into the aic - controller failure (red alarm) has been completed since the original report was filed. imc logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error). the console was sent back fs for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of re-occurrence. after analyzing the logs from the console, it has been determined that the reported issue is a known pattern failure caused by a temporary loss of optical placement signal. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34625 | 1220648-2023-03123 | 17732415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17732415&pc=OZD | 08/15/2023 | 03/04/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  | 05/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368091 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient has an ischemic leg with climbing lactic acid of 10. the impella cp was exchanged for impella 5.5. | [5 empty] | The investigation into the reported limb ischemia reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34626 | 1220648-2023-03124 | 17732495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17732495&pc=OZD | 08/15/2023 | 07/24/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  | 10/31/2024 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285489 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2023 |  |  |  |  | The user facility reported a 63-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. upon insertion of the impella, an aortic dissection was noted. due to this being an elective case, decision was made to abort case and close groin. the groin was accessed using a micro puncture, then 6f, and serial dilated to impella sheath. an angiogram was performed through micro puncture. an abiomed wire had been advanced into the vasculature prior to the dissection being observed. the cause of the dissection was not determined. the dissection was treated by aborting the case to not cause any further issues and risk worsening dissection as it was contained, and the patient was stable. | [5 empty] | The investigation for vascular damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added 925. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03124. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34627 | 1220648-2023-03125 | 17732591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17732591&pc=OZD | 08/16/2023 | 01/16/2024 |  | 09/12/2023 |  |  | 08/16/2023 | 09/12/2023 |  |  | 04/30/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353493 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. once the device was implanted, the patient coded again and was shocked multiple times. percutaneous cardiac insertion was performed and shockwave was used. following this, the patient coded again and was shocked. after the final shock, the optical sensor appeared to fail and the left ventricle and aortic placement wave forms were no longer detected and the ¿placement signal not reliable¿ alarm was displayed across the top of the automated impella controller (aic). checking cable for kinks and use of other aic was attempted but was not successful. the impella pump was removed changed out with another impella cp. the patient was not harmed from this malfunction. | [5 empty] | The investigation into the reported optical signal issue has been completed since the original report was filed. the returned product was investigated and damage to the sensor face was observed. the damage was consistent with shockwave interaction. the 5s parameters were tested prior to sensor removal and were found to be outside of normal range. the lt shows a change is 5s parameters consistent with damage to the sensor occurring when there was a loss of placement signal. the root cause of the optical signal issue was a damaged sensor due to shockwave interaction. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34628 | 1220648-2023-03126 | 17732670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17732670&pc=OZD | 06/04/2023 | 10/26/2023 |  | 09/12/2023 |  |  | 06/04/2023 | 09/12/2023 |  |  | 01/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329159 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2023 |  |  |  |  | The user facility reported an 87-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support during the coronary angioplasty. the pump stopped working, however, the md decided that the patient was in a stable condition, the impella pump was explanted. no medical intervention was required as a result of the pump stopping. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. per the ifu: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency." | The investigation into the pump stop has been completed since the original report was filed. the product was returned for investigation. pump was run to reproduce the failure mode. pump could not start under bench test. no evidence of the patient cable damage or pump plug motor cable disconnect were found. visual inspection found a damage on motor cable inside the catheter. patient cable was able to freely rotate/translate relative to kink protector. kink protector was removed & no glue was seen on the inner surface. the root cause of the pump stop was an open electrical line. the exact location of the electrical open line could not be identified due to accidental damage to catheter during tear down. the cause of the product damage issue was bond failure due to insufficient adhesive at the kink protector of the patient cable.
34629 | 1220648-2023-03127 | 17732983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17732983&pc=OZD | 08/15/2023 | 01/18/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  | 05/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358569 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 49-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had placed a cp for support when there was blood observed between the patient's legs. the blood was from the groin ooze site. the team ordered blood of unknown quantity, placed a femostop, and the fellow placed a stitch. the blood loss occurred on day 1 of the support, and the pump remains on to date today of 8 days of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the pump was not returned for investigation. access site bleeding was reported after transfer to intensive care unit. the doctor successfully managed the bleeding with sheath manipulation, manual pressure, and the femoral stop technique. the cause of the access site bleeding issue was most likely use related.
34630 | 1220648-2023-03128 | 17733040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17733040&pc=OZD | 08/14/2023 | 01/22/2024 |  | 09/12/2023 |  |  | 08/14/2023 | 09/12/2023 |  |  | 01/31/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024316877 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. the patient was reported to be very restless and attempted to stand up. the pump catheter was pulled heavily. the pump stopped with excessive motor current and purge pressure. the impella cp pump could not be restarted. abiomed customer support recommended using a new impella heart pump. the patient was reasonably stable 10 minutes after removal of the impella cp, but a falling mean arterial pressure was reported from 54 to 50mmhg. the patient was already receiving high doses of catecholamines. the patient was transferred to the intensive care unit in critical condition. no new impella heart pump was used. the following day it was reported that the patient died the night after the pump stopped. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. the cause of the pump stop issue was not determined since product and data logs were not returned for investigation.
34631 | 1220648-2023-03129 | 17733111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17733111&pc=OZD | 08/16/2023 | 01/18/2024 |  | 09/12/2023 |  |  | 08/16/2023 | 09/12/2023 |  |  | 04/30/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC |  |  | US |  | 2024346120 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 90-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a impella cp explanted and replaced/escalated to the 5.5 due to access site bleeding from the impella cp and hematoma development. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. according to clinical details, the patient was not bleeding from the puncture site, but rather from blood vessels on the central side. there was no hematoma around the insertion site at the time of removal, and there were no problems with the insertion technique. the relation of the vascular injury/bleeding site to the impella is unknown. the cause of the vascular damage/bleeding issue was most likely patient condition related.
34632 | 1220648-2023-03130 | 17733122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17733122&pc=OZD | 08/15/2023 | 09/12/2023 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  |  |  | 09/12/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1274732 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the impella aic purge disc holder is broken off the automated impella controller (aic). | [5 empty] | An investigation into the purge disc retainer was completed and the flag was confirmed to be broken. during data analysis, the automated impella controller (aic) ((b)(6)) had the case started on (b)(6), 2023 and did not progress past step one before the console is requested to shutdown. the root cause of the issue was a missing purge flag.
34633 | 1220648-2023-03131 | 17733133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17733133&pc=OZD | 08/15/2023 | 07/31/2024 |  | 09/12/2023 |  |  | 08/15/2023 | 09/12/2023 |  |  | 04/30/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353491 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella cp kinked distal to outlet when transitioning from aortic arch to root, during delivery. a new impella was used. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for delivery issues-use has been completed. product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to use issue. instructions for use. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system section: warnings and cautions: warnings to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. h6 medical device problem code added d6a and d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03131: g1 reporting contact fax number missing.
34634 | 1220648-2023-03132 | 17733139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17733139&pc=OZD | 08/16/2023 | 02/02/2024 |  | 09/12/2023 |  |  | 08/16/2023 | 09/12/2023 |  |  | 04/30/2025 |  | 09/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP- EU |  |  | ABIOMED, INC. |  |  | US |  | 2024349312 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2023 |  |  |  |  | The user facility reported a 66-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support to unload the left ventricle. the patient was also under ecmo support. there were no problems during preparation. once the impella cp started, "high purge pressure" and "purge system blocked" alarms triggered. there was no leak in purge tubing, no visible thrombus in lv or aortic root, good position across the aortic valve, no structure on outlet. they decided to change the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure has been completed since the original report was filed. the console was returned for investigation. the cause of the high purge pressure issue was due to biomaterial ingestion in the purge gap near the bearing side per field investigation and data log review.
34635 | 1220648-2023-03087 | 17720683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17720683&pc=PYX | 08/12/2023 | 09/11/2023 |  | 09/11/2023 |  |  | 08/12/2023 | 09/11/2023 |  |  | 07/31/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023281616 |  | 1000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 69-year-old male in st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. the complainant reported that the impella rp had a spike in motor current and the pump exhibited low flows. the impella rp was removed. | [5 empty] | The investigation into the reported low pump flow has been completed. product was returned for investigation. the device was visually inspected and found biomaterial adherred. there were no other abnormality observed. as per clinical investigation, the impella rp had a spike in motor. the pulmonary artery pressure was severely elevated, and the waveform was lost. the impella flows was 0 on p5. the impella rp was removed. the root cause of the low pump flow issue was biomaterial ingestion due to patient condition related.
34636 | 1220648-2023-03088 | 17720735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17720735&pc=OZD | 08/13/2023 | 07/31/2024 |  | 09/11/2023 |  |  | 08/13/2023 | 09/11/2023 |  |  | 04/30/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355062 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the surgeon had a delivery failure of the impella cp. the cp kinked at insertion attempts; therefore, the insertion was aborted. a new impelal cp was placed and supported for 15 days. | [5 empty] | The investigation for delivery issues-use has been completed. complaint device not returned for investigation. the cause of the delivery issue was patient condition related due to patient anatomy. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ h6 medical device problem code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03088: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the ifu: "handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
34637 | 1220648-2023-03089 | 17720747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17720747&pc=OZD | 08/13/2023 | 03/13/2024 |  | 09/11/2023 |  |  | 08/13/2023 | 09/11/2023 |  |  | 06/30/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357365 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 86-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had a 12 day support with a 5.5 pump when the patient wrapped foot around and broke the white cable and caused the pump to stop. the 5.5 pump was replaced and support proceeded with a new pump via the same access site. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the pump issue was an open electrical line at the red handle, resulting from excessive force, as indicated by the provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34638 | 1220648-2023-03090 | 17720897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17720897&pc=OZD | 12/21/2022 | 09/11/2023 |  | 09/11/2023 |  |  | 12/21/2022 | 09/11/2023 |  |  | 05/31/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023228120 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2022 |  |  |  |  | The user facility reported a 68 year-old male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that the purge side arm tubing broke when the patient got the tubing tangled up trying to ride a stationary bicycle. the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm break has been completed. the device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge sidearm break was related to excessive force applied to the sidearm by the patient. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34639 | 1220648-2023-03091 | 17720916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17720916&pc=PYX | 08/14/2023 | 09/23/2024 |  | 09/11/2023 |  |  | 08/14/2023 | 09/11/2023 |  |  | 02/28/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024333407 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2023 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during the procedure and once the pump had started, the impella had a stop alarm, the pump stopped and restarted twice. positioning was verified under fluoroscopy and echo. additionally, there was a high purge pressure that alarmed twice. the decision was made to remove the rp flex and insert a new pump. there was no patient harm. | [5 empty] | The investigation for the pump stop and the purge leak has been completed. the data logs were observed and there were motor current spikes followed by high purge pressure alarms as well as high motor current pump stop alarms, but the pump restarted within seconds. the pump was returned and inspected. there were denatured proteins found in the purge gap as well as around the ball bearing. the root cause of the high purge pressure is most likely due to biomaterial. the pump ran in a bucket of water for 66 hours without a pump stop occurring and it passed all testing. the root cause of the temporary pump stop cannot be determined as the issue was not able to be reproduced. added- d6a/d6b h6 med dev prob code 2954 changed to 1491. f6/f8 should have been left blank in the initial report. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. per the ifu: "unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
34640 | 1220648-2023-03092 | 17720947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17720947&pc=OZD | 08/13/2023 | 01/30/2025 |  | 09/11/2023 |  |  | 08/13/2023 | 09/11/2023 |  |  |  |  | 09/11/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1451015 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Inadequate User Interface | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/13/2023 |  |  |  |  | The user facility reported a patient was being transferred while being supported with an impella cp and an automated impella controller (aic). the aic was dropped and damaged. the user error prompted the team to replace the aic. there was no harm or injury to the patient. | [5 empty] | The investigation for the console (aic) damage has been completed. data logs were reviewed and no anomaly was revealed. the aic was returned for evaluation. inspection of the console revealed a detached front panel optical connector, a crack on the blue purge retainer was noticeable which was unrelated to the complaint. there was no other noticeable damage and console performed to specification when ran with a known good pump and purge line. the root cause for the failure was use issue. added h6 impact code 4629 corrections to the initial mfr report: d2 corrected common device name. d4-updated model number. f6/f8 should have been left blank. g1 mdr reporting contact email. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34641 | 1220648-2023-03093 | 17720961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17720961&pc=OZD | 12/20/2022 | 09/11/2023 |  | 09/11/2023 |  |  | 12/20/2022 | 09/11/2023 |  |  | 09/30/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023257844 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2022 |  |  |  |  | The user facility reported a 55 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after a purge cassette and bag change, there was a cracked yellow luer while swapping the purge system. a purge bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34642 | 1220648-2023-03094 | 17721093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17721093&pc=OZD | 08/13/2023 | 01/16/2024 |  | 09/11/2023 |  |  | 08/13/2023 | 09/11/2023 |  |  | 04/30/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355041 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/13/2023 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. when the physician was adjusting the impella cp, it came out of position and the physician was unable to re-deliver to the left ventricle. the device was removed and replaced with another impella cp. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. based on clinical description, the root cause of the positioning issue is use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34643 | 1220648-2023-03095 | 17721208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17721208&pc=OZD | 11/27/2022 | 09/11/2023 |  | 09/11/2023 |  |  | 11/27/2022 | 09/11/2023 |  |  | 08/31/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023256400 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/27/2022 |  |  |  |  | The user facility reported a 58 year-old female with acute cardiac arrest was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the yellow connector on the purge sidearm was leaking. there was sodium bicarbonate running in the purge with extension tubing in place. patient was previously declining prior to product issue and a decision was made to withdraw care. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the returned product was visually inspected and a crack in the yellow luer was observed. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. per the ifu: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34644 | 1220648-2023-03096 | 17721740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17721740&pc=OZD | 08/12/2023 | 09/30/2025 |  | 09/11/2023 |  |  | 08/12/2023 | 09/11/2023 |  |  | 02/28/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024323375 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/12/2023 |  |  |  |  | The user facility in the eu reported an 80-year-old patient (unknown gender and race) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device alarmed with placement signal issues. the physician verified the placement of the device and the position was correct. in addition, the device had a sudden increase in purge pressure and a drop in purge flow. the values were within range, however, the physician had a concern for a thrombus ingestion. the impella cp was swapped with another cp. there was no patient harm. | [5 empty] | The investigation into the optical signal issue and the high purge pressure has been completed since the original report was filed. data logs were returned and reviewed finding that the placement signal is low throughout the entire case. however, the alarm is triggered at the same time as they are in suction conditions. there are small motor current spikes and a decrease in flow without a p-level change at the time of the suction conditions. additionally, purge pressure begins to rise while purge flow is decreasing. the device was returned for evaluation and was tested after arriving in fs. it was reported that biomaterial was wrapped around. the root cause of the optical signal issue is most likely due to the patient¿s condition, while the high purge pressure is most likely caused from the ingested biomaterial. | The impella device was received from the customer and, an evaluation of the device is pending. upon investigation closure, a supplemental mdr will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the placement signal is low throughout the entire case. however, the alarm is triggered at the same time as they are in suction conditions. there are small motor current spikes and a decrease in flow without a p-level change at the time of the suction conditions. additionally, purge pressure begins to rise while purge flow is decreasing. the product was returned and biomaterial was wrapped around the optical signal issue l is most likely due to the patient¿s condition, while the high purge pressure is most likely caused from the ingested biomaterial. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
34645 | 1220648-2023-03097 | 17721868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17721868&pc=OZD | 01/12/2023 | 09/11/2023 |  | 09/11/2023 |  |  | 01/12/2023 | 09/11/2023 |  |  | 02/29/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023189493 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the purge luer cracked after a purge cassette change. a purge sidearm bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm crack has been completed. the pump and the data logs were returned for evaluation. data logs showed recurrent "purge system open" and "purge pressure low" alarms began following a purge cassette change. additionally, according to the clinical report, the purge solution running was sodium bicarbonate up until 24 hours prior to the purge leak when they started heparin. visual inspection of the returned pump revealed several cracks on the sidearm yellow luer. no other damage or abnormalities were found. therefore, it was determined that the cause of the purge leak was sidearm yellow luer damage due to the use of sodium bicarbonate in the purge solution. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34646 | 1220648-2023-03098 | 17722489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17722489&pc=OZD | 01/17/2023 | 09/11/2023 |  | 09/11/2023 |  |  | 01/17/2023 | 09/11/2023 |  |  | 11/30/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023275520 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/17/2023 |  |  |  |  | The user facility reported a 69 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during a routine purge cassette change by the staff nurse, and removing bicarbonate extension tubing, the yellow luer on the purge side arm cracked, requiring purge bypass. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm crack has been completed. the pump was returned for evaluation. upon inspection, a crack in the yellow luer of the sidearm was confirmed. the clinical report review showed that sodium bicarbonate had been running in the purge solution. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34647 | 1220648-2023-03099 | 17722757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17722757&pc=OZD | 05/23/2023 | 09/11/2023 |  | 09/11/2023 |  |  | 05/23/2023 | 09/11/2023 |  |  | 04/30/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330350 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/23/2023 |  |  |  |  | The user facility reported a 70-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during repositioning, the physician had great difficulty due to the integrated system holding too much rotational tension and it was not translating to repositioning in the body/native anatomy. upon attempting to troubleshoot, he pulled the device into the aorta and then had to advance it back across the valve. the procedure was completed successfully and there was no harm to the patient as a result of this positioning event. | [5 empty] | The investigation into the positioning issue has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the positioning issue was most likely a catheter use issue based on the provided clinical information. the failure modes will be monitored and trended.
34648 | 1220648-2023-03100 | 17725686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17725686&pc=OZD | 08/12/2023 | 01/16/2024 |  | 09/11/2023 |  |  | 08/12/2023 | 09/11/2023 |  |  | 03/31/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357292 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had access site bleeding. nurse held pressure and replaced with a fem stop. anticoagulants were halted as a result of the bleed and support continued with the implanted pump. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34649 | 1220648-2023-03101 | 17726305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726305&pc=OZD | 01/03/2023 | 09/11/2023 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  | 11/30/2023 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023167093 |  | 0550-0008 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2023 |  |  |  |  | The user facility reported a 63 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a purge line leak had occurred on the pump. a crack was noted on the yellow luer. the pump was replaced as a result. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was returned for evaluation finding a damaged yellow luer upon inspection. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34650 | 1220648-2023-03102 | 17726330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726330&pc=OZD | 06/01/2022 | 09/11/2023 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  |  |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022156733 |  | 0550-0008 |  | No |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/01/2022 |  |  |  |  | The user facility reported that a 56-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the leak at the yellow luer was noted while performing a purge cassette. a purge bypass was subsequently performed without issue. of note, the patient had extension tubing and sodium bicarbonate in the purge for a week. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was due to sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34651 | 1220648-2023-03103 | 17726334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726334&pc=OZD | 08/14/2023 | 07/08/2024 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  | 03/31/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344858 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pseudoaneurysm in left subclavian site of the impella cp. a balloon tamponade was performed from groin. a 7.0x60 balloon in subclavian for tamponade, took shot and pseudo-aneurysm in subclavian. then, a covered stent was applied to the subclavian. | [5 empty] | The investigation for the vascular damage has been completed. the pump was not returned for investigation. during cp removal, a pseudoaneurysm was noted at the left subclavian site. a covered stent was placed to resolve the issue. the cause of the vascular damage issue was most likely patient condition related. the instructions for use referenced in the initial report is for the impella cp with smart assist system in section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ d4-corrected udi number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34652 | 1220648-2023-03104 | 17726337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726337&pc=OZD | 02/18/2022 | 09/11/2023 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  | 11/30/2023 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166326 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/18/2022 |  |  |  |  | The user facility reported that a 54-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced malpositioning as the device was pulled into the aorta when the patient was getting out of bed. the sterile sleeve was not connected to the suture pad hub, so the pump had to be replaced. no further information was provided. there was no harm reported to the patient. | [5 empty] | The investigation into the reported malpositioning was completed. the device was not returned for evaluation. based on the information received, the root cause of the positioning issues was determined to be use related as patient movement caused the pump to be pulled into the aorta. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34653 | 1220648-2023-03105 | 17726346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726346&pc=OZD | 02/15/2022 | 09/11/2023 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  | 10/31/2023 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157672 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/15/2022 |  |  |  |  | The user facility reported that a 78-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced malpositioning. while repositioning, the pump was pulled back across the aortic valve. multiple attempts were made to wirelessly recross the valve, but they were unsuccessful. the device was removed as a result and a new pump was placed without issue. there was no harm reported to the patient. | [5 empty] | The investigation into the reported malpositioning was completed. the device was not returned for evaluation. analysis of the data logs confirmed the pump entered the aorta and that the correct positioning was unable to be reestablished. the root cause of the positioning issues was determined to be use related as the device was accidently pulled back across the aortic valve during repositioning. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34654 | 1220648-2023-03106 | 17726351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726351&pc=OZD | 08/14/2023 | 03/15/2024 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  | 05/31/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362355 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported upon an impella cp support for a 57-year-ol male patient who had arrested at home and was admitted with st elevation myocardial infarction. the team intervened upon the coronary arteries with impella support. the team noted recurring ventricular tachycardia (vt) when trying to wean the pump and reposition the pump. the vt was resolved with pump being pulled back to the aorta. | [5 empty] | The investigation into the ventricular tachycardia (vt) issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34655 | 1220648-2023-03107 | 17726372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726372&pc=PYX | 08/14/2023 | 08/09/2024 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  | 02/28/2025 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024333412 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male noted as high risk percutaneous cardiac patient was implanted with an impella rp device for mechanical circulatory support. the impella rp flex team reported there was a self-resolving pump stop alarm, the case continued after event. the pump was restarted. the patient was on impella support for 289.97 hours before the family withdrew care and the patient expired. no allegations were made towards the impella for the cause of death. | [5 empty] | The investigation of pump stop has been completed. the cause of the temporary pump stop was not determined because no product was returned, and not enough clinical information was given. per medical safety officer review use of the device cannot conclusively be associated with the outcome of death g1 reporting contact fax number missing on manufacturer device report 1220648-2023-03107. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34656 | 1220648-2023-03108 | 17726407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726407&pc=OZD | 09/07/2022 | 09/11/2023 |  | 09/11/2023 |  |  |  | 09/11/2023 |  |  | 05/31/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222299 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported that a 47-year-old male patient implanted with the impella 5.5 for mechanical circulatory support was suspected to have heparin-induced thrombocytopenia (hit). there was no heparin in the d5w purge fluid. the patient was administered multiple doses of tissue plasminogen activator (tpa) for low purge flows without success. he was subsequently taken for pump exchange. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure was completed. the device was not returned for evaluation. analysis of the data logs revealed a gradual rise and fall of purge pressure throughout the case. based on all of the information, including the clinical evidence that no heparin was used in the purge; the root cause of high purge pressure is most likely biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34657 | 1220648-2023-03109 | 17726439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17726439&pc=OZD | 07/07/2022 | 09/11/2023 |  | 09/11/2023 |  |  | 07/07/2022 | 09/11/2023 |  |  | 03/31/2024 |  | 09/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023205505 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported that a 21-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced intermittent, and self-resolving, high purge pressure and low flow alarms over several hours. the patient was sleeping, and no movement was recorded when the alarms triggered. additionally, no purge line kinks were observed, and heparin was used in the purge solution. seven days later, purge fluid was observed to be leaking out of the purge cassette, though no obvious damage or loose connections were found. the leaking resolved after the purge cassette was replaced. no further issues were reported. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure and low flow alarms was completed. the purge cassette and data logs were returned for evaluation. analysis of the data logs confirmed several instances of sudden spikes in purge pressure over a 6-hour period. all other pump performance metrics were normal. visual inspection of the purge cassette passed. functional testing of the purge cassette revealed leaking from the piston seal. the root cause of the purge leak was a piston failure. the root cause of the high purge pressure was not determined since the pump was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34658 | 1220648-2023-03075 | 17716637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716637&pc=OZD | 08/10/2023 | 01/24/2024 |  | 09/09/2023 |  |  | 08/10/2023 | 09/09/2023 |  |  | 03/31/2025 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327087 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella was not implanted because the cannula developed a kink when trying to advance into the vessel. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage/cannula kink has been completed. visual inspection found multiple kinks on cannula & catheter below the motor housing. no other issues were found on the rest of the pump. the cause of the delivery issue was patient condition related since resistance was met at anastomosis leading to kink in the cannula.
34659 | 1220648-2023-03076 | 17716638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716638&pc=OZD | 08/10/2023 | 01/22/2024 |  | 09/09/2023 |  |  | 08/10/2023 | 09/09/2023 |  |  | 01/31/2025 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP APAC |  |  | ABIOMED, INC. |  |  | US |  | 2024311907 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the automated impella controller displayed a red alarm stating "high purge pressure (>1000). the purge flow was less than 2. the pump was not removed at that time. then, the patient endured distal limb ischemia after insertion of the impella device. after giving the patient nitroglycerine and applying warmth, the perfusion was partially restored. the device was explanted the following day due to limb ischemia. the patient showed good signs of heart recovery and was maintaining mean arterial pressure. | [5 empty] | The investigation into the limb ischemia ¿ reversible and the high purge pressure has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the root cause of the limb ischemia issue was patient condition related to tortuous vessels. the root cause of the high purge pressure issue was not determined since product was not returned, logs were not returned, and insufficient clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34660 | 1220648-2023-03077 | 17716639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716639&pc=OZD | 08/10/2023 | 09/26/2024 |  | 09/09/2023 |  |  | 08/10/2023 | 09/09/2023 |  |  | 05/31/2025 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352121 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that while attempting to access the right femoral artery a perforation occurred. the 77-year-old female patient expired on the table. | [5 empty] | The investigation into the vascular damage - peripheral vessel has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of vascular damage was most likely due to patient condition. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2 and h1. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34661 | 1220648-2023-03078 | 17716640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716640&pc=OZD | 08/11/2023 | 01/15/2024 |  | 09/09/2023 |  |  | 08/11/2023 | 09/09/2023 |  |  | 11/30/2023 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160227 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 19-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that proper steps were not taken and yellow to yellow were not properly connected in order to purge the impella prior to inserting it in the patient¿s body. the pump was entered through the sheath momentarily into the body and removed without connecting purge line. blood entered the motor gap junction, and the pump could not be used with confidence, and it was replaced. this event did not cause patient¿s harm. | [5 empty] | The investigation for the reported purge leak has been completed. the pump was not returned for evaluation. however, clinical details state that the proper steps were not followed when connecting the purge components. therefore, the cause of the issue was use related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34662 | 1220648-2023-03080 | 17716641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716641&pc=OZD | 07/16/2023 | 09/09/2023 |  | 09/09/2023 |  |  | 07/16/2023 | 09/09/2023 |  |  |  |  | 09/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 37 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/16/2023 |  |  |  |  | The user facility reported that, after a failed attempted to implant impella cp on a patient, when the pump was removed it was noticed that automated impella controller (aic) purge disk holder was broken. there was no harm to the patient as a result of this incident. | [5 empty] | The aic was returned for evaluation and repair. the cause of the issue was a broken purge pressure sensor retainer. it is unknown how the retainer was broken. the failure modes will be monitored and trended.
34663 | 1220648-2023-03079 | 17716642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716642&pc=OZD | 08/11/2023 | 03/12/2024 |  | 09/09/2023 |  |  | 08/11/2023 | 09/09/2023 |  |  | 04/30/2025 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351151 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 49-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a purge leak occurred inside the red plug and could not be resolved by bypass or any other means. recommendation was given to replace the pump due to safety issues and pump functionality being compromised. the pump was replaced. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the device was not returned for investigation. the origin of the leak cannot be confirmed without the returned product. the cause of the purge leak issue was not determined since product was not returned and sufficient clinical details was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34664 | 1220648-2023-03081 | 17716726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716726&pc=OZD | 05/15/2022 | 01/11/2024 |  | 09/09/2023 |  |  | 05/15/2022 | 09/09/2023 |  |  | 01/31/2024 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023185169 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2022 |  |  |  |  | The user facility reported a 31-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was noted to have a hematoma at the access site which required the patient to go in the or to have it surgically removed and have a drain installed. | [5 empty] | The investigation into the access site bleeding-major has been completed. the device was not returned for investigation. the cause of the access site bleeding was most likely due to patient condition since bleeding was also observed from multiple sites along with insertion site. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34665 | 1220648-2023-03082 | 17716728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716728&pc=OZD | 05/16/2022 | 01/25/2024 |  | 09/09/2023 |  |  | 05/16/2022 | 09/09/2023 |  |  | 02/29/2024 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186193 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a hematoma at the access site which required to be washed out. after the procedure the pain and swelling decreased. | [5 empty] | The investigation for the access site bleed has been completed. however, with limited clinical details nor the impella device, the root cause of the reported issue could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34666 | 1220648-2023-03083 | 17716729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716729&pc=OZD | 05/16/2022 | 03/11/2025 |  | 09/09/2023 |  |  | 05/16/2022 | 09/09/2023 |  |  | 10/31/2023 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153878 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a clot which was giving the device purge pressure alarms. tpa (tissue plasminogen activator) protocol was started, pump was doing better and there were no more purge pressure alarms. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency) relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined" this report is being filed as part of a retrospective review of historical records.
34667 | 1220648-2023-03084 | 17716732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716732&pc=OZD | 05/27/2022 | 03/11/2025 |  | 09/09/2023 |  |  | 05/27/2022 | 09/09/2023 |  |  | 06/30/2023 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022105483 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2022 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed clots during impella support. details pertaining to the location of the clots was not provided. the patient was taken to the operating room where the pump was explanted and an intra-aortic balloon pump was placed for ongoing patient support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
34668 | 1220648-2023-03085 | 17716747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716747&pc=OZD | 06/09/2023 | 09/09/2023 |  | 09/09/2023 |  |  | 06/09/2023 | 09/09/2023 |  |  | 04/30/2025 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333567 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Failure to Prime | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/09/2023 |  |  |  |  | The user facility reported that a 79-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak and priming issues. when the purge was connected to perform the initial prime it would not prime past the purge cassette. staff disconnected from the yellow luer but the line still would not prime. a second purge cassette was retrieved and primed successfully. the device was then inserted, and purge flow and pressure were within normal limits. while still in procedural areal however, leaking was noted at the yellow luer connection to the purge sidearm. the purge line was de-aired and visual inspection of the purge sidearm and yellow luer did not reveal any cracks. the purge cassette was changed again, and no further issues were noted. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. two purge cassettes were returned for evaluation. functional testing of purge cassette #1, did not reproduce the priming issue. functional testing of purge cassette #2 passed. no leak was observed near the yellow luer. analysis of the data logs indicated that purge cassette #1 experienced a priming issue and cassette #2 was replaced after a leak was observed. the root cause of the priming issue was use related since the purge cassette #1 functioned/operated to specification. the root cause of the purge leak issue was use related since the purge cassette #2 functioned/operated to specification. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34669 | 1220648-2023-03086 | 17716750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716750&pc=OZD | 09/08/2022 | 09/09/2023 |  | 09/09/2023 |  |  | 09/08/2022 | 09/09/2023 |  |  | 06/24/2024 |  | 09/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237455 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported that a 70-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure 5-10 minutes into the case. the purge cassette was swapped but the high purge pressure remained. the pump was subsequently exchanged. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure was completed. the device was not returned for evaluation. analysis of the data logs revealed a spike in motor current followed by an increase in purge pressure. based on the available information, the root cause is most likely biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34670 | 1220648-2023-03036 | 17710443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710443&pc=OZD | 08/11/2023 | 07/24/2024 |  | 09/08/2023 |  |  | 08/11/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2023 |  |  |  |  | The user facility reported a 79-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. support was ongoing when the patient showed signs of hemolysis. the plasma-free hemoglobin (pfhg) was elevated. once the impella cp was explanted and replaced with an impella 5.5, the pfhg resolved. | [5 empty] | The investigation for hemolysis has been completed. the product was not returned for investigation. data logs show the placement signal trending down at the end of the case run, without any noted position-related abnormalities or motor current spikes. as per the clinical notes, hemolysis was suspected on (b)(6). additionally, on (b)(6), the patient was given a fluid bolus that helped resolve the hemolysis. the cause of the hemolysis issue was most likely patient condition related, based on data log review. h6 component code added 925. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03036: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34671 | 1220648-2023-03037 | 17710484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710484&pc=OZD | 08/11/2023 | 01/16/2024 |  | 09/08/2023 |  |  | 08/11/2023 | 09/08/2023 |  |  | 04/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351151 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2023 |  |  |  |  | The user facility reported a 43-year-old female in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the purge solution visually was leaking slowly from red plug on both sides. dampness was visualized on sheets where red plug sat as well. pressure and flow of purge trending towards low purge pressure alarms. the impella was replaced. the user facility reported there was no patient harm. | [5 empty] | The investigation into the purge leak ¿ pump has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the red pump plug leak was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34672 | 1220648-2023-03038 | 17710532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710532&pc=OZD | 08/10/2023 | 01/17/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 04/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355091 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2023 |  |  |  |  | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had left femoral artery groin site bleeding issues after the impella was implanted for 5 days. hemoglobin (hgb) was 7.6 and a unit of packed red blood cells (prbc¿s) was transfused. patient was being kept heavily sedated to keep flat to help with bleeding problems. at the time of explant a day later the groin was managed by manual compression and femostop placement. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. therefore, the most likely root cause of the access site bleed was user-related due to patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34673 | 1220648-2023-03039 | 17710565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710565&pc=OZD | 08/10/2023 | 10/03/2025 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359249 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | Positioning Problem |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 08/10/2023 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 74-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. there was difficulty with positioning the impella from case start due to patient anatomy. ultimately inserted with pigtail in papillary muscle per echo with continuous suction. attempted to reposition under echo at bedside with improvement in position and better flows, although still not appearing fully free of papillary muscle. shortly after repositioning, the motor current dampened despite a very pulsatile aorta (ao) placement signal. there was a concern for clot, that could be one of the inlet cage windows was blocked by ventricular structure. the decision was made to remove and send the pump back to inspect for clot. |  |  |  |  | The investigation into the reported positioning issues and low pump flow has been completed since the original report was filed. the root cause of the optical signal issue is due to the patient¿s condition as the trending placement signal was still pulsatile when the motor current flattened due to the suction conditions the impella was in. while in the suction conditions, there was low flow due to the patient¿s anatomy. the root cause of the low flows is most likely due to the patient¿s condition. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08108 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was received from the customer and investigation of it is on-going at this time. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. all pertinent information available to abiomed has been submitted at this time. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated.
34674 | 1220648-2023-03040 | 17710593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710593&pc=OZD | 08/10/2023 | 03/01/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 03/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2023 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support after the patient was unstable and "coded". a ct scan of the patient's head revealed a stroke. the patient's family decided to withdraw care after 3 days of impella support. the patient expired. | [5 empty] | The investigation into the reported stroke/cva has been completed since the original report was filed. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34675 | 1220648-2023-03041 | 17710636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710636&pc=OZD | 08/11/2023 | 01/23/2024 |  | 09/08/2023 |  |  | 08/11/2023 | 09/08/2023 |  |  | 01/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024314394 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2023 |  |  |  |  | The user facility reported a 70-year-old female (race unknown) noted as high risk percutaneous cardiac (hrpci) insertion was implanted with an impella cp device for mechanical circulatory support. during pci, the patient bled heavily from the access site. the patient was treated successfully with blood products. when the patient arrived in icu heparin was given. the vascular surgeon declined surgery. the physician could not say whether the patient suffered a vessel perforation or dissection as the patient bled into the thigh. the patient expired during the night. | [5 empty] | The investigation into the access site bleeding issue has been completed. the device was not returned for investigation. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34676 | 1220648-2023-03042 | 17710693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710693&pc=OZD | 08/10/2023 | 12/11/2023 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 06/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356786 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 14x13 peel-away was placed to right femoral artery (rfa). patient with large pannus/bmi. once the dilator was removed, the sheath bowed/kinked. the physician attempted to force impella cp through the peel-away sheath, but it would not advance. the impella device was removed and noted to be damaged from insertion attempt. peel-away exchanged to 14x25. a new impella device was opened and implanted successfully. | [5 empty] | The investigation into the reported delivery issues- use has been completed since the original report was filed. the device was not returned for analysis. the root cause of difficulty inserting pump through introducer was most likely due to use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34677 | 1220648-2023-03043 | 17710718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17710718&pc=OZD | 03/28/2022 | 01/25/2024 |  | 09/08/2023 |  |  | 03/28/2022 | 09/08/2023 |  |  | 11/30/2023 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023167078 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site and that required the patient to receive two units of blood. | [5 empty] | The investigation for the access site bleed has been completed. however, with limited clinical details nor the impella device, the root cause of the reported issue could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34678 | 1220648-2023-03044 | 17711877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17711877&pc=OZD | 04/24/2022 | 09/08/2023 |  | 09/08/2023 |  |  | 04/24/2022 | 09/08/2023 |  |  | 12/31/2023 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023173215 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/24/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that during implant of the impella device, the wire got kinked going into artery due to some calcification. the medical team switched to a new wire, the procedure was completed successfully and there was no harm to the patient reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
34679 | 1220648-2023-03045 | 17711899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17711899&pc=OZD | 11/14/2022 | 09/08/2023 |  | 09/08/2023 |  |  | 11/14/2022 | 09/08/2023 |  |  | 06/30/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023230426 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 60 year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a small leak was noted at the yellow leur. initially the purge pressure was stable but purge pressure suddenly dropped and an attempt was made to purge bypass. after the bypass, the purge pressures increased. bicarbonate was noted to be used in the purge system. there was no patient harm reported. | [5 empty] | The investigation for the leaking sidearm has been completed. only the sidearm was returned. the sidearm was attached to a syringe filled with dyed purge fluid and pressurized. the leak was confirmed at the yellow luer. the yellow luer fracture was imaged. data log review shows the purge pressure increase temporarily after the leak, and then resolve. the purge pressure temporarily increases another three times during the rest of the case, but is resolved and does not impact pump function. thus, the root cause of the purge leak was due to a yellow luer damage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34680 | 1220648-2023-03046 | 17711946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17711946&pc=OZD | 04/14/2022 | 11/13/2023 |  | 09/08/2023 |  |  | 04/14/2022 | 09/08/2023 |  |  | 01/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023178702 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The user facility reported a 19-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team was concerned about an infection at the surgical site and that lead to the patient having the pump exchanged for a new one. | [5 empty] | The investigation into patient's infection/sepsis has been completed. no product was returned. the root cause of the infection/sepsis issue was patient condition related, as per the clinical description provided and the unknown relation to impella. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
34681 | 1220648-2023-03047 | 17711975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17711975&pc=OZD | 04/14/2022 | 03/11/2025 |  | 09/08/2023 |  |  | 04/14/2022 | 09/08/2023 |  |  | 11/30/2023 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023162201 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was found to have a clot in the right axillary sheath. decision was made to pull the graft partially out and clean the clot out. the 5.5 pump remained on for support for 7 days til weaned and explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency) relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined" this report is being filed as part of a retrospective review of historical records.
34682 | 1220648-2023-03048 | 17712095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712095&pc=OZD | 11/12/2022 | 09/08/2023 |  | 09/08/2023 |  |  | 11/12/2022 | 09/08/2023 |  |  | 01/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182454 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/12/2022 |  |  |  |  | The user facility reported a 65 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a fluid leak from the purge sidearm during patient support. a purge sidearm bypass was performed in an attempt to resolve the reported issue, however the purge flow later dropped overnight and the decision was made to remove the pump. there was no patient harm reported. | [5 empty] | The investigation for the purge line leak and the high purge pressures has been completed. a pump without the sidearm and data log were returned for evaluation. the pump was visually inspected. no kink observed. biomaterial observed in impella. no other abnormalities were found. the pump was connected at p-2 and p-9 to reproduce high purge pressure, however no high purge pressure was reproduced. data logs showed that the purge pressure high and purge system blocked alarms occurred after the purge pressure suddenly spiked over 1000 mmhg. the cause of the purge leak was not determined since the sidearm was not returned. the cause of the high purge pressure was most likely ingested biomaterial. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34683 | 1220648-2023-03049 | 17712112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712112&pc=OZD | 11/16/2022 | 09/08/2023 |  | 09/08/2023 |  |  | 11/16/2022 | 09/08/2023 |  |  | 06/30/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237096 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 57 year-old male in heart failure/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a red alarm for purge system open and there was a leak at the junction of the yellow connectors. a purge bypass was performed. bicarbonate was being used in the purge. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34684 | 1220648-2023-03050 | 17712505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712505&pc=OZD | 08/10/2023 | 01/22/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 03/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212693 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. there were concerns for decreased perfusion to left lower extremity (lle) though patient still had doppler able pulses. decision was made to remove the impella. vascular surgery removed impella and performed surgical repair of arteriotomy without issue. | [5 empty] | The investigation into the reported access site bleeding ¿ major and limb ischemia has been completed since the original report was filed. product was not returned for review. the root cause of the limb ischemia was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34685 | 1220648-2023-03051 | 17712536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712536&pc=OZD | 08/10/2023 | 07/31/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 04/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350757 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella was inserted and developed a kink due to patient anatomy, the physician chose to explant and proceed with percutaneous coronary intervention (pci) to dominant right coronary artery (rca). patient became hypotensive, neosynephrine was given and the physician decided to place second impella device at that time. | [5 empty] | The investigation for delivery issues-use has been completed. pump was not returned for investigation. doctor reported kink on pump during initial delivery. the cause of the delivery issue was most likely patient condition related. doctor aborted implant due aortic stenosis and hypertrophic cardiomyopathy heart condition. the cause of the positioning issue was most likely patient condition related. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system section: warnings and cautions: warnings to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. d6a and d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03051: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. g1 reporting contact fax number missing. h.6 code 4625, 2627, and 2976 was reported incorrectly. new code was added to health effect - impact code and medical device problem code. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34686 | 1220648-2023-03052 | 17712573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712573&pc=OZD | 11/17/2022 | 09/08/2023 |  | 09/08/2023 |  |  | 11/17/2022 | 09/08/2023 |  |  | 05/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223999 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 39 year-old female with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that there was a yellow luer leak. during this time, there were high purge pressure and purge system blocked alarms. a purge bypass was performed. bicarbonate was being used in the purge. there was no patient harm reported. | [5 empty] | The investigation for the yellow luer leak was completed. only the data logs were returned for investigation. based on the review, no high purge pressure events occurred on the day of the reported alarms. the sidearm and the pump were not returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34687 | 1220648-2023-03053 | 17712666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712666&pc=OZD | 04/28/2022 | 01/10/2024 |  | 09/08/2023 |  |  | 04/28/2022 | 09/08/2023 |  |  | 01/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023184038 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient the pulled the pump out into the aorta. the patient was taken to the or to re-advance across valve. impella was repositioned and the patient was sent back in cvicu. no lasting harm from the malposition. | [5 empty] | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was most likely patient movement. the impella was pulled into the aorta as per the clinical description provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34688 | 1220648-2023-03054 | 17712727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712727&pc=OZD | 04/27/2022 | 03/11/2025 |  | 09/08/2023 |  |  | 04/27/2022 | 09/08/2023 |  |  | 01/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023184022 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. later the patient became febrile and hypotensive. decision was made to explant impella and the patient blood cultures came back positive indicating that the patient was septic. the patient was sent back to the icu. the cause of the sepsis was not known. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | An infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician. -physical examination findings. -results of special investigations : laboratory test results & imaging studies. -microbiological culture and sensitivity results. -response to previous treatments and therapies. -any relevant epidemiological information or exposure history. -concomitant devices in use. -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined. this report is being filed as part of a retrospective review of historical records.
34689 | 1220648-2023-03055 | 17712862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17712862&pc=OZD | 08/11/2023 | 07/01/2024 |  | 09/08/2023 |  |  | 08/11/2023 | 09/08/2023 |  |  | 02/28/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED INC. |  |  | US |  | 2024325789 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2023 |  |  |  |  | The medical team in the eu placed an impella cp to offer hemodynamic support to the patient admitted with cardiomyopathy and myocarditis. after 10 days of support there was an observation made of an impella cp purge sidearm leak. the pump remained on support until the pump stopped 10 days later (after 21 days of support). the pump was explanted and replaced. | [5 empty] | The investigation into the purge leak ¿ pump and the pump stop issues has been completed since the original report was filed. the pump was returned for investigation. product evaluation confirmed damage to the reservoir filter joint as it came back split in half. the pump additionally had a large purge gap. the cause of the purge leak was damage to the reservoir filter joint however the cause of damage is unknown. the cause of the pump stop was bearing wear due to misuse of the pump over the pump life duration. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. removed f6 and f8. | The medical center in the eu has not yet returned for benchtop analysis any impella pump product or data logs. upon the completion of the investigation, a final report will be forthcoming. per the ifu: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34690 | 1220648-2023-03057 | 17713858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17713858&pc=OZD | 05/09/2022 | 01/26/2024 |  | 09/08/2023 |  |  | 05/09/2022 | 09/08/2023 |  |  | 02/29/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023193694 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support was being weaned off the pump and during explant the md identified a hematoma which required surgical intervention to be remove. the pump was successfully explanted. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34691 | 1220648-2023-03058 | 17713924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17713924&pc=OZD | 11/23/2022 | 09/08/2023 |  | 09/08/2023 |  |  | 11/23/2022 | 09/08/2023 |  |  | 07/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023254427 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 52 year-old female with an impella 5.5 device for mechanical circulatory support. it was reported that there was a crack on the luer lock and and leaking noted in the purge sidearm retainer. there was bicarbonate running in the purge. a purge bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. neither the product nor the data logs were returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34692 | 1220648-2023-03059 | 17714016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17714016&pc=OZD | 08/10/2023 | 08/01/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 04/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338730 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The user facility reported that surgeon had a impella 5.5 supporting a surgical 69-year-old female patient with the chest left open postoperatively. the team noted the patient to have a bleed from the impella graft-native anatomy. the bleed was at the site of the purse sting sutures. hemostasis was not achieved and the team consulted the family and allowed the patient to become a dnr/dni and the patient expired. the support had only been for just over 9 hours. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added-b5 mso review, h6 component code 925 added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34693 | 1220648-2023-03060 | 17715546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715546&pc=OZD | 08/10/2023 | 11/14/2025 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  |  |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1736233 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 08/10/2023 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support in the cath lab. when the patient was transferred to the icu, the automated impella controller (aic) developed a purge failure error. the team exchanged the aic for another aic. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D9 device returned to manufacturer date added d2 common device name revised on manufacturer device report 1220648-2023-03060 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-03060. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03060. g1 reporting contact email revised on manufacturer device report 1220648-2023-03060 in accordance with updated procedures.
34694 | 1220648-2023-03061 | 17715553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715553&pc=OZD | 07/29/2023 | 01/11/2024 |  | 09/08/2023 |  |  | 08/09/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024351856 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility |  |  |  |  | 77 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Impairment | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2023 |  |  |  |  | The complainant reported a 77 year old female patient presenting for elective support via the impella 5.5 device. approximately 2 days after explant, the patient endured renal insufficiency and required dialysis as treatment. this information was received via the long-term outcome and quality indicator (loqi) registry as possibly related to the impella device. | [5 empty] | The investigation into the patient's hemolysis/acute renal insufficiency has been completed. the root cause could not be determined since the complaint device was not returned and insufficient clinical data was returned. there is unknown relation with impella. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34695 | 1220648-2023-03062 | 17715561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715561&pc=OZD | 08/11/2023 | 01/15/2024 |  | 09/08/2023 |  |  | 08/11/2023 | 09/08/2023 |  |  | 04/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351502 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2023 |  |  |  |  | The user facility reported a 57-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was noted that the physician placed the impella pump deep due to the catheter "jumping back" across the aortic valve when trying to place the inlet into the mid-left ventricle. the patient was not harmed. | [5 empty] | The investigation into the reported positioning issues has been completed since the original report was filed. product & data were not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34696 | 1220648-2023-03063 | 17715572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715572&pc=OZD | 08/10/2023 | 01/23/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354654 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during impella support, the staff was unable to find pedal pulses in the impella extremity. the ante-grade perfusion was in place. the patient was on impella support for 73.91 hours before the family withdrew care and the patient expired. no allegations were made towards the impella for the cause of death. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the abiomed products were not returned for evaluation .the cause of the limb ischemia issue was most likely patient condition related based on provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34697 | 1220648-2023-03064 | 17715577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715577&pc=OZD | 08/10/2023 | 01/18/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 08/31/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272558 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. there was bleeding from insertion site of the impella. couple of sutures were placed and no further issues were reported. | [5 empty] | The investigation for the reported access site bleed has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34698 | 1220648-2023-03065 | 17715628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715628&pc=OZD | 08/11/2023 | 01/15/2024 |  | 09/08/2023 |  |  | 08/11/2023 | 09/08/2023 |  |  | 06/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367322 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that despite multiple attempts to deliver the impella 5.5 across the left subclavian artery, the physician was unable to pass catheter into aorta. this extensive manipulation of the impella lead to a kink at the link between the motor and the catheter; therefore, the impella 5.5 was removed. | [5 empty] | The investigation into the delivery issues- use has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the delivery issue was calcification resulting in a catheter kink. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34699 | 1220648-2023-03066 | 17715660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715660&pc=OZD | 08/11/2023 | 01/24/2024 |  | 09/08/2023 |  |  | 08/11/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354130 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that after being on impella support, the pump alarmed that it had high motor current and stopped. the staff switched consoles and attempted to restart without success. when pump stopped patient coded and was successfully resuscitated and patient was placed on ecmo and the impella was removed. the physician discussed that the pump stop could have occurred from a clot. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. data logs were returned and intermittent spikes and high motor current spike with saturated flow followed by a pump stop was observed at the end of the case. also, purge system open, purge line click-on not detected alarms with purge flow at 30 ml/hr was observed at the end. the cause of the pump stop issue was most likely biomaterial with saturated flow and motor current spike observed per log review. the cause of the purge leak-cassette issue was not determined since product was not returned.
34700 | 1220648-2023-03067 | 17715734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715734&pc=OZD | 07/26/2023 | 09/19/2024 |  | 09/08/2023 |  |  | 08/09/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024351860 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility |  |  |  |  | 52 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Ischemia Stroke | [11 empty] | Death | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/09/2023 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a 52 year old male patient presenting for hemodynamic support via the impella 5.5 device. during support, the patient endured hypotension prompting a bolus of 500cc+ as rescue intervention. additionally, the patient enduring an ischemia stroke. this information was received via abiomed's long-term outcome and quality indicator registry with a probably relationship to the impella device. |  |  |  |  | The investigation into the stroke/cva issue has been completed. the device was not returned for investigation. as per clinical, the cause of the stroke was most likely due to patient condition and unknown relation to impella. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34701 | 1220648-2023-03068 | 17715739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715739&pc=OZD | 08/10/2023 | 09/08/2023 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 09/30/2024 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273747 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 78-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that after the physician placed the device across the aortic valve and turned the pump on, low flows and suction alarms were noted. the physician attempted to reposition the device; the flows improved but the alarms returned. on x-ray there was a notable kink in the cannula just prior to the outlet. the physician explanted the pump and re-attempted to place the same device but the same kink in the cannula just prior to the outlet was observed. a new impella was used. | [5 empty] | (h3) the investigation into the reported product damage (cannula kink) has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found that a cannula kink was observed near the outlet cage. there were no other abnormalities observed. during functional analysis the returned complaint pump was connected to automated impella controller (aic) and run on different p-levels. there were no alarms observed. the root cause of the low pump flow was determined to be a cannula kink due to patient condition related. should additional relevant information become available, a supplemental report will be submitted.
34702 | 1220648-2023-03069 | 17715752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17715752&pc=OZD | 08/10/2023 | 03/12/2024 |  | 09/08/2023 |  |  | 08/10/2023 | 09/08/2023 |  |  | 01/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319116 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2023 |  |  |  |  | The user facility reported an 83-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the impella was implanted via left axillary and cut down graft was performed by the vascular surgeon. during intervention, patient became hemodynamically unstable. pressors, fluids and blood given. determined left subclavian and vertebral artery perforation. covered stents placed in both areas. decision made to abort the intervention, remove and bring the patient back at a later date. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34703 | 1220648-2023-03070 | 17716131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716131&pc=OZD | 07/23/2023 | 11/22/2024 |  | 09/08/2023 |  |  | 08/09/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344611 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility |  |  |  |  | 77 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2023 |  |  |  |  | The complainant reported a 77 year old male patient presenting for hemodynamic support via the impella 5.5. during support, signs of hemolysis and rising purge pressure developed. as treatment, tpa was infused and support proceeded. this information was received via abiomed's long-term outcome and quality indicator (loqi) database and alleged as a probable relationship with the device. | [5 empty] | The investigation for the hemolysis and high purge pressure (hpp) has been completed.the impella was not returned for evaluation. data logs were reviewed finding no motor current (mc) spikes occurred and no indication observed for hemolysis issue. hpp alarms occurred and started to reduce after new purge bag connected, tissue plasminogen activator (tpa) was infused and the issues resolved. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and insufficient clinical data returned. the cause of the hpp most likely was biomaterial ingestion. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added h6 component code 925 corrections to the initial mfr report: d4 model number added d6a/d6b f6/f8 should have been left blank. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿
34704 | 1220648-2023-03071 | 17716180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716180&pc=OZD | 08/09/2023 | 01/15/2024 |  | 09/08/2023 |  |  | 08/09/2023 | 09/08/2023 |  |  | 04/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349542 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced profound hypotension and dark red urine with a lactate of 8. the patient was brought to the cath lab and underwent pericardiocentesis with drain alongside a 6/35 reperfusion sheath on left femoral artery. with 500 ml of blood drained, the patient instantly gained pulsatility back decreasing pressor requirements while increasing support to p-9. overnight, the patient experienced hypotension requiring pressor to be put back on despite full support and fluid resuscitation. a head ct also displays an acute but stable right parietal lobe bleed thus systemic heparin and anti-platelet therapies were ceased. the patient was still able to follow commands during sedation. vascular at bedside consulted for distal limb ischemia were able to doppler pt and popliteal pulses on both legs in absence of pedals. dark red urine was noted in foley positioning, reexamined with labs were done to monitor this condition. | [5 empty] | The investigation for the reported stroke and limb ischemia has been completed. the impella device was not returned for evaluation. clinical details states that the patient developed left sided weakness and a stroke alert was called and neurology ordered a computed tomography (ct) scan of the head which revealed a meningioma as well as some infarct sites. the cause of the stroke/ cerebral vascular accident (cva) most likely was patient condition related with unknown relation to impella. additionally, distal limb ischemia was able to doppler pt and popliteal pulses on both legs in absence of pedals. ischemia resolved naturally. the cause of the limb ischemia was patient condition related since patient was put back on pressors and ischemia occurred to both legs. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34705 | 1220648-2023-03072 | 17716182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716182&pc=OZD | 08/09/2023 | 01/19/2024 |  | 09/08/2023 |  |  | 08/09/2023 | 09/08/2023 |  |  | 04/30/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355613 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had no distal pulses at the end of first procedure. patient was hypertensive on impella. therefore, surgeon decision was to wean and explant impella device. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34706 | 1220648-2023-03073 | 17716186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716186&pc=OZD | 08/09/2023 | 08/16/2024 |  | 09/08/2023 |  |  | 08/09/2023 | 09/08/2023 |  |  | 05/31/2025 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024337999 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility that the surgeon was unable to successfully deliver the device due to kinking of the catheter resulting from patient's vascular anatomy. | [5 empty] | The investigation of delivery issues- use has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, as the blood vessels stretched and impella 5.5 did not progress any further and cannula kink occurred. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings and cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings and cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with automated impella controller section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h6 medical device problem code added g1 reporting contact fax number missing on manufacturer device report 1220648-2023-03073. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34707 | 1220648-2023-03074 | 17716188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17716188&pc=OZD | 08/09/2023 | 01/18/2024 |  | 09/08/2023 |  |  | 08/09/2023 | 09/08/2023 |  |  | 04/30/2023 |  | 09/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336956 |  | 1000100 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the impella started alarming-red retrograde flow and position in aorta causing the lowering of p-levels and removal of device. staff contacted the supporting team for support and mentioned that it sounded like a sensor failure issue. | [5 empty] | The investigation into the reported positioning issues has been completed since the original report was filed. the pump was returned for investigation. during testing, the 5s parameters were within specification. no other physical abnormalities were observed. data logs showed a barcode trend in the optical signal while the pump was set to p-level 4. based on clinical, impella device was flipped and sensor face was reported against aortic wall. the root cause of the positioning issue was use related since pump was reported to be flipped during support. corrections: the section d9 was left blank in error in the initial report which has been added to this report. d2: corrected brand name. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
34708 | 1220648-2023-03019 | 17700449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17700449&pc=OZD | 09/18/2022 | 09/07/2023 |  | 09/07/2023 |  |  | 09/18/2022 | 09/07/2023 |  |  | 06/30/2024 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229198 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 67 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after a purge cassette change, a leak was noted at the yellow luer connection. a purge bypass was performed. there was sodium bicarbonate running in the purge. there was no patient harm reported. | [5 empty] | The investigation for the purge sidearm leak has been completed. the product was not returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34709 | 1220648-2023-03020 | 17700790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17700790&pc=OZD | 02/25/2022 | 03/01/2024 |  | 09/07/2023 |  |  | 02/25/2022 | 09/07/2023 |  |  | 09/30/2023 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022145848 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support was noted to have ischemic toes on bilateral feet and md concerned about dry gangrene forming. decision was made to wean the patient off impella and the pump was explanted. the patient prognosis was grim. | [5 empty] | The investigation into the reported limb ischemia- reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34710 | 1220648-2023-03021 | 17700888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17700888&pc=OZD | 08/08/2023 | 01/12/2024 |  | 09/07/2023 |  |  | 08/08/2023 | 09/07/2023 |  |  | 05/31/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355569 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male noted for high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. patient had escalated care from the impella cp to the axillary accessed 5.5 pump. after 2 weeks of support the team had to perform a washout to the axillary site due to elevated wbcs. there was a concern for infection and edema at the site. the pump remained on for support for another 3 weeks after the washout. | [5 empty] | The investigation into the infection/sepsis has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the infection/sepsis issue was patient condition related and the relation to impella is unknown. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34711 | 1220648-2023-03022 | 17702500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17702500&pc=OZD | 11/07/2022 | 09/07/2023 |  | 09/07/2023 |  |  | 11/07/2022 | 09/07/2023 |  |  | 07/31/2024 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023249690 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/07/2022 |  |  |  |  | The user facility reported a 60 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella side arm yellow luer fractured during a purge cassette change by staff. an attempt was made to change the cassette a second time before realizing the fracture. tape was applied which maintained some flow and pressure until patient could be taken back to the operating room for impella replacement. there was no patient harm reported. | [5 empty] | The investigation for the sidearm crack/leak has been completed. the product was returned and the side arm was pressurized using a syringe of color-died water. the leak at the yellow luer was confirmed and the crack was seen. the root cause of the purge leak was determined to be a damaged yellow luer due to bicarbonate use in the purge. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. per the ifu: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol".
34712 | 1220648-2023-03023 | 17703511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17703511&pc=OZD | 03/19/2022 | 03/01/2024 |  | 09/07/2023 |  |  | 03/19/2022 | 09/07/2023 |  |  | 12/31/2023 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023177848 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced necrotic finger noted on left hand and it was worsening. the medical staff agreed to have a dopler study on the patient. the pump remained on for support after the necrosis was observed for 4 more days before weaned and explanted. no noted intervention was performed for the necrotic finger. | [5 empty] | The investigation into the limb ischemia ¿ reversible event has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34713 | 1220648-2023-03024 | 17703538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17703538&pc=OZD | 08/09/2023 | 11/13/2023 |  | 09/07/2023 |  |  | 08/09/2023 | 09/07/2023 |  |  |  |  | 09/07/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1244247 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported when patient was taken to ccl for impella placement, the nurses were setting up the pump. when they opened the purge cassette door, they noticed the portion of the purge area that holds the wings of the purge disc was cracked and missing. they grabbed a console from a separate room to proceed with the case. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the device was returned for investigation. the aic - purge disc/flag issues with detecting the purge disc were due to a broken purge pressure sensor retainer. | The aic controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
34714 | 1220648-2023-03025 | 17703594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17703594&pc=OZD | 08/09/2023 | 10/28/2025 |  | 09/07/2023 |  |  | 08/09/2023 | 09/07/2023 |  |  |  |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during an icu nurse training class, the impella console used for demonstration had a battery failure alarm. the alarm persisted despite following instructions on the screen. the console was placed in the cardiac catheterization lab (ccl) in an area where it wouldn't be used with a sign on the console. staff were informed not to use the console until it has been serviced. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-03025 in accordance with updated procedures. d9 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03025. e2 was erroneously entered on the initial manufacturer device report number 1220648-2023-03025. g1 reporting contact email revised on manufacturer device report 1220648-2023-03025 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03025. | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. data analysis: the logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) and a battery 1 communication failure alarm (smbus communication loss or sda short ¿ event set #352) due to low pack voltage. device analysis: the failure mode was reproduced on an engineering lab bench. the batttest result shows a cell imbalance in cell 3 of battery 1. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure.
34715 | 1220648-2023-03026 | 17704656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17704656&pc=OZD | 08/08/2023 | 03/15/2024 |  | 09/07/2023 |  |  |  | 09/07/2023 |  |  | 03/31/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343094 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2023 |  |  |  |  | The user facility reported an 82-year-old female (race unknown) noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the pump was at p2 for approximately 8 minutes to see if the patient could handle pump removal. suddenly the automated impella controller (aic) alarmed "pump stop". two attempts to restart the pump at p2 failed. white power cord removed and replaced and staff was still unable to restart the pump. decision was made by doctor to remove the pump. closure was successful and patient was stable. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. electrical testing revealed an issue with two out of three open motor cable lines. the cause of the pump stop issue was open electrical lines. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34716 | 1220648-2023-03027 | 17704695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17704695&pc=OZD | 08/08/2023 | 03/07/2024 |  | 09/07/2023 |  |  |  | 09/07/2023 |  |  | 06/30/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357255 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2023 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient developed left sided weakness. a ct scan of the head revealed meningioma as well as some infarct sites. the team and neurologist had concern for device related embolic events due to holding the systemic anticoagulation, a decision was made to wean and explant the device. following the explant, the device did not appear to have any thrombus built up on the catheter or in the cannula, however, an analysis seemed warranted given the patient complications. | The user facility reported a 76-year-old male in acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had a large drop in hemoglobin and hematocrit(h/h). a ct scan found a large retroperitoneal bleed from left iliac, team stopped systemic anticoagulation and gave blood products. overnight the patient developed left sided weakness. a ct scan of the head revealed meningioma as well as some infarct sites. the team and neurologist had concern for device related embolic events due to holding the systemic anticoagulation, a decision was made to wean and explant the device. following the explant, the device did not appear to have any thrombus built up on the catheter or in the cannula, however, an analysis seemed warranted given the patient complications. |  |  |  |  | The investigation for the reported stroke has been completed. the impella device was returned for evaluation. during evaluation, the pump was visually inspected and biomaterial was observed in the inlet and outlet cages. no other abnormalities were observed. however, with sufficient clinical details, the root cause of the stroke could not be determined. corrected b5. describe event or problem note: initial mdr reported bleeding and blood transfusion in error. the bleeding event was not related to the impella 5.5 in this mdr. | The impella products have been received for the investigation into bleeding and cerebral vascular accident. the investigation is not complete at this time. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34717 | 1220648-2023-03028 | 17704820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17704820&pc=OZD | 08/06/2023 | 03/12/2024 |  | 09/07/2023 |  |  |  | 09/07/2023 |  |  | 04/30/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024345386 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. the pump failed to start after an attempted initiation. it was discovered that the easy guide lumen was left in place. the physician noted "this was caused by turning on the pump without confirming that the easyguide lumen was completely removed." a new pump was used with no further issue. | [5 empty] | The investigation for the reported pump "did not start" issue. the impella nor the logs were returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the pump did not start issue was a red lumen removal issue since red lumen was not removed during pump start. the instructions for use states ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34718 | 1220648-2023-03030 | 17705069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17705069&pc=OZD | 03/21/2022 | 03/01/2024 |  | 09/07/2023 |  |  |  | 09/07/2023 |  |  | 11/30/2023 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023161447 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | [6 empty] | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient was taken to the operating room for pump removal due to concerns for a potential clot. although no initial signs of thrombosis were identified during pump explant, a doppler assessment noted no distal pulse past the brachial artery. vascular consult was consulted for intervention. the patient was considered hemodynamically stable at the time of the reported incident. | [5 empty] | The investigation into the reported limb ischemia- reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34719 | 1220648-2023-03029 | 17705080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17705080&pc=OZD | 08/09/2023 | 09/20/2024 |  | 09/07/2023 |  |  |  | 09/07/2023 |  |  | 05/31/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356224 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 77-year-old male noted as high risk for coronary angioplasty was implanted with an impella 5.5 device for mechanical circulatory support. there was a left ventricle (lv) perforation after the impella 5.5 insertion. the impella 5.5 was placed wirelessly into graft and aortotomy was used to pull through atrioventricular (av) into empty lv. as the staff filled the lv, a posterior lv perforation was noted. the staff went back on bypass and fixed the perforation. patient experienced quite a bit of bleeding from multiple areas in the chest. several blood products were given and the right heart began to show further decompensation. staff continued to give blood product and they were able to stay relatively stable. after impella support for 27.78 hours the patient expired. no allegations were made towards the impella for patients death. |  |  |  |  | The investigation into the reported ventricle perforation has been completed since the original report was filed. no product was returned. as per clinical, impella 5.5 was placed wirelessly into graft and aortotomy using direct surgical access and the impella measured to be too deep. a posterior lv perforation was noted and repaired lv. device was repositioned accordingly and achieved flows per requirement. data logs were returned, and positioning issue alarms found during the start of case while repositioning. the cause of the ventricle perforation issue was most likely improper positioning of the device per clinical and log review. the section d4 (udi) was incorrect in the initial report which has been corrected to this report. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ corrected information was provided in sections b2, b5, h1, h6, h10. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34720 | 1220648-2023-03031 | 17705114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17705114&pc=OZD | 08/09/2023 | 01/18/2024 |  | 09/07/2023 |  |  |  | 09/07/2023 |  |  | 10/31/2024 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283302 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had a impella cp delivered to support the 58-year-old male patient with cgs and av workup ongoing. the team noted an access site bleed that was given manual pressure alone. the pump did come out of position. the positioning issues prompted the team to replace the pump as it was unable to be re-delivered back across the valve. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for benchtop evaluation and investigation. the cause of the positioning issue was use related since the pump was pulled out due to reported patient movement, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34721 | 1220648-2023-03032 | 17706446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17706446&pc=OZD | 08/04/2023 | 09/11/2024 |  | 09/07/2023 |  |  | 08/09/2023 | 09/07/2023 |  |  | 07/31/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024368147 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had hemolysis while on impella support. the relationship to hemolysis and the impella is unclear. no problem with pump placement or volume and it was assumed that the hemolysis was caused or contributed by an endoleak after endovascular abdominal aortic aneurysm repair (evar), however that could not be verified. blood products were given to the patient. additionally, there was gastrointestinal bleeding that occurred during the impella support. there was melena, intestinal bleeding, and bleeding from other device insertion sites (systemic bleeding tendency). the number of transfusion units is unknown, but pc, ffp, and rbc were being administered. physician confirmed there was a relationship between the impella and the bleeding. | [5 empty] | The investigation has been completed. clinical details stated that hemolysis was reported on day two of support. gastrointestinal bleeding reported on day nine of support. the data logs showed continuous 'suction' and 'impella position unknown' alarms for the first few days of support when hemolysis was reported. no product was returned. the root cause of the hemolysis was most likely improper positioning of the pump. the root cause of the access site bleeding - major was most likely patient condition related as bleeding was reported at multiple sites. the root cause of the gastrointestinal bleed was not determined as no product was returned and limited clinical information was provided. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ a.3 and e.1 address line 2 were revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2023-03032. b.3 date of event and d.4 unique identifier (udi) # were revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-03032. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-03032 in accordance with updated procedures. h.6 component code was added since the initial submission of manufacturer device report number 1220648-2023-03032 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-03032 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
34722 | 1220648-2023-03033 | 17706450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17706450&pc=OZD | 08/10/2023 | 01/19/2024 |  | 09/07/2023 |  |  | 08/10/2023 | 09/07/2023 |  |  | 01/31/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024312004 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender with acute myocardial infarction/cardiogenic shock received extracorporeal membrane oxygenation (va-ecmo) support and an impella cp smartassist heart pump. a pseudoaneurysm in the right femoral artery has been reported. ct scan showed a pseudoaneurysm without venous pooling and a groin hematoma along the right iliac muscle. the incident required an extended hospital stay. we do not currently have any information about further interventions. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the vascular damage has been completed. product was not returned for review. the clinical review determined the root cause of the vascular damage was most likely due to patient condition of diseased/small vessels.
34723 | 1220648-2023-03034 | 17706486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17706486&pc=OZD | 08/10/2023 | 01/30/2024 |  | 09/07/2023 |  |  | 08/10/2023 | 09/07/2023 |  |  | 05/31/2024 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023211901 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender presented with de-novo cardiomyopathy and was to receive hemodynamic support from an impella 5.0 cardiac pump. during pump setup in the operation room, the physician noticed purge fluid leaking from the purge filter. the automated impella controller (aic) showed a low purge pressure alarm immediately. decision has been made to not implant the pump and replace it by a new impella 5.0 pump. the second device has been successfully delivered into the left ventricle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." | The investigation into the purge leak ¿ pump issue has been completed. the pump was not returned for investigation. the root cause of the purge leak - pump is most likely due to a damaged sidearm based on the clinical description. however, no product or logs were returned to confirm location of the damage.
34724 | 1220648-2023-03035 | 17706552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17706552&pc=OZD | 08/09/2023 | 01/19/2024 |  | 09/07/2023 |  |  | 08/09/2023 | 09/07/2023 |  |  | 05/31/2025 |  | 09/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024358904 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Percutaneous coronary intervention was performed in a 63-year-old patient with acute myocardial infarction/cardiogenic shock due to known coronary artery disease. an impella cp smartassist was inserted via the femoral artery to provide hemodynamic support during the procedure. below the access site, ischemia has been reported. the patient's leg was cold and an external bypass was discussed. a bypass was not possible because the patient's blood vessels were calcified. the following day it was reported that the circulation in the leg had improved sufficiently. no intervention was made because of the ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition.
34725 | 1220648-2023-02996 | 17689876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17689876&pc=OZD | 02/24/2022 | 09/06/2023 |  | 09/06/2023 |  |  | 02/24/2022 | 09/06/2023 |  |  | 11/30/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023167077 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Perforation of Vessels | No Clinical Signs, Symptoms or Conditions | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/24/2022 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team was having difficulty delivering the pump. the pump was passed through the aortic valve however the md lost wire access and had to pull the pump back. the md had difficulty removing the pump and in the process damaged the artery which had to be surgically repaired. the md discussed insertion of the impella cp via other access. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
34726 | 1220648-2023-02997 | 17690321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17690321&pc=OZD | 08/08/2023 | 11/12/2023 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  | 03/31/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343211 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 82-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient experienced left leg cool with doppler popliteal and absent pedal pulse. the impella support was for 4 hours and then weaned and explanted. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the device was not returned by the medical facility. no product was returned. as per clinical, patient pulses were not present in the left leg and pump was explanted to resolved ischemia. patient had calcified vessel in left femoral. the cause of the limb ischemia issue was patient condition related with patient having calcified vessels and pump explant resolving ischemia based on clinical information. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ in the initial report the udi was missing a digit in error, the correct udi is as follow section d4: : the correct udi is (b)(4). | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34727 | 1220648-2023-02998 | 17690446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17690446&pc=OZD | 08/08/2023 | 07/02/2024 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  | 08/31/2024 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a impella cp placed for a 61-year-old male patient to have a coronary artery bypass graft (cabg). the pump access site on day two of support was seen to have a retroperitoneal bleeding (rp). the rp bleed was treated by blood product infusion of 3 units of rbcs. | [5 empty] | The investigation into the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ d.6a implant date, d.6b explant date, g.1 reporting contact fax number and h.10 additional manufacturer narrative instructions for use were revised in accordance with updated procedures since the initial manufacturing device report 1220648-2023-02998 was submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34728 | 1220648-2023-02999 | 17690468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17690468&pc=OZD | 02/22/2022 | 01/25/2024 |  | 09/06/2023 |  |  | 02/22/2022 | 09/06/2023 |  |  | 10/13/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022156734 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site. the patient lost 2 liters of blood and had to be taken to the or to stop the bleeding. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the access site bleed has been completed. however, with limited clinical details nor the impella device, the root cause of the reported issue could not be determined.
34729 | 1220648-2023-03000 | 17690519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17690519&pc=OZD | 02/22/2022 | 01/26/2024 |  | 09/06/2023 |  |  | 02/22/2022 | 09/06/2023 |  |  | 08/31/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130655 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2022 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced oozing at the access site. a computed tomography scan (ct) of the area could not confirm if it was an abscess or hematoma but the physician was leaning more towards hematoma. a washout of the area was performed and one unit of blood was given to the patient. hemoglobin showed improvement form 6.8 to 8. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product were returned for investigation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34730 | 1220648-2023-03001 | 17690539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17690539&pc=OZD | 02/20/2022 | 03/10/2025 |  | 09/06/2023 |  |  | 02/20/2022 | 09/06/2023 |  |  | 09/30/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151065 |  | 055-0008 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2022 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support became thrombocytopenic and the platelets dropped to 25 from 53. purge and systemic heparin were removed and systemic argatroban started. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the issue was not determined due to insufficient clinical details. this report is being filed as part of a retrospective review of historical records.
34731 | 1220648-2023-03002 | 17690575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17690575&pc=OZD | 02/19/2022 | 03/11/2025 |  | 09/06/2023 |  |  | 02/19/2022 | 09/06/2023 |  |  | 10/31/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023159327 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2022 |  |  |  |  | The user facility reported a 39-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a low purge pressure from an alleged kink in the line which was loosened after being noticed. the pump remained on for support despite the kink for a 7.7 day support tile removed for lvad placement. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
34732 | 1220648-2023-03003 | 17692476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17692476&pc=OZD | 06/15/2022 | 09/06/2023 |  | 09/06/2023 |  |  | 06/15/2022 | 09/06/2023 |  |  | 02/29/2024 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023193670 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 58 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge sidearm leak after 5 days on support. the team left the pump on for support. no harm or injury to the patient. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined as no product was returned for this investigation so the location of the leak could not be identified. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34733 | 1220648-2023-03004 | 17692521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17692521&pc=OZD | 08/09/2023 | 03/01/2024 |  | 09/06/2023 |  |  | 08/09/2023 | 09/06/2023 |  |  | 11/30/2024 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306756 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 56-year-old male had an impella cp pump inserted for a support of 5 days when explanted. at the time of explant the pump was seen to be thrombosed and pulses lost to the limb. the team had surgery consult for the salvage ability of the limb. | [5 empty] | The investigation into the limb ischemia irreversible event has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34734 | 1220648-2023-03005 | 17692546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17692546&pc=OZD | 04/20/2022 | 09/20/2024 |  | 09/06/2023 |  |  | 04/20/2022 | 09/06/2023 |  |  | 10/31/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023159166 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2022 |  |  |  |  | The user facility reported that a 49-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak while sodium bicarbonate was being used for the purge solution. a purge bypass was successfully performed. on day 54 of support, the patient was noted to have a purge blocked alarm. the patient was administered two doses of tissue plasminogen activator (tpa) without success. the patient was subsequently changed back to sodium bicarbonate due to bleeding issues. no further information was provided. there were no reports of harm to the patient. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported purge leak and purge system blocked alarm was completed. the device was not returned for evaluation. the root cause remains undetermined as the device was not returned for evaluation. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03005 was provided.
34735 | 1220648-2023-03006 | 17692558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17692558&pc=OZD | 06/12/2022 | 09/20/2024 |  | 09/06/2023 |  |  | 06/12/2022 | 09/06/2023 |  |  | 09/30/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149800 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Crack | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a 59 year-old male was implanted with an impella 5.5 device for mechanical circulatory support after a myocardial infarction. it was reported that the yellow luer had a crack and was leaking. a purge bypass was performed. no harm or injury was noted and support continued with the pump. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the yellow luer crack has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was most likely a yellow luer crack. this crack was likely due to the combination of the use of bicarbonate and the potential use of kelly clamps. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03006 was provided in sections b2, b5, h1, and h6.
34736 | 1220648-2023-03007 | 17694191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17694191&pc=OZD | 08/08/2023 | 03/13/2024 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  | 02/28/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341068 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during procedure, the right femoral artery (rfa) access was obtained and the impella was inserted. a hematoma was noted at the rfa access site with angio revealing perforation. contralateral access obtained and impella removed. balloon tamponade followed by covered stent placement performed in superficial femoral artery (sfa) below profundal bifurcation with hemostasis achieved. it was noted that the perforation site was below the arterial access site. | [5 empty] | The investigation into the vascular damage- great vessel issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the vascular damage issue was most likely use related since sheath kink was noted and resistance encountered during insertion of device, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34737 | 1220648-2023-03009 | 17694356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17694356&pc=PYX | 08/08/2023 | 01/15/2024 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  | 06/30/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024367564 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2023 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was to be implanted with an impella rp device for mechanical circulatory support. during initial insertion under fluoro, the team identified a compromised impella intro sheath. recommendation by local support was to remove the intro sheath. intro sheath inspected outside of the patient¿s body and tear was identified on outer sheath. sheath was not used. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported introducer damage -mechanical has been completed since the original report was filed. product was not returned for review. the root cause of the introducer damage was unable to be determined as limited clinical details were provided and no product was returned for review.
34738 | 1220648-2023-03010 | 17694457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17694457&pc=OZD | 08/09/2023 | 03/12/2024 |  | 09/06/2023 |  |  | 08/09/2023 | 09/06/2023 |  |  |  |  | 09/06/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1444507 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the plug for the automated impella controller (aic) was bent. the staff was unsure when the bending occurred. the aic was taken out of patient use. | [5 empty] | The investigation into the aic ¿ damage has been completed since the original report was filed. the console was returned and tested after arriving in fs. visual inspection of the console power cable plug revealed that the ground terminal was bent. the root cause of the issue was most likely due to a damaged power cord ground terminal during normal use of the device when connecting or unplugging to ac mains. | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
34739 | 1220648-2023-03011 | 17694488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17694488&pc=OZD | 08/08/2023 | 03/13/2024 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  | 04/30/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337415 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. after more than 10 days on support the impella developed suction alarms and there was an acute decrease in impella flows. the impella was kept at p 4-5 and flows remained low without suction events. the next morning the device continued to have suction events and the physician attempted to reposition without resolution of low flows and the device appeared to be in the optimal position with no obstruction noted. the physician dropped impella to p-1 and slowly ramped up to attempt to resolve issue. surgeons unsure if the issue is the impella motor related or possibly a clot even though no clot was visible in the echocardiogram. the concern was that the patients heparin vad goal had been ptt therapeutic or supratherapeutic. the device was explanted as a result of the noted issue and a new impella 5.5 was placed for support. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed a motor current spike without a change in p level before the noted low pump flow and suction event. as per the given clinical information, low pump flow and suction were reported on p level higher than p4. doctor confirmed the good pump position, and providing the specified blood volume did not resolve the issue. the cause of the low pump flow issue was most likely biomaterial, based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34740 | 1220648-2023-03012 | 17694503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17694503&pc=OZD | 08/08/2023 | 10/10/2024 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  |  |  | 09/06/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1504720 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility biomedical engineer report the automated impella controller (aic) had a broken purge clip. a loaner onsite was used. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Manufacturer device report (mdr) 1220648-2023-03012 was created in error and is a duplicate report. all investigation results will be found in mdr 1220648-2023-03013.
34741 | 1220648-2023-03013 | 17694584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17694584&pc=OZD | 08/08/2023 | 01/25/2024 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  |  |  | 09/06/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1504720 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) had a damaged purge disc. the team replaced the aic per instruction for use (ifu) recommendations. this was found outside of any patient use. biomed reported. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon visual inspection, the purge retainer was broken, and the flag was missing. the root cause for the purge pressure low is because of the broken purge retainer and the missing flag. | The impella console was not yet received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34742 | 1220648-2023-03008 | 17694979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17694979&pc=OZD | 08/08/2023 | 01/22/2024 |  | 09/06/2023 |  |  | 08/08/2023 | 09/06/2023 |  |  | 02/28/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322120 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 29-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella had suction alarms. echo was done to evaluate pump positioning. a kink was observed in the cannula very near to the motor housing. the physician was convinced that it was due to the patient¿s anatomy. the physician described the guide catheters also not sitting right in the aortic root and thinks that the arch takeoff was coming almost straight from the aortic valve. attempts were made to reposition the cannula, but no amount of repositioning straightened out the kink. the physician then proceeded to remove the impella. the patient tolerated the pump removal well. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. the device was not returned by the medical facility. data logs confirmed the low flow when the pump started and remained to the rest of the case. the device was visually inspected and found a kink on the cannula about 14 mm from the bonding site between the of cage and cannula. no other issues were found on the rest of the pump. the pump ran without issue under bench test. based on clinical description, product analysis and data review, the root cause of low flow was due to canula kink. the cause of cannula kink was due to patient condition. | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
34743 | 1220648-2023-03014 | 17695043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17695043&pc=OZD | 03/05/2022 | 01/26/2024 |  | 09/06/2023 |  |  | 03/05/2022 | 09/06/2023 |  |  | 10/31/2023 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157372 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2022 |  |  |  |  | The user facility reported a 49-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site which required the patient to receive multiple blood products and the bleeding stopped. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
34744 | 1220648-2023-03015 | 17695593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17695593&pc=OZD | 08/08/2023 | 09/22/2024 |  | 09/06/2023 |  |  |  | 09/06/2023 |  |  | 06/30/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024357706 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/08/2023 |  |  |  |  | The user facility reported a male (unknown age and race) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, a red alarm stating "high purge pressure" occurred. the team changed the purge system and there was a short term improvement. finally, purge fluid and purge pressure stabilized. the pump was not explanted. there was no patient harm. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the high purge pressure issue has been completed since the original report was filed. no product was returned for evaluation. based off the clinical information available, red high purge pressure/purge system blocked alarm was observed. data logs were not returned for review. the cause of the high purge pressure issue was not determined since the product and data logs were not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03015 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
34745 | 1220648-2023-03016 | 17695919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17695919&pc=OZD | 08/09/2023 | 11/09/2023 |  | 09/06/2023 |  |  | 08/09/2023 | 09/06/2023 |  |  | 03/31/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024332534 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 65-year-old female patient received extracorporeal membrane oxygenation (va-ecmo) and impella cp smartassist heart pump support for acute myocardial infarction/cardiogenic shock. it has been reported that the impella cp pump was accidentally started after placement prior to removal of the guidewire. the pump stopped immediately. the patient required cardiopulmonary resuscitation. after removing the guide wire, the impella cp pump could be started. after spontaneous blood circulation returned, the position of the impella heart pump was optimized. it has been pointed out by abiomed customer support that the impeller may have been damaged when the pump stopped. a decision was made not to remove the impella cp. the impella support went without further complications. | [5 empty] | The investigation into pump did not start has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the root cause of pump did not start was due to use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
34746 | 1220648-2023-03017 | 17695921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17695921&pc=OZD | 08/09/2023 | 12/07/2023 |  | 09/06/2023 |  |  | 08/09/2023 | 09/06/2023 |  |  | 03/31/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347026 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp had a pump stop during support. there were purge pressure (pp) low alarms and then the pump stopped. the staff viewed the impella connect and could see flow saturation after pp low. the staff attempted to restart, but was unsuccessful. the pump was explanted. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. the device was not returned by the medical facility. no product was returned. as per clinical, physician observed purge pressure low alarm followed by high motor current pump stop. data logs were returned and impella high motor current pump stop alarms were observed after purge pressure low alarm. bag change was done and then purge pressure low alarm occurred, later purge cassette change was also done. flow saturation was present after purge pressure low at p-4 but without any prior motor current spike. the cause of the pump stop was most likely due to purge leak per log and clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34747 | 1220648-2023-03018 | 17695942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17695942&pc=OZD | 08/09/2023 | 07/09/2024 |  | 09/06/2023 |  |  | 08/09/2023 | 09/06/2023 |  |  | 06/30/2025 |  | 09/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359702 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male was implanted with an impella 5.5 device for mechanical circulatory support during a high risk coronary intervention. the patient became hypotensive and poor pump performance included decreased performance levels before suctioning. echo showed large pericardial effusion, thus brought back down to the cardiac lab for a pericardiocentesis. intensive team was unable to drain given location of effusion. after tapped in lab patient pressor requirement decreased and allowance of increasing p levels with proper waveforms and flows. physician was unsure if the dissection of left anterior descending artery (lad) or pump placement was responsible for the pericardial effusion. | [5 empty] | The investigation for the ventricle perforation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ventricle perforation could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (unites states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ d4-corrected udi number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34748 | 1220648-2023-02977 | 17682611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17682611&pc=OZD | 12/02/2021 | 10/16/2023 |  | 09/05/2023 |  |  | 12/02/2021 | 09/05/2023 |  |  | 07/31/2023 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022126634 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2021 |  |  |  |  | The user facility reported that a 68-year-old male implanted with the impella 5.5 for mechanical circulatory support experienced low purge pressure and high purge flow alarms and purge system blocked alarms occurred. a yellow luer leak was also suspected due to saturation of the cloth around the sidearm. there were no reports of harm to the patient. no further information was provided. | [5 empty] | The investigation into the reported low purge pressure and high purge flow alarms and purge system blocked alarms and purge leak was completed. the pump and data logs were returned for evaluation. analysis of the data logs show a sudden, sharp increase in purge pressure followed by purge system blocked alarms. visual inspection of the pump revealed a large clot in the outflow cage, likely occluding the purge gap. functional testing of the pump reproduced the high purge pressure and visual inspection revealed biomaterial in the purge gap until the biomaterial was removed. the root cause of the high purge pressure was ingested biomaterial occluding the purge gap. additionally, a yellow luer leak was confirmed. the root cause of the yellow luer leak was due to sodium bicarbonate weakening the plastic. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | The investigation into the reported low purge pressure and high purge flow alarms and purge system blocked alarms and purge leak was completed. the pump, purge cassette and data logs were returned for evaluation. analysis of the data logs show a sudden, sharp increase in purge pressure followed by purge system blocked alarms. visual inspection of the pump revealed a large clot in the outflow cage, likely occluding the purge gap. functional testing of the pump reproduced the high purge pressure. based on the available information, the root cause of the high purge pressure was ingested biomaterial. additionally, a yellow luer (purge) leak was confirmed. the root cause of the yellow luer leak was due to sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34749 | 1220648-2023-02978 | 17682770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17682770&pc=OZD | 05/29/2022 | 09/22/2023 |  | 09/05/2023 |  |  | 05/29/2022 | 09/05/2023 |  |  |  |  | 09/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1413896 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The user facility had an automated impella controller (aic) alarming for a purge issue. the purge cassette (pc) did not seat and detect. the team replaced the pc, the pump and the aic. the replacement of the products caused no harm or injury. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of purge issues could not be determined. after successfully passing all pm tests, the console operated within specification. there were no further complaints against it after that day. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34750 | 1220648-2023-02981 | 17684344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17684344&pc=OZD | 08/07/2023 | 01/12/2024 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  | 05/21/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024356586 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2023 |  |  |  |  | The user facility reported a 19-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during insertion of the device, the catheter shaft prolapsed on itself upon advancement. when removed, a kink was present in the catheter shaft. impella was replaced successfully. upon removal, bleeding and vascular complications were noted, and vascular surgery came to aid in closing the left femoral groin site. distal flows were restored and vascular repair of left femoral artery was deemed successful. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) has been completed since the original report was filed. the impella pump was returned for investigation. it was reported that the pump moved out of position from left ventricle (lv) due to continued cpr and was unable to reposition with catheter being kinked. upon device return, visual inspection revealed that the catheter shaft had been kinked at multiple locations before 20 cm marking of the catheter. no other abnormalities were found. the cause of the positioning issue was use related with continued cpr leading the pump being pushed out from the lv and was unable to reposition the pump with the catheter being kinked.
34751 | 1220648-2023-02982 | 17684911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17684911&pc=OZD | 05/04/2022 | 09/05/2023 |  | 09/05/2023 |  |  | 05/04/2022 | 09/05/2023 |  |  | 01/31/2024 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183209 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a 76 year-old male was implanted with an impella 5.5 device for cardiogenic shock for mechanical circulatory support. it was reported that the purge sidearm had leakage. the team did a bypass to the 5.5 pump and support proceeded with the pump without harm or injury. | [5 empty] | The investigation into the leaking sidearm has been completed. the clinical site did not return a device for evaluation. however, the clinical report was evaluated. the cause of the purge leak was most likely sidearm yellow luer damage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. per the ifu: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol."
34752 | 1220648-2023-02983 | 17685638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17685638&pc=OZD | 02/02/2022 | 03/10/2025 |  | 09/05/2023 |  |  | 02/02/2022 | 09/05/2023 |  |  | 07/31/2023 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118205 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the md was having difficulty positioning the pump properly in the patient. the support was started and flows of 4.5 lpm were achieved, however patient did not tolerate any manipulation of heart due to the positioning of the impella and decision was made to remove the device and proceed with on-pump bypass. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The cause of the delivery issue was use issue since the wire was exchanged while removing pigtail cath leading to wire malfunction. this report is being filed as part of a retrospective review of historical records.
34753 | 1220648-2023-02984 | 17685950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17685950&pc=OZD | 08/07/2023 | 01/24/2024 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  | 04/30/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024349655 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 54-year-old male patient with acute myocardial infarction and cardiogenic shock was treated with percutaneous coronary intervention. the patient went into cardiac arrest and had to be resuscitated. extracorporeal membrane oxygenation (va-ecmo) was used and then an impella cp smartassist heart pump was introduced. two days after insertion of the impella cp, ischemia was reported in the left leg, below the impella access site. the attending physicians planned the use of an external bypass. the blood circulation in the leg was suboptimal at first, but eventually improved so much that no external bypass was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's limb ischemia has been completed. complaint device not returned for investigation. the cause of limb ischemia - reversible issue was patient condition related as the patient had calcified vessels condition.
34754 | 1220648-2023-02985 | 17685976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17685976&pc=OZD | 02/17/2022 | 09/18/2024 |  | 09/05/2023 |  |  | 02/17/2022 | 09/05/2023 |  |  | 08/31/2023 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130668 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2022 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced massive bleeding from the impella access site. the patient had a hematoma removed and had pressure and dressing applied. in addition, at a later date the patient was found to have a gastro-intestinal (gi) bleed. decision was made for (do not resuscitate) dnr and the patient expired that morning. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02985 was provided in sections b2, b5, h1, and h6.
34755 | 1220648-2023-02986 | 17685994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17685994&pc=OZD | 08/07/2023 | 01/23/2024 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  |  |  | 09/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1477997 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the purge disc holder of an automated impella controller (aic) was cracked and unable to detect the purge disc. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for evaluation and was tested after arriving in fs. the issue with properly detecting the purge disc was reproduced, and purge retainer was confirmed to be broken. the root cause of the issues detecting the purge disc were due to a broken purge pressure sensor retainer. | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
34756 | 1220648-2023-02987 | 17686007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686007&pc=OZD | 08/07/2023 | 07/23/2024 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  | 04/30/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330390 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2023 |  |  |  |  | The user facility reported a 58-year-old male in non-st elevated myocardial infarction was implanted with an impella 5,5 device for mechanical circulatory support. the impella was prepped and inserted in normal fashion with graft anastomosed to anterior ascending aorta and tunneled out right supraclavicular. impella inserted wirelessly across valve, but it was noted that patient had a horizontal heart. impella couldn't be maneuvered away from the mitral valve or the sub valvular space. after the second attempt to position the impella the doctor noticed blood around the heart and found that the impella had gone through the ventricle wall. impella was taken out, wall was repaired, and an intra-aortic balloon pump was inserted from the groin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The pump was not returned for investigation. the patient was diagnosed with a thinned heart wall condition and a horizontally positioned heart. during the second attempt to position the pump using wireless access techniques, the doctor discovered a ventricle wall perforation. the cause of the ventricle perforation issue was most likely patient condition related, based on given clinical details. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank in the initial report.
34757 | 1220648-2023-02989 | 17686036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686036&pc=OZD | 08/07/2023 | 11/26/2025 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  |  |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1374547 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Mechanical Problem | Noise, Audible | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the aic was making a grinding noise. the loaner was sent to initiate the return for investigation. there was no noted patient harm related to the reported issue. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-02989 in accordance with updated procedures. e2 and e3 were erroneously entered on the initial manufacturer device report number 1220648-2023-02989. g1 reporting contact email revised on manufacturer device report 1220648-2023-02989 in accordance with updated procedures. | The investigation has been completed. data analysis: on aug 06, 2023, imc log shows that the cassette was changed from s/n (b)(6) to s/n (b)(6) an hour before the pump was unplugged from the aic. the pump sn: (b)(6) was then moved to the second console (b)(6) and the support was provided for 1 day 15 hours. there was no evidence in the logs to confirm the reported claim ¿pump is grinding¿. device analysis: the console was tested after arriving in fs. upon visual inspection, the disk-detect flag was dislodged and the retainer was not seated flush with the aic. no dextrose residue was observed. root cause: the root cause of the reported failure mode, 'pump is grinding,' was not determined. additionally, there is no log evidence to confirm this failure mode. the root cause for the broken retainer and the dislodged disk detect flag was defective component (5500-0036).
34758 | 1220648-2023-02990 | 17686045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686045&pc=OZD | 07/17/2023 | 05/16/2025 |  | 09/05/2023 |  |  | 07/17/2023 | 09/05/2023 |  |  | 01/31/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327953 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient was taken to evor for cutdown to remove impella due to possible limb ischemia. patient did have hematoma under skin. a fogerty catheter used to remove clot from external bypass sheath and the removal was successful. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the patient was taken to evor for cutdown to remove impella due to possible limb ischemia. patient did have hematoma under skin. a fogerty catheter was used to remove clot from external bypass sheath and the removal was successful. a vascular surgery treatment resoled the limb ischemia. the cause of the limb ischemia issue was biomaterial since issue solved with the removal of clot instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information for the initial mfr 1220648-2023-02990 was provided in sections d6a, d6b, h1, and h6.
34759 | 1220648-2023-02991 | 17686056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686056&pc=OZD | 08/07/2023 | 02/05/2024 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  | 05/31/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361971 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2023 |  |  |  |  | The user facility reported a 49-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump suddenly alarmed for continuous suction that was not resolved by normal trouble shooting which included decreasing p-level and repositioning under echo guidance. decision was made to remove the pump for fear of clot ingestion. when impella was removed there was some thrombus that came out with it, but none on the catheter. on echo post removal it did appear that there was some thrombus in the left ventricle apex. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus and low pump flow issues has been completed since the original report was filed. the device was returned for investigation. as per clinical, continuous suction alarms were observed with low flow and thrombus was found during pump explant. visually inspected and thrombus was observed to be ingested in the outflow cage and the impeller. no other abnormalities were found. data logs were reviewed and suction event with motor current spike was observed causing the low flow issues. the cause of the low pump flow issue was due to biomaterial based on visual inspection and data log review.
34760 | 1220648-2023-02992 | 17686115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686115&pc=OZD | 08/07/2023 | 01/23/2024 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  |  |  | 09/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1218783 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during prepping of patient for a impella 5.5 placement and coronary artery bypass graft (cabg) in operating room (or), staff turned on automated impella controller (aic) to make sure it was going to boot up without any issues. unfortunately, it booted up with controller error alarm and saying to consider back up controller. the staff replaced the aic successfully. | [5 empty] | The investigation into the reported optical signal issue has been completed since the original report was filed. the device was returned by the medical facility. device analysis console (b)(6) was sent back to field service for further investigation and was tested on a lab bench. testing confirmed that the controller error was caused by a defective/nonoperational lamp in the optical bench, which was identified by both visual inspection and comprehensive testing of the optical system. testing confirmed that replacing the original lamp in the optical bench with a known operational lamp resolved the controller error. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿ alarm header displayed in the left column; window is color-coded red for critical alarms, yellow for serious alarms, white for advisory notifications, gray for resolved alarms.¿ overlaid with a translucent yellow when the controller cannot determine catheter position, position is wrong, and placement monitoring is suspended or disabled. when the impella catheter is operating in a low flow range, placement monitoring may be suspended and the flow rate in the lower left corner of the controller display screen will turn yellow to indicate that impella position is unknown. abnormal situation. prompt action needed. 3 beeps every 15 seconds alarm header on yellow background. | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿alarm header displayed in the left column; window is color-coded red for critical alarms, yellow for serious alarms, white for advisory notifications, gray for resolved alarms.¿ overlaid with a translucent yellow when the controller cannot determine catheter position, position is wrong, and placement monitoring is suspended or disabled. when the impella catheter is operating in a low flow range, placement monitoring may be suspended and the flow rate in the lower left corner of the controller display screen will turn yellow to indicate that impella position is unknown. abnormal situation. prompt action needed. 3 beeps every 15 seconds alarm header on yellow background.
34761 | 1220648-2023-02994 | 17686207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686207&pc=OZD | 08/07/2023 | 09/21/2024 |  | 09/05/2023 |  |  | 08/07/2023 | 09/05/2023 |  |  | 04/30/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/07/2023 |  |  |  |  | The user facility reported a 75-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had severe vascular disease, but the patient was coding. decision was made to place impella. after the impella placement, staff could not find a pulse. attempted to wean pressors with hopes of reducing vaso constriction with no success in icu. decision was made to remove the impella several hours later. once impella was removed, a runoff of the leg was performed and showed no issues with flow at the impella site. patient expired shortly after explant. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's limb ischemia has been completed. product was not returned for review. the clinical investigation reads: before placing device physician was aware of the patient having severe peripheral vascular disease. after placement, could not find a pulse. discussed external bypass. per physician, he didn¿t think it would work in this case. attempted to wean pressors with hopes of reducing vaso constriction with no success in icu. decision was made to remove impella several hours later. once impella was removed, a runoff of the leg was performed and showed no issues with flow at the impella site and similar vasculature in the lower leg. based on clinical description, the root cause of limb ischemia was most likely due to patient condition. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34762 | 1220648-2023-02995 | 17686215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686215&pc=OZD | 08/08/2023 | 12/12/2023 |  | 09/05/2023 |  |  | 08/08/2023 | 09/05/2023 |  |  | 06/30/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356778 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during the peel away of a sheath for the impella cp patient, the 76-year-old male patient moved his leg, and caused the peel away intoroducer sheath to pull out unexpectedly, causing a moderate bleed requiring immediate manual pressure. the patient became hypotensive and levo was started. the stat echo showed moderate effusion. the physician decided of getting 4 units of prbcs and repeat echo in one hour. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's access site bleeding has been completed. product was not returned for investigation. per given clinical, the patient leg movement, causing the peel-away sheath to pull out, resulting in moderate bleeding. the root cause of the access site bleeding issue was patient movement based on given clinical details.
34763 | 1220648-2023-02980 | 17686821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17686821&pc=OZD | 08/06/2023 | 09/20/2024 |  | 09/05/2023 |  |  |  | 09/05/2023 |  |  |  |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/06/2023 |  |  |  |  | The user facility reported a 53-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient did not have dopplerable pulse since pump placed. nurse increased p-level on automated impella controller (aic) from p-2 back to p-6 and patient had brief return of circulation. patient eventually expired after 14 hours of impella support. no allegations were made for patient expiration due to impella. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the limb ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based off the clinical info available, the cause of the limb ischemia issue was patient condition related to small vessel based on provided clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34764 | 1220648-2023-02979 | 17687011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17687011&pc=OZD | 08/07/2023 | 09/21/2024 |  | 09/05/2023 |  |  |  | 09/05/2023 |  |  | 06/30/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364339 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 40-year-old male patient had a impella 5.5 support that was ongoing for 2 weeks. the patient suffered a cva during the support. the team made choice to withdraw care and allow the patient to expire. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into a patient stoke/cva has been completed. the impella product was not returned. the data logs were not made available. per the clinical review, the cause of the stroke was most likely due to patient condition and unknown relation to impella. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34765 | 1220648-2023-02993 | 17687131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17687131&pc=OZD | 05/25/2022 | 09/05/2023 |  | 09/05/2023 |  |  |  | 09/05/2023 |  |  | 02/29/2024 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196471 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a 39 year-old male was implanted with an impella 5.5 device for mechanical circulatory support due to cardiogenic shock. it was reported that there were purge system alarms. the staff found a potential leak in the extension tubing while running sodium bicarbonate. after replacing the purge cassette and extension tubing, the alarms resolved. pump remained on for support and there was no harm or injury. | [5 empty] | The investigation into the purge leak has been completed. the device was not returned for evaluation. however, the clinical report was evaluated. the root cause of the purge leak was determined to be a damaged extension set as a pin hole leak was observed in the set and replacing the extension set resolved the issue. this is an off the shelf extension set. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34766 | 1220648-2023-02988 | 17687161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17687161&pc=OZD | 08/07/2023 | 01/24/2024 |  | 09/05/2023 |  |  |  | 09/05/2023 |  |  | 04/30/2025 |  | 09/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355665 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2023 |  |  |  |  | The user facility reported a 65-year-old male (race unknown) in acute decompensated cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced limb ischemia with thrombosis. thrombectomy performed and distal reperfusion catheter placed. | [5 empty] | The investigation into the patient's limb ischemia with thrombosis has been completed. the product was not returned for evaluation. based on the clinical description, the patient's vessel conditions were peripheral artery disease/peripheral vascular disease. the root cause of limb ischemia was most likely thrombosis patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34767 | 1220648-2023-02955 | 17669321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17669321&pc=OZD | 09/24/2021 | 09/01/2023 |  | 09/01/2023 |  |  | 09/24/2021 | 09/01/2023 |  |  | 03/31/2023 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060374 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2021 |  |  |  |  | The user facility reported that a 62-year-old male implanted with the impella 5.5 pump experienced high purge pressure during the last 12 hours of support. the automated impella controller (aic) did not produce any low purge flow or high purge pressure alarms. the pump was subsequently explanted. there were no reports of harm to the patient from the console failing to alarm for the purge pressure issues. | [5 empty] | The investigation into the reported low purge flow and high purge pressure was completed. the device and data logs were returned for evaluation. analysis of the data logs confirmed an increased purge pressure and decreased purge flow on the last day of support. functional testing of the pump was able to reproduce the high purge pressure. the root cause of the high purge pressure was biomaterial buildup occluding the purge gap. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34768 | 1220648-2023-02956 | 17670589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17670589&pc=OZD | 11/27/2021 | 09/01/2023 |  | 09/01/2023 |  |  | 11/27/2021 | 09/01/2023 |  |  |  |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130653 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/27/2021 |  |  |  |  | The user facility reported that a 34-year-old male implanted with the impella 5.5 noted that high purge pressure/low purge flow was observed 13 days into support. the purge pressure was noted to be at 1080 mmhg and the purge flow was 1.7 ml/hr. the "purge pressure high" alarm did not trigger. the decision was then made to infuse tissue plasminogen activator (tpa) into the purge solution, which successfully resolved the issue. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low purge flow and no high purge pressure alarm was completed. the device was not returned for evaluation. after review of the clinical information, it was determined that the cause of the high purge pressure was most likely biomaterial buildup occluding the purge gap. the root cause of the high purge pressure was most likely biomaterial buildup occluding the purge gap. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34769 | 1220648-2023-02957 | 17670608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17670608&pc=OZD | 09/05/2021 | 09/01/2023 |  | 09/01/2023 |  |  | 09/05/2021 | 09/01/2023 |  |  | 06/30/2023 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022104964 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/05/2021 |  |  |  |  | The user facility reported that a 59-year-old male implanted with the impella 5.5 pump had low purge flow which was treated with tpa for several hours. the subsequent minimum and maximum flows were equal and pump replacement was recommended due to pump saturation. the pump was removed hours later when the issue did not resolve. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported flow saturation was completed. analysis of the returned data logs confirmed a sharp increase in purge pressure and saturated flows correlated with a high motor current. this spike in purge pressure suggests a kink in the purge system. the pump was returned, and the high purge pressure and saturated flows were reproduced. blood was found in the purge line and a kink was found in the catheter just above the kink protector. the root cause of the high purge pressure was a kink in the purge line resulting in blood ingress into the purge line causing a blockage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34770 | 1220648-2023-02958 | 17670629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17670629&pc=OZD | 11/25/2021 | 09/01/2023 |  | 09/01/2023 |  |  | 11/25/2021 | 09/01/2023 |  |  |  |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022126653 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2021 |  |  |  |  | The user facility reported that a 65-year-old male implanted with the impella 5.5 noted that the pump did not alarm when purge flow dropped below limits. the rn recognized when flow was 0.5 ml/hr and called for help. the purge cassette was changed and tpa protocol was then initiated, but neither method resolved the issue. the next day the patient received purge outflow blocked alarms. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low purge flow and no high purge pressure alarm was completed. the data logs were received for evaluation. analysis of the data logs revealed a gradual increase in purge pressure throughout the third day of support with a concurrent decrease in purge flow. the "purge pressure high" alarm did not trigger, however the "purge system blocked" alarm occurred repeatedly the following day. additionally, the "suction" alarm triggered throughout support. no other abnormalities were noted. the device was not returned; therefore, the root cause of the high purge pressure was unable to be determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34771 | 1220648-2023-02959 | 17670679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17670679&pc=OZD | 10/31/2021 | 09/01/2023 |  | 09/01/2023 |  |  | 10/31/2021 | 09/01/2023 |  |  | 06/30/2023 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022117799 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/31/2021 |  |  |  |  | The user facility reported a 62 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that they are having leaking issues from purge side. there was no patient injury reported.the pump was run until planned vad placement with no adverse impact to the patient. | [5 empty] | The pump was returned for evaluation. visual inspection confirmed a crack in the yellow luer. sodium bicarbonate was used as a purge solution. the pump was run in the lab without issue. data logs were provided and reviewed and found that 'purge pressure low' alarms had been captured. the logs showed the pump operating as expected. the root cause of the reported purge sidearm leak was determined to be a damaged yellow luer. the root cause of the reported low flow was unable to be determined as no flow issues were reproduced. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. per the ifu: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34772 | 1220648-2023-02960 | 17673123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673123&pc=OZD | 04/19/2023 | 09/01/2023 |  | 09/01/2023 |  |  | 08/03/2023 | 09/01/2023 |  |  | 12/31/2024 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024318894 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | [5 empty] | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/03/2023 |  |  |  |  | The complainant reported a 76 year old male patient presenting for hemodynamic support via the impella 5.5 device. new information received, via abiomed's long-term outcome and quality indicator (loqi) registry, revealed a potential relationship of vt requiring defibrillation. the impella device was implicated as a possible causal/contributory factor. | [5 empty] | The impella device was not received and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34773 | 1220648-2023-02961 | 17673710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673710&pc=OZD | 08/04/2023 | 02/05/2024 |  | 09/01/2023 |  |  | 08/04/2023 | 09/01/2023 |  |  | 06/30/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024360861 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 57-year-old patient received support from an impella 5.5 smartassist heart pump and veno-arterial extra-corporeal membrane oxygenation (va-ecmo) after coronary artery bypass surgery. on day 7 after inserting the impella 5.5, a sudden pump stop occurred. the pump could be restarted first. the customer was advised by abiomed customer support that the pump could stop again and it was recommended that the pump be removed. later the impella 5.5 stopped again and could not be restarted. the pump was removed and the patient was subsequently stable on a low dose of catecholamines. the va-ecmo was explanted two days before the pump stop. the day before the incident, the patient was diagnosed hit positive, so the purge fluid was switched from 5% glucose solution with heparin to 5% glucose solution with sodium bicarbonate. according to the doctor treating him, there were a large number of thrombi in the patient's body. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, data logs were returned and multiple motor current spikes with saturated flow followed by pump stop was found. the cause of the pump stop issue was most likely biomaterial with motor current spike observed followed by saturated flow and several pump stop alarms per log review.
34774 | 1220648-2023-02962 | 17673741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673741&pc=OZD | 08/04/2023 | 09/20/2024 |  | 09/01/2023 |  |  | 08/04/2023 | 09/01/2023 |  |  | 04/30/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported admitted to priming the impella device inside of the patient prior to abiomed on-site support. this occurred independent of abiomed on-site support. the physician questioned the amount of air that could have been introduced as a result of going against instruction for use (ifu) in this manner. the patient stabilized after insertion of the impella, however, became profoundly hypotensive and coded approximately 1.5 hours later and expired. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into embolism issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the embolism issue was most likely patient condition related and determined to be unknown relation to the pump. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
34775 | 1220648-2023-02963 | 17673756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673756&pc=OZD | 08/04/2023 | 01/23/2024 |  | 09/01/2023 |  |  | 08/04/2023 | 09/01/2023 |  |  |  |  | 09/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1320441 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 96 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/04/2023 |  |  |  |  | The user facility reported that during preparation of an impella cp device, the automated impella controller (aic) did not prime the purge tubing. the user facility used a backup aic. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. product was returned for evaluation. console inspected and evaluated on an engineering lab bench. the console was found to be within spec. according to pm ((b)(6)) 12 procedure. purge pressure low issue was not reproducible during investigation. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34776 | 1220648-2023-02964 | 17673762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673762&pc=OZD | 08/05/2023 | 09/01/2023 |  | 09/01/2023 |  |  | 08/05/2023 | 09/01/2023 |  |  | 05/31/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024349818 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/05/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the pump stopped and the team was unable to restart. the patient was dependent on the pump. the patient decompensated, cpr was initiated and the patient expired shortly after the pump stoppage. | [5 empty] | The investigation into the patient death has been completed. the data logs show an impella #119 alarm. the pump is needed to confirm the cause of the #119 alarm. the cause of the pump stop was not determined since product was not returned. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34777 | 1220648-2023-02965 | 17673782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673782&pc=OZD | 04/21/2022 | 09/01/2023 |  | 09/01/2023 |  |  | 04/21/2022 | 09/01/2023 |  |  | 01/31/2024 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183581 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 61 year-old male was implanted with an impella 5.5 device for bridge to transplant. it was reported there was a small leak at the purge side arm after rolling the patient. the side arm retainer was on and physician removed it to visualize the leak, which could be seen at side arm itself. bone wax applied over leak. there was no patient harm reported. | [5 empty] | The facility did not return the device nor data logs for evaluation. however, the clinical review was done and reported that the leak occurred after rolling the patient and that the location of the leak was at the sidearm itself. the root cause of the sidearm leak could not be determined as the product was not returned for evaluation. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34778 | 1220648-2023-02966 | 17673845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673845&pc=OZD | 08/05/2023 | 09/20/2024 |  | 09/01/2023 |  |  | 08/05/2023 | 09/01/2023 |  |  | 04/30/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352831 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Material Deformation | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Vascular Dissection | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/05/2023 |  |  |  |  | Additional information received that blood products were required and vascular damage was observed. manual pressure held, difficulty achieving hemostasis. pump was pulled out of position while attempting to hold pressure. impella explanted to perform contralateral access balloon tamponade and cutdown and was able to control the bleeding. | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient's intra-aortic balloon pump was removed and the impella cp was exchanged over the wire without difficulty. after correct impella position was verified by fluoro and waveforms, the impella peel-away sheath was removed completely out of the body before peeling away. the repositioning sheath was then advanced, but md met some resistance and pulled back and readvanced. impella placement was again verified by fluoro, however, reposition sheath was noted to be kinked. the impella cp was explanted. | The vascular surgery team explanted the impella and performed a contralateral access balloon tamponade and cutdown. they were able to control the bleeding; however, there was a delay in blood products from the blood bank and a drop in mean arterial pressure (map). the patient went into arrest and cardiopulmonary resuscitation was administered. the patient coded for 40 min before the patient was pronounced dead. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  | The investigation into the reported delivery issue - use, hematoma and vascular damage has been completed since the original report was filed. per the clinical information provided, the root cause of the vascular damage - peripheral vessel issue was most likely use related. the repo unit was found kinked after insertion which caused vascular damage and hematoma at the access site. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrected information for the initial mfr 1220648-2023-02966 was provided in sections b2, b5, d6a, d6b, h1, and h6.
34779 | 1220648-2023-02967 | 17673852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673852&pc=OZD | 08/05/2023 | 07/08/2024 |  | 09/01/2023 |  |  | 08/05/2023 | 09/01/2023 |  |  | 03/31/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346243 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was bleeding and had a hematoma at the impella cp insertion site. three units of blood were administered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations section: entering cardiac output section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ component code 92.
34780 | 1220648-2023-02968 | 17673863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673863&pc=OZD | 04/18/2023 | 09/15/2024 |  | 09/01/2023 |  |  | 04/18/2023 | 09/01/2023 |  |  | 08/31/2024 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259596 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 73-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. a small dissection was noted on x-ray at the impella cp inserted femoral artery. staff is not sure when it happened and not sure what caused it. while still in case patient began to decompensate and sadly despite maximal effort and after significant advanced cardiovascular life support (acls), patient passed. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported vascular damage to the peripheral vessel has been completed since the original report was filed. no product was returned for investigation. it was noted that the patient had pad/pvd and was coding during access. the root cause of the vascular damage issue was most likely patient condition related. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02968 was provided.
34781 | 1220648-2023-02969 | 17673878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673878&pc=OZD | 04/14/2022 | 09/01/2023 |  | 09/01/2023 |  |  | 04/14/2022 | 09/01/2023 |  |  | 01/31/2023 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023178682 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2022 |  |  |  |  | The user facility reported a 36 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported there was leaking from the purge sidearm and flow saturation during support. a purge side arm bypass was performed. the pump was ultimately replaced. there was no patient harm reported. | [5 empty] | The impella 5.5 was returned without the sidearm so the exact location of the original purge leak was unable to be determined. the pump was evaluated and biomaterial was found to be built up in the purge gap. the data logs were returned and confirm the purge leak that was resolved using a purge bypass. the logs also confirm that purge pressure dropped at the end of the case briefly and then purge pressure increased to higher than expected values. when purge pressure increased, both motor current and flow increased resulting in saturated flows. the root cause of the original purge leak was unable to be determined as the sidearm was not returned for analysis. the root cause of the saturated flows was determined to be biomaterial buildup in the purge gap likely as a result of the purge leak. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34782 | 1220648-2023-02970 | 17673892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673892&pc=OZD | 08/03/2023 | 02/20/2024 |  | 09/01/2023 |  |  | 08/03/2023 | 09/01/2023 |  |  | 12/31/2024 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024317533 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Separation | Inadequate User Interface | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2023 |  |  |  |  | The user facility reported a 60-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the sterile sleeve of an impella 5.5 pump detached during patient support. it was believed that the sterile sleeve may have come loose when the patient was exercising. the sleeve was reattached in sterile fashion and support continued. there were no reports of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - great vessel) issue has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the sterile sleeve damage issue was most likely use related to excessive force supposedly happened while patient was exercising.
34783 | 1220648-2023-02971 | 17673895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673895&pc=OZD | 01/16/2022 | 03/10/2025 |  | 09/01/2023 |  |  | 01/16/2022 | 09/01/2023 |  |  | 09/30/2023 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022144172 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2022 |  |  |  |  | The user facility reported a 42-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical team performing an impella implant had difficulty with anastomosis resulting in vessel dissection and oozing. the vessel dissection was sutured and patient was put on frequent checks for hematoma. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34784 | 1220648-2023-02972 | 17673908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673908&pc=OZD | 10/22/2022 | 09/19/2024 |  | 09/01/2023 |  |  | 10/22/2022 | 09/01/2023 |  |  | 09/30/2023 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143115 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/22/2022 |  |  |  |  | The user facility reported a 63-year-old male (race unknown) in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the team escalated the patient to an impella 5.5 device for increased cardiac support. the patient experienced access site bleeding issues and was given one unit of packed red blood cells. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported access site bleeding has been completed since the original report was filed. no device was returned for investigation. the cause of the access site bleeding was most likely due to anticoagulation management since last result of acts reported was 297 which is above the range of 160-180. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34785 | 1220648-2023-02973 | 17673932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17673932&pc=OZD | 08/04/2023 | 01/15/2024 |  | 09/01/2023 |  |  | 08/05/2023 | 09/01/2023 |  |  | 05/31/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352055 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2023 |  |  |  |  | The user facility reported a 49-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while inserting the peel away sheath, the patient was moving aggressively. nothing else noted at the time of insertion. peel away removed and repo sheath advanced. the patient's groin would not stop bleeding after that. the patient was given blood products. the patient also had bleeding from another access site not related to the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident | The investigation for the vascular damage has been completed. the purge cassette was returned for evaluation but was not related to the complaint. however, clinical details provided are sufficient to determine the root cause. the most likely root cause of the reported vascular damage was most likely due to use related.
34786 | 1220648-2023-02974 | 17674738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17674738&pc=OZD | 08/06/2023 | 01/11/2024 |  | 09/01/2023 |  |  | 08/06/2023 | 09/01/2023 |  |  | 02/28/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338810 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient's right groin impella insertion site had a new lateral swelling/purple color that the clinician initially thought it was from a known and resolved hematoma however, the discoloration and edema appear to have moved from initial medial right groin. the medical team was made aware, and area of concern outlined. the dressing was clean, and the angle matched. 2 units of prbcs were given overnight. the patient¿s hemoglobin(hg) dropped again to 6.9 and the team was likely going to transfuse 1 more unit. after several days the hematoma was successfully managed with stitches around site. the insertion site is now showing clean, dry, and intact. slight old bruising evidence around site but with no active swelling noted and groin was soft. it was further reported that the team was planning to remove the impella cp when left ventricle thrombus was noted and therefore did not feel comfortable with any repositioning or movement of device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding and thrombus has been completed since the original report was filed. product was not returned for review. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review and the root cause of thrombus was most likely due to patient condition. the failure modes will be monitored and trended.
34787 | 1220648-2023-02975 | 17674740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17674740&pc=OZD | 08/06/2023 | 01/17/2024 |  | 09/01/2023 |  |  | 08/06/2023 | 09/01/2023 |  |  | 04/30/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349765 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the nurse stated there was a lot of bleeding during impella support. sales rep. walked the nurse through changing the dressing and noted a significant amount of clot. all of this was removed and the angle of entry was maintained and a new dressing was applied. no hematoma noted. groin was soft. the rep. encouraged the nurse to change the dressing and address any bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported thrombus has been completed since the original report was filed. no product was returned. as per clinical patient had a lot of bleeding at access site, dressing changed and later noted a significant amount of clot with calcified vessel condition. clot was removed and new dressing changed to resolve bleeding. during weaning and explant of impella heparin stopped due to concerns for heparin induced thrombocytopenia with platelet drop of 50% noted. the cause of the access site bleeding issue was most likely patient condition related per clinical. the cause of the thrombocytopenia issue was most likely patient condition related and unknown relation to impella per clinical.
34788 | 1220648-2023-02976 | 17674742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17674742&pc=OZD | 08/05/2023 | 04/12/2024 |  | 09/01/2023 |  |  | 08/05/2023 | 09/01/2023 |  |  | 04/30/2025 |  | 09/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338159 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Nervous System Injury | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2023 |  |  |  |  | The user facility reported a 34-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient had right arm discomfort. vascular was consulted and stated patient had likely brachial plexus injury with 5.5 placement. no additional details are available. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the patient's nervous system injury has been completed since the initial report was filed. no impella products were returned for investigation. the root cause of the nerve compression injury cannot be determined due to insufficient clinical information provided and lack of product returned.
34789 | 1220648-2023-02920 | 17660698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17660698&pc=OZD | 11/16/2021 | 08/31/2023 |  | 08/31/2023 |  |  | 11/16/2021 | 08/31/2023 |  |  | 07/31/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022121282 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2021 |  |  |  |  | The user facility reported a consistent purge leak at the impella 5.5 yellow luer which triggered a purge system open alarm. the purge cassette was replaced without resolution. the pump was subsequently replaced. of note, sodium bicarbonate was used in the purge. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was complete. the device was not returned for evaluation. the site reported a purge leak at the yellow luer and that sodium bicarbonate was used as a purge solution. therefore, the root cause of the purge leak was most likely a damaged yellow luer on the purge sidearm. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. of note the ifu states: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34790 | 1220648-2023-02921 | 17660731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17660731&pc=OZD | 12/10/2021 | 01/26/2024 |  | 08/31/2023 |  |  | 12/10/2021 | 08/31/2023 |  |  | 08/31/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022137425 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2021 |  |  |  |  | The user facility reported a 33-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a hematoma which required the patient to go to the or to have it surgically addressed. seven days later the patient had developed ischemia in the right arm and had to have the pump explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the access site bleed and limb ischemia has been completed. however, with limited clinical details nor the impella device, the root cause of the reported could not be determined. the root cause of the limb ischemia was most likely patient condition related since the pump was not providing enough support.
34791 | 1220648-2023-02922 | 17660759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17660759&pc=OZD | 11/05/2021 | 08/31/2023 |  | 08/31/2023 |  |  | 11/05/2021 | 08/31/2023 |  |  | 07/31/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113610 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/05/2021 |  |  |  |  | The user facility reported that the patient experienced an impella 5.5 pump purge leak and that the purge system was "cracked". this had occurred on day 11 of the pump support. a purge system bypass was performed without issues. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the pump was returned with a severed catheter as well as evidence that a purge sidearm bypass had been performed. the purge sidearm was not returned for evaluation. based on the clinical details, the root cause of the reported purge sidearm leak was most likely a damaged filter. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34792 | 1220648-2023-02923 | 17660777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17660777&pc=OZD | 11/04/2021 | 08/31/2023 |  | 08/31/2023 |  |  | 11/04/2021 | 08/31/2023 |  |  | 08/31/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022125474 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/04/2021 |  |  |  |  | The user facility reported that a 51-year-old male patient implanted with the impella 5.5 for management of cardiomyopathy experienced a purge fluid leak after the "purge system" cracked during a purge cassette exchange. the pump was replaced with another impella 5.5. of note, sodium bicarbonate was used in the purge solution throughout support. no other cleansing agents were reported to be in use. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. neither the device nor the data logs were returned for investigation. therefore, the cause of the purge leak is undetermined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34793 | 1220648-2023-02924 | 17660786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17660786&pc=OZD | 01/01/2022 | 08/31/2023 |  | 08/31/2023 |  |  | 01/01/2022 | 08/31/2023 |  |  | 06/30/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022105414 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/01/2022 |  |  |  |  | The user facility reported that the plastic on the pump side of the purge sidearm broke. initially, a purge side arm bypass was performed but the pump was ultimately replaced with a new impella 5.5. additional information indicated that the sidearm had broken between the infusion filter and the red impella plug. of note, sodium bicarbonate was used as a purge solution for a few days at the beginning of the case. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge sidearm break and purge system open alarms was completed. the device was not returned for evaluation, but pictures of the device and data logs were evaluated. the data logs confirmed multiple purge system open alarms that were associated with a drop in purge pressure. evaluation of the submitted image confirmed a complete break of the sidearm proximal to the filter. the root cause of the low purge pressure was the damaged purge sidearm. this damage most likely occurred due to excessive force during use. of note, sodium bicarbonate was used in the purge solution. the root cause of the low purge pressure was the damaged purge sidearm. this damage most likely occurred due to excessive force during use. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. of note the ifu states: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34794 | 1220648-2023-02925 | 17660857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17660857&pc=OZD | 12/23/2021 | 08/31/2023 |  | 08/31/2023 |  |  | 12/23/2021 | 08/31/2023 |  |  | 07/31/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022115139 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2021 |  |  |  |  | The user facility reported a 31 year-old male was implanted with an impella 5.5 pump for bridge to transplant. it was reported that there was a leak at the luer after 5 weeks of support. bicarbonate was used as the purge fluid. the pump was ultimately explanted as a result. there was no patient injury reported. | [5 empty] | The investigation into the reported luer leak has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. based on the use of sodium bicarbonate as a purge solution, the root cause of the purge leak was most likely a damaged yellow luer on the purge sidearm. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34795 | 1220648-2023-02926 | 17660886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17660886&pc=OZD | 09/21/2021 | 08/31/2023 |  | 08/31/2023 |  |  | 09/21/2021 | 08/31/2023 |  |  | 06/30/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022106334 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2021 |  |  |  |  | The user facility reported that the purge system of the impella 5.5 was leaking at the yellow luer after the purge cassette was changed. purge bypass was performed. no alcohol or hemostats were used, and sodium bicarbonate was used in the purge. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the clinical information was reviewed. the root cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate in the purge solution. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. of note the ifu states: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34796 | 1220648-2023-02927 | 17661839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17661839&pc=OZD | 12/15/2021 | 11/29/2023 |  | 08/31/2023 |  |  | 12/15/2021 | 08/31/2023 |  |  | 09/30/2022 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021003859 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2021 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient undergoing impella implant experienced bleeding from the insertion site during the procedure. multiple blood products had to be administered. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into access site bleeding has been completed. the cause of the access site bleeding cannot be determined as patient insufficient clinical information was provided to determine a root cause and no product was returned.
34797 | 1220648-2023-02928 | 17661890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17661890&pc=OZD | 01/07/2022 | 08/31/2023 |  | 08/31/2023 |  |  | 01/07/2022 | 08/31/2023 |  |  | 07/31/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022119971 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/07/2022 |  |  |  |  | The user facility reported a 50 year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge sidearm leak during routine tubing change. they were using bicarbonate in the purge. there was no patient injury reported. | [5 empty] | The investigation into the reported purge system leak has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. based on the use of sodium bicarbonate as a purge solution, the root cause of the purge leak was most likely a damaged yellow luer on the purge sidearm.. no product or data logs were provided. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34798 | 1220648-2023-02929 | 17662044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17662044&pc=OZD | 12/13/2021 | 01/11/2024 |  | 08/31/2023 |  |  | 12/13/2021 | 08/31/2023 |  |  | 09/30/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022140634 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2021 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had experienced bleeding in the chest and required surgery and a washout to have the bleeding stopped, and in addition also required 2 units of blood to be given. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding- major has been completed. the product was not returned. as per the clinical information provided, the root cause of the bleeding is patient condition related.
34799 | 1220648-2023-02930 | 17662804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17662804&pc=OZD | 08/03/2023 | 11/29/2023 |  | 08/31/2023 |  |  | 08/03/2023 | 08/31/2023 |  |  | 02/28/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338757 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2023 |  |  |  |  | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed a groin hematoma at the access site after removal of the peel-away sheath. physician decided to remove the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's hematoma has been completed. no product was returned. as per clinical, hematoma was observed when peel away sheath was removed and pump was explanted. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information.
34800 | 1220648-2023-02931 | 17663339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17663339&pc=OZD | 08/04/2023 | 01/12/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  | 08/31/2024 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023257456 |  | 0048-0014 EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2023 |  |  |  |  | The user facility was to implant a patient (unknown age, race, gender) with an impella cp for circulatory support. when the cp was connected to the automated impella controller (aic) there was an error message stating the optical sensor was defective. the team exchanged the aic but that did not resolve the error message. the team then used a new impella cp and the error message was eliminated. there was no patient harm. | [5 empty] | The device has been received from the user facility. the investigation has not been completed at this time. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical signal issue has been completed. the impella device and the data logs were returned for evaluation. upon visual inspection of the device, a kink was observed. the light continuity test was done, and the light test was passed. performed the optical bench service and verified that the 5s parameters were in nominal range. the device functioned/operated to specification with flow of 3.9 l/min at auto. no other abnormalities were found. data logs were reviewed which revealed pump stop alarms and a significant increase in gain, signal-to-noise (snr) and lamp being maxed out and contrast greater than 40%. the cause of the optical signal failure was most likely an air gap from between the aic and the patient cable plug during the case run. the cause of the pump did not start issue was not determined with high motor current pump stop observed per log review and the issue not being reproduced during field investigation.
34801 | 1220648-2023-02932 | 17663400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17663400&pc=OZD | 04/11/2023 | 01/23/2024 |  | 08/31/2023 |  |  | 08/02/2023 | 08/31/2023 |  |  | 11/30/2024 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297359 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2023 |  |  |  |  | Per proctect iv: a 41-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp. the patient had a vascular access site wound infection. no other details have been provided. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the infection/sepsis issue has been completed. the device was not returned for investigation. product was not returned for review. based on clinical description, the cause of the infection was most likely patient condition related since the device is sterilized under validated conditions and there was no evidence to indicate a break in the sterile barrier or pump contamination, unknown relation to impella.
34802 | 1220648-2023-02933 | 17663416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17663416&pc=OZD | 08/03/2023 | 01/22/2024 |  | 08/31/2023 |  |  | 08/03/2023 | 08/31/2023 |  |  | 04/30/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024345390 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2023 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump stopped during support on day 15. the impella was removed with no patient hemodynamic issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the of pump stop issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs were not available for review of pump stop issue. as per clinical information, the pump was used for more than 14 days, and the pump stop alarm was observed after design life period of the impella cp. the cause of the pump stop issue was not determined since the product and data logs were not returned for review.
34803 | 1220648-2023-02934 | 17663447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17663447&pc=OZD | 07/29/2023 | 02/08/2024 |  | 08/31/2023 |  |  | 08/03/2023 | 08/31/2023 |  |  | 02/28/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024320751 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2023 |  |  |  |  | The user facility reported a patient (unknown age and gender) noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during support, there was a pump stop and the motor current increased after surgical mode. evaluation showed that the pump was placed 2-3 cm deeper, and the surgeon was concerned that the tip of the pigtail might damage the left ventricular myocardium during prolapse. after adjusting the position and reconfirming the echo evaluation before blocking, aorta blocking was performed. it was difficult to obtain an intracardiac bloodless visual field, because the blockage site was poor. the blockage site was adjusted twice and ventricular septal defect closure surgery was completed. although the surgical mode was canceled, the motor current increased and the support could not be resumed, so the impella was removed. | [5 empty] | The impella device has been received for investigation, however the investigation is incomplete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issues and pump stop has been completed. the impella device was returned and visually inspected finding biomaterial observed around impella. the pump was connected to the console and a "pump stop" alarm was reproduced. the pump was soaked in water and cleaning enzymes to remove the biomaterial. once the biomaterial was removed, the pump ran successfully in water without stopping within a specific p-level and flow rate. no other abnormality observed. the root cause of the pump stop was biomaterial ingestion due to patient condition. the root cause of the positioning issues was use related issue as the pump was placed 2-3 cm deeper due to improper technique. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34804 | 1220648-2023-02935 | 17663482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17663482&pc=OZD | 08/04/2023 | 04/02/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  | 04/30/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350513 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 81-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient had a hematoma develop during transplant from cath lab to room. the hematoma was expressed, redressed and a femstop applied. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. per the provided clinical details, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided.
34805 | 1220648-2023-02936 | 17663721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17663721&pc=OZD | 08/03/2023 | 09/20/2024 |  | 08/31/2023 |  |  | 08/03/2023 | 08/31/2023 |  |  | 12/31/2024 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315324 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Difficult to Advance | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/03/2023 |  |  |  |  | The user facility reported a 60-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. once the impella was in the aorta, the physician had difficulty with crossing valve. ventricle filled with attempts to make valve incompetent. decision made to try diagnostic catheter and wire. 0.035" and al1 advanced across valve without issue. placement wire advanced and catheter was removed. impella advanced without issue. slowly ramped up as bypass flows reduced. able to get to p9 with >5 l flow. eventually encountered suction alarms. flows reduced, volume given back to patient. then the patient had significant bleeding issues further impacting volumes and then right ventricle began to fail. the team was unable to maintain much flow, ended up maxed out on chemical support. patient expired after 1.2 hours of impella support. surgeon does not believe the cause of death was related to the impella. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. it was reported that the patient has heavy anastomosis near the valve area and the physician had difficultly crossing the valve. the root cause of the delivery issues is due to the patient¿s condition. the root cause of the low pump flow issue was patient condition related to diseased vessels. the root cause of the access site bleeding issue cannot be determined as product was not returned and significant clinical information was not provided. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34806 | 1220648-2023-02937 | 17664483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17664483&pc=OZD | 11/01/2021 | 08/31/2023 |  | 08/31/2023 |  |  | 11/01/2021 | 08/31/2023 |  |  | 07/31/2023 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022126633 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2021 |  |  |  |  | The user facility reported that while attempting to change the purge tubing per hospital protocol, a leak was observed at the yellow-to-yellow connection of the impella 5.5 pump. purge flow (pf) and purge pressure (pp) remained below normal limits. the surgeon performed a purge bypass to prevent a possible pump stop. the pump was successfully weaned and removed on (b)(6) 2021. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak and low pf and pp was completed. the device was returned for evaluation. the returned product was visually inspected and a purge bypass was confirmed. the pump sidearm was not returned for review. the cause of the purge leak was most likely sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34807 | 1220648-2023-02938 | 17665146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665146&pc=OZD | 08/02/2023 | 09/01/2023 |  | 08/31/2023 |  |  | 08/02/2023 | 08/31/2023 |  |  | 03/31/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024333678 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2023 |  |  |  |  | The user facility reported a 51-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed left lower limb ischemia which was resolved with an antegrade sheath insertion. the patient was provided va-ecmo blood infusion tube. the patient remained on impella support for 8 additional days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. the cause of the limb ischemia was most likely patient condition related based on the provided clinical details that the patient was ischemic in both legs.
34808 | 1220648-2023-02939 | 17665244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665244&pc=OZD | 08/02/2023 | 08/31/2023 |  | 08/31/2023 |  |  | 08/02/2023 | 08/31/2023 |  |  |  |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1650098 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the loaner automated impella controller (aic) returned from field for functional check and when turned on a flickering screen was observed. this issue happened at the manufacturing site. | [5 empty] | The investigation into the reported aic - display issue" has been completed, however, the device failure was not reproduced during testing. because screen flickering can be caused by both the 4-in-1 and lcd screen, and this failure mode happened in emfg, both will be replaced as a precaution. during data analysis, there was nothing in the logs that indicate the screen flickering. the cause of the issue was not determined since failure was not reproduced.
34809 | 1220648-2023-02941 | 17665352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665352&pc=OZD | 08/04/2023 | 01/15/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  | 05/31/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360995 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2023 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella cp for escalation of care from an intra-aortic balloon pump. the patient had a massive blood loss from the impella insertion site. patient hemoglobin was 9.1 and after receiving 2 units of blood it was 6.2. impella was removed and 16 fr gore sheath placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. therefore, the root cause of the access site bleed was due to patient condition, lack of vessel recoil after repo unit insertion due to stenosis.
34810 | 1220648-2023-02940 | 17665355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665355&pc=OZD | 08/02/2023 | 01/17/2024 |  | 08/31/2023 |  |  | 08/02/2023 | 08/31/2023 |  |  | 06/30/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359700 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 prep for insertion. a new surgical tech was prepping introducer/dilator and during handling of the sheath the sheath cracked. axillary kit from another impella 5.5. kit was opened to replace the cracked sheath. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. clinical details stated that a new surgical tech was prepping the introducer/dilator and during handling of the sheath, the sheath cracked. the cause of the issue was an introducer sheath split along the scorelines.
34811 | 1220648-2023-02942 | 17665410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665410&pc=OZD | 08/02/2023 | 02/02/2024 |  | 08/31/2023 |  |  | 08/02/2023 | 08/31/2023 |  |  | 04/30/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347933 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/02/2023 |  |  |  |  | The user facility reported a 53-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. shortly after the device was implanted, the pump stopped. the device was removed and replaced with no injury to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. the device was not returned by the medical facility. per the clinical information provided, the root cause of the pump stop/did not start issue was not determined since product was not returned and data logs were not returned as well.
34812 | 1220648-2023-02943 | 17665566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665566&pc=OZD | 08/02/2023 | 09/19/2024 |  | 08/31/2023 |  |  | 08/02/2023 | 08/31/2023 |  |  | 05/31/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024358375 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/02/2023 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pigtail/inlet was in the patient's papillary muscle and the physician was unable to optimize performance. the patient had ongoing hemolysis and expired the following day. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the positioning issues and the hemolysis issue has been completed since the original report was filed. neither the product nor data logs were returned for investigation. the root cause of the hemolysis issue was most likely improper positioning of the device against the mitral apparatus. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
34813 | 1220648-2023-02945 | 17665688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665688&pc=OZD | 08/03/2023 | 03/01/2024 |  | 08/31/2023 |  |  | 08/03/2023 | 08/31/2023 |  |  | 08/31/2024 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2023256741 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2023 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on impella support for 208 hours before the patient expired. the patient expired due to cerebral hemorrhage. additionally, there was additional ecmo treatment. there were no issues that occurred with the operation of the impella. the relationship of the cerebral hemorrhage to the impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined.
34814 | 1220648-2023-02946 | 17665795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665795&pc=OZD | 07/29/2023 | 08/31/2023 |  | 08/31/2023 |  |  | 08/03/2023 | 08/31/2023 |  |  | 02/28/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341773 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 64-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the attending physician reports patient's aortic valve was damaged s/p implant or d/t 2 hour explant of the impella cp. the patient was explanted on (b)(6) 2023. patient was stable post explant and was sitting up in the chair post explant on no support . per report, the patient decompensated the weekend of the (b)(6) 2023 and was reintubated and tee performed and showed a damaged aortic valve. patient was taken to the operating room (or) for aortic valve replacement. | [5 empty] | (h3) the investigation into the reported ¿heart valve damage¿ was completed, despite not receiving the device from the customer. per cts doctor patient aortic valve was damaged s/p implant or d/t 2 hour explant of the impella cp. the patient was explanted on (b)(6) 2023. the patient was stable post explant and was sitting up in the chair post explant on no support . per report, the patient decompensated the weekend of the (b)(6) 2023 and was reintubated and tee performed and showed a damaged aortic valve. the patient was taken to the operating room (or) for aortic valve replacement. the cause of the heart valve damage issue was not determined since product was not returned and sufficient clinical information was not provided. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34815 | 1220648-2023-02947 | 17665940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665940&pc=OZD | 08/04/2023 | 01/24/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  |  |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1197229 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/04/2023 |  |  |  |  | The user facility reported automated impella controller (aic) failing to recognize a purge cassette during startup of an impella case. a new console was used without further issue. | [5 empty] | The impella device was received and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the automated impella controller purge disc/flag issues has been completed. data review: imc logs from the complaint date revealed that the clinical staff was following the case wizard when a purge-related alarm was triggered. the purge cassette was detected successfully, the alarm was triggered by a purge disc not detected event. the clinical staff attempted to repeat the case wizard, however the purge disc not detected events reoccurred. device analysis: the console was tested after arriving in field service. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer).
34816 | 1220648-2023-02948 | 17665952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665952&pc=OZD | 08/04/2023 | 01/19/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  | 06/30/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359698 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported in a bridge to transplant 65-year-old male patient a hematoma was noted and increasing at pocket site of impella 5.5, physician decision to stop all heparin until signs of bleeding resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. no product was returned. based off the clinical information available, the patient had hematoma at the impella 5.5 access site and other clinical information are unknown. the cause of the access site bleeding issue was not determined since the product was not returned and due to insufficient clinical information.
34817 | 1220648-2023-02944 | 17665970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17665970&pc=OZD | 08/03/2023 | 01/23/2024 |  | 08/31/2023 |  |  | 08/03/2023 | 08/31/2023 |  |  |  |  | 08/31/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1133151 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Unknown |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/03/2023 |  |  |  |  | The user facility reported their automated impella controller (aic) was not reading the purge disc. a loaner was sent to the user facility. no patient harm related to this event. | [5 empty] | The aic has been received for repair and investigation. the investigation is not complete at this time. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. product was returned for evaluation. the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken (missing at blue disc retainer). the failure mode will be monitored and trended.
34818 | 1220648-2023-02949 | 17666115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17666115&pc=OZD | 08/04/2023 | 01/24/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  | 05/31/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345496 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2023 |  |  |  |  | The user facility reported a 43-year-old male was implanted with an impella 5.5 as a bridge to lvad. the patient started having nerve pain around the axillary cutdown site. the site appeared slightly swollen but similar to previous assessments. the ct scan showed some kind of material that could be causing impingement. the patient underwent an incision and drainage procedure. it was found that the patient had a fibrin/clot like substance that was causing the impingement. it was removed and closed with no signs of any infection. the patient was given prophylactic antibiotics given the fact the patient will soon undergo durable lvad placement. after removal of fibrin/clot, the patient had significant relief of symptoms. the patient was transitioned to lvad. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the issue of thrombus has been completed. the device was not returned for investigation. based on the clinical investigation, the root cause of thrombus was most likely due to patient condition.
34819 | 1220648-2023-02950 | 17667574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17667574&pc=OZD | 08/05/2023 | 10/09/2023 |  | 08/31/2023 |  |  | 08/05/2023 | 08/31/2023 |  |  |  |  | 08/31/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old male patient was presented for impella placement. while in the operating room (or) escalating care, it was noted that the aic experienced purge disk errors with two consecutive purge cassettes. the staff switched the aic's and the purge cassette was recognized appropriately. the original aic was taken out of service after event. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the product was returned for investigation. the console was tested after arriving in field service. the issue with not properly detecting the purge disc was reproduced, and housing where the screw that fastens the purge retainer to the purge pressure sensor assembly was found to be broken. the root cause of the aic - purge disc/flag was related to a broken purge retainer.
34820 | 1220648-2023-02951 | 17667584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17667584&pc=OZD | 08/05/2023 | 02/28/2024 |  | 08/31/2023 |  |  | 08/05/2023 | 08/31/2023 |  |  | 05/31/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354088 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Frayed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 0.018 wire shredded while forming j tip prior to insertion for impella implant. a replacement guidewire was subsequently utilized and the pump was placed for patient support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue -use has been completed since the original report was filed. per the clinical information provided, the root cause of the delivery issue - guidewire damage was most likely use related.
34821 | 1220648-2023-02952 | 17667586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17667586&pc=OZD | 08/04/2023 | 08/30/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  | 05/31/2025 |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024342562 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2023 |  |  |  |  | The user facility reported a 64-year-old male with congestive heart failure was implanted with an impella 5.5 as a bridge to transplant. during support, a leak was noticed at the purge luer of the impella 5.5 pump. the purge cassette was changed with no resolution of the leak. at this time, the leak is minimal and is not impacting the purge pressures or purge flows. the attending physician is aware and has expressed not to change the purge cassettes as the patient is getting offers for orthotopic heart transplant and he does not want to exacerbate the problem. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported purge leak-pump has been completed. the pump was returned for evaluation. upon visual inspection, no abnormalities were observed. the pump was run through de-air and no leaks were reproduced. the retainer cover was removed, a syringe with colored fluid was connected to the yellow luer, high pressure was applied and no leaks were reproduced. the data logs showed no abnormalities. the root cause of the purge leak-pump was a damaged sidearm yellow luer but the cause of the damage was not determined as the issue could not be reproduced.
34822 | 1220648-2023-02953 | 17667589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17667589&pc=OZD | 08/04/2023 | 01/23/2024 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  |  |  | 08/31/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1211525 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Component Missing | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/04/2023 |  |  |  |  | The abiomed service representative reported the blue purge retainer on the automated impella controller (aic) was found to be missing while preventive maintenance was being performed. the aic will be returned for repair. no patient involvement. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the automated impella controller missing component has been completed. the device was returned for evaluation. data analysis: no evidence of the issue was found within the logs. device analysis: the console was tested after arriving in field service. the purge retainer was confirmed to be broken (missing at blue disc retainer). field service was instructed on what repairs were required. the cause of the issue was broken purge disc retainer.
34823 | 1220648-2023-02954 | 17667597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17667597&pc=OZD | 08/04/2023 | 10/24/2025 |  | 08/31/2023 |  |  | 08/04/2023 | 08/31/2023 |  |  |  |  | 08/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1486033 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a purge disc was broken on their aic. a loaner was sent to the site to initiate the return for investigation. there was no reported patient harm resulting from the issue. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-02954 in accordance with updated procedures. e2 and e3 were erroneously entered on the initial manufacturer device report number 1220648-2023-02954. g1 reporting contact email revised on manufacturer device report 1220648-2023-02954 in accordance with updated procedures. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. data analysis: the logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. re-insertion of purge cassettes/purge disc was observed, yet the console still had problem in detecting purge disc. device analysis: the console was tested and the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. there are some residue in purge insert and the purge motor shaft is not immobilized. the purge disc retainer was observed to be not broken, but a small crack is present. it must be replaced, even though this issue is not related to the reported complaint. the cause of the issue was a defective component.
34824 | 1220648-2023-02898 | 17649403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17649403&pc=OZD | 11/25/2021 | 01/24/2024 |  | 08/30/2023 |  |  | 11/25/2021 | 08/30/2023 |  |  | 07/31/2023 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130624 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2021 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient inadvertently pulled on the pump and it came across the valve when walking in the hospital. the patient was stable, however, decision was made to replace the pump due to the malposition. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was most likely patient movement. as per the clinical description provided.
34825 | 1220648-2023-02899 | 17649559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17649559&pc=OZD | 07/16/2021 | 08/30/2023 |  | 08/30/2023 |  |  | 07/16/2021 | 08/30/2023 |  |  | 11/30/2022 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022020862 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/16/2021 |  |  |  |  | The user facility reported that the impella 5.5 purge sidearm was covered in blood following surgery, after the purge retainer had been cleaned with betadine containing isopropyl alcohol. a leak was also reported to be coming from the purge filter. purge sidearm bypass was performed. it was also reported that the optical placement signal was reading 100 mmhg higher than the patient's arterial line. there was no harm to the patient. no further information was provided. | [5 empty] | The investigation into the reported purge leak and optical signal issue was completed. the device was returned for evaluation. the purge side arm was noted to be detached from the red pump handle. the retainer was partially over the sidearm. a crack in the filter was observed and the leak was able to be reproduced. the root cause of the purge leak was sidearm filter damage as a result of cleaning the filter with betadine and isopropyl alcohol. data log analysis revealed no issues with the placement signal (ps) until 5 days into the case, when suddenly the ps began to fluctuate. device analysis did not reveal any physical sensor malfunctions. the root cause of the optical sensor inaccuracies could not be determined. impella 5.5 ifu states: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34826 | 1220648-2023-02900 | 17649593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17649593&pc=OZD | 08/20/2021 | 08/30/2023 |  | 08/30/2023 |  |  | 08/20/2021 | 08/30/2023 |  |  | 04/30/2023 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022074272 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | False Alarm | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/20/2021 |  |  |  |  | The user facility reported an unknown leak in the catheter near the side arm of a 68-year-old, black male patient implanted with an impella 5.5 for management of cardiomyopathy. the user facility additionally reported "false in aorta alarms". pump position was confirmed via ultrasound and placement signal monitoring was ultimately disabled. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak and "false in aorta alarms" was completed. the device was returned for evaluation. a leak reproduction test was performed which confirmed a purge fluid leak at the yellow luer connection on the purge sidearm. visual inspection of the yellow luer revealed a crack. the root cause of the reported purge sidearm leak was determined to be a damaged yellow luer. of note, sodium bicarbonate was used as a purge solution. based on the clinical information provided that the pump was properly positioned, the root cause of the false in aorta alarms was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34827 | 1220648-2023-02901 | 17650683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17650683&pc=OZD | 10/09/2021 | 08/30/2023 |  | 08/30/2023 |  |  | 10/09/2021 | 08/30/2023 |  |  | 04/30/2023 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071812 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/09/2021 |  |  |  |  | The user facility reported that the yellow purge connection cracked on the impella 5.5 being used for a cardiomyopathy patient. staff reported that the luer never became stuck, no tools were used, and the purge cassette was changed every four days. the cassette had been changed the day prior to the luer crack. the pump was not replaced due to the yellow luer crack. there were no reports of harm to the patient. the pump support continued til transplant accomplished. | [5 empty] | The investigation into the reported purge leak was completed. the pump, purge cassette sidearm, and device stick were returned for evaluation. the returned sidearm was pressurized with colored fluid and a leak was reproduced at the base of the yellow luer. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34828 | 1220648-2023-02902 | 17650799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17650799&pc=OZD | 10/12/2021 | 08/30/2023 |  | 08/30/2023 |  |  | 10/12/2021 | 08/30/2023 |  |  | 07/31/2023 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022117719 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2021 |  |  |  |  | The user facility reported that the patient¿s sidearm on the impella 5.5 cracked overnight while sleeping. purge bypass was completed without issue. staff stated that the 3-point fixation and case remained in place and intact and there were no reports of use of alcohol products. sodium bicarbonate was used in the purge. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported the cracked purge sidearm was completed. the device was not returned for evaluation. the cause of the sidearm break was likely due to the bicarb weakening the plastic of the yellow luer. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol."
34829 | 1220648-2023-02903 | 17651457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17651457&pc=OZD | 08/01/2023 | 06/24/2024 |  | 08/30/2023 |  |  | 08/01/2023 | 08/30/2023 |  |  | 06/30/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361160 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported an 87 year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site bleed. the bleed was treated by femostop and one unit of red blood cells. there was no major bleed that would have prompted further intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the product was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
34830 | 1220648-2023-02904 | 17651581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17651581&pc=OZD | 10/23/2021 | 08/30/2023 |  | 08/30/2023 |  |  | 10/23/2021 | 08/30/2023 |  |  | 05/31/2023 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022084077 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/23/2021 |  |  |  |  | The user facility reported that 17 days into support, after changing the purge cassette, the rn noticed that the yellow luer was leaking on the impella 5.5 pump. the clinical support call center team member was able to assist the nurse via phone to perform a purge system bypass without issue. it was also reported that sodium bicarbonate was used in the purge solution throughout support. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. after reviewing the clinical information, the cause of the purge leak was determined to be most likely related to sidearm yellow luer damage due to sodium bicarbonate in the purge solution. the root cause of the purge leak was most likely sidearm yellow luer damage due to sodium bicarbonate in the purge solution. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34831 | 1220648-2023-02905 | 17651858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17651858&pc=OZD | 08/01/2023 | 01/11/2024 |  | 08/30/2023 |  |  | 08/01/2023 | 08/30/2023 |  |  | 05/31/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345926 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility had a 55-year-old male with impella 5.5 support ongoing for a patient one year out from transplant, but due to underlying condition is in need of another transplant, and possible renal as well. the 5.5 was placed and has supported for over 2 weeks to date. the patient has an ooze at the access site and arm swelling/pain. the team has treated with one unit of blood, imaging to confirm the limb has flow, and a jackson pratt drain pulling from the access site/hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. product was not returned for review. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
34832 | 1220648-2023-02906 | 17652689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17652689&pc=OZD | 10/11/2021 | 08/30/2023 |  | 08/30/2023 |  |  | 10/11/2021 | 08/30/2023 |  |  | 07/31/2023 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022112656 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/11/2021 |  |  |  |  | The user facility reported a 45-year-old male patient implanted with the impella 5.5 for management of acute myocardial infarction (ami) and cardiogenic shock, received a false "impella in aorta" alarm. correct pump position was verified under echocardiogram (echo). hospital staff disabled the placement signal (ps) alarm and was advised to use the motor current. additionally, a leak at the yellow luer was reported on the 21st day of support. a purge sidearm bypass was performed and the sidearm was discarded. the rn stated that no tools had been used on the connection and no patient movement occurred that would have caused the leak. sodium bicarbonate was used in the purge solution. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported "false impella in aorta" alarms and low purge pressure/purge system leak was completed. the device was returned for evaluation. when the "false impella in aorta" alarms were triggered, the pump's position was confirmed via echo, so the ps alarm was disabled, and the case continued. the root cause of the optical placement signal alarm issue was the patient condition. the cause of the low purge pressure/purge system leak was determined to be yellow luer damage. the root cause of the crack was most likely a result of the use of sodium bicarbonate in the purge fluid. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34833 | 1220648-2023-02907 | 17652777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17652777&pc=OZD | 08/03/2023 | 07/31/2024 |  | 08/30/2023 |  |  | 08/03/2023 | 08/30/2023 |  |  | 03/31/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343118 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the orange hub on the 14 x 13cm peel away sheath was oozing intraoperatively. staff used a 2nd person to pinch the orange hub and prevent further bleeding. the product malfunction occurred despite all best practices being followed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Product was not returned for investigation. per the clinical information provided, the root cause of the was most likely patient condition related to patients calcified vessels. d6a and d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-02907: g1 reporting contact fax number missing. h.6 code 1888, 4627, and 2993 were reported incorrectly. new code was added to health effect - clinical code, medical device problem code, and component code.
34834 | 1220648-2023-02908 | 17652788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17652788&pc=OZD | 07/28/2022 | 01/19/2024 |  | 08/30/2023 |  |  | 07/28/2022 | 08/30/2023 |  |  | 03/31/2024 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203076 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient had access site bleeding. blood products were given to the patient to make the patient more hemodynamically stable. decision was made to explant impella and place the patient on va ecmo in axillary artery graft. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the impella was not returned for evaluation. however, the data logs revealed impella position unknown alarms with recorded low flow during the case. according to clinical notes, the patient had a decompressed left ventricle (lv), and blood products were administered during treatment. additionally, bleeding was reported from the sutured graft site after observing the blue hub move into the graft. the cause of the access site bleeding issue was most likely use related based on provided clinical information. the cause of the low pump flow issue was most likely patient condition related to lv decompression, based on provided clinical.
34835 | 1220648-2023-02909 | 17653868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17653868&pc=OZD | 08/03/2023 | 09/26/2024 |  | 08/30/2023 |  |  | 08/03/2023 | 08/30/2023 |  |  | 06/30/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364772 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 75-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. patient expired of a cerebral hemorrhage during support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/cva has been completed since the original report was filed. the device was not returned by the medical facility. data logs was not provided. the cause of the stroke issue was most likely patient condition related. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections h1 and h6.
34836 | 1220648-2023-02910 | 17655932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17655932&pc=OZD | 08/01/2023 | 07/31/2024 |  | 08/30/2023 |  |  | 08/01/2023 | 08/30/2023 |  |  | 05/31/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352554 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. the cannula kinked during attempted pump insertion. a second pump was subsequently prepped and placed due to the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for delivery issues-use has been completed. product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to use issue. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). impella cp with smartassist system section: warnings and cautions: warnings to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels. physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. d6a and d6b added. h6 medical device problem code added 2682 the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-02910: f6/f8-should have been left blank.
34837 | 1220648-2023-02911 | 17655960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17655960&pc=OZD | 08/01/2023 | 01/23/2024 |  | 08/30/2023 |  |  | 08/01/2023 | 08/30/2023 |  |  |  |  | 08/30/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1218792 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported an 83-year-old female undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. prior to insertion of the impella cp, the automated impella controller (aic) had an error message and a sensor failure. the staff was going to switch to a backup aic, but could not find one. the impella cp insertion was aborted and an intra-aortic balloon pump was placed. when the field service representative arrived onsite, the aic was located in the adjacent room and had been overlooked. there was no patient harm. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into optical sensor issues on the automated impella controller (aic) has been completed. data analysis: logs were searched for optical sensor issues. on day of complaint, ¿optical sensor not supported - alarm issued¿ was found. before the complaint date field service set to have a hardware version revision 1 instead of the expected revision 257. device analysis: console was inspected and tested on an engineering lab bench. system acted normal except for ¿optical sensor not supported¿ alarm present when optical pump was inserted. this confirmed the complaint. the cause of the optical sensor not support alarm was due to the wrong hardware revision set in software.
34838 | 1220648-2023-02912 | 17656060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656060&pc=OZD | 08/01/2023 | 01/23/2024 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  |  |  | 08/30/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1429644 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The abiomed field service representative reported that at the user facility during setup of an impella pump, the automated impella controller (aic) did not recognize the purge cassette. replacing the purge cassette did not resolve the issue. the aic was exchanged and the issue was resolved. no patient involvement at the time of this event. | [5 empty] | The aic has been returned for evaluation, however the investigation is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. product was returned for evaluation. the console was tested after arriving in fs. purge shaft could not be turned even with the motor dismounted and failed torque test step 11 of pm procedure (b)(6). dried glucose was not the root cause of this failure. the failure mode will be monitored and trended.
34839 | 1220648-2023-02913 | 17656118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656118&pc=OZD | 08/03/2023 | 01/15/2024 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  | 05/31/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344637 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male patient had pfo closure, avr and impella 5.5 placement. device was inserted, and while rewarming, the left ventricle (lv) perforation was noted in the posterior wall. staff patched the perforation and the patient was weaned off the pump without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the ventricle perforation has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the ventricle perforation issue was not determined since product was not returned and insufficient clinical information was provided.
34840 | 1220648-2023-02914 | 17656208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656208&pc=OZD | 08/02/2023 | 10/27/2025 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  |  |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1493604 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 08/02/2023 |  |  |  |  | The user facility reported a male (age unknown) undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller gave a yellow error alarm. all connections were secure, and the patient was sedated and not moving. the error message went away but the impella was switched to the backup controller. all wave forms were within normal limits. no patient consequences were reported as result of this event. | [5 empty] | The device has been received for evaluation, however the investigation is incomplete. upon investigation closure, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-02914 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-02914. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-02914. g1 reporting contact email revised on manufacturer device report 1220648-2023-02914 in accordance with updated procedures. g2 report source; foreign was missing from manufacturer device report 1220648-2023-02914.
34841 | 1220648-2023-02915 | 17656305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656305&pc=OZD | 08/03/2023 | 02/14/2024 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  | 03/31/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347882 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported 69-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp dislodged out of the ventricle on day 4 while logrolling the patient. decision was made to place a new impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the positioning issue (cp dislodged out of ventricle) was patient movement related due to improper technique while logrolling the patient.
34842 | 1220648-2023-02916 | 17656653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656653&pc=OZD | 08/01/2023 | 01/15/2024 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  | 05/31/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357192 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported an 89-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. during the implant of impella cp, he cannula was unable to advance through the right subclavian artery into the innominate. after trying several times, a decision was made to remove and exchange the guidewire. upon the removal, a kink was noticed in the coil of the cannula. the patient was escalated to right axillary placement of an impella 5.5. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the clinical details states that the patient had calcified and tortuosity vessels condition. the cause of the delivery issue was patient condition related due to calcified and tortuosity vessels condition.
34843 | 1220648-2023-02917 | 17656710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656710&pc=PYX | 08/01/2023 | 02/05/2024 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  | 02/28/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024333413 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. two days after implant, the physician decided to remove the device due to low pump flows, patient hemolysis and a blood clot. the patient was put on va ecmo immediately after explant. | [5 empty] | The impella device was received and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the low pump flow, the thrombus and the hemolysis issues has been completed. the device was returned for benchtop investigation. returned complaint rp flex pump was visually inspected. biomaterial observed around impella. no other abnormality observed. the field lt log was reviewed and a drop in pump flow with a concurrent spike in motor current and increase in motor current thereafter despite no change in p-level confirms the low pump flow. no positioning issue or trending placement signal to indicate the hemolysis was patient condition or use related. the root cause of the thrombus and the hemolysis issues was biomaterial.
34844 | 1220648-2023-02918 | 17656737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656737&pc=OZD | 08/01/2023 | 03/15/2024 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  | 02/28/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337862 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella was placed without any issue. the patient became bradycardic after placement of impella and a transvenous pacemaker (tvp) was placed. the case was unremarkable after placement of tvp. the left main (lm) was shockwave with 1 stent placed with multiple post-inflations. the patient was weaned from the impella without any issue. attempted was made to wean pacer after impella removal but the patient was still bradycardic. the impella site closed with manta and the patient when to icu because of tvp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported bradycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the bradycardia could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34845 | 1220648-2023-02919 | 17656760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17656760&pc=OZD | 08/01/2023 | 01/11/2024 |  | 08/30/2023 |  |  |  | 08/30/2023 |  |  | 06/30/2025 |  | 08/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356203 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Swelling/ Edema | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing escalation of care was implanted with an impella 5.5 device for mechanical circulatory support. the team went back to the operation room to check for bleeding since the jackson pratt drain yielded around 3l, but nothing was found. the impella site was clean and dry, no bleeding or hematoma noted. the clinician noted right arm is slightly swollen and will monitor. since the implant the day before, the patient has received 19 packed red blood cells (prbc), 10 fresh frozen plasma (ffp), 7 platelet(plt), 20 cryo, 750 albumins, 1.5 of lr. the plan is to diurese and rest on impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major has been completed. the product was not returned. as per the clinical information provided, the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
34846 | 1220648-2023-02882 | 17641468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17641468&pc=OZD | 08/15/2021 | 08/29/2023 |  | 08/29/2023 |  |  | 08/15/2021 | 08/29/2023 |  |  | 05/31/2023 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088796 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 63 year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was leaking inside the sidearm retainer. there was no patient harm as a result of the issue. | [5 empty] | Only the purge sidearm was returned. a crack was observed on the yellow luer. leak reproduction testing yielded leaking from only the yellow luer. no other damage or abnormalities were found. the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate in the purge solution. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
34847 | 1220648-2023-02883 | 17643747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17643747&pc=OZD | 08/10/2021 | 08/29/2023 |  | 08/29/2023 |  |  | 08/10/2021 | 08/29/2023 |  |  | 05/31/2023 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088337 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2021 |  |  |  |  | The user facility reported that during an impella 5.5 support there was a crack in the yellow purge luer and that purge pressure (pp) was low. staff troubleshoot with purge bypass. bicarb was utilized as the purge solution. the purge leak occurred after just under 1 month of support. the pump support remained on for 6 days after the purge leak was observed. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge system leak was completed. the device was returned for evaluation. based on the clinical report it can be confirmed that the cause of the purge system leak was due to damage of the purge luer with the use of sodium bicarbonate. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol."
34848 | 1220648-2023-02884 | 17643909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17643909&pc=OZD | 09/12/2021 | 08/29/2023 |  | 08/29/2023 |  |  | 09/12/2021 | 08/29/2023 |  |  | 05/31/2023 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022091940 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2021 |  |  |  |  | The user facility reported that after changing the purge cassette per hospital protocol. the automated impella controller (aic) displayed a "purge system open" alarm¿ and leaking around the yellow-to-yellow connection was noted. the purge cassette was changed again without resolution. the nurse denied use of alcohol or hemostats and reported that the impella 5.5's sidearm protector was in place and 3-point fixation was used. the purge cassette was changed a third time, and a very slow leak was still visible. the physician hoped to wean the patient from support to avoid emergent sidearm bypass or pump exchange. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported "purge system open" alarm¿ and purge sidearm leak was completed. the device was not returned for evaluation. the cause of the yellow luer damage was due to bicarbonate weakening the plastic of the yellow luer. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34849 | 1220648-2023-02885 | 17643953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17643953&pc=OZD | 09/12/2021 | 08/29/2023 |  | 08/29/2023 |  |  | 09/12/2021 | 08/29/2023 |  |  | 04/30/2023 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022072746 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2021 |  |  |  |  | The user facility reported that the purge sidearm of the impella 5.5 leaked 24 days into support. sodium bicarbonate was used in the purge solution throughout the duration of support. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported leaking purge sidearm was completed. the device was not returned for evaluation. it was determined that the cause of the purge leak was most likely due to sidearm yellow luer damage since sodium bicarbonate was used in the purge solution. the root cause of the purge leak was most likely sidearm yellow luer damage since sodium bicarbonate was used in the purge solution. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34850 | 1220648-2023-02886 | 17644092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17644092&pc=OZD | 09/11/2021 | 08/29/2023 |  | 08/29/2023 |  |  | 09/11/2021 | 08/29/2023 |  |  | 07/31/2023 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022110421 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/11/2021 |  |  |  |  | The user facility reported that a 47-year-old male, unknown race, implanted with an impella 5.5 for management of cardiogenic shock post acute myocardial infarction (ami), noted a cracked purge sidearm and "open purge" alarm displayed on the device. the alarm was mitigated, and normal function was observed. the physician subsequently replaced the pump on day 5 of the support. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported cracked purge side arm and "open purge" alarm was completed. the device was returned for evaluation. the returned purge sidearm was visually inspected and a crack on the male portion of the yellow luer was observed. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34851 | 1220648-2023-02887 | 17645311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17645311&pc=OZD | 07/31/2023 | 07/26/2024 |  | 08/29/2023 |  |  | 07/31/2023 | 08/29/2023 |  |  | 04/30/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351460 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/31/2023 |  |  |  |  | The user facility reported a 47-year-old male in st-elevated myocardial infarction was implanted with an impella cp as an upgrade from intra-aortic balloon pump. while on support, the patient developed a left ventricle thrombus which was identified under echo. the patient was weaned and the cp was explanted the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the potential adverse events section, which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
34852 | 1220648-2023-02888 | 17645371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17645371&pc=OZD | 08/02/2023 | 08/29/2023 |  | 08/29/2023 |  |  | 08/02/2023 | 08/29/2023 |  |  | 02/28/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024311730 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 58-year-old male patient on impella support had experienced blood leaking out of the catheter plug. there were no serious adverse event to the patient as a result of this incident. the patient was weaned off from the impella on the next day of the issue discovery since the heart function recovered. | [5 empty] | (h3) the investigation into the reported "damage product (hole in catheter)" has been completed since the original report was filed. the medical center did not return any impella products. there was bleeding noticed near the red pump plug. the physician wondered if the blood had adhered to it and wiped it off. however, the blood leaked out again while fixing it in a different position. the impella was removed since the heart function recovered. the root cause of the hole in the catheter cannot be determined due to insufficient clinical details and no product returned.
34853 | 1220648-2023-02889 | 17645412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17645412&pc=OZD | 08/02/2023 | 01/15/2024 |  | 08/29/2023 |  |  | 08/02/2023 | 08/29/2023 |  |  | 01/31/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024315908 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 33-year-old female patient was on impella support, the impella pump stopped and an alarm sounded. there was no increase in purge pressure or decrease in purge flow rate. they tried to restart the pump, but it did not restart. the impella was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. the root cause of the pump stop could not be determined due to insufficient clinical information and no product or logs were returned.
34854 | 1220648-2023-02890 | 17646222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17646222&pc=OZD | 10/06/2021 | 08/29/2023 |  | 08/29/2023 |  |  | 10/06/2021 | 08/29/2023 |  |  | 04/30/2023 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022079869 |  | 0550-0008 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2021 |  |  |  |  | The user facility reported low purge pressure and that purge fluid was leaking at the spot where the purge line connected to the device twenty days into support with the impella 5.5. a crack was noted on the yellow portion of the purge. staff "tightened hard" and purge pressures returned to normal. a purge bypass was performed several days later. the clinical team also reported that hospital staff had switched the purge solution from heparin to sodium bicarbonate early in support. there was no harm to the patient. | [5 empty] | The investigation into the reported low purge pressure and purge fluid leak was completed. the device was not returned for evaluation. after review of the clinical information, it was determined that the cause of the purge leak was due to sidearm yellow luer damage. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34855 | 1220648-2023-02891 | 17646642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17646642&pc=OZD | 08/01/2023 | 01/11/2024 |  | 08/29/2023 |  |  | 08/01/2023 | 08/29/2023 |  |  | 04/30/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male patient developed hemolysis during impella support. patient experienced decreased renal function and was unable to wean off pressor support. due to the noted condition, the physician planned to escalate to an impella 5.5 for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, hemolysis was observed during impella support with impella in aorta alarms seen and the pump was repositioned with echocardiogram. data logs were reviewed. improper position of pump was observed with several aorta alarms and alarms resolved after repositioning. the cause of the hemolysis issue was most likely improper positioning of the device.
34856 | 1220648-2023-02892 | 17646765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17646765&pc=PYX | 08/02/2023 | 07/31/2024 |  | 08/29/2023 |  |  | 08/02/2023 | 08/29/2023 |  |  | 11/30/2024 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024289426 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon insertion of the introducer for impella rp delivery , the sheath was damaged as a result of the patient's thick skin. a new sheath was utilized from an unopened rp kit and the implant was successful. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ impella rp: ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for delivery issues-use has been completed. the impella device was not returned for evaluation. the cause of the introducer damage issue was most likely use related since second sheath was inserted without issue.
34857 | 1220648-2023-02893 | 17646793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17646793&pc=OZD | 07/31/2023 | 01/18/2024 |  | 08/29/2023 |  |  | 07/31/2023 | 08/29/2023 |  |  | 04/30/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348466 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/31/2023 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma around the impella access site during pump support. it was noted by the physician that the patient was thrashing around leading up to the development of the hematoma. as the patient showed improvement in their hemodynamic status, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. product was not returned for review. based on clinical description, the cause of access site bleeding was most likely due to patient movement.
34858 | 1220648-2023-02894 | 17646803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17646803&pc=OZD | 07/31/2023 | 06/25/2024 |  | 08/29/2023 |  |  | 07/31/2023 | 08/29/2023 |  |  | 04/30/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350504 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/31/2023 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. post procedure, reported there was significant bleeding noted around repo sheath post pannus release. when pannus was secured during the procedure, there was no bleeding. all issues occurred once patient's panis was released. while in icu, post procedure patient continued to have significant bleeding from access site which required patient to have blood transfusion and subsequent early explant of the impella per physician request to stop bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
34859 | 1220648-2023-02895 | 17646909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17646909&pc=OZD | 08/02/2023 | 02/07/2024 |  | 08/29/2023 |  |  | 08/02/2023 | 08/29/2023 |  |  | 06/30/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367305 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Flush | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that staff were unable to flush saline through the side-port of an inserted introducer during preparations for impella implant. it was noted by the physician that they were unable to compress their saline syringe when placed in the side port for the flush. the introducer was replaced with an onsite spare and the subsequent implant was successful. there was no patient harm. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported introducer damage has been completed since the original report was filed. product was returned and investigated. the three way introducer valve was found to be damaged and did not pass reverse flow. upon further investigation, it was observed that there was no flow path in the valve to begin with. the product was manufactured incorrectly. per the clinical information and the investigation, the root cause of the introducer issue was a damaged valve related to manufacturing operator error. the product is a malfunction.
34860 | 1220648-2023-02896 | 17646925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17646925&pc=OZD | 08/02/2023 | 01/18/2024 |  | 08/29/2023 |  |  | 08/02/2023 | 08/29/2023 |  |  | 05/31/2025 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354075 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient still has reasonable bleeding from leg fasciotomies and chest tubes. patient's was ecmo 6l/min. the impella cp suction decreased from p9 to p5. patient's labs have hemolyzed by last ptt 46. the physician discussed that p9 with ecmo and suction contributing to hemolysis. th physician states due to extremely dilated heart, physician wants to unload at p9 as soon as they are able to do so again. giving 1 unit of prbc shortly. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's hemolysis has been completed. product was not returned for investigation. the data log shows the several suction and improper position alarm during time of hemolysis. the cause of the hemolysis issue was most likely improper positioning of the device.
34861 | 1220648-2023-02897 | 17647661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17647661&pc=OZD | 11/08/2021 | 03/01/2024 |  | 08/29/2023 |  |  | 11/08/2021 | 08/29/2023 |  |  | 07/31/2023 |  | 08/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118235 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2021 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support was taken to perform a computed tomography(ct) scan and it showed that the patient had experienced a frontal lobe infarct. cause and severity were undetermined at the time. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported stroke/cva has been completed since the original report was filed. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
34862 | 1220648-2023-02850 | 17630628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17630628&pc=OZD | 07/29/2023 | 09/19/2024 |  | 08/28/2023 |  |  | 07/29/2023 | 08/28/2023 |  |  | 02/28/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324713 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Survival outcome at explant noted the patient coded and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a impella cp support ongoing for a 52-year-old female patient post stemi, vt code, and iabp placement for air lift to tertiary center. the tertiary center placed the cp and there was signs of hemolysis that went untreated by any intervention or confirmation by pfhg draw. additionally there was bilateral groin site bleeding that was not treated by any form of intervention. the pump remained on for a support of just over 30 hours. |  |  |  |  | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis and the access site bleeding issues was not determined since product was not returned and data log investigation was inconclusive, and relation to impella is unknown. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34863 | 17630637 | 17630637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17630637&pc=OZD | 07/29/2023 | 08/10/2023 | 08/10/2023 | 08/28/2023 | 08/10/2023 | 08/28/2023 |  | 08/28/2023 |  |  |  |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 2024356588 |  | 0048-0003 |  | No |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 16790 DA |  | Male | 97 KG |  |  | Malfunction | No | Yes | [7 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella cp with smartassist heart pump failed to purge and detect on console. trouble shooting was unable to improve equipment malfunction. representative was consulted as well as the 1-800 number. we were advised to open a new box and start over.
34864 | 1220648-2023-02851 | 17631725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17631725&pc=OZD | 07/13/2021 | 08/28/2023 |  | 08/28/2023 |  |  | 07/13/2021 | 08/28/2023 |  |  | 04/30/2023 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022080692 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/13/2021 |  |  |  |  | An impella 5.5 was placed for support of a patient suffering from cardiomyopathy. the user facility reported that the purge was leaking close to the purge sidearm after 18 days of support. purge sidearm bypass was discussed and instruction given. there were no reports of patient harm. the patient was bridged to transplant. | [5 empty] | The investigation into the reported high purge flow and purge leak was completed. the device was not returned for evaluation. based on the clinical account the root cause of the purge sidearm damage is most likely the use of bicarbonate weakening the plastic of the yellow luer. the root cause of the purge leak was sidearm yellow luer damage as a result of bicarbonate weakening the plastic. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34865 | 1220648-2023-02848 | 17634140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17634140&pc=PYX | 07/28/2023 | 10/09/2023 |  | 08/28/2023 |  |  | 07/28/2023 | 08/28/2023 |  |  | 12/31/2024 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024320047 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2023 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. the impella flows dropped from 4.0 to 3.5 l/m at p9 and purge fluid dropped to 6.0. the medical staff suspected of an issue with the pump so they opted for removal and bridge to protecduo, but ended up just placing a second rp instead. no patient harm was reported. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, the root cause of the low pump flow issue was not determined since neither product nor logs were returned and sufficient clinical information was not provided. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34866 | 1220648-2023-02852 | 17635518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17635518&pc=OZD | 07/22/2023 | 01/23/2024 |  | 08/28/2023 |  |  | 07/22/2023 | 08/28/2023 |  |  | 02/28/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322144 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2023 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the rn reported that flow dropped and had suction alarms. suspect the low flow issues to be due to clot ingestion. it was decided the patient met criteria for wean to explanation. | [5 empty] | The investigation into the thrombus has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the low pump flow issue was most likely biomaterial ingestion base on recorded data log events of motor current spike followed low pump flow. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34867 | 1220648-2023-02853 | 17638261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638261&pc=OZD | 08/02/2023 | 01/15/2024 |  | 08/28/2023 |  |  | 08/02/2023 | 08/28/2023 |  |  | 11/30/2024 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient was bleeding from the groin access site. to treat the patient the medical team made choice to administer 4 units of blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | No product was returned. as per clinical, the patient had bleeding at the impella groin site and four units of packed red blood cells were given to resolve the bleeding. the cause of the access site bleeding issue was not determined since the product was not returned and due to insufficient clinical information.
34868 | 1220648-2023-02854 | 17638270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638270&pc=OZD | 08/02/2023 | 01/24/2024 |  | 08/28/2023 |  |  | 08/02/2023 | 08/28/2023 |  |  | 12/31/2024 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024300511 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 30-year-old male patient that had a impella 5.5 support for over eight days when the pump was explanted due to concerns of infection. the pocket had purulent drainage and the patient had cultures sent for analysis and antibiotics ordered. | [5 empty] | The investigation into the issue of infection/sepsis has been completed. the device was not returned for investigation. based on the clinical investigation, the cause of the infection was most likely patient condition related since the device is sterilized under validated conditions and there was no evidence to indicate a break in the sterile barrier or pump contamination. unknown relation to impella. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34869 | 1220648-2023-02855 | 17638306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638306&pc=OZD | 08/02/2023 | 07/30/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  | 10/31/2024 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285793 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 84-year-old male patient was placing a impella cp across the valve and into the left ventricle (lv) for support of a surgical turndown patient. the placement was causing the patient to experience bradycardia. as such, the team intervened and placed a temporary pacer and medicated with atropine. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported arrhythmia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of arrhythmia could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
34870 | 1220648-2023-02856 | 17638582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638582&pc=OZD | 07/29/2023 | 01/22/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  | 02/28/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311643 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 experienced purge system blocked/high purge pressure alarms with use of sodium bicarb as purge solution during support. tpa given to resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the reported high purge pressure has been completed since the original report was filed. no product was returned for evaluation. as per clinical, several high purge pressure alarms were observed during support with low purge flow. data logs were reviewed and observed that there is stepwise rise in purge pressure due to kinked purge line at the sidearm with gradual reduction in purge pressure during unkinking and normal purge flow returned. during kinking the purge flow ticks were stalling based on the log review. the cause of the high purge pressure issue was most likely a kinked purge line in the sidearm tubing based on the data log review.
34871 | 1220648-2023-02857 | 17638587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638587&pc=OZD | 07/29/2023 | 07/25/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Cardiogenic Shock | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/29/2023 |  |  |  |  | The user facility reported a 75-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. once peel away sheath was removed and repo sheath inserted significant bleeding noted from insertion site. the bleeding would not subside even with perfect angle matching, forward tension mattress sutures and manual pressure. md shot angio and verified no retroperitoneal bleed but there was a small hematoma. the cp was explanted the following day. | It was further reported that the care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | Additional informations received via a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening of cardiogenic shock with a "possible" relationship to the impella device used. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the cardiogenic shock and access site bleed has been completed. the impella was not returned for evaluation. however, clinical details were sufficient to determine the root cause of the access site bleed. the root cause of access site bleeding was determined to be use issue because the issue by application of forward tension sutures and manual pressure. the root cause of the cardiogenic shock could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- b5 mso review, h6 component code 925. f6/f8 should have been left blank in the initial report.
34872 | 1220648-2023-02858 | 17638590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638590&pc=OZD | 07/29/2023 | 02/07/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  | 12/31/2023 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178355 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2023 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. white the patient was on support, the device reported air in purge system on the console. the team attempted to de-air purge, but fluid would not move through the tubing. the purge cassette and tubing were replaced. purge flow returned to normal ranges. within an hour of patient arrival, the motor current and left ventricle signal started increasing outside normal limits. later, there were failure controller alarms. pump was in good position across valve on echo. there was eventually a pump stop. decision was to remove pump. pump removed and there was a large clot visible covering outlet cage. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed since the original report was filed. the device was returned for investigation. the impella cp pump was visually inspected and soaked the pump for suspected clot ingestion. electrical resistance test was done, and all three motor phase lines resistance values were within normal range. after 20-25 mins of soaking, significant biomaterial was observed to be ingested at the outflow cage and wrapped around impeller. the cause of the pump stop issue was biomaterial wrapped around impeller and ingested over the outflow cage per field investigation and log, clinical review. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ h.10 instructions for use were inadvertently left off the original manufacturer device report. they can be found in h.10 in this report.
34873 | 1220648-2023-02859 | 17638591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638591&pc=OZD | 07/29/2023 | 02/02/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  | 11/30/2024 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311725 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2023 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. upon weaning and explant of the impella, staff noticed that ci/co, co2, and pressures started decreasing quickly. hgb had dropped. staff observed the iliac and noticed that the iliac had dissected. iliac was stented. hemostasis was achieved by perclose. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vascular damage - peripheral vessel issue has been completed. the product was not returned. as per clinical review, upon impella explant, the iliac was dissected and then stented. hemostasis was achieved using perclose. the cause of the vascular damage issue most likely patient condition and unknown relation to impella per clinical information.
34874 | 1220648-2023-02860 | 17638595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638595&pc=OZD | 07/29/2023 | 01/19/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350799 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during patient support, the pump stopped and after several attempts, the pump was unable to restart. the patient was taken to cardiac catheterization lab for pump replacement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed. product nor data logs were returned for review. the root cause of the pump stop was most likely a bearing wear as this pump stop occurred after 12 days of run time, which is well beyond the design life of 8 days.
34875 | 1220648-2023-02862 | 17638632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638632&pc=OZD | 07/30/2023 | 01/15/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  | 05/31/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354610 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2023 |  |  |  |  | The user facility reported a 45-year-old male in myocardial infarction was implanted with an impella cp device for mechanical circulatory support. once the impella cp was in place, the patient began to flail limbs on the cardiac catheterization lab (ccl) table and hematoma was noted to the impella site. manual pressure applied. after the procedure the patient continued to decompensate, and manual pressure still being applied to the hematoma. decision was made to explant the impella and exchange for va ecmo. the impella cp was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the cause of the bleeding issue was patient movement as reported that patient began to flail limbs on the cardiac catheterization laboratory (ccl) table and hematoma noted to impella site. manual pressure was applied for a long time and ended explanting the impella.
34876 | 1220648-2023-02863 | 17638649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638649&pc=OZD | 08/01/2023 | 01/23/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  |  |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1127807 |  | 042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the purge holder on the aic broke while prepping the device. | [5 empty] | The impella device was received from the customer and an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. product was returned for evaluation. the cause of the aic issues was broken/ missing purge flag.
34877 | 1220648-2023-02864 | 17638657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638657&pc=OZD | 08/01/2023 | 01/19/2024 |  | 08/28/2023 |  |  |  | 08/28/2023 |  |  |  |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1232682 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender received hemodynamic support from an impella cp smartassist cardiac pump due to acute myocardial infarction / cardiogenic shock. when the purge cassette was changed, the cassette was not detected by the automated impella controller (aic). a second purge cassette was also not recognized. the problem was resolved by switching to the backup controller. the attending physician assessed the situation as critical, but the patient survived the incident. | [5 empty] | The impella device was received and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge issue has been completed since the original report was filed. the product was returned for investigation. the console was tested on an engineering lab bench. the issue was reproduced with the original parts and configuration; upon connecting the pump, unable to start priming, no piston blockage was found; also, there were no traces of glucose on the purge unit. it was determined that the root cause of the issue was a defective component.
34878 | 1220648-2023-02865 | 17638963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638963&pc=OZD | 08/01/2023 | 03/11/2024 |  | 08/28/2023 |  |  | 08/01/2023 | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353451 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the purge leaked at the yellow connector. the tubing was changed, and it was resolved. impella support continued following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. clinical details were insufficient to determine the root cause. therefore, the root cause of the purge leak was unable to be determined.
34879 | 1220648-2023-02866 | 17638967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638967&pc=OZD | 07/30/2023 | 01/18/2024 |  | 08/28/2023 |  |  | 07/30/2023 | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336984 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2023 |  |  |  |  | The user facility reported a 68-year-old male was going to have an escalation of therapy from va ecmo and impella cp to an impella 5.5 device. during support, the patient was hemodynamically unstable and bleeding at the impella site. hematoma noted at site as well. 250 of albumin given and patient is doing better. the team is managing the site. | [5 empty] | The investigation for the reported access site bleed has been completed. the impella device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. clinical details states that the patient was given albumin and their condition improved thereafter. the patient was also reported to have had calcification. the cause of the bleeding issue most likely was patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34880 | 1220648-2023-02867 | 17638971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638971&pc=OZD | 08/01/2023 | 01/24/2024 |  | 08/28/2023 |  |  | 08/01/2023 | 08/28/2023 |  |  | 02/28/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325512 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Rupture | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 80-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient experienced an artery rupture leading to bleeding and ischemia that required surgical repair. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's access site bleeding (with rupture of artery) and limb ischemia has been completed. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. the cause of the limb ischemia issue was most likely patient condition related with ischemia observed on impella limb.
34881 | 1220648-2023-02868 | 17638986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638986&pc=OZD | 07/30/2023 | 01/19/2024 |  | 08/28/2023 |  |  | 07/30/2023 | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2023 |  |  |  |  | The user facility reported a 55-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during support, the device had low flow and suction alarms. the physician repositioned the pump with echo at bedside. the flows were unchanged and suction alarms continued. it was decided to remove the pump as the patient was stable enough to do so rather than exchange for a new pump. the pump was explanted, and patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the cause of the low pump flow cannot be determined since the complaint pump not returned for investigation and lack of the clinical details. data logs reviewed no mc spikes, low flow occurred in p-9, but nothing indicated as reason for the issue.
34882 | 1220648-2023-02869 | 17638993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638993&pc=OZD | 07/30/2023 | 03/05/2024 |  | 08/28/2023 |  |  | 07/30/2023 | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2023 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the pump was inadvertently pulled back into the aorta following placement. repositioning was attempted; however, the pump became kinked at the cannula. difficulty was met when trying to straighten the pump, and a snare was subsequently needed to straighten and then remove the device. another pump was placed for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was due to wireless re-access which is outside product claims.
34883 | 1220648-2023-02870 | 17638995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17638995&pc=OZD | 08/01/2023 | 06/17/2024 |  | 08/28/2023 |  |  | 08/01/2023 | 08/28/2023 |  |  | 05/31/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358281 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  |  | Injury | Yes | Yes | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 59-year-old patient was treated for acute myocardial infarction/cardiogenic shock and received hemodynamic support from an impella cp smartassist heart pump. immediately after the pump was inserted into the left ventricle, a sudden increase in purge pressure and a decrease in purge flow were observed. the values were still within specification and the pump continued to function. the doctor treating her suspected that a thrombus had been sucked in. the heparin dosage was increased. after 42 hours, the impella cp was removed and life support was discontinued due to the patient's poor condition. | The patient expired. upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 codes 4644 and 2964 were reported incorrectly on the previously submitted report. new codes has been added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
34884 | 1220648-2023-02871 | 17639000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639000&pc=OZD | 07/29/2023 | 10/09/2023 |  | 08/28/2023 |  |  | 07/29/2023 | 08/28/2023 |  |  |  |  | 08/28/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/29/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the display on the automated impella controller (aic) had a blackout during patient support. the screen was restored immediately and there was no noted effect on the support. it was advised that the aic should be replaced; however, the physician stated that they would only replace the aic under certain situations. as there was no effect on the patient's hemodynamics, support continued with the aic and implanted pump. | [5 empty] | The aic has been received for investigation, however, the investigation is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - shutdown or restart issue has been completed since the original report was filed. the product was returned for investigation. the data logs confirm the failure mode described from the field, however the failure mode from the field was not reproduced during investigation. it was determined that the root cause of the aic issue was the 3rd party hardware/software that runs io-graphics.
34885 | 1220648-2023-02872 | 17639014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639014&pc=OZD | 07/29/2023 | 01/22/2024 |  | 08/28/2023 |  |  | 07/29/2023 | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349722 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2023 |  |  |  |  | The user facility reported a 59-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during support, the pump was alarming high motor current and then the pump stopped. several attempts made to restart pump but they were unsuccessful. decision was made to replace pump and the pump was successfully explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the of pump stop issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the pump was used for more than 19 days, and the pump stop high motor current alarm was observed after design life period of the impella cp. the cause of the pump stop issue was not determined since the product and data logs were not returned for review.
34886 | 1220648-2023-02873 | 17639024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639024&pc=OZD | 07/29/2023 | 07/23/2024 |  | 08/28/2023 |  |  | 07/29/2023 | 08/28/2023 |  |  | 05/31/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356588 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 46-year-old male patient requiring impella support had encountered issues with the impella plug to aic malfunctioning. the procedure was aborted and a new pump used. there was no adverse event reported as a result of this incident. | [5 empty] | The investigation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34887 | 1220648-2023-02874 | 17639029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639029&pc=OZD | 07/30/2023 | 08/28/2023 |  | 08/28/2023 |  |  | 07/30/2023 | 08/28/2023 |  |  | 03/31/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339517 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that impella cp experienced purge blocked alarms during support. the physician declined all other options of tpa or pump exchange and felt it would be better to take patient to the operating room (or) and remove with surgical closure. | [5 empty] | The investigation into the reported "high purge pressure" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; however, logs were returned and evaluated and numerous purge system alarms were found along with a few motor current spikes. this combined with the given clinical data suggests that the root cause of the high purge pressure issue was biomaterial related. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34888 | 1220648-2023-02875 | 17639043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639043&pc=OZD | 07/31/2023 | 01/19/2024 |  | 08/28/2023 |  |  | 07/31/2023 | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337956 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old female patient on impella cp support had experienced a subarachnoid hemorrhage. a definitive cause is unknown, but it was believed to have been caused by anticoagulant therapy. per medical staff, they could not disassociate the use of impella cp from the adverse event. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. no product or data logs were returned. based on the clinical information provided, subarachnoid hemorrhage occurred on the patient with the cause and the date of original occurrence being unknown. the cause of the stroke issue was most likely patient condition related and unknown relation with insufficient clinical/log review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34889 | 1220648-2023-02876 | 17639046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639046&pc=OZD | 07/31/2023 | 04/19/2024 |  | 08/28/2023 |  |  | 07/31/2023 | 08/28/2023 |  |  | 06/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357102 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Malfunction | Yes | Yes | Activation, Positioning or Separation Problem | Infusion or Flow Problem | [5 empty] | Swelling/ Edema | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male patient was taken to cvor for cabg on impella cp. plan was to leave impella in place and cross clamp in surgical mode. impella was deep on tee at start of case so the doctor repositioned under imaging. ekg was silent but after a few minutes, and they had concerns there was movement in the heart and an additional antegrade (impella back up to p2 during that) yielded a distended left ventricle (lv). doctor decided they needed to pull the cp into the descending (ripped the sterile sleeve while doing that) and they replaced cross clamp, but realized shortly after that the lv was still distended and that the doctors root vent tubing had gotten kinked on the field. doctor was trying to move the blood from lv into the root vent and they thought the problem was actually the vent not the impella. after the case, they attempted to recross the valve several times, they came completely off bypass flows and still could not cross the valve even with maneuvers. patient was completely normotensive off of cbp with an impella in the aorta unable to flow above p2. the rep. expressed concern for clot in cannula since the cp should be flowing way more than 0.4 lpm in the aorta without suction, so the doctor decided it needed to be removed. impella was removed using guidewire re-access port and perclose sutures tied down with 5-10 min of manual pressure and they dermabonded the site. | [5 empty] | The investigation into the product damage (sterile sleeve damage), the positioning issues and the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the cause of the sterile sleeve damage issue was use related due to excessive force during the repositioning the pump. the root cause of the positioning issues causing the low flow is most likely due to use related as well. | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34890 | 1220648-2023-02878 | 17639066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639066&pc=OZD | 07/31/2023 | 05/28/2025 |  | 08/28/2023 |  |  | 07/31/2023 | 08/28/2023 |  |  | 03/31/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient was suspected to have a gastrointestinal bleed and exhibited bleeding from the right groin impella closure sites. vascular surgery assessed the patient that morning and placed additional sutures. | The user facility reported a 66-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. patient ct angiogram revealed a complex dissection in the distal transverse arch extending into the descending thoracic aorta. there was a large 11 cm thrombus associated with the impella as well as associated with the dissection flap. the thrombus began in the distal transverse arch and extended into the descending aorta. the pump was subsequently removed and an impella 5.5 was placed for ongoing support. |  |  |  |  | The investigation for vascular damage and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component coded added 925. | Additional information was provided in sections b5 and h6. pump was not returned for investigation. a complex dissection was observed in the distal transverse arch extending into the descending thoracic aorta. additionally, there was an 11 cm thrombus associated with both the impella and the dissection flap. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. as the necessary information was not provided, the root cause of the thrombus and access site bleeding was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34891 | 1220648-2023-02880 | 17639082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17639082&pc=OZD | 07/31/2023 | 01/22/2024 |  | 08/28/2023 |  |  | 07/31/2023 | 08/28/2023 |  |  | 04/30/2025 |  | 08/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351143 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had a large gi bleed and heparin was removed. ten days later the impella suddenly stopped and did not restart. it was further reported that upon explanting, the impella had a large clot around the pigtail. | [5 empty] | The investigation for the reported pump stop and thrombus has been completed. the device was not returned for evaluation. however, data logs and clinical information was reviewed and evaluated. the data logs confirmed the failure mode. data showed the pump ran without issue for 325.5 hours then stopped. multiple attempts to restart the pump failed. the pump then was disconnected from the console. the root cause of the pump stop was unable to be determined as limited clinical details were provided and no product was returned for review. this pump stop occurred after 14 days of run time, which is well beyond the design life of 8 days. the root cause of the thrombus was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34892 | 1220648-2023-02830 | 17622472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17622472&pc=OZD | 07/27/2023 | 07/23/2024 |  | 08/25/2023 |  |  | 07/27/2023 | 08/25/2023 |  |  | 04/30/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351190 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was noted that there was a perforation after angiogram was performed to the right femoral artery (rfa). a stent was used to cover the perforation. | [5 empty] | The investigation for vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. according to the clinical, access to the right femoral artery (rfa) was made using best practices and high bifurcation was noted. after performing serial dilations, the impella cp was implanted with flows of 3.5l/min. an attempt was made to wire the right common artery (rca), but was aborted due to complexity. later the same day, the patient was weaned off of support and the impella cp was explanted. the patient remained hemodynamically stable at this time. however, a completion angiogram performed on the rfa revealed a perforation requiring a covered stent. no product was returned for a visual inspection. the most likely root cause of the vascular damage - peripheral vessel was patient condition related. instructions for use: section: warnings and cautions: physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. d6a/d6b added. h6 component code added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34893 | 1220648-2023-02831 | 17622525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17622525&pc=OZD | 07/28/2023 | 02/06/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 05/31/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352351 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 48-year-old female on impella support experienced a low pump flow. the pump was repositioned, however, that did not resolve the low pump flow issue and the physician suspected of a clot formation. the following day the patient was partially disoriented and the ct scan noted a subacute infarct. there was no intracranial hemorrhage noted. since the patient was scheduled a heart transplant the medical team determined that it would be best to remove the pump when performing the heart transplant. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the given clinical information, the root cause of the low pump flow issue was biomaterial. clinical information points towards biomaterial clot which was not confirmed until two days after. data logs show suction at the start of the low pump flow event corresponding to motor current spike further confirming clot presence. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34894 | 1220648-2023-02833 | 17624336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17624336&pc=OZD | 07/07/2021 | 08/25/2023 |  | 08/25/2023 |  |  | 07/07/2021 | 08/25/2023 |  |  | 04/30/2023 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022076922 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. staff reported a purge cassette leaking around the connection. this patient had sodium bicarb running in the purge for less than 24 hours. no reports of patient harm. | [5 empty] | The investigation into the reported purge cassette leak has been completed. the medical center did return impella products for benchtop analysis. the cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
34895 | 1220648-2023-02834 | 17625128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17625128&pc=OZD | 06/23/2023 | 09/20/2024 |  | 08/25/2023 |  |  | 07/26/2023 | 08/25/2023 |  |  | 03/31/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024344676 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/26/2023 |  |  |  |  | Additional information noted that the patient survived the explant procedure but subsequently expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a 72 year old patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during the first day of support, the patient endured sepsis. antibiotics (vancomycin) were administered as a result. the relationship was noted as "possible," therefore, could not be ruled out. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the infection/sepsis has been completed since the original report was filed. the cause of the sepsis issue was most likely patient condition related since the pump was sterilized under validated condition and determined to be unknown relation to the pump. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34896 | 1220648-2023-02836 | 17625917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17625917&pc=OZD | 07/28/2023 | 10/09/2023 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 04/30/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339400 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2023 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device experienced some suction problems and a thrombus like substance was observed. impella had to be removed under surgical thoracotomy. the patient was in a stable condition at the end of the pump explant. | [5 empty] | The investigation into low pump flow issue has been completed since the original report was filed. the medical center in japan did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the low pump flow issue was biomaterial found on the inlet cage based on provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
34897 | 1220648-2023-02838 | 17626113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17626113&pc=OZD | 07/27/2023 | 09/22/2024 |  | 08/25/2023 |  |  | 07/27/2023 | 08/25/2023 |  |  | 03/31/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345246 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/27/2023 |  |  |  |  | The patient became hemodynamically unstable and went into cardiac arrest. the procedure outcome was that patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | The user facility reported a 43-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient had a family history of heart disease and was not a surgical candidate. the patient was admitted for a high-risk percutaneous coronary intervention (hrpci) without impella support. upon ballooning left anterior descending (lad) patient became hemodynamically unstable and grossly hypertensive. the circulator gave three times ordered dose. the patient subsequently went into cardiac arrest. the impella was placed fifteen minutes into active cpr. fluoro images taken to verify et tube placement. endotracheal tube (ett) was found to be in the esophagus. the patient was reintubated. the impella appeared to be kinked distal to outlet based on fluoroscopy images. it was communicated to the team that repositioning would be required to attain adequate flows. cpr continued for a total of 1.5 hours, but the patient had target organ damage (tod) and the patient expired. |  |  |  |  | The investigation into the low pump flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the cause of the low pump flow was a cannula kink due to insertion during cardiopulmonary resuscitation (cpr). | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02838 was provided in sections b1, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34898 | 1220648-2023-02839 | 17626124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17626124&pc=OZD | 07/27/2023 | 01/22/2024 |  | 08/25/2023 |  |  | 07/27/2023 | 08/25/2023 |  |  | 04/30/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337427 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The user facility reported a 62-year-old male (race unknown) was to be implanted with an impella 5.5 for circulatory support. while attempting to place the impella 5.5, the procedure was aborted. a perforation of the axillary artery occurred. | [5 empty] | The investigation into vascular damage has been completed. the pump was not returned for investigation. the first impella 5.5 was removed due to a positioning issue. the second impella was stuck in the body and there were bleeding noted due perforation of the axillary artery. the cause of the vascular damage issue was most likely use related based on provided clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
34899 | 1220648-2023-02840 | 17626143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17626143&pc=OZD | 07/28/2023 | 01/18/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  |  |  | 08/25/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1332033 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/28/2023 |  |  |  |  | The abiomed field service represented reported that the automated impella controller (aic) at the user facility did not recognize the purge cassette. the aic was successfully swapped out. there was no patient harm related to this event. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the product was returned for investigation. console was tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge disc flag on the aic was broken which would issue the "purge disc not detected" alarm when attempting to set up the purge assembly. the root cause of the aic - purge disc/flag issues were due to a broken purge pressure sensor flag. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34900 | 1220648-2023-02841 | 17626148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17626148&pc=OZD | 07/28/2023 | 01/18/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  |  |  | 08/25/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1097390 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/28/2023 |  |  |  |  | The abiomed field service representative reported when the user facility power on their automated impella controller (aic), a battery failure alarm was noted. the console had been plugged in but was showing a battery level of 50%. troubleshooting was initiated making sure that the bottom of the console was turned on and it was. aic was removed from use. there was no patient harm associated with this event. | [5 empty] | The investigation for the reported console battery failure alarm has been completed. the console and the data logs were returned for evaluation. the console was tested in the lab and found that the blinking pattern of the lcd of battery two is indicative of cell imbalance in the battery. additionally, there was no fringe pattern on the optical bench. there was voltage and current waveform across the lamp was roughly linear and not the expected pulse width modulated waveform from the feedback loop. data logs were reviewed which confirmed the reported alarms. the root cause of the battery failure was due to a defective battery. the root cause of the defective battery was single cell failure, a known pattern failure. additionally the defective optical bench was due to light leakage along fiber. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
34901 | 1220648-2023-02842 | 17626165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17626165&pc=OZD | 07/28/2023 | 01/17/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 03/31/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343812 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male patient on impella cp support that had been upgraded to an impella 5.5. after the 5.5 was placed, the previous pump's femoral access site was closed using perclose without any noted issues. the following day, a large hematoma was observed with leaking from impella insertion site. the medical team applied a covered stent, resolving the issue. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34902 | 1220648-2023-02844 | 17626378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17626378&pc=OZD | 07/28/2023 | 03/12/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 11/30/2024 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303999 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. when the impella was weaned, explanted and attempting to post-close the site with the pre-close sutures, the physician had difficulty with hemostasis and elected to add an 8fr angio-seal. physician believes that the pre-closed sutures snapped while attempting to deliver the 8fr angio-seal. the decision was made to stabilize the patient using contralateral access (pci site) to balloon tamponade and upsize sheath on the left femoral artery (lfa). patient received x2 units packed red blood cells and intubated. the patient was brought to the operation room for vascular repair. | [5 empty] | The investigation into the access site bleeding ¿ major and the vascular damage- peripheral vessel issues has been completed since the original report was filed. according to the clinical details, bleeding was reported after the removal of the peel away sheath and the use of perclose sutures, necessitating surgery to achieve hemostasis. the patient had a known peripheral artery disease (pad) condition. the cause of the vascular damage issue was most likely patient condition related to small vessel, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34903 | 1220648-2023-02845 | 17626916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17626916&pc=OZD | 07/28/2023 | 12/11/2023 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 03/31/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346259 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient experienced access site bleeding caused by introducer damage and vessel dissection. there was a surgical cutdown to repair vessel damage. | [5 empty] | The investigation into the patient's vascular damage-peripheral vessel has been completed. no product was returned. as per the clinical investigation, the impella introducer sheath kinked during removal of dilator due to obese patient condition. due to reintroduction of dilator, right femoral artery got dissected with patient anatomy complication. the cause of the introducer damage issue was an introducer sheath kink during removal of the dilator due to obese patient condition. the cause of the vascular damage issue was patient condition related with patient having obese condition and reintroduction of dilator causing a right femoral artery dissection. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34904 | 1220648-2023-02835 | 17627589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17627589&pc=OZD | 06/23/2023 | 09/20/2024 |  | 08/25/2023 |  |  | 07/26/2023 | 08/25/2023 |  |  | 03/31/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024326593 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/26/2023 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a 72 year old patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed anisoconria. a relationship with the impella device could not be ruled out. |  |  |  |  | The investigation into the stroke/cva has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34905 | 1220648-2023-02847 | 17628374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17628374&pc=PYX | 07/28/2023 | 09/26/2025 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 08/31/2024 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023250922 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/28/2023 |  |  |  |  | The user facility reported a 70-year-old female with right heart failure was implanted with an impella rp device for mechanical circulatory support. the pump started to fail on the impella. it was no longer showing adequate flows. the mc spiked on (b)(6), 2023 and lost pa pressures and flows a day later. a clot was suspected and was found on the pump when the impella was removed. | [5 empty] | The investigation for the reported pump stop and thrombus issue has been completed. impella rp product was returned for evaluation. visually inspected the pump and found the biomaterial to be present in the outflow and inflow area of the pump. the pump was connected to the console and tried running at p-9, but it was unable to flow due to complete biomaterial blockage at the outflow. data logs were reviewed, and multiple motor current spikes were observed with saturated flows after the event. placement signal drift, loss of pa pressures observed during chained event of biomaterial ingestion with saturated flows. additionally, placement signal drift and loss of pa pressures were observed during the chained event of biomaterial ingestion with saturated flows. the root cause of the pump stop issue was due to placement signal drift which occurred during ingested biomaterial per field investigation and data log review. the instructions for use states the following :¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ it also states ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The impella device was received from the customer and an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | D6a/d6b updated with additional information corrections that were made from the initial manufacturer device report number 1220648-2023-02847.
34906 | 1220648-2023-02846 | 17628499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17628499&pc=OZD | 07/28/2023 | 01/24/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  |  |  | 08/25/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1148921 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/28/2023 |  |  |  |  | The abiomed field service representative reported that the user facility's automated impella controller (aic) had a clip sensor that was not working for purge infusion. a loaner was sent to the user facility. there was no patient involvement in this event. | [5 empty] | The console was returned. visual inspection of the console revealed that the purge flag to be missing. there was also noticeable residue at the behind the purge unit. the root cause of the clip sensor not working was a missing purge flag. | The impella device was received and an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
34907 | 1220648-2023-02843 | 17628524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17628524&pc=OZD | 07/28/2023 | 08/10/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 06/30/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024361274 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported left ventricle (lv) access attempted using angled glide 0.035", multiple attempts with different wires and catheters were all unsuccessful. the physician attempted to cross and was successfully with an al1 and straight guidewire. however, the physician was unable to get catheter to advance into the lv without a lot of resistance. tee showed catheter was on top of the leaflet causing paravalvular leak, catheter and wire retracted. tee showed type a dissection. decision made to abort and place patient on ecmo with iabp 1:1. blood products were given to the patient. | [5 empty] | The investigation for ventricle perforation has been completed. no product was returned for a visual inspection. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the ventricle perforation was use related. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2023-02843. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-02843 in accordance with updated procedures. e.1 first and last name were revised as previously entered incorrectly on manufacturer device report number 1220648-2023-02843. h.6 added code 3038 since initial manufacturer device report number 1220648-2023-02843 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-02843 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34908 | 1220648-2023-02837 | 17628681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17628681&pc=OZD | 07/28/2023 | 01/22/2024 |  | 08/25/2023 |  |  | 07/28/2023 | 08/25/2023 |  |  | 05/31/2025 |  | 08/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356721 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2023 |  |  |  |  | The user facility reported a 65-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in their impella leg during pump support. it was noted by the staff that the patient's leg was cold to the touch. an external fem-fem bypass was performed coinciding with a fasciotomy to treat the noted condition. the patient was subsequently taken to the or to remove the impella cp and to have an impella 5.5 placed via axillary access for ongoing support. | [5 empty] | The investigation into ischemia has been completed. product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition. vessel conditions: calcification. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34909 | 1220648-2023-02809 | 17611530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17611530&pc=PYX | 07/26/2023 | 02/06/2024 |  | 08/24/2023 |  |  | 07/26/2023 | 08/24/2023 |  |  | 12/31/2024 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024299080 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. the team checked the impella's position with x-ray and reviewed impella connect. via chest x-ray, it appeared outlet was in position but the rp device was more linear compared to natural shape. when arriving on site, the rep. also noticed significant changes in purge flow/pressure. intubated/drips were added. rp reposition attempted to no improvement of flow. rp was removed. clot was noted in outlet cage. physician did not elect to replace rp device at this time. there was no patient harm reported. | [5 empty] | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. per log review and clinical review, multiple small motor current spikes were observed at p-6 with low pump flows. no positioning issue alarms were found. also, high purge pressure trend was found after motor current spikes with low pump flows, waveforms and pump flow noted to change with significant changes in purge pressure, flows later. the cause of the low pump flow issue was most likely biomaterial ingestion with several intermittent motor current spikes observed at p-6 with low flows. the cause of the high purge pressure issue was most likely biomaterial ingestion with purge trend noticed after several intermittent motor current spikes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34910 | 1220648-2023-02810 | 17614862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17614862&pc=OZD | 07/27/2023 | 08/24/2023 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 11/30/2024 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302343 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 75-year-old female had just begun impella cp support, experienced low pump flow. the medical team attempted to reposition the pump in an attempt to restore the normal flow, however, they were unable to do so. decision was made to explant impella. the patient was left in a stable condition. | [5 empty] | The investigation into the reported low pump flow has been completed. product was returned for investigation. the device was visually inspected and found the impeller blades to be broken. also, biomaterial was found to be present over impeller. no other abnormalities were found during case review. data logs were returned for review and motor current spike, suction alarms were found. the pump flow reduces abruptly after suction alarm. because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. per the ifu: handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
34911 | 1220648-2023-02811 | 17617918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17617918&pc=OZD | 07/25/2023 | 01/19/2024 |  | 08/24/2023 |  |  | 07/25/2023 | 08/24/2023 |  |  | 05/31/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2023 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at both the impella access site and the ecmo site. red blood cells and fresh frozen plasma were given to the patient. significant oozing near the impella site. will re-do angle matching to help minimize oozing. | [5 empty] | The root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34912 | 1220648-2023-02812 | 17617946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17617946&pc=OZD | 07/26/2023 | 01/18/2024 |  | 08/24/2023 |  |  | 07/26/2023 | 08/24/2023 |  |  | 04/30/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338785 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male patient had a hematoma with pain on the impella 5.5 access site. patient had a washout due to the hematoma the following day. | [5 empty] | The investigation into the access site bleeding. major has been completed since the original report was filed. the product was returned for investigation and no physical abnormalities were found. product was returned for investigation and no physical abnormalities were found. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical information regarding the bleeding was given. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34913 | 1220648-2023-02813 | 17617954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17617954&pc=OZD | 07/27/2023 | 01/24/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 08/31/2024 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Unspecified Eye / Vision Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The user facility reported an 81-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a lack of dorsalis pedis (dp) and posterior tibial (pt) pulses in the patient's feet prior to explant. both feet were swollen. the right foot was more swollen than the left one. the patient was also exhibiting a lack of bilateral pupil activity. the device was removed. | [5 empty] | The investigation into the reversible limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the cause of the limb ischemia - reversible was patient condition related since both legs were ischemic. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34914 | 1220648-2023-02814 | 17620291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620291&pc=OZD | 07/27/2023 | 07/24/2024 |  | 08/24/2023 |  |  |  | 08/24/2023 |  |  | 06/30/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359950 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old female patient went into ventricular tachycardia (vt) upon crossing the atrioventricular (av) valve with the impella 5.5. this was treated by cardioverting with a shock of 200 joules. cardioverted successfully and patient back into normal sinus rhythm (nsr). the 5.5 pump remained on for support with the 5.5 for 5 days til family choice was made to explant and provide palliative care only. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined. d.4 revised unique identifier (udi) # from the initial submission of manufacture device report number 1220648-2023-02814 as it was previously entered incorrectly. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-02814 in accordance with updated procedures. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-02814 was submitted. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report 1220648-2023-02814 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34915 | 1220648-2023-02815 | 17620830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620830&pc=OZD | 07/26/2023 | 01/15/2024 |  | 08/24/2023 |  |  | 07/26/2023 | 08/24/2023 |  |  | 12/31/2024 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315442 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that due to a challenging 76-year-old male patient's anatomy several attempts were done to pass the catheter through the aorta, but the catheter kink in the ascending aorta. another impella 5.5 was prepped and inserted into place. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. clinical details provided were sufficient to determine the root cause. according to clinical details, the catheter kinked due to patient challenging anatomy because of tortuosity vessel condition. therefore, the delivery issue was patient condition related due to patient tortuosity vessels. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
34916 | 1220648-2023-02817 | 17620849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620849&pc=OZD | 07/26/2023 | 01/19/2024 |  | 08/24/2023 |  |  | 07/26/2023 | 08/24/2023 |  |  | 03/31/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343203 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the doctor began with right femoral artery (rfa) access and angio, vessel of smaller caliber, but looked appropriate for use. sheaths swapped and impella locked into place. foot examined and shown to be cold and no dopplerable pulse. doctor then placed external fem to fem bypass. very red urine post impella placement. the rep encouraged repositioning, but not adjusted. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient had no flow down insertion site leg due to small and calcified artery. fem-fem bypass was done to restore flow on the right femoral leg access. the cause of the limb ischemia issue was patient condition related with patient having small and calcified vessels in right femoral and fem-fem bypass resolved ischemia. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34917 | 1220648-2023-02818 | 17620859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620859&pc=OZD | 07/26/2023 | 02/21/2024 |  | 08/24/2023 |  |  | 07/26/2023 | 08/24/2023 |  |  | 04/30/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355068 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Frayed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2023 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the staff stated that the wire was removed prior to starting impella. the impella .018 wire was frayed upon removal from the impella catheter but was not noticed until after the impella was turned on. shortly after impella was started, the staff member noticed the wire was frayed and the decision was made to turn the p level to p0 and remove the device as a precaution. a new impella pump was then successfully inserted over a new .018 wire. there was no patient harm related to this event. | [5 empty] | The investigation into the product damage (guidewire damage - great vessel) issue has been completed. the device was not returned for investigation. the root cause of the guidewire damage issue was not determined since product was not returned for investigation and the clinical information provided is inconclusive. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34918 | 1220648-2023-02819 | 17620862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620862&pc=OZD | 07/26/2023 | 01/24/2024 |  | 08/24/2023 |  |  | 07/26/2023 | 08/24/2023 |  |  | 01/31/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333012 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Unintended Movement | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2023 |  |  |  |  | The user facility reported a 68-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the patient vomited violently and dislodged the impella. furthermore, the patient began bleeding profusely from the groin. a hematoma was formed in the groin. two units of blood and femstop were applied. the pump was explanted that night and the patient was considered to be stable. | [5 empty] | The investigation into the positioning issues and the major access site bleeding issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the root cause of access site bleeding and positioning issues were due to patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34919 | 1220648-2023-02820 | 17620867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620867&pc=OZD | 07/27/2023 | 08/05/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 06/30/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024359663 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump failed to advance through the patient's vasculature during attempted implant. upon removal of the pump, it was discovered that the shaft of the pump had kinked. the patient was subsequently supported with an intra-aortic balloon pump. there was no noted harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the cause of delivery issue was most likely was use related due to excessive force to kink the device. complaint device not returned for investigation. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34920 | 1220648-2023-02821 | 17620880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620880&pc=OZD | 07/27/2023 | 06/06/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 05/17/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336748 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Excess Flow or Over-Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The patient expired after care was withdrawn. the patient's outcome was reviewed by the post market surveillance team and it was determined that this reportable event will be classified as a death. | The user facility reported a 70-year-old male was implanted with an impella 5.5 for escalation of therapy. on day 13 of support, the impella pump had shown signs of saturated flow and requested the product representative to come and troubleshoot it. the patient on this pump was reported to be hemodynamically stable. |  |  |  |  | The investigation for the saturated flow on the pump has been completed. the product was not returned for investigation. data logs confirm saturated flow but no motor current spike or other alarms related to that were seen. per the given clinical information, the root cause of the saturated flow issue was not determined since product was not returned and data logs were inconclusive. b1-selected adverse event b2-selected death. b5-added pms review determination. h6-impact code changed to 1806. | The impella device was not received from the customer and follow up with the field service representative has not been completed therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34921 | 1220648-2023-02822 | 17620892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620892&pc=OZD | 07/27/2023 | 01/22/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  |  |  | 08/24/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1606415 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The abiomed field service representative was performing preventive maintenance at the user facility. the automated impella controller (aic) failed when the support level was set to p1. there was no patient involvement. | [5 empty] | The investigation into the reported console issue has been completed. the device and data logs were returned for evaluation. the data logs were reviewed and confirmed that an impella stopped error occurred. this error was coupled with an error of ¿motor controller stopped the pump due to excessive differences in phase currents¿. the device was evaluated by comprehensive analysis and testing of the console¿s pump motor driver circuitry performed on a lab bench. during evaluation, an electrical component was found to have a voltage gain that was lower than expected, which resulted in an inaccurate measurement of the current delivered to phase a of the impella motor. other relevant components in the pump motor driver circuitry were tested and confirmed to be non-defective. replacing the impellatronic board in the console resolved the issue. the cause of the aic issue was a defective impellatronics board. the defect was specifically due to a defective operational amplifier ic (u308.2) that had a lower-than-expected gain. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34922 | 1220648-2023-02823 | 17620910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620910&pc=OZD | 07/27/2023 | 01/19/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 04/30/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351185 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The user facility reported a 84-year-old male (race unknown) in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma and persistent bleeding from the access site during impella support. although the bleeding was initially controlled with a femostop, the decision was ultimately made to explant the pump when the bleeding continued. | [5 empty] | The investigation into access site bleeding has been completed. the product was not returned for review. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34923 | 1220648-2023-02825 | 17620982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620982&pc=OZD | 07/27/2023 | 01/19/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 02/28/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024325820 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male patient had the pump stop working while on impella support. the defective pump was removed and a new one was placed. there was no harm to the patient. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the product was returned for investigation. as per clinical, impella cp had a high motor current pump stop alarm within 78 hours of support and did not resolve by replacing the controller. visually inspected the returned pump and there was damage noted to the cannula and the inflow cage was found to be deformed. it appears that the pump was likely to be clamped at the cannula inlet. a large impeller purge gap was observed during visual inspection. bearing wear was confirmed with large purge gap. no other abnormalities were found. the cause of the pump stop was due to bearing wear occurred within design life based on the field investigation results. | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34924 | 1220648-2023-02826 | 17620999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17620999&pc=OZD | 07/27/2023 | 03/06/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 03/31/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024329834 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during implantation of the impella, the sleeve on the sheath ripped. there was no patient harm related to this event. despite the torn sleeve, support with the implanted pump proceeded as planned. | [5 empty] | The investigation into the product damage (sterile sleeve damage) issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the sterile sleeve damage issue was not determined due to insufficient clinical information and no product returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34925 | 1220648-2023-02827 | 17621016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17621016&pc=OZD | 07/27/2023 | 01/18/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 03/31/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343846 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The user facility reported a 59-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient went into ventricular tachycardia and code was called. cpr performed, and the patient intubated and return of spontaneous circulation achieved. upon pulse check, it was noted that patient had no pulse in the right leg and it was cold. decision was made to consult cardio-thoracic surgeon and remove femoral-placed cp and place an axillary cp instead. | [5 empty] | The investigation into the reported limb ischemia ¿ reversible and vf/vt/af/at has been completed since the original report was filed. no product was returned. data logs were returned and no positioning issue relative to vt event was found. suction alarms were observed, with low flows at high p-level. the auto suction response was not allowing pump to achieve higher flows by reducing p-level. the cause of the ventricular tachycardia issue was most likely patient condition related and unknown relation to impella based on clinical and log review. the cause of the limb ischemia issue was most likely patient condition related based on clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34926 | 1220648-2023-02828 | 17621041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17621041&pc=OZD | 07/27/2023 | 08/24/2023 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 03/31/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340021 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 74-year-old female who needed impella support had the left leg become molted and a hematoma was observed. this required the medical team to place a percutaneous coronary intervention (pci) on the left main artery. the left femoral artery (lfa) impella was removed. the right femoral artery (rfa) access dilated and new impella placed. | [5 empty] | (h3) the investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the vascular damage was most likely use related and related to access being obtained too low based on the provided clinical information. a definitive cause for the ischemia could not be established. should additional relevant information become available, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34927 | 1220648-2023-02829 | 17621068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17621068&pc=OZD | 07/27/2023 | 01/24/2024 |  | 08/24/2023 |  |  | 07/27/2023 | 08/24/2023 |  |  | 03/31/2025 |  | 08/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343142 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp started to go into suction and could not be broken. when reduced to p2 it was giving false aortic alarms. proper position was confirmed by tte. measuring 3.6 cm. regardless of proper position, the physician still attempted repositioning. nothing would stop the alarms. decision was made to start epi and remove the device. | [5 empty] | The investigation into low pump flow and positioning issues has been completed. the device was returned for evaluation. the data logs were not received. visual inspection found biomaterial on the sensor face, inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. conclusion: the root cause of low flow was due to biomaterial ingestion. the positioning issue was also a result of the biomaterial on the device. | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
34928 | 1220648-2023-02804 | 17605739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17605739&pc=OZD | 07/26/2023 | 11/09/2023 |  | 08/23/2023 |  |  | 07/26/2023 | 08/23/2023 |  |  | 05/31/2025 |  | 08/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024347076 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult or Delayed Positioning | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a impella 5.5 that failed delivery due to the anatomy and the shaft of the pump kinking. the kink prompted the team to remove and replace with a impella cp pump that could be delivered. | [5 empty] | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the root cause of the delivery issue was the tortuosity on brachiocephalic artery. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34929 | 1220648-2023-02805 | 17607365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17607365&pc=OZD | 07/26/2023 | 09/19/2024 |  | 08/23/2023 |  |  | 07/26/2023 | 08/23/2023 |  |  | 06/30/2025 |  | 08/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357375 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demi. | The user facility reported a difficulty in placement a impella 5.5 on the 76-year-old male patient. as a result of the noted difficulty, the patient experienced bleeding and required a blood transfusion of unknown quantity. support continued with the implanted pump. |  |  |  |  | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. no device has been returned for investigation. per the clinical information provided, the root cause of the access site bleeding issue was patient condition related to patient anatomy. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34930 | 1220648-2023-02806 | 17607536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17607536&pc=OZD | 06/27/2023 | 09/04/2024 |  | 08/23/2023 |  |  | 06/27/2023 | 08/23/2023 |  |  | 04/30/2025 |  | 08/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333553 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2023 |  |  |  |  | The complainant reported that a connected pump was not recognized by the automated impella controller. a second pump was subsequently utilized for patient support. there was no patient harm. no patient demographics are available. | [5 empty] | Upon review of the provided logs, it was determined that this was a known but rarely occurring console issue, as such, no further investigation is warranted in regards to the utilized pump. there is currently no allegation of a malfunction or serious injury associated with the implanted pump. a regulatory report is no longer necessary please see manufacturer device report 1220648-2023-04594 for the investigation against the automated impella controller (aic) pertaining to this event. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34931 | 1220648-2023-02807 | 17607588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17607588&pc=OZD | 06/22/2023 | 01/24/2024 |  | 08/23/2023 |  |  | 06/22/2023 | 08/23/2023 |  |  |  |  | 08/23/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1349165 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/22/2023 |  |  |  |  | The abiomed field service representative reported the user facility had a battery failure on the automated impella controller (aic) during set-up. there was no patient involvement. | [5 empty] | The investigation into the aic - battery failure (red alarm) has been completed since the original report was filed. the device was returned for investigation. console was inspected and tested on an engineering lab bench. system functions properly when transport protection is engaged. when the transportation switch was disengaged, the system was not able to run without ac power- as expected and issued a battery failure alarm. the cause of the battery alarm failures in the field was likely due to a momentarily disengaged transport connection switch (use error). | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34932 | 1220648-2023-02808 | 17608483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17608483&pc=OZD | 07/26/2023 | 09/18/2023 |  | 08/23/2023 |  |  |  | 08/23/2023 |  |  | 02/28/2025 |  | 08/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338585 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had a 53-year-old male had an impella cp pump placed for hemodynamic support. the pump was being repositioned within the left ventricle (lv) when the team visualized a kink. the pump was replaced due to the kink. the pump was explanted without harm or injury and was replaced by a new impella pump via the other leg's femoral artery. | [5 empty] | The investigation into the reported ¿product damage (cannula kink)¿ has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found cannula kink observed near inlet cage. as per clinical investigation, the impella was pulled into aorta. during attempted repositioning under echo guidance in the unit, pump became kinked. kink was discovered when patient was brought to ccl for repositioning. the physician was able to straighten the pump in the ccl and removed without incident. the most likely root cause of the cannula kink was use related issue due to improper technique while attempted to reposition impella. | The impella device was received from the customer and an evaluation of the device on the benchtop is ongoing at this time. upon investigation closure, a supplemental mdr will be filed. per the ifu: handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
34933 | 1220648-2023-02777 | 17595005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595005&pc=OZD | 07/24/2023 | 04/11/2024 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 03/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345201 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella's pump had suction alarms. the staff attempted to reposition the pump as it was trapped in the mv apparatus. the impella was freed into the left ventricle (lv) and proper position could not be achieved. the pump flow could not pass 2 liters and a new impella device was inserted. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. cannula kink near outlet cage observed and the pigtails was tilted up from original shape. the root cause of the positioning issue and occurred cannula kink was use related issue due to improper technique as the cannula became trapped in the mitral valve apparatus. low pump flow was also discovered during investigation. the root cause of the low pump flow was due to the cannula kink. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34934 | 1220648-2023-02778 | 17595540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595540&pc=OZD | 07/05/2023 | 09/21/2024 |  | 08/22/2023 |  |  | 07/05/2023 | 08/22/2023 |  |  | 03/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343490 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/05/2023 |  |  |  |  | The user facility reported a 71-year-old male in cardiogenic shock and left main coronary artery bifurcation was implanted with an impella cp device for mechanical circulatory support. the patient experienced a left ventricle thrombus and positive blood cultures. the patient was being treated for sepsis. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the thrombus and the infection/sepsis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the thrombus issue was most likely patient condition related. the root cause of the infection issue is most likely due to patient condition related. the relation to impella is unknown for infection. the failure modes will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34935 | 1220648-2023-02779 | 17595549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595549&pc=OZD | 07/23/2023 | 01/19/2024 |  | 08/22/2023 |  |  | 07/23/2023 | 08/22/2023 |  |  | 02/29/2024 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023187531 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/23/2023 |  |  |  |  | The user facility reported a 13-year-old male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. during support, the pump encountered a fluid lead from what appeared to be a small hairline crack on the yellow luer. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation into the purge leak ¿ pump has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the purge leak- pump was due to sidearm yellow luer damage with bicarbonate use per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. information for use for the related event: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34936 | 1220648-2023-02780 | 17595602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595602&pc=OZD | 07/24/2023 | 04/19/2024 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 01/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024294359 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2023 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the initial flows were low until the impella was advanced further into the ventricle to 4.8cm under transesophageal echocardiography guidance. closure of chest cavity started, surgeon was called back approximately 40 minutes later after significant bleeding in the chest cavity noted. multiple blood products, factor vii and fluids given while identifying and treating areas of bleeding. impella motor current waveform became flat, and multiple suction alarms occurred. volume status addressed and position was confirmed, as flows steadily decreased and then fluctuated between 0.5l and 2l. multiple suction and impella position unknown alarms were noted while having low flow. flows eventually decreased and sustained 0.0 flow. no pump stop occurred, and no issues noted with motor current values. upon removal of the pump, biomaterial noted in the outlet cage. | [5 empty] | The investigation into the low pump flow and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the low pump flow issue was unable to be determined based on data log, clinical review and since product was not returned. the cause of the thrombus issue was most likely patient condition related per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34937 | 1220648-2023-02781 | 17595615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595615&pc=OZD | 05/05/2022 | 01/22/2024 |  | 08/22/2023 |  |  | 05/05/2022 | 08/22/2023 |  |  | 10/31/2023 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153078 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2023 |  |  |  |  | The user facility reported a 61-year-old male was implanted with an impella 5.5 for escalation of care from an impella cp. while the staff was pulling the patient over to the or table, the impella 5.5 was moved and continuous suction alarms occurred. transesophageal echocardiography was performed in or to reposition the pump, which was stuck in the mitral valve with the tip in the papillary muscle. multiple attempts were made to reposition the impella and they were unsuccessful. staff replaced the impella. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. per the provided clinical details, the root cause of the positioning issue was most likely due to patient movement. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34938 | 1220648-2023-02782 | 17595705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595705&pc=OZD | 07/24/2023 | 10/09/2023 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 11/30/2024 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024289421 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male on impella support had developed some obstruction on blood flow on right lower limb of impella insertion site due to very small artery. a femoral to femoral crossover bypass was done to maintain sufficient flow to the right leg and right lower extremities. | [5 empty] | The investigation into limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the limb ischemia issue was patient condition related with patient having small vessels in right femoral and fem-fem bypass resolved ischemia. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34939 | 1220648-2023-02783 | 17595842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595842&pc=OZD | 07/24/2023 | 10/17/2023 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  |  |  | 08/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AIC CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1584993 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the aic will not power up. this console was removed from use. | [5 empty] | The investigation into the aic - shutdown or restart issue has been completed since the original report was filed. the product was returned for investigation, and it was determined that the root cause of the failures were corrupt compact flash cards. | The impella device was returned and an evaluation is pending. upon investigation closure, a supplemental mdr will be filed.
34940 | 1220648-2023-02784 | 17595933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595933&pc=OZD | 07/24/2023 | 01/17/2024 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 03/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024329673 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2023 |  |  |  |  | The user facility reported a 67-year-old male was implanted with an impella 5.5 as a bridge to transplant. during support, a fluid leak was identified from the purge side arm during impella support. despite the noted leak, the purge flow and purge pressure were both still within normal limits. as a result of the leak, the site was advised to replace the implanted pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. the root cause of the purge leak was unable to be determined because the product was not returned and limited clinical details were provided.
34941 | 1220648-2023-02785 | 17595978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17595978&pc=OZD | 07/25/2023 | 01/24/2024 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  | 06/30/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364345 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a difficult axillary insertion of impella 5.5. there was a low pump flow after pump insertion. a removal of pump decision was made by physician due to suspected clot ingestion, after implantation tissue found on impella, appearing to be intima. the patient has been monitored for right arm pulses possible development of aneurysm in artery. a new pump was inserted without issue. | Patient had symptoms of occipital stroke. stroke alert activated. ct showed also subarachnoid hemorrhage. stroke team recommended no intervention. no impella concerns. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vascular damage- peripheral vessel and stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the root cause of the vascular damage was due to use issue. the cause of the stroke was not determined due to insufficient clinical details and unknown relation to impella. correction: section b5: the information was inadvertently left out in the initial report which has been updated in this report. correction: section h6: the health effect-clinical code was inadvertently left out in the initial report which has been updated in this report.
34942 | 1220648-2023-02786 | 17596236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17596236&pc=OZD | 07/24/2023 | 03/01/2024 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 02/28/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024319595 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2023 |  |  |  |  | The user facility reported an 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient endured a cerebral hemorrhage during support and expired. the physician could not rule out a relationship with the impella device as a potential factor. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke has been completed since the original report was filed. as the necessary clinical information was not provided, the true root cause of the stroke/cva could not be determined.
34943 | 1220648-2023-02787 | 17597550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17597550&pc=OZD | 07/25/2023 | 10/30/2025 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  |  |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024311706 |  | 0042-0000-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a that a 50-year-old male patient on impella support had a alarm indicating that the aic was not recognizing the cassette. a new aic was used and that resolved the issue. there was no harm to the patient as a result of this incident. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella pump and cassette were discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation of automated imeplla controller (aic): purge issue: other has been completed. data analysis: the logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. the alarm reported in complaint was consistent with defective disk detect flag. device analysis: the console was tested and the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. the cause of the inability to detect the purge disc was due to a broken flag. d1 brand name was erroneously entered on the initial manufacturer device report number 1220648-2023-02787. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-02787. d9 device returned to manufacturer date was inadvertently omitted on the initial manufacturer device report number 1220648-2023-02787. g1 reporting contact email revised on manufacturer device report 1220648-2023-02787 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-02787. h5 was erroneously selected on the initial manufacturer device report number 1220648-2023-02787 . h8 was erroneously selected on the initial manufacturer device report number 1220648-2023-02787.
34944 | 1220648-2023-02788 | 17597580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17597580&pc=OZD | 07/25/2023 | 09/20/2024 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  | 11/30/2023 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 JP |  |  | ABIOMED, INC. |  |  | US |  | 2022154240 |  | 05066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Death | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/25/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the patient was also on ecmo support. the impella support exceeded the 5.0 indication when the pump stopped on day 21. the pump was explanted and not replaced per the condition of the patient. the patient remained on ecmo support alone. | The patient's overall condition was poor and recovery was difficult. outcome of the procedure is patient passed away. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the data logs showed a gradual rise in purge pressure and spikes in motor current indicative of biomaterial ingestion causing the high purge pressure and the eventual blood ingress pump stop due to a lack of sufficient purge. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02788 was provided in sections b2, b5,d6, h1, and h6.
34945 | 1220648-2023-02789 | 17597868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17597868&pc=OZD | 07/25/2023 | 02/06/2024 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  | 03/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344812 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old female had a impella cp with low flows and suction during support. the suboptimal flow was troubleshot by repositioning and right heart assessment, but the flows did not improve. the team replaced the pump and support proceeded. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the low pump flow issue was biomaterial ingestion due to patient condition related and biomaterial was observed in the outflow after explant.
34946 | 1220648-2023-02790 | 17597905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17597905&pc=OZD | 07/25/2023 | 11/09/2023 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  | 01/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2024342716 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2023 |  |  |  |  | The user facility reported a male (age and race unknown) was implanted with an impella cp device for mechanical circulatory support of an angioplasty procedure. the impella cp crossed the aortic valve and the delivery across the valve caused complete heart block. the harm was treated by atropine and pacing. the support was for just 85 minutes, explanted and the groin site closed by perclose. | [5 empty] | The investigation into the heart block has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the cause of the complete heart block was patient condition related and it solved after medicine used. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34947 | 1220648-2023-02791 | 17599121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17599121&pc=OZD | 07/25/2023 | 12/04/2023 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  | 11/30/2024 |  | 08/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1417541 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Run on Battery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the aic battery would only charge less than 59% and not to a full charge, like it was supposed to. there was no harm to the patient as a result of this incident. | [5 empty] | The impella pump and console were not yet received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the automated impella controller (aic) battery failure has been completed. the logs for ic7035 showed a battery level low alarm on the complaint date. the lt logs on the complaint date show that battery 1 requested charging consistently whereas battery 2 is requesting charge intermittently as it consumes power as intende the root cause was battery charging issues was most likely due to a defective battery.
34948 | 1220648-2023-02792 | 17599385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17599385&pc=OZD | 07/02/2023 | 01/15/2024 |  | 08/22/2023 |  |  | 07/02/2023 | 08/22/2023 |  |  | 05/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345918 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2023 |  |  |  |  | The user facility reported a 76-year-old female decompensating after multiple surgeries was implanted with an impella 5.5 device for mechanical circulatory support. the patient bled after graft was inserted and required blood products. physician reported that patient was on anticoagulants. family withdrew care after 73 hours and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the patient was on anti-platelet use which was a contributing factor for bleeding at the graft site. the cause of the access site bleeding issue was most likely anticoagulation management related due to anti-platelet use being a key factor for bleeding.
34949 | 1220648-2023-02793 | 17599498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17599498&pc=OZD | 07/25/2023 | 08/22/2023 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  | 10/31/2024 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276782 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient that had a impella cp support that was ongoing but, was exhibiting low flows and suction. the impella cp was removed and replaced with escalated support to the 5.5 pump. no lasting harm or injury from the issue/malfunction. upon investigation of the returned product, there was an observation made of impeller blade damage. | [5 empty] | (h3) the investigation into the reported low pump flow has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and broken blades were observed. there was no cannula kink, no biomaterial and no other abnormality observed. the cause of the low pump flow was fractured impeller blades. the fracture was characteristic of an interaction with another device but that was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿
34950 | 1220648-2023-02794 | 17599592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17599592&pc=OZD | 07/24/2023 | 08/22/2023 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  |  |  | 08/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1121971 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Unable to Obtain Readings | Device Sensing Problem | Failure to Shut Off |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/24/2023 |  |  |  |  | The user facility reported a patient (demographics not provided) was implanted with an impella 5.5 as a bridge to transplant. the impella was placed without complications. upon arriving to the unit, a controller error appeared on the controller recommending switching the controller. the aortic and left ventricle waveform was no longer present in the controller. once the controller was exchanged, it sensed an impella even though it was not connected and the console could not be turned off. per the healthcare team this event did not affect the patient. | [5 empty] | The impella device has been received and is pending evaluation. upon investigation closure, a supplemental mdr will be filed.
34951 | 1220648-2023-02795 | 17600356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600356&pc=OZD | 07/24/2023 | 01/15/2024 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 02/28/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320374 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed a left ventricle thrombus during impella support. the thrombus was identified via echo and the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported thrombus has been completed since the original report was filed. the root cause of the thrombus is most likely due to the patient¿s condition as the clinical states thrombus was observed upon removal.
34952 | 1220648-2023-02796 | 17600365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600365&pc=OZD | 07/23/2023 | 10/10/2023 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  |  |  | 08/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1340487 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | No Display/Image | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/24/2023 |  |  |  |  | The user facility reported the automated impella controller (aic) blacked out during the priming step. the team switched to another aic and started support. the patient was unaffected. | [5 empty] | The investigation into the aic - power on issues/blank screen has been completed since the original report was filed. the product was returned for investigation. the console was tested after arriving in field service. the cause of the aic - power on issues/blank screen was not determined since failure mode was unable to be reproduced, failure mode will be monitored and trended. | The impella device was returned from the customer and an evaluation is ongoing. upon investigation closure, a supplemental mdr will be filed.
34953 | 1220648-2023-02797 | 17600368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600368&pc=OZD | 07/24/2023 | 02/06/2024 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 12/31/2024 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318893 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 07/24/2023 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding and a hematoma at the impella access site. patient was given blood products. the patient remained on support for three days and was successfully weaned. | [5 empty] | The investigation into the access site bleeding ¿ major and the thrombocytopenia issues has been completed. the device was not returned for investigation. as per clinical, patient started to bleed while introducer was in place with no pre-dilation prior to sheath delivery being followed. the cause of the access site bleeding issue was most likely use related due to no serial pre-dilation done prior to introducer insertion per clinical. the cause of the thrombocytopenia issue was most likely patient condition related and unknown relation to impella per clinical review. | Additional information was received via abiomed's long-term outcome and quality indicator (loqi) registry indicating the patient had endured thrombocytopenia and required platelet transfusion. purge solution was also switched to sodium bicarbonate. the allegation is stated to have a probable relationship with the device. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34954 | 1220648-2023-02798 | 17600370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600370&pc=OZD | 07/25/2023 | 01/25/2024 |  | 08/22/2023 |  |  | 07/25/2023 | 08/22/2023 |  |  | 11/30/2023 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166246 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 21-year-old male patient on impella support had encountered some issues with purge flow and pressure allegedly from a kink in the white catheter. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the reported high purge pressure issue has been completed since the original report was filed. pump was not returned for investigation. data logs shows the several purge system block alarm with barcoding trend in purge pressure, which indicate the sign of purge line kinking and unkinking. per clinical details, kink was noted near kink protector and team was trying to keep white cable unkinked with available hospital supplies. the cause of the high purge pressure issue was most likely a kinked purge line, based on data log review and provided clinical details.
34955 | 1220648-2023-02799 | 17600419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600419&pc=OZD | 07/24/2023 | 01/18/2024 |  | 08/22/2023 |  |  | 07/24/2023 | 08/22/2023 |  |  | 05/31/2025 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024346443 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old female with cardiac index and rhythm issues was implanted with an impella 5.5 device for support. during the procedure via cutdown to right subclavian, the impella sheath was inserted, and graft locks were secured. left ventricle (lv) access was obtained using standard 0.035 j and al1 catheter. the 0.035 was exchanged for abiomed 0.018 and impella was backloaded over the wire. the impella was inserted into graft, but there was difficulty crossing subclavian with what appeared to be a high take off. after multiple attempts, inlet finally crossed the atrioventricular (av) node into the lv. the wire was removed and pump started on p2. a small circular matter not previously identified on echo at the inlet was noted. pump moved up to p6 with suction. the pump pointed towards mitral valve and was repositioned. successfully torqued mid lv. there were questionable clot calcification notes. therefore, the pump turned off and removed. decision was made to place the patient on peripheral va ecmo, but physician was unable to get vascular access. the patient¿s chest was opened for central va ecmo. during cannulation, the physician also noted significant calcification at aortic root. | [5 empty] | The impella device was received from the customer and an evaluation is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported ¿low pump flow¿ has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found cannula kink and delaminated coils observed near outlet cage. there was no biomaterial and other abnormalities observed. as per clinical investigation, the during the procedure via cutdown to right subclavian, the impella sheath was inserted, and graft locks were secured. left ventricle (lv) access was obtained using standard 0.035 j and al1 catheter. the 0.035 was exchanged for abiomed 0.018 and impella was backloaded over the wire. the impella was inserted into graft, but there was difficulty crossing subclavian with what appeared to be a high take off. after multiple attempts, inlet finally crossed the atrioventricular (av) node into the lv. the wire was removed and pump started on p2. a small circular matter not previously identified on echo at the inlet was noted. pump moved up to p6 with suction. the pump pointed towards mitral valve and was repositioned. successfully torqued mid lv. there were questionable clot calcification notes. therefore, the pump turned off and removed. decision was made to place the patient on peripheral va ecmo, but physician was unable to get vascular access. the patient¿s chest was opened for central va ecmo. during cannulation, the physician also noted significant calcification at aortic root. the most likely root cause of the low pump flow was the cannula kink that happened due to patient condition related because the patient had significant calcification and multiple attempted to cross the av to lv that caused the kink.
34956 | 1220648-2023-02800 | 17600456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600456&pc=OZD | 07/26/2023 | 01/18/2024 |  | 08/22/2023 |  |  | 07/26/2023 | 08/22/2023 |  |  |  |  | 08/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1168797 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/26/2023 |  |  |  |  | The user facility reported the screen of an automated impella controller (aic) was glitching and notifications were flickering in and out. each time the pump was started, the unit asked to visualize the connector plug otherwise it would give an error message. the device was beeping intermittently with the audio disabled. there was no harm to a patient reported as a result of this incident. the device was removed from service as a result of the noted issue. | [5 empty] | The device investigation has not been completed. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported controller failure (red alarm) has been completed since the original report was filed. the console was returned for evaluation. during testing of the console, the failure mode was unable to be reproduced. additionally, the optical pump connector showed no signs of residue, and the placement signal was steady with a known good pump. the data logs were reviewed which revealed multiple placement signal unreliable and controller error alarms that went on for an hour until the pump was unplugged. these particular alarms are characterized by a temporary loss of optical data and have a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be managed monitored with imaging. the cause of the issue was not determined since the failure mode was not reproduced. added g2-user facility.
34957 | 1220648-2023-02802 | 17600500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600500&pc=OZD | 07/26/2023 | 08/22/2023 |  | 08/22/2023 |  |  | 07/26/2023 | 08/22/2023 |  |  |  |  | 08/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1142508 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/26/2023 |  |  |  |  | The abiomed field service representative reported the user facility's automated impella controller (aic) has a battery failure alarm even though the battery switch was on. the user replaced the unit with another aic. there are no patient details related to this case. no harm was reported. | [5 empty] | An investigation into the reported battery failure alarm has been completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. the failure modes will be monitored and trended.
34958 | 1220648-2023-02801 | 17600550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17600550&pc=OZD | 07/26/2023 | 01/18/2024 |  | 08/22/2023 |  |  | 07/26/2023 | 08/22/2023 |  |  | 10/31/2024 |  | 08/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2023276474 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2023 |  |  |  |  | The user facility reported a 56-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the pump purge flow gradually decreased before the pump stopped unexpectedly. as the patient was deemed hemodynamically stable at the time of the failure, the decision was made to explant the pump. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the pump was returned and investigated. data logs around the time of the failure were not returned. upon investigation it was found that the root cause of the pump stop issue was biomaterial found in the purge gap. | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
34959 | 1220648-2023-02754 | 17583919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17583919&pc=OZD | 07/23/2023 | 10/27/2025 |  | 08/21/2023 |  |  | 07/23/2023 | 08/21/2023 |  |  |  |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 130129 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | No Pressure | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the aic was unable to recognize the purge disk. the aic was replaced. there was no patient harm. | Additional information received that the procedure outcome was patient did not survive. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue b5 additional information received regarding the patient outcome and mso statement d1 model number was erroneously entered on the initial manufacturer device report number. d4 model number was erroneously entered on the initial manufacturer device report number. d9 device returned date was inadvertently omitted on the initial manufacturer device report number g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h3 device not returned was erroneously selected yes on the initial manufacturer device report number. h5 was erroneously selected yes on the initial manufacturer device report number. h6 medical device problem code 3270 and type of investigation 4114 was erroneously entered on the initial manufacturer device report number. h8 was erroneously selected on the initial manufacturer device report number. h10 statement erroneously stated the device was not returned on the initial manufacturer device report number.
34960 | 1220648-2023-02755 | 17583934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17583934&pc=OZD | 07/22/2023 | 10/30/2023 |  | 08/21/2023 |  |  | 07/22/2023 | 08/21/2023 |  |  | 04/30/2025 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP - EU |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024342970 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the purge system of the impella cp is leaking during patient's support. the yellow luerlock is broken; however, it is unclear if leak is at purge cassette or pump. | [5 empty] | The investigation into the purge leak has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical info provided, the root cause of the purge leak was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
34961 | 1220648-2023-02757 | 17584424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584424&pc=OZD | 05/13/2021 | 03/10/2025 |  | 08/21/2023 |  |  | 05/13/2021 | 08/21/2023 |  |  | 02/28/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022052485 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2021 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the yellow purge connector was stuck. it was reiterated to the medical team that it should not be forced and an instrument should not be used. there was no harm to the patient as a result of this event. support continued with the implanted pump for one more day before it was eventually explanted. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34962 | 1220648-2023-02758 | 17584461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584461&pc=OZD | 04/19/2021 | 01/24/2024 |  | 08/21/2023 |  |  | 04/19/2021 | 08/21/2023 |  |  | 10/31/2022 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021010655 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2021 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient who had just received an impella pump experienced some oozing from the access site. a jackson-pratt drain was used and fresh frozen plasma was provided to treat the oozing. | [5 empty] | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation. | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34963 | 1220648-2023-02759 | 17584656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584656&pc=OZD | 05/07/2021 | 09/20/2023 |  | 08/21/2023 |  |  | 05/07/2021 | 08/21/2023 |  |  | 02/28/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022052508 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2021 |  |  |  |  | The user facility reported a 46-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced a clot in the left atrium and required multiple blood transfusions throughout the course of support. care was eventually withdrawn following this event. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the thrombus issue was not determined since product was not returned and sufficient clinical information was not provided. the failure mode will be monitored and trended. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34964 | 1220648-2023-02760 | 17584682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584682&pc=OZD | 05/17/2021 | 10/16/2023 |  | 08/21/2023 |  |  | 05/17/2021 | 08/21/2023 |  |  | 09/30/2022 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1501813 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2021 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that upon planned explant of the impella 5.5 pump, a clot was noted within utilized graft. forceps and manual manipulation were utilized to expel the clot from the graft. bleeding resulted from the intervention and the graft was subsequently clamped, trimmed, and over-sewn to treat the condition. the patient was stable following the intervention. | [5 empty] | The investigation into the thrombus has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical account, the root cause of the thrombus was most likely patient condition related. it was reported that the clot was noted from the graft and have no issue with the working of the impella at the location of the graft and they were finally able to expel the clot. the clot was not observed at any position that has the direct contact with impella 5.5 during the insertion or removal. so it can be considered as the patient condition because of which the clot was formed and no other issues with were reported. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34965 | 1220648-2023-02761 | 17584740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584740&pc=OZD | 05/19/2021 | 01/22/2024 |  | 08/21/2023 |  |  | 05/19/2021 | 08/21/2023 |  |  | 01/31/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022040023 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2021 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient experienced bleeding from multiple sites during impella support. a blood transfusion of unknown quantity was provided to the patient to counteract the condition. additional information pertaining to the involved sites was not provided. support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the device was not returned by the medical facility. the root cause of the patient injury was patient condition. it was reported that there was massive bleeding that had occurred and there was no additional information that was provided that identifies as the potential cause for bleeding. bleeding from multiple sites does not relate the impella 5.5. the bleeding sites are not particular to the impella 5.5 site. the root cause was patient condition. unknown relation to the impella. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34966 | 1220648-2023-02762 | 17584777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584777&pc=OZD | 05/21/2021 | 01/23/2024 |  | 08/21/2023 |  |  | 05/21/2021 | 08/21/2023 |  |  | 09/30/2022 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1505056 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2021 |  |  |  |  | The user facility reported a 37-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, a thrombus developed in the patient's left ventricle. due to the noted clot, tissue plasminogen activator was administered and the issue resolved. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported thrombus has been completed since the original report was filed. the device was not returned by the medical facility. it was reported that there were concerns for clot in the left ventricle. there was no mentioning of any other clinical details that the injury was related to the functioning of the pump. later, the issue was resolved after the administration of tissue plasminogen activator (tpa) and the pump function was stable and no other problems was reported. the root cause of the patient injury was patient condition.
34967 | 1220648-2023-02764 | 17584849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584849&pc=OZD | 06/11/2021 | 11/10/2023 |  | 08/21/2023 |  |  | 06/11/2021 | 08/21/2023 |  |  | 03/31/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022067477 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2021 |  |  |  |  | The user facility reported a 64-year-old male patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced "impella in the aorta" alarms. the position was verified and the medical staff was walked through disabling alarm. there was no harm reported as a result of this incident. | [5 empty] | Correction: upon review of the provided information, this is no longer being considered a reportable malfunction. the optical signal that was reported to be faulty, is not required for patient support and the functionality of the pump can be monitored via other means without risk to the patient. however, the investigation into the allegation has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based off the clinical info available, this is a false impella in aorta alarm. based on the provided information, the cause for optical placement signal issue was believed to be related to the patient condition. | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34968 | 1220648-2023-02765 | 17584905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584905&pc=OZD | 06/12/2021 | 01/24/2024 |  | 08/21/2023 |  |  | 06/12/2021 | 08/21/2023 |  |  | 03/31/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022067496 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2021 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient experienced active bleeding from the access site during impella support. as a result of the noted condition, the patient was taken to the operating room for a washout and pump explant. the patient was stable following the intervention. | [5 empty] | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
34969 | 1220648-2023-02766 | 17584954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584954&pc=OZD | 06/12/2021 | 03/11/2025 |  | 08/21/2023 |  |  | 06/12/2021 | 08/21/2023 |  |  | 04/30/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022074349 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2021 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a thrombus and experienced persistent suction issues during impella support. the ongoing alarms were silenced and the patient was given a bolus of heparin to counteract the thrombosis. support continued with the implanted pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Thrombus - the cause of the issue is patient condition as thrombus was observed in the imaging during process. low pump flow - 1200+ suction alarms were observed on the logs occurring intermittently. the cause of the low pump flow was patient condition since there was in decline in the rv function as well as the logs indicated flatline in the motor current. purge leak (cassette) - the root cause of the purge cassette leak was unable to be determined as no product was returned for investigation. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
34970 | 1220648-2023-02767 | 17584978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17584978&pc=OZD | 07/23/2023 | 08/21/2023 |  | 08/21/2023 |  |  | 07/23/2023 | 08/21/2023 |  |  | 03/31/2025 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024333863 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/23/2023 |  |  |  |  | The user facility reported a patient (unknown age and gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had a blood sample taken which showed to be hemolytic. some adjustments to the pump were made, however, there was no improvement. the decision was made to replace impella cp with impella 5.5 due to concerns of hemolysis. | [5 empty] | The investigation into the reported hemolysis has been completed. the cause of the issue was not determined since product was not returned, and sufficient clinical information and data logs were not returned. the failure modes will be monitored and trended. instructions for use follow: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
34971 | 1220648-2023-02756 | 17585542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17585542&pc=OZD | 05/19/2021 | 03/07/2024 |  | 08/21/2023 |  |  | 05/19/2021 | 08/21/2023 |  |  | 03/31/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022058032 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2021 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that as the pump was inserted for delivery, the patient went into ventricular tachycardia (vt). the patient stabilized and the pump was successfully implanted. however, upon attempting to turn the patient to change the bedding, the patient again went into vt. their rhythm again stabilized and support with the pump continued. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the vt/vf issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
34972 | 1220648-2023-02768 | 17585674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17585674&pc=OZD | 07/22/2023 | 01/23/2024 |  | 08/21/2023 |  |  | 07/22/2023 | 08/21/2023 |  |  | 01/31/2025 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311960 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 45-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella was explanted due to a pulseless right foot. after explant the patient's right groin site was stable under pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia - reversible has been completed since the original report was filed. product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition.
34973 | 1220648-2023-02769 | 17585718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17585718&pc=OZD | 07/21/2023 | 01/19/2024 |  | 08/21/2023 |  |  | 07/22/2023 | 08/21/2023 |  |  | 12/31/2024 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325175 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2023 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient suffered an embolic stroke during impella support. an embolectomy was successfully performed to treat the occlusion. the patient showed signs of short-term memory difficulty following the stroke. impella support remains ongoing with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed. the device was not returned for benchtop evaluation and investigation. the relationship between the cva/stroke and the use of impella was unable to be determined. the root cause of the stroke/cva cannot be determined due to insufficient clinical information.
34974 | 1220648-2023-02770 | 17585724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17585724&pc=OZD | 07/23/2023 | 01/15/2024 |  | 08/21/2023 |  |  | 07/23/2023 | 08/21/2023 |  |  | 02/28/2025 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338536 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old male patient was oozing from access groin site during impella support and received two units of blood products. support continued with the implanted pump following the transfusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
34975 | 1220648-2023-02763 | 17585735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17585735&pc=OZD | 06/08/2021 | 09/20/2024 |  | 08/21/2023 |  |  | 06/08/2021 | 08/21/2023 |  |  | 03/31/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071483 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2021 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient developed persistent position induced hemolysis during impella support. the patient was provided multiple units of prbc throughout the course of support and the pump was repositioned on multiple occasions due to the persistent condition. support with the pump continued despite the hemolysis. | It was reported the patient was placed on continuous renal replacement therapy while on impella support, due to decreased urinary output, hemolysis, and rising creatinine. it was noted patient care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02763 was provided in sections b2, b5, d6a, d6b, h1, and h6.
34976 | 1220648-2023-02771 | 17590092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17590092&pc=OZD | 07/05/2023 | 10/20/2023 |  | 08/21/2023 |  |  | 07/05/2023 | 08/21/2023 |  |  | 01/31/2025 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334534 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2023 |  |  |  |  | The user facility reported 46-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced poor distal perfusion while on support. the impella cp was explanted. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient had ischemia on impella leg and pump was explanted due to poor flow down the leg. the cause of the limb ischemia issue was most likely patient condition, based on the clinical review. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34977 | 1220648-2023-02773 | 17591696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17591696&pc=OZD | 07/23/2023 | 01/24/2024 |  | 08/21/2023 |  |  | 07/23/2023 | 08/21/2023 |  |  | 04/30/2024 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 31-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella was placed but groin started oozing. the physician proceeded with repo sheath positioning, but the oozing worsened and was not resolved with extended compression. the decision was made to remove the impella as the patient was going to be transferred to another facility. the impella was removed without incident. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the device was not returned by the medical facility. the root cause of the access site bleeding was most likely due to calcified vessel patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34978 | 1220648-2023-02772 | 17591740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17591740&pc=OZD | 07/22/2023 | 03/08/2024 |  | 08/21/2023 |  |  | 07/22/2023 | 08/21/2023 |  |  | 04/30/2025 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350782 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50-year-old male was implanted with an impella cp device for mechanical circulatory support. patient was brought back to the cath lab after multiple attempts on the unit under echo to reposition impella, the pump was facing posteriorly towards mitral valve, upon trying to reposition the impella in cath lab, a pigtail caught and it was unable to get optimal position, patient plasma free hemoglobin 430 and having red urine. a decision to pull impella back into aorta and attempt to reposition; however, physician was unable to get pump into optimal position, a plan to use re-access port and place 0.035 wire to maintain access and place a new impella. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. data logs show alarms for impella position in the aorta during the 30-minute repositioning of the device at p level 2. as per the clinical details, doctor encountered difficulties while repositioning the impella. despite efforts, optimal positioning couldn't be achieved, leading to the insertion of a new impella due to red urine and plasma free hemoglobin was given. the cause of the hemolysis issue was most likely improper positioning of the device, based on given clinical details and data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34979 | 1220648-2023-02774 | 17591761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17591761&pc=OZD | 07/23/2023 | 01/31/2024 |  | 08/21/2023 |  |  | 07/23/2023 | 08/21/2023 |  |  | 04/30/2025 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355672 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/23/2023 |  |  |  |  | The user facility reported a 75-year-old male patient in ventricular fibrillation prior to arriving in cath lab, return of spontaneous circulation achieved in around 30 minutes. upon arrival to the cath lab, impella cp in place in right fem; patient received drug eluting stent to the left anterior descending artery. the decision was made to leave impella in patient. while removing peel away sheath, the impella was pulled back into the aorta. the physician was unable to advance the impella back across the aortic valve. an 0.035" wire was inserted through the guide wire access port, the impella was removed, the unused 14f 13cm sheath was inserted over the 0.035" wire. the impella was successfully replaced using a bs v18 wire. | [5 empty] | The investigation into the positioning issues has been completed. the pump was not returned for investigation. as per clinical, pump was pulled back into the aorta while removing peel away sheath. the cause of the positioning issue was most likely use related, based on clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34980 | 1220648-2023-02775 | 17591774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17591774&pc=OZD | 04/26/2021 | 01/26/2024 |  | 08/21/2023 |  |  | 04/26/2021 | 08/21/2023 |  |  | 03/31/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022057801 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2021 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient inadvertently pulled back the pump as the patient was in a [confused] state. patient was taken to or for explant and exchange. new pump was inserted without any complications. the patient expired on a later date, there were no allegations against the pump. | [5 empty] | The investigation for the positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. the root cause of the positioning issue was most likely patient movement due to the pump inadvertently pulled back by patient in confessed state as per the clinical description provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
34981 | 1220648-2023-02776 | 17591865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17591865&pc=OZD | 05/13/2021 | 03/10/2025 |  | 08/21/2023 |  |  | 05/13/2021 | 08/21/2023 |  |  | 01/31/2023 |  | 08/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022037403 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2021 |  |  |  |  | The user facility reported a 43-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that upon planned explant of the impella 5.5 pump, a clot was noted within the right axillary graft. the graft was subsequently cleaned and a catheter was utilized to remove the clot. the patient was considered stable following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
34982 | 1220648-2023-02729 | 17571858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17571858&pc=OZD | 07/19/2023 | 10/26/2023 |  | 08/18/2023 |  |  | 07/19/2023 | 08/18/2023 |  |  |  |  | 08/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1273705 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/19/2023 |  |  |  |  | The biomedical engineer contacted abiomed field service representative to report an automated impella controller (aic) yellow alarm. the alarm will prompt the team to utilize a backup console per ifu instructions. no patient harm or injury was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) issue has been completed since the original report was filed. the aic product was returned for investigation. console was inspected and tested on an engineering lab bench. system self-check failed on startup. the cause of the system self-check failure was determined to be the corrosion of the pbm.
34983 | 1220648-2023-02730 | 17571893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17571893&pc=OZD | 07/19/2023 | 01/15/2024 |  | 08/18/2023 |  |  | 07/19/2023 | 08/18/2023 |  |  | 11/30/2024 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308425 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2023 |  |  |  |  | The user facility reported a 58-year-old female who had cardiac surgery was implanted with an impella cp device for mechanical circulatory support post-op due to cardiac event. the impella cp was having persistent suction and low flow issues. the team attempted to troubleshoot but decided to replace the cp with a new cp pump. support proceeded without harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported low pump flow has been completed since the original report was filed. the cause of the low pump flow issue cannot be determined since the complaint device not returned for investigation and data logs reviewed and many suction alarm occurred and low flow confirmed but no indication for issue cause.
34984 | 1220648-2023-02731 | 17571939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17571939&pc=OZD | 07/18/2023 | 09/26/2025 |  | 08/18/2023 |  |  | 07/19/2023 | 08/18/2023 |  |  |  |  | 08/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1276242 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/19/2023 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the automated impella controller (aic) had the alarm "battery temperature high" occur. the alarm lasted 10 seconds and the patient reports that the patient does not believe the device stopped at all. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported aic - battery charging/other issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the logs from the console confirm that battery hot temperature alarm was issued by the aic but almost immediately and cleared. a visual inspection of the internal circuitry of the console didn't reveal any issues with the console. the issue could not be reproduced in the lab when a known good pump and purge was run with the console for over an hour. the root cause of the temperature spike is a glitch of the power battery management (pbm) communication triggered by the console being plugged into ac mains. this report is being filed as part of a retrospective review of historical records.
34985 | 1220648-2023-02733 | 17573166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17573166&pc=OZD | 07/20/2023 | 10/21/2025 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  |  |  | 08/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1095114 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/20/2023 |  |  |  |  | The abiomed field service representative reported that while removing the purge cassette and disc from the automated impella controller (aic) console, the blue portion the disc slides into broke off. this event occurred during training and there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34986 | 1220648-2023-02734 | 17573699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17573699&pc=OZD | 07/19/2023 | 01/19/2024 |  | 08/18/2023 |  |  | 07/19/2023 | 08/18/2023 |  |  | 12/31/2024 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024304118 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2023 |  |  |  |  | The user facility reported a 68-year-old male with low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. on the second day after the introduction, there was ischemia in the lower limb, and an antegrade sheath was inserted to bypass blood from the unaffected side. after bypassing the blood from the antegrade sheath lead, the ischemia was resolved. additionally, when the impella cp was removed, roughly 30 cm of intima adhered to the pigtail. the iliac artery was dissected from the iliac artery to the periphery, and the intima of the blood vessel was peeled off. the intima was cleanly detached, and there was blood flow in the lower extremities when angiography was performed. after removal of impella, blood flow in the medial malleolar artery was confirmed by doppler, and there was no ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible and vascular damage- peripheral vessel issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, incorrect introducer 16fr medikit sheath was used causing the vascular damage. patient had calcified vessels on femoral access and antegrade sheath was used to resolve ischemia. the cause of the limb ischemia issue was patient condition related due to patient having calcified vessels. the cause of the vascular damage issue was use related due to 16 fr medikit introducer being used. impella access is not indicated for apart from the specified manufacturer introducer listed per ifu 0042-9010-jp.
34987 | 1220648-2023-02735 | 17573982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17573982&pc=OZD | 07/19/2023 | 11/14/2025 |  | 08/18/2023 |  |  | 07/19/2023 | 08/18/2023 |  |  |  |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1415135 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Operating System Version or Upgrade Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/19/2023 |  |  |  |  | The abiomed field service representative was performing preventive maintenance on the automated impella controller (aic) when a system error occurred after successfully uploading software version 12.1.1. upon rebooting the aic, another system error was triggering. there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. console logs were not available for analysis, as they're being erased during aic sw upgrade. analysis of the device and its file structure identified that after failed installation, the systemconfig file (which is supposed to list all fw versions compatible with current aic sw) showed "err", which is expected after sw installation fails. the content of systemconfig file, however, is not updated after consecutive aic sw installation attempts. after the file was manually deleted from aic's cf card during our troubleshooting, the next aic sw installation was successful, and console operated within specification. the cause of the first aic sw upgrade failure was not determined d2 common device name updated. d3 fax number updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated.
34988 | 1220648-2023-02736 | 17574152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17574152&pc=OZD | 07/20/2023 | 08/18/2023 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 01/31/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327997 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/20/2023 |  |  |  |  | The user facility reported a 57-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, after that nurse turned the patient, the purge flow dropped and, upon inspection noted that it was fluid was coming out from the filter area. there was no serious harm to the patient as a result of this incident. the patient was escalated to an impella 5.5 device with no issues. | [5 empty] | The impella cp and purge cassette were received from the user facility. pump was investigated and the purge leak-pump issue reproduced during the investigation. leak was observed on the reservoir adhesive joint to the filter. no other abnormalities were found. the cause of the purge leak issue was due to damage to the reservoir adhesive joint to filter joint. the purge leak failure was observed close to 12 days after implant. the impella cp optical device is not indicated for support more than 8 days of design life period according to design per information for use. the failure modes will be monitored and trended.
34989 | 1220648-2023-02737 | 17574501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17574501&pc=PYX | 07/20/2023 | 10/10/2023 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 06/30/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024359021 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Movement | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 52-year-old male patient on impella support had experienced ventricular fibrillation while at home with cpr was administered. the pump was in good position, but came out of ventricular positioning when suturing at the groin. pump unable to re-positioned. decision was made to explant the impella and implant a new one. the patient was in a stable condition after the procedure. | [5 empty] | The investigation into the positioning issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0052-9013. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
34990 | 1220648-2023-02738 | 17574536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17574536&pc=OZD | 07/20/2023 | 01/22/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 04/30/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024344488 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a leakage of purgeline between the reservoir and filter coinciding with impella support. the leak seemingly resolved itself when purge bypass placed, and no purge cassette change was needed. the pump purge sidearm was broken/damaged during the support. the pump was explanted a week post implantation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿. | (h3) the investigation into the reported "purge leak - pump" has been completed since the original report was filed. product was returned for investigation. the impella cp pump complete product was not returned, only the sidearm was returned. during visual inspection found that the sidearm filter to be cracked and damaged on the surface. there were no other abnormalities found. the cause of the purge leak was due to the sidearm filter damage with possible ipa interaction.
34991 | 1220648-2023-02739 | 17574620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17574620&pc=OZD | 07/20/2023 | 01/16/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 06/30/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356782 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 81-year-old male patient experienced an embolic stroke during impella support. following the initiation of support, the patient became hypertensive and experienced shortness of breath. upon exhibiting difficulty with some extremity movement, the patient was sent for a ct scan. it was then verified that the patient had experienced a stroke. a thrombectomy was subsequently performed to treat the condition. the patient was considered to be stable following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the stroke was likely due to patient's condition. it is unknown if the relationship to the stroke is due to heparin used in the purge.
34992 | 1220648-2023-02740 | 17574752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17574752&pc=OZD | 07/20/2023 | 07/23/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 02/28/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342239 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Signal Artifact/Noise | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was a constant alarm stating impella flow low, despite multiple repositioning efforts under echo. the impella was almost pulled out of the patient¿s heart and the alarm continuously kept on alarming. a decision was made to silence the alarm as it was suspected that the alarm was false. furthermore, suction alarms also occurred. the suction alarm was resolved with volume. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for optical signal issue has been completed. no product was returned for a visual inspection. data log analysis confirmed suction alarms requiring volume, multiple placement signal low alarms throughout the case, no low flow, and no other abnormalities related to the 5s parameters. the most likely cause of the optical signal issue was patient condition related. d6a/d6b added. h6 medical device problem code 2917 added the following have been reported incorrectly or missing from manufacturer device report 220648-2023-02740: g1 reporting contact fax number missing.
34993 | 1220648-2023-02741 | 17575062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17575062&pc=OZD | 07/20/2023 | 09/21/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 10/31/2024 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287630 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 79-year-old female with a history of chronic obstructive pulmonary disease (copd), aspergillosis, hypothyroidism, and htn presented to ed yesterday afternoon with shortness of breath and chest pressure. the patient was admitted and was brought down to the cath lab for impella placement and possible angioplasty. the patient was taken emergently to or for cabg procedure. during the off-pump cabg, the physician lifted and turned heart, which resulted in a perforation from impella cannula. the patient was placed on cardiopulmonary bypass (cpb). the perforation of the left ventricle was successfully closed. due to the extent of dissection down lateral wall, no other options were available to restore flow. the patient was unable to tolerate coming off cpb and patient expired. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the ventricle perforation issue has been completed since the original report was filed. no product was returned. as per clinical review, while attempting to get to circulation doctor lifted and turned heart which resulted in a perforation from impella cannula. the cause of the ventricle perforation issue was use related based on clinical review. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
34994 | 1220648-2023-02742 | 17575077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17575077&pc=OZD | 07/20/2023 | 10/07/2023 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  |  |  | 08/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/20/2023 |  |  |  |  | The user facility reported a 60-year-old male patient was implanted with an impella cp for circulatory support. during support, the automated impella control (aic) had a failure alarm that resolved immediately. the aic function was not affected. the aic was replaced as a precaution. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, the failure mode could not be reproduced during testing and the console functioned as expected. the root cause of the issue was not determined since failure mode was not reproduced. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34995 | 1220648-2023-02743 | 17575247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17575247&pc=OZD | 07/20/2023 | 09/07/2023 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  |  |  | 08/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1566931 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/20/2023 |  |  |  |  | The user facility reported the blue plastic piece that fits in purge disc slot of the automated impella controller (aic) broke off during set up for a high risk percutaneous cardiac insertion case. the medical center removed the aic from patient use. the team was able to successfully support the patient with another console. no harm or injury from delay. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the device subcomponent was returned and investigated. the benchtop analysis and investigation determined that the root cause of the aic - purge disc/flag issues was a broken purge disc retainer and purge flag. failures of this type will be monitored for trends. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
34996 | 1220648-2023-02744 | 17575332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17575332&pc=OZD | 07/20/2023 | 03/08/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 10/31/2024 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2023 |  |  |  |  | The abiomed contact reported that a patient on impella support was getting low flows. imaging showed that the device was in proper position; despite that, the pump was still giving low flow alarms. there was no harm to the patient reported as a result of this incident. patient demographics are unknown at this time. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed which confirmed low flow throughout the case but there were no indications of motor current spikes, placement signal trends, or positioning issues during the case run. the doctor was unable to achieve proper pump flow despite a good pump position and adequate fluid given. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided.
34997 | 1220648-2023-02745 | 17575393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17575393&pc=OZD | 07/20/2023 | 01/12/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 05/31/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356214 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2023 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella 5.5 device for mechanical circulatory support for bridge to transplant. during support, the patient experienced internal bleeding probably in the 5.5 axial access pocket, no external access site bleeding noted. blood transfusion and heparin resolved the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleed has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine root cause. the cause of the access site bleeding issue was most likely high patient act values since the heparin stop solved the issue, based on provided clinical information.
34998 | 1220648-2023-02748 | 17576754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17576754&pc=OZD | 07/21/2023 | 01/24/2024 |  | 08/18/2023 |  |  | 07/21/2023 | 08/18/2023 |  |  | 01/31/2024 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023178277 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility in the eu reported patient experienced access site bleeding requiring blood transfusion on day 3 of support with the impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of access site bleeding issue cannot be determined since the complaint device was not returned for investigation and no sufficient data returned.
34999 | 1220648-2023-02749 | 17577331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17577331&pc=OZD | 07/21/2023 | 01/12/2024 |  | 08/18/2023 |  |  | 07/21/2023 | 08/18/2023 |  |  | 05/31/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024342274 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male in post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 device for mechanical circulatory support. however, the impella 5.5 showed a "purge-pressure low¿ alarm, so the staff checked the purge-side arm and noticed a leak. after a short time alarm was resolved, but purge-side arm is still leaking. the impella 5.5 was explanted on (b)(6) 2023. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g.,infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak - pump has been completed since the original report was filed. the device was returned for investigation. it was determined that the root cause of the purge leak is due to a crack in the yellow luer. the leak was able to be reproduced using a syringe.
35000 | 1220648-2023-02750 | 17577349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17577349&pc=OZD | 07/21/2023 | 07/23/2024 |  | 08/18/2023 |  |  | 07/21/2023 | 08/18/2023 |  |  | 06/30/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357338 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male patient that had a positioning issue involving an impella cp pump during support. following initial placement, it was stated that the pump was pulled into the aortic arch when the guidewire was being removed. the pump prolapsed on itself and, in the process of trying to remove it, the physician lodged the device in the iliac artery. a snare was unsuccessful in retrieving the pump and surgical intervention was subsequently performed to remove the pump. there was no lasting harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | No product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issue was use related. the cause of the removal issue was not determined because no product was returned and not enough clinical information was given. h6 added medical device problem code g1 reporting contact fax number missing.
35001 | 1220648-2023-02751 | 17577564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17577564&pc=OZD | 07/21/2023 | 10/16/2023 |  | 08/18/2023 |  |  | 07/21/2023 | 08/18/2023 |  |  | 03/31/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344609 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump was placed in the catheterization lab to allow for angioplasty and then sent to the unit (icu) for continued monitoring. the patient presented to the icu with low pulses in the leg and a mottled left leg. the decision was made to explant the impella after just 2 1/2 hours of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the cause of limb ischemia most likely patient condition related.
35002 | 1220648-2023-02752 | 17578464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17578464&pc=OZD | 07/19/2023 | 02/14/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 03/31/2025 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346332 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2023 |  |  |  |  | The user facility reported a 56-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support during coronary artery bypass grafting. during the bypass surgery, the patient was found to have severe mitral disease. the physician performed mitral valve replacement and decided on escalation of care for the patient. an impella 5.5 was inserted. during the exchange of cp for 5.5, the right femoral artery had massive blood loss and a total of 4 packed red blood cells and platelets were infused. the blood loss was at the impella cp access site. the team was unable to obtain good echo views due to artifact from mitral valve replacement. there was no suction or hemolysis noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the access site bleeding issue was patient condition related due to patient calcified vessels.
35003 | 1220648-2023-02753 | 17578483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17578483&pc=OZD | 07/20/2023 | 01/24/2024 |  | 08/18/2023 |  |  | 07/20/2023 | 08/18/2023 |  |  | 09/30/2024 |  | 08/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274149 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2023 |  |  |  |  | The user facility reported a male (age unknown) undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding from the groin site. the sheath was advanced, and the site was re-sutured to mitigate the bleeding. the bleeding successfully subsided. the patient remained on support for one day and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the device was not returned by the medical facility. the root cause of the access site bleeding was patient movement during transfer which was resolved with re-suturing.
35004 | 1220648-2023-02726 | 17568712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17568712&pc=OZD | 07/20/2023 | 12/03/2024 |  | 08/17/2023 |  |  | 07/20/2023 | 08/17/2023 |  |  | 06/30/2025 |  | 08/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024360351 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 38-year-old male patient was having access site bleeding during impella 5.5 support. the patient returned from the operating room (or) having the axillary site cleaned out from bleeding/hematoma. the wash resolved and controlled the access site bleeding. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: d4-udi number. added d6a/d6b.
35005 | 1220648-2023-02727 | 17568751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17568751&pc=OZD | 07/20/2023 | 01/25/2024 |  | 08/17/2023 |  |  | 07/20/2023 | 08/17/2023 |  |  | 01/31/2025 |  | 08/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315321 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old female patient on impella support had experienced a access site bleed. the bleeding improved slightly after the sheath adjustment and the patient was administered and unknown amount of units of blood. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding issue has been completed since the original report was filed. product was not returned for investigation. the cause of the access site bleeding issue was most likely use related since bleeding was improved after angle match of sheath, based on provided clinal information.
35006 | 1220648-2023-02728 | 17569253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17569253&pc=OZD | 07/18/2023 | 08/17/2023 |  | 08/17/2023 |  |  | 07/18/2023 | 08/17/2023 |  |  |  |  | 08/17/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1319826 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | Unexpected Shutdown | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/18/2023 |  |  |  |  | The complainant reported a 67 year old female patient presenting for hemodynamic support using the impella device. during support, the console appeared to be rebooting while patient was on support. once the console came back on, they received an alert of "unexpected console shutdown". there did not appear to be a disruption in hemodynamic support during the reboot. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
35007 | 1220648-2023-02687 | 17556064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17556064&pc=OZD | 07/18/2023 | 10/16/2023 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  | 09/30/2024 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273637 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 52-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp accessed limb had inadequate flow down it. the team explanted and ballooned the iliac and the vessel was seen to have a dissection in the common femoral. later the team additionally performed a fasciotomy. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the root cause of limb ischemia was most likely due to patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35008 | 1220648-2023-02688 | 17556187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17556187&pc=OZD | 07/17/2023 | 02/02/2024 |  | 08/16/2023 |  |  | 07/17/2023 | 08/16/2023 |  |  | 04/30/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350289 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2023 |  |  |  |  | The user facility reported an 83-year old female with non-st elevated myocardial infarction had impella cp support ongoing for 4 days when there was persistent suction events with the device. the team made choice to explant the pump. at the time of explant the pump, there was clot burden observed. the clot was seen in the limb and the cannula had clot observed adhered to the pump. | [5 empty] | The investigation into the reported thrombosis and low pump flows has been completed since the original report was filed. the device was not returned by the medical facility. however, data logs were reviewed which revealed many suction alarms, motor current spikes, and purge pressure spikes that indicated biomaterial ingestion. the cause of the low pump flow issue most likely was the clot and biomaterial ingestion. the cause of thrombus issue was patient condition related as clot observed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ instructions for use for the related event are as follows ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35009 | 1220648-2023-02689 | 17556197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17556197&pc=OZD | 07/17/2023 | 12/12/2023 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  |  |  | 08/16/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1409579 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | The user facility in japan had an impella 5.5 support ongoing for a patient with cardiomyopathy. on day 37 of the support the automated impella controller (aic) failed with "purge line click-on not detected alarm" the team attempted to troubleshoot by replacing the purge cassette, but when the cassette replacement did not remedy the issue the team replaced the aic itself. this allowed for one day more of support until the patient was weaned and explanted. no harm or injury noted. on 20-jul-2023, additional information was provided from the user facility. "visual inspection of the subject control unit revealed that the transmitter was not recognized when attached because the lever was broken." | [5 empty] | The investigation into the reported aic purge disc/flag issues has been completed since the original report was filed. console logs from the reported day of event are consistent with complaint and show that after 37 days of uninterrupted support, console suddenly presented with purge disk not detected alarm (#402) that did not resolve after purge cassette reinsertion or purge cassette swap. inspection of the console performed at japan field service confirmed broken (missing) purge disk detect flag, which explained the alarm observed during a case. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35010 | 1220648-2023-02690 | 17557712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17557712&pc=OZD | 07/18/2023 | 02/05/2024 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  | 03/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347229 |  | 048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump experienced a drop in flow during impella support. it was noted that the measured flow would not exceed 0.5 l/min while attempting to obtain p-6 support levels. as positioning verification and clot searches were inconclusive under tee, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. data logs confirmed the failure mode. data showed low flow occurred when pump started on p-9 that triggered auto suction response & suction alarms. pump then was only able to run on p-6 on auto mode during support. about 241 hours into the case, p level was turned down to p-1 for cabg then resume to p-6 after cabg, but flow was only 0.5 l/min. pump then was disconnected from the console. the root cause of the low flow was unable to be determined as no product was returned for review.
35011 | 1220648-2023-02692 | 17557859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17557859&pc=OZD | 07/18/2023 | 09/21/2023 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  | 12/31/2024 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309853 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 87-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. however, patient has horrible peripheral vascular disease (pvd) and the sheath was occlusive in the artery. the physician attempted an external fem-fem bypass and was unable to find a suitable vessel for the return sheath and the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was patient condition related to pvd vessel condition.
35012 | 1220648-2023-02693 | 17558053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558053&pc=OZD | 07/18/2023 | 07/01/2024 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  | 12/31/2024 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293186 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Unfold or Unwrap | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. upon delivery of the pump there was a pump-guidewire interaction. the guidewire removal would pull the pump back as well. the team observed that the wire had become knotted. the team removed the damaged wire and pump. the product was replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - peripheral vessels) and the product damage (cannula kink) issues has been completed since the original report was filed. the pump was returned for investigation. product evaluation confirmed the presence of guidewire tip damage and a catheter kink on the pump. the cause of the guidewire product damage was use related. the cause of the cannula kink product damage was a catheter kink.
35013 | 1220648-2023-02694 | 17558056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558056&pc=OZD | 07/17/2023 | 01/10/2024 |  | 08/16/2023 |  |  | 07/17/2023 | 08/16/2023 |  |  | 05/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354418 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2023 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during attempts to place the perclose for preclosure, there were multiple failed attempts attributed to peripheral vascular disease. there were at least two failed attempts after only one successful deployment. while the impella implant was completed, the next physician started the ep ablation procedure by crossing the aortic valve in a retrograde fashion, and during this, the team believed that there was a left ventricle perforation from the ablation catheter. the patient experienced some instability and required a pericardial centisis. the stability of the patient was questionable with the hemoglobin drop. prior to leaving the cath lab, there was also a notable hematoma from the impella site. this was addressed with a femstop with moderate pressure. the patient finally began to stabilize, but required multiple high dose drips with no significant drainage at the pericardial site, and the hematoma site was under control. it was after the ct that the patients impella site began bleeding again, and the call was made to explant impella and explore failed perclose and impella site. that evening the device was weaned and removed by surgical closure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the delivery issues- anatomy, the ventricle perforation, and the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. per the provided clinical information and doctors comments, the root cause of the delivery issues was due to the patient's peripheral vascular disease, the root cause of the ventricle perforation issue was use related which happened by ablation catheter during the ep procedure while crossing the aortic valve, and the root cause of the access site bleeding issue was use related to unsuccessful closing of the access site using failed preclose.
35014 | 1220648-2023-02696 | 17558148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558148&pc=OZD | 07/18/2023 | 09/18/2023 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  |  |  | 08/16/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1690016 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | Abiomed field service representative reported the user facility had a console in storage at the hospital when the automated impella controller (aic) was pulled to setup impella connect. upon powering up, the aic had a battery failure, a red alarm. the team allowed for a 45 minute charge but still the aic battery failed. the ifu was followed and aic removed from use, a backup console will be used for patient use. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus has been completed since the original report was filed. the product was returned; the benchtop analysis and the clinical report evaluation revealed that the root cause of the battery alarm failures in the field was likely due to a momentarily disengaged transport connection switch (use error).
35015 | 1220648-2023-02697 | 17558169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558169&pc=OZD | 07/19/2023 | 01/16/2024 |  | 08/16/2023 |  |  | 07/19/2023 | 08/16/2023 |  |  | 04/20/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352834 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male patient experienced access site bleeding, leading to removal of pump. impella placed pre-procedure for vt ablation via rfem successfully. vt ablation performed. throughout patient required total of 10 defib. successful ablation. patient hemodynamically stable post ablation, peel away sheath removed and repo sheath positioned well. after approximately 5 min access site began to bleed, attempted angle matching with no decrease in bleeding. decision made to explant impella in procedure room with manual hold and protamine reversal d/t act >300. the impella weaned and explanted successfully and manual pressure held for hemostasis for 40min. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. as per clinical, access site began to bleed and decision made to explant impella due to act values higher than 300. manual pressure was held for 40 mins to achieve hemostasis. the cause of the access site bleeding issue was anticoagulation management related due to high patient act values (>300) per clinical review.
35016 | 1220648-2023-02698 | 17558216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558216&pc=OZD | 04/03/2021 | 01/19/2024 |  | 08/16/2023 |  |  | 04/03/2021 | 08/16/2023 |  |  | 10/31/2022 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021018094 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2021 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient had a bleeding complication in the chest after surgery and impella insertion. bleeding was hard to control. the purge solution was then changed to bicarb per abiomed protocol. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu, bleeding is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical info provided, the root cause of the access site bleeding was unable to be determined as no product and insufficient clinical details were provided.
35017 | 1220648-2023-02699 | 17558391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558391&pc=OZD | 04/07/2021 | 09/28/2023 |  | 08/16/2023 |  |  | 04/07/2021 | 08/16/2023 |  |  | 02/28/2023 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022052207 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Air/Gas in Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2021 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the pump assembly sidearm had a fluid leak prompting a sidearm bypass. a few days later, a simulation of a purge pressure alarm was attempted on the console by kinking the purge tubing. after releasing the kink the console alarmed air in purge system. after multiple attempts to de-air the alarm would not clear. cassette change performed which resolved the alarm. purge bypass is in place and advised to monitor closely as the time we were changing the cassette there was no purge pressure and blood could have backed up into motor. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the purge leak has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. per the provided clinical information, the root cause of the purge leak was a leak from the sidearm, but the exact location of the leak was unknown. hence, the root cause is undetermined since the problem cannot be evaluated without the product.
35018 | 1220648-2023-02700 | 17558512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558512&pc=OZD | 04/12/2021 | 03/10/2025 |  | 08/16/2023 |  |  | 04/12/2021 | 08/16/2023 |  |  | 11/30/2022 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021020019 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2021 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. site has been bleeding so plan is to change to the other arm tomorrow. pt became septic after 3 mos. surgeon wanted to pulled the device and placed new at a new site. device explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
35019 | 1220648-2023-02701 | 17558584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558584&pc=OZD | 07/19/2023 | 01/22/2024 |  | 08/16/2023 |  |  | 07/19/2023 | 08/16/2023 |  |  | 01/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024307825 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old female patient had known peripheral arterial disease (pad) and despite placing a functioning reperfusion catheter down the right leg there was no pulse to the right foot. patent was taken to the cath lab and angiography performed down antegrade sheath. angio revealed 100% popliteal occlusion with small collaterals filling tibeal territory. no flow was present in the foot. patient was requiring additional cardiac support, so va ecmo was placed on left side, and the team ran into a similar problem with distal limb perfusion. the left leg had a 100% occlusion in the mid-superficial femoral artery (sfa) and the only reperfusion catheter that could be placed was in the left profunda artery in hopes that it would help supply collaterals to where the sfa reconstituted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the limb ischemia issue has been completed. the device was not returned for investigation. per the provided clinical information, the cause of the limb ischemia issue was patient condition related to pad.
35020 | 1220648-2023-02702 | 17558603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17558603&pc=OZD | 07/18/2023 | 11/14/2025 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  |  |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420044 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported an 86-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the team lifted up the power cord for a ventilator to pass underneath. the cord came out of the automated impella controller (aic) plug in the outlet. the impella pump and purge system were switched over to a new aic without issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35021 | 1220648-2023-02703 | 17559017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17559017&pc=OZD | 07/19/2023 | 09/22/2024 |  | 08/16/2023 |  |  | 07/19/2023 | 08/16/2023 |  |  | 02/28/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347725 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 61-year-old male patient on impella device experienced some flow inconsistencies which the medical staff was unable to resolve by repositioning the device. decision was made to explant the pump and insert a new one. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02703 was provided in section h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the positioning issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, data logs show the spike in motor current followed by the low pump flow. also, several motor current spikes were reported at the end of the case while running on low p-level. the cause of the low pump flow issue was most likely biomaterial.
35022 | 1220648-2023-02704 | 17559359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17559359&pc=OZD | 07/12/2023 | 09/13/2023 |  | 08/16/2023 |  |  | 07/12/2023 | 08/16/2023 |  |  | 05/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355507 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 52-year-old female patient had a impella 5.5 support. the impella 5.5 was successful for 6.7 days. there was an observation made of thrombus at the inlet of the pump at explant and on echo. there was no noted harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported thrombus has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. data review analysis showed about 9 hours into the case, the suction alarms were triggered with low flow. based on clinical description and data review, the root cause of thrombus was most likely due to patient condition.
35023 | 1220648-2023-02705 | 17561037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561037&pc=OZD | 07/18/2023 | 02/02/2024 |  | 08/16/2023 |  |  | 07/18/2023 | 08/16/2023 |  |  | 03/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338632 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, it was suspected that the patient had developed hemolysis. as plasma free hemoglobin testing was not available at the facility, hemolysis could not be definitively confirmed. the pump was subsequently weaned and explanted as a result of the noted condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis issue has been completed. the data logs were returned for evaluation only. review of the logs did not show any motor current spikes or any other abnormalities which gives concrete evidence. the root cause of the hemolysis issue was not determined.
35024 | 1220648-2023-02706 | 17561323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561323&pc=OZD | 07/17/2023 | 09/17/2024 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 03/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 54-year-old male patient was implanted with an impella cp device for mechanical circulatory support. the impella was upgraded from a cp to a 5.5 due to the patient going into cardiogenic shock (cgs) and needing full support. the patient was also hemolyzing. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported hemolysis has been completed since the original report was filed. the device was not returned by the medical facility. data review: data logs showed pump was in proper position & run without issues for entire case. product was not returned for review. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. it was unknown relation to impella because it was reported that the injury was not attributed to impella despite occurring during support. instructions for use states: "¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
35025 | 1220648-2023-02707 | 17561334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561334&pc=OZD | 07/17/2023 | 03/08/2024 |  | 08/16/2023 |  |  | 07/17/2023 | 08/16/2023 |  |  | 03/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346252 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during cross clamp the impella stopped. it was presumed that the clamp was close to the motor. the impella started without difficulty when coming off pump and clamp was removed. the patient was hemodynamically stable. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not returned for investigation. data logs were reviewed which showed the impella stop 115-motor current high alarms with intermediate loss of placement signal after 5 hours of case start. according to clinical records, a pump stop was reported after clamping the device, and the issue was resolved after unclamping it. the cause of the pump stop issue was use related since the problem occurred while the device was clamped, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35026 | 1220648-2023-02708 | 17561338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561338&pc=OZD | 07/17/2023 | 01/24/2024 |  | 08/16/2023 |  |  | 07/17/2023 | 08/16/2023 |  |  | 04/30/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351086 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp failed to start. the cannula was completely kinked. the pump was removed and replaced with a new impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure. which triggers the ¿purge pressure high¿ alarm message. could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into cannula kink has been completed. the product was not returned for evaluation. the data logs showed low flow occurred when pump started & remained to the rest of the case. the root cause of low flow was most likely due to cannula kink, technique/use.
35027 | 1220648-2023-02709 | 17561346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561346&pc=OZD | 07/17/2023 | 01/24/2024 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 07/31/2024 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2023 |  |  |  |  | The user facility reported complainant reported male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was intubated and sedated showing signs of hemolysis, dark red urine, and several hemolyzed lab samples. the physician attempted to reposition the impella at bedside, but it was unsuccessful. the patient was taken to the cardiac catheterization lab (ccl) for another attempt. the impella waveforms and flows are appropriate, and the outcome was successful. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported hemolysis has been completed since the original report was filed. the device was not returned by the medical facility. data logs reviewed: many impella position unknown alarms occurred that indicated to improper impella positioning during support. the cause of hemolysis issue was improper positioning related due to improper use technique.
35028 | 1220648-2023-02710 | 17561353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561353&pc=OZD | 07/19/2023 | 01/18/2024 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 01/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326485 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male patient experienced impella access site bleed that prompted removal of the pump during support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the access site bleeding issue was most likely patient condition related.
35029 | 1220648-2023-02711 | 17561358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561358&pc=OZD | 07/18/2023 | 02/06/2024 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 11/30/2024 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278015 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported a 69year-old female in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. during attempted manipulation of the impella pump's position, a ventricular rupture occurred. intervention was required in the form of plegeted sutures to repair the injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the reported ventricle perforation has been completed since the original report was filed. no product was returned. as per clinical imaging revealed suboptimal placement of 5.5 pump appearing to contact mv. multiple attempts to torque cannula clockwise were unsuccessful resulting in lv lateral wall rupture. data logs were reviewed, and no positioning issue alarms were observed. the cause of the ventricle perforation issue was most likely patient condition related per clinical review.
35030 | 1220648-2023-02712 | 17561375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561375&pc=OZD | 07/19/2023 | 01/18/2024 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 03/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344856 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported concern for ischemia in the patient's impella inserted limb. it's unclear if any intervention was required, however, we were informed that a fellow from the hospital had consulted the patient for this issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient had right lower extremity ischemia, and introducer was not removed from body. the root cause of the limb ischemia issue was use related due to introducer left in place and not peeled away leading to no pulses on right leg.
35031 | 1220648-2023-02713 | 17561379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561379&pc=OZD | 07/18/2023 | 01/19/2024 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 03/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347270 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. due to patient's anatomy, the physician was unable to advance the impella's repositioning sheath during an attempted delivery of the pump. the decision was made to remove the pump. additionally, bleeding from the access site could not be controlled, therefore, a balloon tamponade was required after device removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the major access site bleeding has been completed since the original report was filed. per the clinical information provided, the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical information provided.
35032 | 1220648-2023-02714 | 17561381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561381&pc=OZD | 07/18/2023 | 09/12/2023 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 03/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024319988 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2023 |  |  |  |  | The user facility reported a 61-year-old male with a history of heart failure was implanted with an impella 5.5 device for mechanical circulatory support. the patient required an embolectomy in the impella-inserted limb during support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the thrombus issue was not determined since product was not returned.
35033 | 1220648-2023-02715 | 17561383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561383&pc=OZD | 04/13/2021 | 09/28/2023 |  | 08/16/2023 |  |  | 04/13/2021 | 08/16/2023 |  |  | 01/31/2023 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022040605 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/13/2021 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. an impella position in aorta alarm occurred during support. while repositioning of the pump, it became dislodged completely in the aorta. doctor attempted (unsuccessfully) to cross the valve under transesophageal echocardiography guidance. due to patient having transesophageal echocardiography in 70¿s on low dose milrinone and levo, decision was made for surgical removal of impella and repair at bedside. | [5 empty] | The investigation into the positioning issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical details and data logs, the root cause of the positioning issue was use related since the pump pulled out during repositioning. wireless re-access was unsuccessful. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
35034 | 1220648-2023-02716 | 17561406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561406&pc=OZD | 07/18/2023 | 01/17/2024 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 01/31/2025 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326907 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male patient had an impella site hematoma developing during support. in addition to applying a femostop, 1 unit of packed red blood cells were delivered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that the bleeding occurred after insertion. therefore, the root cause of the access site bleed was due to patient condition.
35035 | 1220648-2023-02717 | 17561414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561414&pc=OZD | 05/14/2021 | 03/10/2025 |  | 08/16/2023 |  |  |  | 08/16/2023 |  |  | 12/31/2022 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022030501 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2021 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support was found by echo to have a thrombus on the impella inlet. impella was weaned and removed. in addition, impella was contaminated on the back table and a new impella was opened. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Thrombus - no product was returned. the root cause of the thrombus issue was most likely patient condition related. after the impella 5.5 was removed, it was contaminated on the back table and cause the sterile sleeve damage to use it and they have opened an new one to continue. it was just a user mistake and it have no failure mode that was related to the pump. therefore, the root cause was determined to be use related.
35036 | 1220648-2023-02718 | 17561449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561449&pc=OZD | 04/06/2021 | 01/23/2024 |  | 08/16/2023 |  |  | 04/06/2021 | 08/16/2023 |  |  | 02/28/2023 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022049826 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2021 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient returned to or today for coronary artery bypass grafting x3 and mitral ring. impella was in a posterior position toward mitral valve. when the md attempted to reposition, the catheter would go toward the apex and then turn abruptly back to the mitral valve. another md tried to help transition the torque but repositioning was unsuccessful. md decided to pull the pump back (not in optimal position) and run at p5. | [5 empty] | The investigation into the positioning issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the cause of the positioning issue was determined to be use issue as no product malfunction had occurred. patient had no harm, and the outcome of the device was successful. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
35037 | 1220648-2023-02719 | 17561469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561469&pc=OZD | 04/12/2021 | 09/26/2023 |  | 08/16/2023 |  |  | 04/12/2021 | 08/16/2023 |  |  | 02/28/2023 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022049813 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/12/2021 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient became acutely agitated, thrashing in bed, grabbing & pulling the impella. impella console noted with red alarm: impella position in aorta, dual arterial waveforms & flat motor current. impella was decreased to p2; picture received with locking ring/tuohy lock pulled away from the suture hub. doctor immediately notified and requested to verify device position due to direct ascending aorta insertion with subxyphoid exit from the body. the patient¿s blood pressure was 148/57/78 on no impella or vasopressor support. noted with impella motor & outlet in the graft & inlet remaining in the aorta. md decided to have rn retract impella, with aid impella completely back into the graft and turned off. device explanted. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. per the provided clinical information, the root cause of the positioning issues was determined to movement of the patient. the outcome of the device was unsuccessful. patient was taken to the or and the device was surgically explanted without issue. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
35038 | 1220648-2023-02720 | 17561479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561479&pc=OZD | 10/28/2021 | 10/27/2025 |  | 08/16/2023 |  |  | 10/28/2021 | 08/16/2023 |  |  | 06/30/2023 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022105508 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2021 |  |  |  |  | The user facility reported a 56-year-old male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. some oozing noted at insertion site. clot mobilization to the right upper extremity post impella removal. recommended vessel loops were not used during removal. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. this report is being filed as part of a retrospective review of historical records. | Thrombus: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. bleeding: the root cause of the access site bleeding minor was determined to be patient condition as evidenced by oozing beginning after movement d6a/d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated.
35039 | 1220648-2023-02722 | 17561480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561480&pc=OZD | 09/16/2021 | 09/17/2024 |  | 08/16/2023 |  |  | 09/16/2021 | 08/16/2023 |  |  | 09/30/2022 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1505056 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2021 |  |  |  |  | The user facility reported delivery issued with impella 5.5 device. some difficulty crossing atrioventricular (av). once across av and pump turned on, unable to achieve flows 3lpm when above p4 when extracorporeal membrane oxygenation was being weaned. pump then removed for better positioning. upon removal of pump, noted kinked catheter behind motor housing. decision was made to remove pump and open a new 5.5 for implant instead. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02722 was provided.
35040 | 1220648-2023-02721 | 17561482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561482&pc=OZD | 09/25/2021 | 01/11/2024 |  | 08/16/2023 |  |  | 09/25/2021 | 08/16/2023 |  |  | 05/31/2023 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022091947 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 09/25/2021 |  |  |  |  | The user facility reported a 74-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. during support, the device had a pump flow issue. suction occurred whenever the team went above p6 in the settings. multiple drips were required. the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported positioning issues has been completed since the original report was filed. the root cause of the positioning issue was most likely use related as the pump was in a suboptimal position leading to low flow. correction: section h6 :the health effect- impact code was inadvertently left out in the initial report which is been added to this report.
35041 | 1220648-2023-02723 | 17561483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561483&pc=OZD | 11/10/2021 | 09/18/2024 |  | 08/16/2023 |  |  | 11/10/2021 | 08/16/2023 |  |  | 09/30/2022 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1501813 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2021 |  |  |  |  | The user facility reported an 72-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. during support, patient suffered ventricular tachycardia (vt) / ventricular fibrillation (vf) requiring defibrillation x4 no cpr was done, medications were administered and device was adjusted. patient is mildly sedated on precedex following commands. echo has been reviewed by both physicians, explained concerns of impella causing the vt/vf as it is interacting with posterior wall and deep: advised to reposition. both physicians stated at this point will reduce the adrenergic stimulus from current medications and reassess impella position in am. will maintain adequate filling pressure to prevent a suck-down event which could trigger a malignant arrhythmia. a third physician reviewed echo and electrocardiogram tracing; he stated he doesn't think the impella is causing the vt/vf given the premature ventricular contractions morphology before the vt/vf episodes patient had, however impella as the cause can't be fully ruled out. shortly after initiation of support impella noted to be interaction with posterior wall (statements above), eventually needing to be rewired and reinserted. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
35042 | 1220648-2023-02724 | 17561485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17561485&pc=OZD | 11/09/2021 | 10/30/2025 |  | 08/16/2023 |  |  | 11/09/2021 | 08/16/2023 |  |  | 09/09/2023 |  | 08/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022117759 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 11/09/2021 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. could not get impella to cross aortic valve, so pigtail and j wire were then placed. it was very challenging to cross the valve as well. finally did cross valve and exchanged to impella wire. the 5.5 was inserted and still had difficulty crossing valve: transesophageal echocardiography imaging was sub-optimal and visualization of wire was compromised. as further manipulation was attempted, patient had ventricular fibrillation arrest. defib x4, then cardiac massage. decision to abort impella attempt(s). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2023-02724 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-02724 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-02724 in accordance with updated procedures. h6 health effect - impact code 4647 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-02724. h6 medical device problem code 2993 and component code 3038 was erroneously entered on the initial manufacturer device report number 1220648-2023-02724.
35043 | 17547970 | 17547970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17547970&pc=OZD | 08/05/2023 | 08/10/2023 | 08/10/2023 | 08/15/2023 | 08/10/2023 | 08/15/2023 |  | 08/15/2023 |  |  |  |  | 08/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  | 2024364436 |  | 1000100 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [8 empty] | Unknown |  |  |  | Death | Yes | Yes | [7 empty] | Cardiac Arrest | Loss of consciousness | [11 empty] | Death | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella implanted, approximately a month later the patient reported feeling uneasy 2130, repositioned per request after sitting on side of bed for comfort. after patient repositioned, the impella device stopped working (@2154), abiomed (device company) was called as well as the md called to bedside. attempted restarting impella with low flow rates multiple times per abiomed request, however would not start. patient began to feel more uneasy. patient lost consciousness w/ pulseless electrical activity (pea) arrest.
35044 | 1220648-2023-02680 | 17549419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17549419&pc=OZD | 07/16/2023 | 01/16/2024 |  | 08/15/2023 |  |  | 07/16/2023 | 08/15/2023 |  |  | 04/30/2025 |  | 08/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337405 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2023 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. during support, there was a very slow purge filter leak from the device. the leak was contained and no intervention was required. | [5 empty] | The investigation into the purge sidearm leak has been completed. no product was returned for evaluation. as per clinical, a slow leak was observed from the purge filter. the exact location of the leak associated to the sidearm filter is unknown. the cause of the purge leak issue was most likely a leak from the sidearm, but the exact location of the leak is unknown since the product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
35045 | 1220648-2023-02681 | 17550746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17550746&pc=OZD | 07/18/2023 | 02/08/2024 |  | 08/15/2023 |  |  | 07/18/2023 | 08/15/2023 |  |  | 10/31/2024 |  | 08/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269285 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while driving cpsa with ecpella, the icu determined that the pump position was deep and pulled it out using tte. the doctor pulled it out while confirming the tte, and entered the aorta with the shaft and cannula bent. it was decided to drive p1 in the descending aorta and remove it. in hybrid or, the extracardiac doctor retracted the femoral insertion site, and the circulatory doctor inserted the sheath through the brachial and inserted the snare catheter. when the tip of the cpsa was pulled while it was held with a snare, the flexion was released, but the descending aorta ruptured. the doctor stopped the bleeding with a balloon, reversed protamine, administered blood transfusion, and managed to maintain hemodynamics with ecmo support. the previous doctor had suspected sarcoidosis and had administered steroids, so it is possible that the vascular wall was fragile. it was recognized that re-insertion of the pump was prohibited when it was pulled out from the lv to the pigtail, except for the physician who performed the position adjustment. | [5 empty] | The investigation into the ventricle perforation and the positioning issues has been completed since the original report was filed. the product was returned for investigation. the device was visually inspected and found that the cannula was kinked/bent, and the catheter was kinked at multiple locations due to excessive force. the sterile sleeve was found to be damaged during support. no other abnormalities were found. as per clinical review, the impella was placed deep and was tried to pull back and slipped out of left ventricle with the cannula being bent. the pigtail was pulled while it was held with a snare. the flexion was released, and the descending aorta ruptured. the cause of the vascular damage - great vessel issue was use related based on field investigation and clinical review. the cause of the positioning issue was use related based on field investigation and clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35046 | 1220648-2023-02682 | 17552572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17552572&pc=OZD | 07/17/2023 | 09/26/2023 |  | 08/15/2023 |  |  | 07/17/2023 | 08/15/2023 |  |  |  |  | 08/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1216754 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/17/2023 |  |  |  |  | The user facility reported a 73-year-old male was on impella support for percutaneous cardiac insertion. during setup of the device, the automated impella controller (aic) failed as a purge disk holder was found to be broken. the team followed the instructions for use and a backup aic was used. the backup did not have the same failure. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the product was returned for investigation. per the provided clinical information, the root cause of the aic - purge disc/flag issues was a broken purge pressure sensor retainer. the failure mode will be monitored and trended. | The device has been received for investigation and repair. a supplemental mdr will be submitted when the investigation is complete.
35047 | 1220648-2023-02683 | 17552620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17552620&pc=OZD | 07/18/2023 | 10/31/2025 |  | 08/15/2023 |  |  | 07/18/2023 | 08/15/2023 |  |  | 01/31/2025 |  | 08/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316568 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 82-year-old male patient had a impella cp support ongoing when the team had to replace the aic due to a controller error and ps loss. there was no harm or injury from the replacement, the team followed the instructions for use (ifu) and placed the pump on a 2nd aic. the 2nd aic additionally had purge issues and so the team utilized a 3rd console. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and confirm the reported failure mode given there were ¿algs suspended opm crc error ¿events (1045), and a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. console ic3515 was sent back fs for further investigation. the controller error and loss of placement signal were due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated. d4 updated. d6a / d6b updated. g1 reporting contact email and reporting contact fax number updated. h4 device manufacture date updated.
35048 | 1220648-2023-02684 | 17552668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17552668&pc=OZD | 07/16/2023 | 09/11/2023 |  | 08/15/2023 |  |  | 07/16/2023 | 08/15/2023 |  |  | 01/31/2025 |  | 08/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024307950 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2023 |  |  |  |  | The user facility reported a 63-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had limb circulation issues and the impella was removed. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was most likely patient condition related since issue was resolved after pump explant. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35049 | 1220648-2023-02685 | 17552739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17552739&pc=OZD | 07/16/2023 | 09/17/2024 |  | 08/15/2023 |  |  | 07/16/2023 | 08/15/2023 |  |  | 05/31/2025 |  | 08/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352097 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Other | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/16/2023 |  |  |  |  | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 74-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient endured continuous suction events with suspected hemolysis. the decision was made to refrain from repositioning, due to patient's worsening status. |  |  |  |  | The investigation for the reported low pump flow and hemolysis has been completed. the device was not returned for evaluation. data logs were reviewed which revealed several instances of suction conditions and low pump flow. the downward placement signal trend was reported without affecting optical signal parameters. no motor current spike or improper position-related alarms were reported during the case. according to clinical records, suction alarm was temporarily resolved after administering 1l of normal saline. the cause of the low pump flow issue was most likely patient condition related based on data log review. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35050 | 1220648-2023-02686 | 17552748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17552748&pc=PYX | 07/16/2023 | 08/02/2024 |  | 08/15/2023 |  |  | 07/16/2023 | 08/15/2023 |  |  | 02/28/2025 |  | 08/15/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024333389 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2023 |  |  |  |  | The user facility reported a 55-year-old female in st elevated myocardial infarction was implanted with an impella rp flex device and an impella cp for mechanical circulatory support. . the patient went into ventricular fibrillation/ventricular tachycardia multiple times while trying to advance the swan for wire placement. the flex was advanced but kept pushing the wire out of the pulmonary artery. the impella rp flex was aborted. the impella cp was implanted without issue. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. per medical safety officer review use of the device cannot conclusively be associated with the outcome of death. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
35051 | 1220648-2023-02667 | 17540600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17540600&pc=OZD | 07/17/2023 | 09/05/2023 |  | 08/14/2023 |  |  | 07/17/2023 | 08/14/2023 |  |  | 12/31/2024 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024305289 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | As reported, during the procedure the aic driving the impella cp pump reported multiple faulty impella systems yellow alarm (no pump detected). there was no reported patient harm. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the device was returned and investigated. the pump functioned/operated to specification. there was no issue found during the case. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35052 | 1220648-2023-02668 | 17545749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17545749&pc=OZD | 07/11/2023 | 03/01/2024 |  | 08/14/2023 |  |  | 07/17/2023 | 08/14/2023 |  |  | 04/30/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337449 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male patient, on impella support, was returned to the intensive care unit (icu) extubated and doing well, but showed signs of stroke soon after. patient was taken to the cardiovascular catheterization lab (ccl) for neuro intervention where ischemic infarct found. extent of intervention unknown, but patient was still intubated and on ventilator support. the pump was subsequently explanted. | [5 empty] | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35053 | 1220648-2023-02669 | 17545835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17545835&pc=OZD | 07/17/2023 | 10/14/2024 |  | 08/14/2023 |  |  | 07/17/2023 | 08/14/2023 |  |  | 03/31/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344844 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient exhibited bleeding from all sites and heparin was withheld until oozing subsided. | The user facility reported a 47-year-old female patient on impella support was unable to move hand/fingers and arm post implant. arterial flow assessed and low/no flow visualized to lower upper extremity. the physician opened the impella access site and manually pulled back repositioning sheath within graft to ensure sheath was not entering axillary artery. minimal improvement to flow reported. on exam, the patient showed to have very small vessels throughout and nothing further to be done at this time, improvement in passive range of motion reported. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the given clinical information, the root cause of the ischemia issue was patient condition related to diseased/ small vessels. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35054 | 1220648-2023-02670 | 17545996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17545996&pc=OZD | 07/17/2023 | 02/02/2024 |  | 08/14/2023 |  |  | 07/17/2023 | 08/14/2023 |  |  | 04/30/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349659 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 83-year-old female had a drop in pump flow and a spike in motor current 15 minutes after impella implant. the flow rates were measured to be as low as 0.5 l/min and suction was encountered at support levels above p-1. the pump was subsequently explanted and a clot was discovered around the impeller blades. | [5 empty] | The investigation into the low pump flow has been completed. the cause of the low pump flow most likely was biomaterial (thrombus) ingested. complaint device not returned for investigation. data log was reviewed; flow dropped then motor current spike that indicated to biomaterial was ingested by the pump. h1, b1-updated to reflect malfunction. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35055 | 1220648-2023-02671 | 17546198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546198&pc=OZD | 07/17/2023 | 01/25/2024 |  | 08/14/2023 |  |  | 07/17/2023 | 08/14/2023 |  |  |  |  | 08/14/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | IMPELLA CONTROLLER EU |  |  | OZD |  |  | IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1282817 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a ¿replace the console¿ error showed in the aic. the device was removed from service as a result of the noted issue. there was no reported patient harm. | [5 empty] | Updated sections d9, g3, g6, h2, h3 & h6: (h3) the investigation into the reported system automated impella controller (aic) - self-check error and purge issue have been completed since the original report was filed. product was returned for investigation. the device was visually inspected, and the console was tested after arriving at the manufacturer. the complaint "the front door is broken" turned out to be a mistranslation and the part that was cracked is the upper right front of the console. the failure mode was reproduced when the console was turned on and immediately got a system self-check fail. the super cap on the pbm was confirmed to have corroded the pbm communication trace next to it (ic4987 pbm front corruption.png and ic4897 pbm back corruption). after the pbm was replaced, the system self-check failure went away but the batteries were no longer charging. replacing the batteries alleviated this issue. during data analysis the automated impella controller (aic) logs for ic4987 show erroneous pbm1 communication that started on 4jul2023 and remained consistent throughout the case. eventually, the pbm1 communication crashed causing the "replace the console" message. as per clinical investigation, a ¿replace the console¿ error showed in in the aic. the front door was also noted to be partially broken. there was no reported patient harm. the root cause of the automated impella controller (aic) issue was contamination of a communication line on the pbm. | The impella device has been received and an evaluation/investigation is ongoing. upon investigation closure, a supplemental mdr will be filed.
35056 | 1220648-2023-02672 | 17546210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546210&pc=OZD | 07/15/2023 | 01/19/2024 |  | 08/14/2023 |  |  | 07/15/2023 | 08/14/2023 |  |  | 04/30/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332840 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male patient experienced several ventricular tachycardia associated with the impella during impella 5.5 support. the impella was repositioned four times under echo. ultimately, a decision was made to explant the impella. | [5 empty] | The investigation into the vf/vt/af/at issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the field lt log was reviewed, and motor current was pulsatile, so no positioning issues were confirmed. the root cause of the ventricle tachycardia was most likely due to patient condition. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35057 | 1220648-2023-02673 | 17546221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546221&pc=OZD | 07/07/2023 | 10/01/2024 |  | 08/14/2023 |  |  | 07/07/2023 | 08/14/2023 |  |  |  |  | 08/14/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1545370 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | Abiomed field service representative reported that during a presentation of demo pump impella cp sa, the automated impella controller (aic) was shutting down. restarts and shows error message to use a different controller (red alarm). no patient was harmed, but aic could have been used as a loaner for patient cases. eventually, after restart demo was continued and no error shows in the following hours. | [5 empty] | The investigation for the console (aic) controller failure-red has been completed. data logs from the date of occurrence confirm unexpected shutdown alarm. the console was returned for evaluation. while in the lab under investigation, console was power cycled 10 times while a similar cp type was run for 5 minutes in between each cycle without any observable anomaly. unrelated to the complaint; during the functional check of the console, cell imbalance was observed within battery 1 with differences in cell voltages not up-to specification. replacing the batteries showed that cell voltage difference where then within specification. after battery replacement, emergency shutdown of the console was utilized due to regular shutdown sequence not being available. the root cause for the unexpected controller shutdown could not be determined since the issue could not be reproduced from investigation. added- d9 device return date d2-corrected common device name d4-corrected model number e2-changed to no f6/f8 should have been lleft blank. h6 med dev prod code 4042 changed to 2880. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35058 | 1220648-2023-02674 | 17546239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546239&pc=OZD | 07/16/2023 | 08/14/2023 |  | 08/14/2023 |  |  | 07/16/2023 | 08/14/2023 |  |  | 05/31/2024 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220154 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the first implant was aborted by the physician during the pump insertion into the sheath. the physician noticed that the easy guide lumen was still in place and removed the pump completely. afterwards physician had concerns to use this same pump (physician was not sure if the pump was damaged during the procedure) and decided to implant another new pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should additional relevant information become available, a supplemental report will be filed. (h3) the investigation into the reported event of the red lumen removal issue has been completed since the original report was filed. the physician explained that he had not removed the red easy guide lumen when implanting the impella but had advanced the impella into the airlock as it was. unfortunately, he cannot get the impella to retract and asks if the easy guide lumen must be removed. he was strongly advised to remove the lumen before advancing the pump further. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
35059 | 1220648-2023-02675 | 17546297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546297&pc=OZD | 07/15/2023 | 11/19/2025 |  | 08/14/2023 |  |  | 07/15/2023 | 08/14/2023 |  |  |  |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/15/2023 |  |  |  |  | The user facility reported the purge pressure was getting low every ten minutes on the automated impella controller (aic). purge fluid went up to 30 and then went back to the regular rate that is has been infusion (10cc/hr). there were no changes in patient or patient position. the cassette was changed and the issue was still ongoing. patient details are not available at this time. the operator was going to swap to a new aic. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35060 | 1220648-2023-02676 | 17546304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546304&pc=OZD | 07/15/2023 | 09/17/2024 |  | 08/14/2023 |  |  | 07/15/2023 | 08/14/2023 |  |  | 05/31/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356198 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/15/2023 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 50-year-old female was escalated from an impella cp to an impella 5.5. while transferring patient to bed, the patient went into ventricular fibrillation. compressions started, direct current cardioversion converted into ns 90¿s, amiobolus given, dobutamine, global thrombosis test started. suction event, p5 with volume. transesophageal echocardiography (tee) revealed underfilled left ventricle w/ slightly dilated right ventricle. after volume p level increased to p8, improvements noted on tee. pt taken to cvicu 10. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | The investigation for the reported ventricular fibrillation (vf) has been completed. the impella device nor flouroscopic imaging and data logs were returned for evaluation. however, clinical information provided is sufficient to determine a root cause. the reported root cause of the vf was most likely patient condition related and unknown relation to impella. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35061 | 1220648-2023-02677 | 17546439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546439&pc=OZD | 07/15/2023 | 03/08/2024 |  | 08/14/2023 |  |  | 07/15/2023 | 08/14/2023 |  |  | 01/31/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP APAC |  |  | ABIOMED, INC. |  |  | US |  | 2024333489 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2023 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. since the patient's native heart contractility has improved significantly after 48+ hours of support with impella, and patient's hemoglobin was dropping because of significant hemolysis, operator has decided to slowly wean and explant impella on day three. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35062 | 1220648-2023-02678 | 17546595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546595&pc=OZD | 07/15/2023 | 08/14/2023 |  | 08/14/2023 |  |  | 07/15/2023 | 08/14/2023 |  |  | 05/31/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024356773 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/15/2023 |  |  |  |  | The complainant reported an 83 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, there was concerns of hemolysis as the patient's urine was tinged red and labs were hemolyzed. the clinician stated that urinary output (uop) had been tinged red since a foley catheter was inserted the previous day. the physician attempted to reposition the pump with unintentional retraction into the aorta. the pump was unable to advance back into ventricle, despite exhausted efforts. the decision was then made to remove the pump at bedside with manual pressure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35063 | 1220648-2023-02679 | 17546631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17546631&pc=OZD | 07/15/2023 | 08/20/2024 |  | 08/14/2023 |  |  | 07/15/2023 | 08/14/2023 |  |  | 03/31/2025 |  | 08/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2023 |  |  |  |  | Updated patient outcome. the patient did not survive the procedure. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a 50-year-old female was implanted with an impella cp for circulatory support. while on support, the patient had a large hematoma at insertion site. the team was unable to find pulses on right foot, lactate increasing and cardiac output was down trending. a subarachnoid hemorrhage found on head ct scan. neuro surgeon and ct surgeon consulted. neuro ok with heparin bolus for possible ecmo and or impella 5.5 placement. plan is for patient to go to or for escalation of care. |  |  |  |  | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information. h6 component code added. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank on manufacturer device report 1220648-2023-02679. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35064 | MW5138716 | 17531805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17531805&pc=OJE |  | 04/01/2022 |  | 08/12/2023 |  |  |  | 08/12/2023 |  |  |  |  | 08/12/2023 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA |  |  | OJE |  |  | LMPELLA DEVICE |  |  | ABIOMED | [8 empty] | No answer provided |  |  | Not Available | [6 empty] | Voluntary report |  | Other | [8 empty] | Unknown |  |  |  | Injury | No | No | Microbial Contamination of Device | [6 empty] | Abscess | Hematoma | [40 empty] | Initial submission | [9 empty] | Hematoma over impella site at right axillary, post-procedure. : exploration of right chest wall impella wound. cultures obtained. wound washed out and vac dressing applied note states r pectoralis intramuscular hematoma wo discrete abscess identified. (b)(4). this report reflects information received by fda in the form of a notification per 803.22 (b)(2).
35065 | MW5144201 | 17537322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17537322&pc=OJE |  | 04/05/2019 |  | 08/12/2023 |  |  |  | 08/12/2023 |  |  |  |  | 08/12/2023 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE | [5 empty] | ABIOMED | [8 empty] | No answer provided |  |  | Not Available | [6 empty] | Voluntary report |  | Other | [8 empty] | Unknown |  |  |  | Injury | No | No | Insufficient Information | [6 empty] | Thrombosis/Thrombus | [41 empty] | Initial submission | [9 empty] | Medtronic received information that during removal of the lmpella rp device from the hemostatic sheath in right groin, a thrombus was noted with this associated device. the thrombus was mechanically evacuated with sequelae and the device was successfully removed. no additional adverse patient effects were reported. (b)(4). this report reflects information received by fda in the form of a notification per 803.22 (b)(2).
35066 | 1220648-2023-02647 | 17510851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17510851&pc=OZD | 07/14/2023 | 01/22/2024 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 04/30/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old male post impella placement, the patient had no flow distal to sheath in right leg. staff placed antegrade sheath in right leg and retrograde sheath in left groin to perform fem-fem bypass temporarily. this was successful under fluoroscopy. all sheaths sutured in place. | [5 empty] | The investigation into limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. per the provided clinical information, patient had lack of flow over right femoral leg due to diseased/small vessels and fem-fem bypass was done from right femoral leg to left femoral leg to restore flow. the cause of the limb ischemia issue was patient condition related with patient having diseased vessels in right femoral and fem-fem bypass resolved ischemia. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35067 | 1220648-2023-02649 | 17510993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17510993&pc=OZD | 07/15/2023 | 10/21/2025 |  | 08/11/2023 |  |  | 07/15/2023 | 08/11/2023 |  |  |  |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129860 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that their aic displayed a controller failure alarm during patient support. coinciding with failure, it was noted that there was a loss in ao and lv waveforms for roughly 15 seconds; however, there was no loss in motor current. the alarm resolved, however the decision was made to replace the console. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: alarm header displayed in the left column; window is color-coded red for critical alarms, yellow for serious alarms, white for advisory notifications, gray for resolved alarms. overlaid with a translucent yellow when the controller cannot determine catheter position, position is wrong, and placement monitoring is suspended or disabled. when the impella catheter is operating in a low flow range, placement monitoring may be suspended and the flow rate in the lower left corner of the controller display screen will turn yellow to indicate that impella position is unknown.
35068 | 1220648-2023-02651 | 17511062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17511062&pc=OZD | 06/20/2023 | 10/13/2024 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 10/31/2023 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154780 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 8 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the pump experienced persistent suction issues throughout the 8 year-old male patient support. repositioning was attempted to troubleshoot the issue; however, the flows remained low. it was noted that the pump looked to be bent on fluoroscopy at the time of the reduced flow. the pump was eventually explanted as the patient's overall condition worsened. there was no harm as a result of the reported issue. | It was also reported there was concern for limb ischemia in the patient's leg distal to the impella access site. |  |  |  |  | The investigation into low pump flow has been completed. returned complaint impella cp pump was visually inspected. cannula kinks were observed near the pump inlet and the bent area. there was no biomaterial, no broken blades, and no other abnormality found in the pump. per the clinical review, the root cause of the low pump flow was patient condition related due to right ventricle dysfunction and a small left ventricle. added section d9. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop*. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
35069 | 1220648-2023-02652 | 17511493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17511493&pc=PYX | 07/14/2023 | 10/06/2025 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 02/28/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337526 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a drop in pump flow coinciding with a spike in motor current during 60 year-old female patient support. the pump was repositioned during troubleshooting; however, the flow remained low. the pump was subsequently replaced as a result of the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the low pump flow has been completed. the clinical team reported a drop in pump flow coinciding with a spike in motor current during patient support. it was noted that the patient was not on any systemic anticoagulation and suggested act values was not maintained. the pump was inspected and biomaterial was found wrapped beneath the impeller. the data logs confirm motor current spike before reported drop in flow. the root cause of the low pump flow was biomaterial. added d6a/d6b corrections to the initial mfr report: h6 med dev prob code changed to 1248.
35070 | 1220648-2023-02653 | 17511506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17511506&pc=OZD | 07/14/2023 | 03/12/2024 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 01/31/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293953 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. support was ongoing when on day three of support, the pump was explanted due to the pump having flow saturation. the patient endured ventricular tachycardia (vt), which was suspected due to the decrease in performance level of the impella. the patient was also on ecmo at the time of the vt. the 5.5 was explanted and replaced. | [5 empty] | The investigation for the reported saturated pump flow and ventricular tachycardia has been completed. the impella device was not returned for evaluation. data logs were reviewed which show a motor current spike and placement signal spike followed by saturated pump flow. the cause of the saturated flow issue was most likely biomaterial, based on data log review. as per the clinical details, cpr was initiated due to sustained ventricular tachycardia (vt) that was unresponsive to shocks. the pump was exchanged due to concerns regarding saturated pump flow and elevated left ventricular (lv) waveform. it was determined that decreased left ventricular unloading at lower p-levels, not physical interaction, likely contributed to the observed increase in vt. the cause of the ventricular tachycardia issue was most likely patient condition related and its relation to impella is unknown. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35071 | 1220648-2023-02650 | 17513540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17513540&pc=OZD | 05/15/2021 | 09/16/2025 |  | 08/11/2023 |  |  | 05/15/2021 | 08/11/2023 |  |  | 01/31/2023 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047923 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male patient developed a retroperitoneal bleed during impella support. due to the noted condition, the impella cp pump was explanted and the left femoral artery was surgically repaired. an impella 5.5 pump was subsequently placed to support the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for retroperitoneal bleed has been completed. pump was not returned for investigation. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the retroperitoneal bleed could not be determined. g1 mdr reporting contact email updated corrections to the initial mfr report: h6 med dev prob code 4008 changed to 2993. updated component code 3038 to 925.
35072 | 1220648-2023-02654 | 17515858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515858&pc=OZD | 07/16/2023 | 03/08/2024 |  | 08/11/2023 |  |  | 07/16/2023 | 08/11/2023 |  |  | 04/30/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351465 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient's labs were all hemolyzing during support. nurse reported there was no echo. the impella cp pump was explanted (b)(6) 2023. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. d4-updated udi. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35073 | 1220648-2023-02655 | 17515859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515859&pc=OZD | 07/16/2023 | 01/30/2024 |  | 08/11/2023 |  |  | 07/16/2023 | 08/11/2023 |  |  | 09/30/2024 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella pump being stuck in the patient's papillary muscle, while the patient exhibited lack of pulse and coolness in the impella inserted limb. the patient was ultimately explanted 3 days later. | [5 empty] | The investigation into the limb ischemia ¿ reversible and the positioning issues has been completed. the pump was not returned for investigation. per the clinical information provided, the root cause of the positioning issue was use related. the root cause of the limb ischemia issue was patient condition related to diseased/small vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
35074 | 1220648-2023-02656 | 17515861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515861&pc=OZD | 07/14/2023 | 09/26/2025 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  |  |  | 08/11/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1164960 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The abiomed field service representative reported a user facility's automated impella controller (aic) battery failure, where the aic was stating to plug the controller to power even though the controller was plugged in. multiple attempts were made to troubleshoot this issue, however, they were unsuccessful. there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported aic - battery failure (red alarm) issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that after reviewing the imc logs, the reported battery failure issue was confirmed. there were numerous instances of battery failure alarms (transport connection blown/missing) observed on the reported complaint date. the console was tested after arriving in field service. the battery failure issue was able to be reproduced. results of running the batttest utility revealed that cell 4 in battery 1 had a voltage that was significantly less than the rest of the cells in the battery. lcd screen of battery one was blank. the reported battery failure issue was determined to be caused by internal battery cell imbalance (found in cell 4 of battery 1). this report is being filed as part of a retrospective review of historical records.
35075 | 1220648-2023-02658 | 17515863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515863&pc=OZD | 07/14/2023 | 09/26/2023 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 05/31/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352063 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The user facility reported a 51-year-old male noted as high risk percutaneous cardiac insertion (hrpci) was implanted with an impella cp device for mechanical circulatory support. upon completion of the hrpci, the pump was explanted and groin site seen to be ischemic. the limb ischemia prompted the impella limb to be treated by the vascular surgeon. | [5 empty] | The investigation into the limb ischemia and the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. per the provided clinical information, the root cause of the limb ischemia was use related issue due to improper technique and closure complication in which resulted in a lack of perfusion down the impella leg. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35076 | 1220648-2023-02659 | 17515864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515864&pc=OZD | 07/14/2023 | 10/27/2023 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  |  |  | 08/11/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1550078 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The user facility reported that the automated impella controller (aic) had a broken purge clip. a loaner was sent to the user facility and return of aic with broken clip is anticipated. there was no patient involvement related to the event. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the medical center returned the console for analysis. the root cause of the reported complaint was a defective purge pressure sensor. the faulty component was replaced. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35077 | 1220648-2023-02657 | 17515865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515865&pc=OZD | 07/14/2023 | 10/27/2025 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 12/31/2024 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292228 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 07/14/2023 |  |  |  |  | It was further reported that the patient expired at the time the pump was explanted. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that the use of the device cannot conclusively be associated with the outcome. | The user facility report a 61-year-old male was being upgraded from an impella cp to an impella 5.5 device. the patient had the 5.5 and ecmo support ongoing when placement signal was lost. the team troubleshot by automated impella controller replacement. this did not remedy the issue and so the impella 5.5 pump itself was replaced and the issue was resolved. there was no patient harm. |  |  |  |  | The investigation for the optical signal issue has been completed. according to the clinical, shortly after inserting the impella 5.5 pump placement signal was lost and a controller error alarm was recorded. staff then performed a controller swap, which reproduced the same error and did not fix the issue. the decision was then made to replace the pump with another impella 5.5 and the support proceeded without any other issues. no product was returned for a visual inspection. data log analysis confirmed a fiber break after the pump ran for two minutes which triggered the ps not reliable alarm and caused all 5s parameter values to fall out of specification. the most likely cause of optical signal issue was due to a fiber break. the location of the fiber break is unknown as no product was returned for review. added- b5 mso review added d6a/d6b. f6/f8 should have been left blank in the initial report. h4 should have been initial use of device in the initial report. h5 should have been yes in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B1 and b2: added information h1: corrected reportable event type. h6: health effect code: corrected the clinical code and health effect code. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
35078 | 1220648-2023-02660 | 17515878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515878&pc=PYX | 07/16/2023 | 07/16/2024 |  | 08/11/2023 |  |  | 07/16/2023 | 08/11/2023 |  |  | 02/28/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024348544 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. | The user facility reported that the purge sidearm of an impella rp flex pump was found to be broken in two pieces while in the packaging. a replacement pump was subsequently implanted for patient support. there was no patient harm. |  |  |  |  | B.5 additional information was received, and abiomed's medical safety officer review was completed. h.6 add code 2907 to medical device problem codes. the investigation for the product damage has been completed. according to the clinical description, before beginning impella rp flex support, the team noticed that the purge sidearm was broken into two pieces. at the time, the pump had not even been removed from the packaging. the decision was made to use a replacement pump. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the cause of the packaging issue was not determined because no product was returned. d.1 and d.4 revised model number from manufacturer device report 1220648-2023-02660 in accordance with updated procedures. d.6a and d.6b revised from manufacturer device report 1220648-2023-02660 in accordance with updated procedures. g.1 revised first name as it was report incorrectly on manufacturer device report 1220648-2023-02660. revised reporting contact fax number from manufacturer device report 1220648-2023-02660 in accordance with updated procedures. h.6 added code 4629 as it was inadvertently left off manufacturer device report 1220648-2023-02660. h.10 additional manufacturer narrative instructions for use (ifus) were reported on manufacturer device report 1220648-2023-02660 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿ do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿
35079 | 1220648-2023-02661 | 17515889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515889&pc=OZD | 07/14/2023 | 01/19/2024 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  |  |  | 08/11/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1643465 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Run on Battery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The user facility reported a 62-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a better level low alarm. it was noted that the registered nurse had accidentally pulled the cord instead of the adapter plug and that the battery would no longer charge when plugged in. as a result of the reported issue, the decision was made to replace the console with another aic. there was no patient harm. | [5 empty] | The investigation into the aic - battery charging/other issues has been completed since the original report was filed. the product was returned for investigation. the console was inspected and tested on an engineering lab bench. the end of the power cable appeared to be missing the strain relief, which suggests that it was compromised in the field. even though charging was within specification, the cause of the battery charging issues was most likely a defective power cable. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35080 | 1220648-2023-02662 | 17515892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515892&pc=OZD | 07/14/2023 | 01/24/2024 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 05/31/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024351422 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The user facility reported a 57-year-old male (race unknown) in acute myocardial infarction / cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding. two units of blood products were administered. on (b)(6) 2023, the physician reported the bleeding occurred due to initial high activated clotting time (act) / partial thromboplastin time (ptt) and poor repo-sheath handling. after adjustment of act / ptt and improvement of repo-sheath, the bleeding stopped. | [5 empty] | The investigation into the access site bleeding issue has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding issue was high patient act values due to improper anticoagulation technique used. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35081 | 1220648-2023-02664 | 17515893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515893&pc=OZD | 07/14/2023 | 01/25/2024 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 03/31/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342987 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The user facility reported a 45-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the patient was taken to the or for an intra-aortic balloon pump placement (iabp). the medical staff determined that the patient had developed hemolysis and leg ischemia. pump was explanted and iabp was placed. | [5 empty] | The investigation into the reversible limb ischemia and the hemolysis issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, patient had hemolysis and leg ischemia during impella support. data logs were reviewed, and no positioning issues or motor current spike were observed. the cause of the limb ischemia issue was most likely patient condition related with ischemia observed on impella leg per clinical review; however, the cause of the hemolysis issue was not determined since product was not returned and with inconclusive evidence from log and clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35082 | 1220648-2023-02663 | 17515894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515894&pc=OZD | 07/16/2023 | 01/16/2024 |  | 08/11/2023 |  |  | 07/16/2022 | 08/11/2023 |  |  | 08/25/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338520 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 22-year-old male patient was oozing form the impella site during support. doctor replaced the sutures. | [5 empty] | Incorrect year for date of occurrence and aware date used. revised b(3) and g(4) from year of 2022 to 2023. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the access site bleeding issue was most likely use related as staff was educated on maintaining angle of entry and hemostasis achieved with forward tension sutures after changing angle per clinical review.
35083 | 1220648-2023-02665 | 17515895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515895&pc=OZD | 07/16/2023 | 09/19/2024 |  | 08/11/2023 |  |  | 07/16/2023 | 08/11/2023 |  |  | 05/31/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351912 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 70-year-old male patient experienced an embolic stroke during impella 5.5 support. systemic heparin stopped and purged. it was further reported that blood products were administered for drop in hematocrit (hct). |  |  |  |  | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr1220648-2023-02665 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva issue has been completed. the device was not returned for investigation. per the provided clinical information, the cause of the stroke was patient condition related. heparin was used in purge but unknown if it was the cause. unknown relation to impella.
35084 | 1220648-2023-02666 | 17515901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17515901&pc=OZD | 07/14/2023 | 08/11/2023 |  | 08/11/2023 |  |  | 07/14/2023 | 08/11/2023 |  |  | 03/31/2025 |  | 08/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024347302 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/14/2023 |  |  |  |  | The complainant reported a 65 year old male patient presenting with cardiomyopathy for impella support. during support, the patient exhibited limb ischemia. the physician advised the patient would be taken to the operating room for impella explant due to lost distal pulse. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35085 | 1220648-2023-02635 | 17502329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502329&pc=OZD | 07/10/2023 | 01/19/2024 |  | 08/10/2023 |  |  | 07/10/2023 | 08/10/2023 |  |  | 11/30/2024 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312828 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2023 |  |  |  |  | The user facility reported a 59-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient demonstrated to have no pedal pulses and also showed cold distal limbs. after several days of support, the patient coded multiple times over night and the family decided to withdraw care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient did not demonstrate positive pedal pulses and shows cold distal limbs, and both the legs were ischemic. the cause of the limb ischemia issue was patient condition related and unrelated to impella with both legs being ischemic per clinical review.
35086 | 1220648-2023-02637 | 17502581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502581&pc=OZD | 05/03/2023 | 09/25/2024 |  | 08/10/2023 |  |  | 07/05/2023 | 08/10/2023 |  |  | 12/31/2024 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/05/2023 |  |  |  |  | Stemi-dtu notification revealed the 64-year-old male patient endured thrombocytopenia prompting transfusion of platelets and packed red blood cells. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the thrombocytopenia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, there was no evidence indicated that impella caused thrombocytopenia. therefore, the root cause of thrombocytopenia was most likely due to the patient¿s condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02637 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
35087 | 1220648-2023-02636 | 17502673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502673&pc=OZD | 06/15/2023 | 01/19/2024 |  | 08/10/2023 |  |  | 07/11/2023 | 08/10/2023 |  |  | 02/28/2025 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325027 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2023 |  |  |  |  | Loqi study notification reported a 61-year-old male post cardiac artery bypass grafting surgery with an impella cp in place endured multiple episodes of ventricular tachycardia (vt). the device was removed and reinserted but the vt continued and required multiple cardioversions. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the ventricular tachycardia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. during case stat, data log recorded the improper position of the pump in ventricle followed by several suction for the remaining case run. the cause of the ventricular tachycardia issue was use related since pump position was compromised from beginning of the case start.
35088 | 1220648-2023-02639 | 17502680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502680&pc=OZD | 07/13/2023 | 11/19/2025 |  | 08/10/2023 |  |  | 07/13/2023 | 08/10/2023 |  |  |  |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1309664 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Low Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 59-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during impella support, the device encountered an alarm sound during suction. the alarm was barely audible and was muffled. there was no harm to the patient as a result of this incident. | It was later reported that care was withdrawn and that the patient expired. |  |  |  |  | Follow up with the field service representative has not been completed. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: the console log did not confirm the reported muffled alarm but showed communication issues with the pbm and opm module (unrelated to the complaint). device analysis: the console was tested in lab. all alarms are audible, and there were no communication issues reproduced during normal operation. root cause: the cause of the muffled/inaudible alarm was not determined since issue wasn't reproduced. b1/b2 updated. b5 updated. h1 updated. d2 common device name updated. d4 model number, lot number, and unique identifier (udi) # updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated. h4 updated.
35089 | 1220648-2023-02641 | 17502737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502737&pc=OZD | 07/14/2023 | 01/22/2024 |  | 08/10/2023 |  |  | 07/14/2023 | 08/10/2023 |  |  |  |  | 08/10/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1690147 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 07/14/2023 |  |  |  |  | The user facility reported that upon removal of the catheter from controller, the plastic dust cover and plug was broken of the controller. there was no staff or patient harm. | [5 empty] | The investigation into the console "aic ¿ damage" has been completed since the original report was filed. the console was returned for investigation. the console was tested on an engineering lab bench. upon visual inspection, the front panel optical connector on the aic was broken, which confirms the reported failure that the plastic dust cover and the plug were broken off the controller. the root cause for the broken front panel optical connector was not determined. | The device is going to be returned to abiomed for repair. the investigation will be performed once device is received. upon investigation closure, a supplemental mdr will be filed.
35090 | 1220648-2023-02642 | 17502806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502806&pc=OZD | 07/13/2023 | 01/16/2024 |  | 08/10/2023 |  |  | 07/13/2023 | 08/10/2023 |  |  | 04/30/2025 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350826 |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 70-year old male with persistent access site bleeding coinciding with impella support. due to the noted condition, the decision was made to explant the pump from the right femoral artery and a second pump was implanted in the existing left femoral artery sheath. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp ¿assess access site for bleeding and hematoma.¿ introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the cause of the access site bleeding was anticoagulation management related due to improper anticoagulation: high acts.
35091 | 1220648-2023-02640 | 17502850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502850&pc=OZD | 07/13/2023 | 01/25/2024 |  | 08/10/2023 |  |  | 07/13/2023 | 08/10/2023 |  |  | 03/31/2025 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324345 |  |  |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 48-year-old female had a 5.5 pump inserted for hemodynamic support. the patient had known multiple cardiac comorbidities. the 5.5 was placed via the axillary and had a persistent access site ooze at the site. the site bleed was treated with multiple units of blood products, rbcs and fresh frozen plasma (ffp). | [5 empty] | The investigation into the major access site bleeding issue has been completed. the device was not returned for investigation. based on the clinical information, the cause of the access site bleeding- major most likely was use related as there weren't any hemashield platinum grafts available and opted to use hemashield gold even though they know they have much higher risk for oozing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
35092 | 1220648-2023-02643 | 17502871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502871&pc=OZD | 07/14/2023 | 10/09/2023 |  | 08/10/2023 |  |  | 07/14/2023 | 08/10/2023 |  |  | 02/28/2025 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333102 |  |  |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male was implanted with an impella cp device for hemodynamic support. the complainant reported that the physician was uncomfortable leaving impella cp in patient during transport due to inability to position impella in lv without creating ectopy. physician showed lv gram and lv cavity visibly looked anatomically abnormal with ¿waste band¿ in mid ventricle that possibly prevented impella from being advanced into ventricular free-space. the product rep facetimed physician who was then able to reposition the pump and he felt like it was in a good position. however, the next update was received 20 minutes later indicating that the team was explanting the impella and placing patient on intra-aortic balloon pump (iabp) support due to continuous ectopy caused by impella catheter. | [5 empty] | The investigation into positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the positioning issue was patient condition related to abnormal left ventricle. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35093 | 1220648-2023-02638 | 17502933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502933&pc=OZD | 07/14/2023 | 09/06/2023 |  | 08/10/2023 |  |  | 07/14/2023 | 08/10/2023 |  |  | 04/30/2025 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351471 |  |  |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male was implanted with an impella cp device for circulatory support. when the patient was taken to the or for a cabg and a left main occlusion, it was noted that there was no flow to the right leg. the patient was therefore put on bypass, and the impella was explanted. the impella site was surgically closed prior to completion of the cabg. staff had no additional concerns. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the limb ischemia was unable to be determined as no product was returned and insufficient clinical details were provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35094 | 1220648-2023-02644 | 17502941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502941&pc=OZD | 04/30/2023 | 08/31/2023 |  | 08/10/2023 |  |  | 04/30/2023 | 08/10/2023 |  |  | 12/31/2024 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320298 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2023 |  |  |  |  | A 64-year-old male enrolled in stemi dtu trial presented as st elevated myocardial infarction. during support the impella displayed suction alarms coinciding with the presence of a thrombus in the left ventricle. additionally, new information information was received on 11-jul-2023 regarding a retroperitoneal hematoma that formed prompting intervention in the form of packed red blood cell delivery, platelet delivery, which resolved after replacing the pump. | [5 empty] | The investigation into the low pump flow, thrombus, major access site bleeding, and pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the low pump flow and the pump stop was biomaterial, the root cause of the thrombus was patient condition related, however the root cause of the major access site bleeding was not determined since product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be file.
35095 | 1220648-2023-02645 | 17502994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17502994&pc=OZD | 07/13/2023 | 01/23/2024 |  | 08/10/2023 |  |  | 07/13/2023 | 08/10/2023 |  |  | 12/31/2024 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 26-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, that there was no distal flow to the patient's lower extremity following impella implant. an antegrade reperfusion catheter was subsequently placed to resolve the noted condition. support continued with the implanted plant following the intervention. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. per the clinical information provided. the device was not returned by the medical facility. the root cause of the limb ischemia issue was patient condition related to disease/ small vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35096 | 1220648-2023-02646 | 17503023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17503023&pc=OZD | 07/13/2023 | 07/23/2024 |  | 08/10/2023 |  |  | 07/13/2023 | 08/10/2023 |  |  | 01/31/2025 |  | 08/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338586 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | User facility reported a 72-year old male with a past medical history of coronary artery disease, hypertension, congestive heart failure (chf), amyloidosis, and ablation with reoccurring angina chf symptoms arrived to the hospital and was placed on impella support. during the procedure, access was obtained through the left femoral artery (lfa), a 14 f sheath was inserted and the pigtail was advanced into the left ventricle (lv). however, the 0.018" guidewire was accidentally bent upon insertion by the operator. staff replaced the bs pp 0.018" wire and the impella was implanted without complications. | [5 empty] | The investigation for delivery issues-delivery has been completed. the impella device was not returned for evaluation. according to the clinical information provided, the root cause of the delivery issue was kinked guidewire due to improper technique. instructions for use: section: warnings and cautions: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. d4 catalog number d6a/d6b. d8-no. g1reporting contact fax number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time¿.
35097 | 1220648-2023-02619 | 17491204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17491204&pc=OZD | 07/11/2023 | 02/29/2024 |  | 08/09/2023 |  |  | 07/11/2023 | 08/09/2023 |  |  | 04/30/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333536 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2023 |  |  |  |  | The user facility reported an 80-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. during support, the impella was alarming ¿placement signal low¿ throughout the night. echo was done and the 5.5 was in good position. placement signal was disabled overnight. it is suspected optical sensor may have been damaged during placement of the device. | [5 empty] | The investigation into the reported delivery issue -use has been completed since the original report was filed. the pump was not returned for investigation. the pump was then used on the console ic4713 and the data logs also show a consistent placement signal low alarm from the beginning of the case. none of the optical 5s parameters were affected during the case run. according to clinical records, the doctor disabled the placement signal after continuously observing a low placement signal alarm despite the pump being in a good position. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35098 | 1220648-2023-02620 | 17491256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17491256&pc=OZD | 07/13/2023 | 01/22/2024 |  | 08/09/2023 |  |  | 07/13/2023 | 08/09/2023 |  |  |  |  | 08/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1302104 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility called the abiomed field service to report an automated impella controller (aic) error message. a loaner was onsite and was provided to the user facility for use. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned to the lab for investigation. the console was tested after arriving. the purge disc retainer was confirmed to be broken. the cause of the issue was due to a broken blue purge retainer. d.9 the date the device was returned to manufacturer was incorrect on the previous manufacturer device report and was updated to reflect the correct date the device was returned. | The impella device has been received from the customer. upon investigation closure, a supplemental mdr will be filed.
35099 | 1220648-2023-02621 | 17491342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17491342&pc=OZD | 07/11/2023 | 07/26/2024 |  | 08/09/2023 |  |  | 07/11/2023 | 08/09/2023 |  |  | 11/30/2024 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297116 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male patient developed a thrombus in their left ventricle during impella support. the pump was subsequently removed as a result of the noted thrombus. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is for the impella cp with smart assist system, which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
35100 | 1220648-2023-02622 | 17491367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17491367&pc=OZD | 07/11/2023 | 01/17/2024 |  | 08/09/2023 |  |  | 07/11/2023 | 08/09/2023 |  |  | 01/31/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024307992 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male was implanted with an impella cp device for mechanical circulatory support. however, the pump experienced persistent suction issues during patient support. support was reduced and the pump was subsequently explanted as a result of the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported low pump flow has been completed since the original report was filed. the cause of the low pump flow cannot be determined since no complaint device for investigation and no data logs for review returned.
35101 | 1220648-2023-02623 | 17491452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17491452&pc=OZD | 07/10/2023 | 09/28/2023 |  | 08/09/2023 |  |  | 07/10/2023 | 08/09/2023 |  |  | 11/30/2024 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024311980 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/10/2023 |  |  |  |  | The complainant reported a 76 year old female patient presenting with acute myocardial infarction and cardiogenic shock for independent impella insertion. a notification was captured via long-term outcome and quality indicator (loqi) impella registry on (b)(6) 2023 that reported limb ischemia in the impella inserted limb. vascular consultation found no emergent intervention required, but recommended outpatient support for a possible limb amputation (if deemed necessary). | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, patient had faint pulses over the right leg (rle), and amputation was done to treat ischemia. the root cause of the limb ischemia issue was not determined due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35102 | 1220648-2023-02624 | 17493889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17493889&pc=OZD | 07/12/2023 | 01/11/2024 |  | 08/09/2023 |  |  | 07/12/2023 | 08/09/2023 |  |  | 12/31/2024 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024304103 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient developed red urine while driving at p9 during support. haptoglobin was administered (number of units unknown). after that, when the supplement level was lowered, hemolysis was improved, but renal function did not recover. impella support continued and discussions were initiated regarding the potential need for hemodialysis. | [5 empty] | The investigation for the reported hemolysis has been completed. without the impella device being returned nor was the clinical data provided sufficient. additionally, there was no data logs to review. therefore, the root cause of the reported issue could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35103 | 1220648-2023-02625 | 17493991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17493991&pc=OZD | 07/12/2023 | 02/07/2024 |  | 08/09/2023 |  |  | 07/12/2023 | 08/09/2023 |  |  | 05/31/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352453 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the operators were unable to increase support levels beyond p4 due to persistent suction alarms. support continued at the reduced level until planned explant. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. data confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response & suction alarms. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows:
35104 | 1220648-2023-02626 | 17494010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494010&pc=OZD | 07/13/2023 | 07/19/2024 |  | 08/09/2023 |  |  | 07/13/2023 | 08/09/2023 |  |  | 02/28/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342002 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/13/2023 |  |  |  |  | New information was received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case and complaint. in addition to hemolysis prompting escalation to higher support via impella 5.5, the patient endured aki prompting crrt. | The user facility reported a 54-year-old male (race unknown) in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support while undergoing coronary artery bypass grafting. during support, the patient experienced hemolysis. according to the complainant the pump was running at p5 due to hemolysis. an attempt was made to reposition the pump, however, it still appeared to be angled toward the mitral structures and the urine was dark red. the impella cp was explanted during escalation to an impella 5.5 device. |  |  |  |  | New information received and added to b5 & h6. the investigation is still in process. a supplemental mdr will be filed at the completion of our analysis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ . ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. | The investigation for the hemolysis has been completed. no product returned. the data logs show "impella position unknown" and "impella position in aorta" alarms which resolved during the case. pump ran at p5 due to hemolysis with pump angled towards mitral structures. the root cause of the hemolysis was most likely improper positioning of the pump as mentioned in the clinical summary. the root cause of the renal failure is most likely due to improper positioning of the device. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
35105 | 1220648-2023-02627 | 17494208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494208&pc=OZD | 07/12/2023 | 09/21/2024 |  | 08/09/2023 |  |  | 07/12/2023 | 08/09/2023 |  |  | 02/28/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342023 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 57-year-old female had an impella cp support ongoing for over two weeks when the pump stopped. the patient condition deteriorated and chemical code was necessary with falling bp. the team then transitioned the patient to care/comfort only as the patient was not a candidate for lvad or heart transplant. the patient expired. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the pump stop has been completed since the original report was filed. the product was returned for investigation. returned pump was visually inspected. larger impella purge gap observed. pump connected to aic and pump stop alarm reproduced. the root cause of the pump stop issue was bearing wear and the failure occurred beyond design life as the pump stopped on day 13th. cannula kink near outlet observed it may have occurred during pump removing process. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02627 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
35106 | 1220648-2023-02628 | 17494211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494211&pc=OZD | 07/13/2023 | 10/31/2025 |  | 08/09/2023 |  |  | 07/13/2023 | 08/09/2023 |  |  |  |  | 08/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Operating System Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility's biomedical engineer called abiomed field service representative to report a controller error on the automated impella controller (aic). there was a loaner on site for customer use. the customer returned the aic for investigation. there was no mention of patient involvement during this error message. | The complainant reported that a patient on impella support had encountered a controller error. the loaner automated impella controller (aic) was on site and the old aic was shipped for investigation. there was no harm reported to the patient as a result of this incident. |  |  |  |  | The automated impella controller has been received. the investigation is incomplete at this date. upon investigation closure, a supplemental mdr will be filed. | Data analysis: imc logs confirm that the defective optical sensor lamp controller error alarm was issued on initial plugging of pump as reported on the date of occurrence. device analysis: the console was inspected on the bench. the failure mode as reported was reproduced upon boot up in the lab. the front optical connector was inspected and cleaned before performing section 23 of the preventative maintenance procedure with a test cavity of specified length. the results were not within specification. conclusion: the root cause for the controller error (defective optical bench) was a defective lamp. fs instructed to replace optical bench (0042-1012). perform section 23 of the pm and full functional check for both console and optical bench.
35107 | 1220648-2023-02629 | 17494216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494216&pc=OZD | 07/11/2023 | 10/16/2023 |  | 08/09/2023 |  |  | 07/12/2023 | 08/09/2023 |  |  | 02/29/2024 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193121 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old female patient who required impella support and underwent the impella pump implant, was found to have a retroperitoneal bleed, for which the medical team had to administer four (4) units of blood. the medical staff was unsure what caused the bleed and think that it could maybe gastrointestinal (gi) related. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical information, ultrasound was not used for imaging. the cause of the vascular damage issue was related to the improper use of imaging to assess the artery. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35108 | 1220648-2023-02630 | 17494243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494243&pc=OZD | 07/12/2023 | 07/05/2024 |  | 08/09/2023 |  |  | 07/12/2023 | 08/09/2023 |  |  | 12/31/2024 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024304216 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2023 |  |  |  |  | The user facility reported a 57-year-old male awaiting heart transplant was implanted with an impella 5.5 device for mechanical circulatory support. while doing routine patient checks, the rn noticed a kink in catheter just proximal to the red plug. there were no alarms reported and rn said they hadn¿t noticed it. there was no harm to the patient as a result of this incident. the patient successfully received heart transplant on (b)(6) 23. the impella device was removed after transplant. patient is currently in stable condition. | [5 empty] | The investigation for the product damage has been completed. the impella was not returned for evaluation. according to the clinical, on the 67th day of impella 5.5 support a kink in the catheter was found close to the red plug. data log analysis confirmed there were no other abnormalities. no issue found in logs. no attempts were made to resolve the issue. the most likely cause of the cannula kink product damage was a catheter kink.(technique/use). the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock in section: warnings & cautions: cautions states ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ h6 impact code 4627 added- d6a/d6b. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35109 | 1220648-2023-02631 | 17494250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494250&pc=OZD | 07/11/2023 | 01/19/2024 |  | 08/09/2023 |  |  | 07/11/2023 | 08/09/2023 |  |  | 04/30/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350814 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. a tee showed that the cannula had migrated towards mitral valve. the physician attempted to torque cannula, but was unsuccessful. fluoroscopy was utilized to better visualize impella placement which showed good position across atrioventricular node (av). while troubleshooting, patient cardiac arrested requiring re-initiation of cpb. the decision was made to rewire impella cp in order to explant and place new pump. new impella cp was inserted successfully via right femoral artery (rfa) and optimal position confirmed under tee. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible. only the clinical report was evaluated. based on the clinical information provided, transesophageal echocardiography (tee) confirmed that the cannula had migrated towards the mitral valve and the physician attempted to torque cannula but was unsuccessful. the cause of the positioning issue was use related due to incorrect repositioning and a new impella cp was replaced without any issues per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35110 | 1220648-2023-02632 | 17494301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494301&pc=OZD | 07/11/2023 | 03/06/2024 |  | 08/09/2023 |  |  | 07/11/2023 | 08/09/2023 |  |  | 11/30/2024 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308434 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2023 |  |  |  |  | The user facility reported a 62-year-old female undergoing surgery to remove balloon pump and escalate to ecpella. an impella cp device was implanted for mechanical circulatory support. the medical team experienced pump flow issues which required administration of several blood products for more hemodynamic support. the cp was explanted on (b)(6) 2023 and patient was escalated to an impella 5.5 device. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35111 | 1220648-2023-02633 | 17494330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494330&pc=OZD | 07/12/2023 | 09/18/2023 |  | 08/09/2023 |  |  | 07/12/2023 | 08/09/2023 |  |  | 04/30/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351110 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had an out of box observation made of the 14fr sheath being bent/kinked. the surgeon still made choice to implant the sheath and under imaging of the iliofemoral saw the kink. the team made choice to replace (with a new 14fr x13cm) and deliver the impella cp to the left ventricle (lv). the impella cp had low flows after insertion and was explanted after the conclusion of the percutaneous coronary intervention (pci). | [5 empty] | The investigation into the reported ¿low pump flow and product damage¿ has been completed since the original report was filed. product was returned for investigation. the 14fr peel away introducer sheath and dilator was visually inspected and found a kink observed in peel away sheath and dilator tip bent observed. the device data logs were reviewed and found suction alarms occurred. as per clinical investigation, the complainant had an out of box observation made of the 14fr sheath being bent/kinked. the surgeon still made choice to implant the sheath and under imaging of the iliofemoral saw the kink. the team made choice to replace (with a new 14fr x13cm) and deliver the cp to the lv. the cp had low flows after insertion and was explanted after the conclusion of the pci. the cause of the low pump flow was most likely a kinked cannula from a difficult delivery through tortuous vessels. the most likely root cause of the introducer damage was sheath kink due to patient condition related as patient had tortuous vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows:
35112 | 1220648-2023-02634 | 17494559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17494559&pc=OZD | 07/13/2023 | 09/14/2023 |  | 08/09/2023 |  |  | 07/13/2023 | 08/09/2023 |  |  | 04/30/2025 |  | 08/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332335 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the md was experiencing a difficult insertion of the pump due to the patient's tortuous anatomy. ultimately the md was unable to be advanced despite many tries and great effort. | [5 empty] | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the delivery issue was patient condition related to tortuous vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35113 | 1220648-2023-02607 | 17481286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17481286&pc=OZD | 07/08/2023 | 09/11/2023 |  | 08/08/2023 |  |  | 07/11/2023 | 08/08/2023 |  |  | 01/31/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024314710 |  | 0048-0024 JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2023 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had shown signs of hemolysis. the medical team explanted the impella and the hemolysis was resolved. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis was not determined since product was not returned, and sufficient clinical information and data logs were not provided. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes; may also play a role in patient susceptibility to hemolysis.¿
35114 | 1220648-2023-02608 | 17484085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484085&pc=OZD | 07/11/2023 | 02/29/2024 |  | 08/08/2023 |  |  | 07/11/2023 | 08/08/2023 |  |  | 04/30/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024335029 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the 87-year-old female patient's vessel was small in diameter and was damaged during impella insertion. the vessel was surgically repaired and support with the implanted pump continued. the patient later passed away during support, however, there was no allegation or evidence of association with the impella pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported vascular damage - peripheral vessel has been completed since the original report was filed. the pump was not returned for investigation. according to clinical details, the vessel was damaged during insertion. the cause of the vascular damage issue was most likely use related, as excessive force was applied during insertion. subsequently, the doctor discovered that the patient had a small vessel diameter.
35115 | 1220648-2023-02610 | 17484132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484132&pc=OZD | 07/11/2023 | 07/25/2024 |  | 08/08/2023 |  |  | 07/11/2023 | 08/08/2023 |  |  | 04/30/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351039 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2023 |  |  |  |  | The user facility reported a 46-year-old male in cardiac arrest was implanted with an impella cp device for mechanical circulatory support. the patient experienced an access site bleed around the repositioning sheath of an impella cp pump after an impella rp 23 fr introducer was utilized for implant. vascular surgery was called for pump removal and site closure. a new pump was subsequently implanted via left femoral artery after it was discovered that the originally implanted impella cp pump had a clot. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. product was not returned for review. based on clinical description, the root cause of access site bleeding was due to misuse of rp introducer for cp impella. the root cause of the thrombus could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
35116 | 1220648-2023-02611 | 17484196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484196&pc=OZD | 07/11/2023 | 09/21/2024 |  | 08/08/2023 |  |  | 07/11/2023 | 08/08/2023 |  |  | 04/30/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024347390 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Failure to Power Up | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/11/2023 |  |  |  |  | The user facility reported a male of unknown age and race in critical condition with no pulsatility was implanted with an impella cp device for mechanical circulatory support after the first impella cp device malfunctioned. the impella cp was implanted successfully but would not start. the patient expired due to their condition and the device was left in the patient for autopsy. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Pump and data logs were not returned for investigation. the cause of the pump did not start issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02611 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
35117 | 1220648-2023-02612 | 17484212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484212&pc=OZD | 07/11/2023 | 08/01/2024 |  | 08/08/2023 |  |  | 07/11/2023 | 08/08/2023 |  |  | 03/31/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2023 |  |  |  |  | The user facility reported a 79-year-old female in rapid atrial fibrillation was implanted with an impella cp device for mechanical circulatory support. while on impella support, there was oozing out the the impella access site. manual pressure was held and eventually increased pressure and mattress sutures were required. additionally, pink-red urine was observed by the staff and there was also decreasing urine output. the patient was successfully weaned off support the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for access site bleeding and hemolysis has been completed. the product was not returned for investigation. the data logs did not indicate any abnormalities related to hemolysis. according to the given clinical details, the pump was in a good position. hemolysis was diagnosed with tinged foley/urine, and the removal of the pump resolved the hemolysis issue. additionally, the impella access site was closed and dry, but the right internal jugular was still oozing. the cause of the access site bleeding issue was most likely patient condition related since oozing was reported from multiple access site. the cause of the hemolysis issue was not determined since the product was not returned and sufficient clinical information was not provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-02612 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-02612 and should not have been. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2023-02612. a new code has been added to type of investigation codes. h.6 code 925 was added since the initial submission of manufacturer device report number 1220648-2023-02612 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-02612 in accordance with updated procedures. additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2023-02612 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
35118 | 1220648-2023-02613 | 17484219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484219&pc=OZD | 07/13/2023 | 09/15/2023 |  | 08/08/2023 |  |  | 07/13/2023 | 08/08/2023 |  |  | 02/28/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024311698 |  | 1000211 JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella cp device for mechanical circulatory support for escalation of therapy. the patient developed a thrombus during impella support. at the time of planned explant, it was discovered that a thrombus had adhered to the inlet of the pump. despite the thrombus almost entirely occluding the inlet, there was no noted deficiencies with the function of the pump. no further intervention was necessitated as a result of the thrombus. | [5 empty] | The investigation into the thrombus has been completed since the original report was filed. the device was returned and investigated. the benchtop analysis and investigation determined that the root cause of the thrombosis is most likely due to the patient's condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35119 | 1220648-2023-02614 | 17484269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484269&pc=PYX | 07/13/2023 | 07/23/2024 |  | 08/08/2023 |  |  | 07/13/2023 | 08/08/2023 |  |  | 10/31/2024 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023274766 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 79-year-old male undergoing coronary artery bypass grafting and aorta repair was implanted with an impella rp device for mechanical circulatory support. while on support, the impella rp triggered an error message upon start up. the automated impella controller (aic) was swapped in an attempt to troubleshoot the issue, however the pump still would not start up. the pump was replaced with another impella rp device to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation for pump did not start has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35120 | 1220648-2023-02615 | 17484314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484314&pc=OZD | 07/13/2023 | 10/07/2023 |  | 08/08/2023 |  |  | 07/13/2023 | 08/08/2023 |  |  | 05/31/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024355912 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a patient (unknown age, gender, race) was implanted with an impella 5.5 for mechanical circulatory support. the device generated a high purge pressure (pp) alarm coinciding with impella support. pump was explanted the following day due to reported malfunction. there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported limb ischemia and vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that high purge pressure alarm with 1060 mm/hg pressure and purge flow of 1.7 ml/hr was observed. the patient has bleeding problems and ptt levels were kept at 35s due to bleeding risk. the impella data logs were not returned to perform log review for high purge pressure event. the cause of the high purge pressure issue was not determined since product, data logs were not returned and due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35121 | 1220648-2023-02616 | 17484388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484388&pc=OZD | 07/12/2023 | 01/18/2024 |  | 08/08/2023 |  |  | 07/12/2023 | 08/08/2023 |  |  | 06/30/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356987 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2023 |  |  |  |  | The user facility reported a 75-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the patient lost pulse in the leg. the medical team removed the impella cp and escalated treatment to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported limb ischemia ¿ reversible has been completed since the original report was filed . the root cause of the limb ischemia issue was patient condition related to small diseased vessels. correction : the clinical code in section h6, was submitted incorrect inadvertently in the initial report which has been corrected in this report.
35122 | 1220648-2023-02617 | 17484404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484404&pc=OZD | 07/10/2023 | 01/24/2024 |  | 08/08/2023 |  |  | 07/10/2023 | 08/08/2023 |  |  | 04/30/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337437 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male patient had a stroke during impella support, the care team seems to believe it was related to impella. the patient expired. | [5 empty] | The investigation into stroke/cva issue has been completed. the device was not returned for investigation. per the provided clinical information, the cause of the reported stroke issue was most likely patient condition related and determined to be unknown relation to impella since the heparin was not given in purge fluid. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35123 | 1220648-2023-02618 | 17484416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17484416&pc=OZD | 07/11/2023 | 08/09/2023 |  | 08/08/2023 |  |  | 07/11/2023 | 08/08/2023 |  |  | 04/30/2025 |  | 08/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024345080 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Disconnect | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2023 |  |  |  |  | The user facility reported a male of unknown age and race was implanted with an impella cp for circulatory support. during implant, when the pump was placed, the 0.18 abiomed wire could not be removed as normal and was stuck at the pump. the pump was explanted and a new impella cp was implanted. the patient expired during unsuccessful attempt of implantation. | [5 empty] | The investigation into the reported guidewire malfunction has been completed since the original report was filed. the medical center returned the impella pump for benchtop analysis, but did not returned the guidewire which was the component that malfunctioned. visual inspection found no issues on the pump. the root cause of delivery issue was unable to be determined as guidewire was not returned for review. the failure modes will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35124 | 1220648-2023-02597 | 17475179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17475179&pc=OZD | 07/08/2023 | 09/25/2025 |  | 08/07/2023 |  |  | 07/08/2023 | 08/07/2023 |  |  | 05/31/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355484 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | Positioning Problem |  |  |  |  | Hemolysis | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/08/2023 |  |  |  |  | The user facility reported a 53-year-old male in cardiac arrest was implanted with an impella 5.5 device for mechanical circulatory support. the device reported a position unknown alarm on (b)(6) 2023. on (b)(6) 2023, the patient had discolored urine indicating hemolysis. two days later, on (b)(6) 2023, the device reported multiple suction alarms. the medical team decided to withdraw care on (b)(6) 2023. | It was further reported that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that the use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the positioning issue, low pump flow, and optical issues has been completed. no product was returned for a visual inspection. data log analysis confirmed several placement signal spikes up to 3000 during impella position unknown alarms. additionally, the logs did not indicate any issues related to the position of the device or spikes in motor current during the time of discolored urine. finally, several suction alarms were recorded without any motor current spikes or placement signal trends. the cause of the optical signal issue was not determined because no product was not returned and not enough clinical information was given. the cause of hemolysis was not determined because no product was not returned and not enough clinical information was given. the cause of the low pump flow was not determined because no product was not returned and not enough clinical information was given. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ added- b5 mso review added- d6a/d6b. f6/f8 should have been left blank in the initial report. | H1 updated to death. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-02597. b2 updated death date. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
35125 | 1220648-2023-02598 | 17475351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17475351&pc=OZD | 07/08/2023 | 01/19/2024 |  | 08/07/2023 |  |  | 07/08/2023 | 08/07/2023 |  |  | 05/31/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352565 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2023 |  |  |  |  | The user facility reported a 41-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had pink-tinged urine which is a concern for hemolysis. the impella cp was explanted after one day of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no issues were found on the data logs. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
35126 | 1220648-2023-02599 | 17475442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17475442&pc=OZD | 07/08/2023 | 01/26/2024 |  | 08/07/2023 |  |  | 07/08/2023 | 08/07/2023 |  |  | 03/31/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331062 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2023 |  |  |  |  | The user facility reported an 85-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site bleed during impella support. the site was re-dressed and the patient was transfused one unit packed red blood cells to treat the condition. support continued with the implanted pump following the treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleeding has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. the cause of the access site bleeding issue was most likely patient movement with pump position moved back despite knee immobilizer use and bleeding occurred due to persistent patient movement.
35127 | 1220648-2023-02601 | 17476232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17476232&pc=OZD | 06/13/2023 | 08/07/2023 |  | 08/07/2023 |  |  | 06/13/2023 | 08/07/2023 |  |  | 12/31/2024 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325227 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 76-year-old male patient suffering from acute myocardial infarction and cardiogenic shock was presented for impella placement. during support, it was noted that the patient experienced low pump flows. the pump was replaced as a result of the noted issue. following explant, an evaluation was completed on the pump and it was discovered that the blades of the impeller were damaged. there was no patient harm resulting from the noted issue. | [5 empty] | (h3) the product was returned for investigation. the device was visually inspected and found the impeller blades to be damaged. as per clinical investigation, the patient had undergone tavr with medtronic corevalve. the impella pump impeller blades would be damaged with interaction of tavr. the cause of the low pump flow issue was due to damaged impeller blade(s). the damage was characteristic of an interaction with tavr (medtronic corevalve). the investigation determined that the root cause of the low pump flow issue was found to be damaged impeller blades. should additional relevant information become available, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
35128 | 1220648-2023-02602 | 17476400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17476400&pc=OZD | 06/16/2023 | 01/18/2024 |  | 08/07/2023 |  |  | 06/13/2023 | 08/07/2023 |  |  | 01/31/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316448 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Fracture | Incorrect, Inadequate or Imprecise Result or Readings | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a male was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that there was a sudden loss in placement signal coinciding with a placement signal not reliable alarm. a second aic was utilized in an attempt to troubleshoot the issue, but failures persisted on the second console. due to the ongoing issue, it was suspected that the sensor on the pump was faulty. the pump was explanted as a result of the failure. following explant, an evaluation was initiated and it was additionally discovered that the impeller blades were damaged. there was no resulting harm to the patient. | [5 empty] | The impella device was received and an evaluation is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿. | The investigation into the optical signal issue has been completed since the original report was filed. the medical center has returned for analysis the impella pump and the benchtop analysis revealed that the root cause of the optical sensor issue was a damaged sensor due to tavr interaction with the optical sensor being broken and not available in the pump along with the damaged impeller blades.
35129 | 1220648-2023-02603 | 17476766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17476766&pc=OZD | 07/08/2023 | 01/22/2024 |  | 08/07/2023 |  |  |  | 08/07/2023 |  |  | 03/31/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  |  |  | 2024327074 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Computer Operating System Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/10/2023 |  |  |  |  | The user facility reported a 70-year-old male with history of heart failure was implanted with an impella 5.5 device for mechanical circulatory support. the nurse was performing a cassette change and received an alarm stating "automated impella controller (aic) will not detect the purge disc". the device (serial number (b)(6)) was swapped with another aic. simultaneously, the patient had an axillary hematoma, at the site of the impella 5.5 pump delivery. patient was complaining of tingling in their fingers resulting in a wash out of site. this mdr is being submitted for both the durable console ((b)(6)) and disposable impella pump ((b)(6)). | [5 empty] | Follow up with the field service representative has not been completed. upon investigation closure, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the failure mod will be monitored and trended.
35130 | 1220648-2023-02604 | 17476797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17476797&pc=OZD | 07/08/2023 | 09/26/2025 |  | 08/07/2023 |  |  |  | 08/07/2023 |  |  | 01/31/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316617 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 70-year-old male on impella support experienced a low placement signal and the medical team was unable to resolve it by repositioning the device. the following day the patient was experiencing bleeding from multiple sites, heparin had to be stopped and device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issues has been completed. the device was not returned for benchtop investigation. data logs show a downward trend in placement signal, with several suction alarms and low placement signal alarms, but without affecting 5s optical parameters or causing low pump flow. the data log does not indicate any abnormalities related to improper pump positioning. according to the clinical notes, the low placement signal alarm persisted even after repositioning and confirming a good pump position, leading the doctor to disable the placement monitoring. the cause of the optical signal issue was most likely patient condition related based on data log review. the cause of the access site bleeding issue was patient condition related since bleeding was reported from multiple access site, based on given clinical details. | B1 product problem updated. d6a / d6b updated. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-02604 b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
35131 | 1220648-2023-02605 | 17476803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17476803&pc=OZD | 07/08/2023 | 01/12/2024 |  | 08/07/2023 |  |  |  | 08/07/2023 |  |  | 04/30/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349750 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male had ongoing hemolysis concerns during impella cp support. the impella cp was repositioned twice during the day. also, physician discussed right heart failure (rhf) as possible cause of hemolysis. they are giving lasix to reduce right heart numbers. the patient remained on impella cp support until explanted on (b)(6) 2023. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported hemolysis has been completed since the original report was filed. given the timing of the hemolysis, the positioning alarms occurred on the same day. when considering they had multiple positioning issues. the root cause of hemolysis testing was most likely improper positioning due to improper technique. complaint device not returned for investigation.
35132 | 1220648-2023-02606 | 17476807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17476807&pc=OZD | 07/10/2023 | 09/20/2023 |  | 08/07/2023 |  |  |  | 08/07/2023 |  |  | 02/28/2025 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337751 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during the impella support the patient moved and impella cp pulled back into aorta and was not able to be repositioned. a new impella (5.5) was placed. there was no harm to the patient. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. data log review confirmed that the pump was in aortic area (about 4 hours into the case). the pump then was stopped manually & disconnected from the console. based on clinical description & data review, the root cause of positioning issue was due to patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows:
35133 | 1220648-2023-02600 | 17477052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17477052&pc=OZD | 07/11/2022 | 03/28/2025 |  | 08/07/2023 |  |  |  | 08/07/2023 |  |  | 04/30/2024 |  | 08/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023215047 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Difficult to Advance | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the medical staff was attempting to implant impella 5.5 and encountered some delivery issues due to the patient's anatomy. the medical staff had to abort the procedure and also had to ask for assistance to the vascular surgery team to close the damaged patient arteries. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. no product was returned for evaluation. per the given clinical information, the root cause of the access site bleeding issue was most likely patient condition related. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
35134 | 1220648-2023-02583 | 17461604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461604&pc=OZD | 07/10/2023 | 01/18/2024 |  | 08/04/2023 |  |  | 07/10/2023 | 08/04/2023 |  |  | 02/28/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320368 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/10/2023 |  |  |  |  | The user facility reported, a 61-year-old male undergoing percutaneous cardiac insertion surgery. was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella support, the device experienced low pump flow. from what was alleged to be a thrombus. the device involved with this incident had to be explanted. and a new one had to be implanted in its place. | [5 empty] | The impella device was not received from the customer. and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the thrombus has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, suction alarms observed at higher p-level with low flows and p-level reduced per auto suction response. no motor current spike was found per logs. the cause of the low pump flow issue was most likely patient condition related with suction alarms observed at higher p-levels and reduced p-level with low flows during entire case run per log review.
35135 | 1220648-2023-02584 | 17461608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461608&pc=OZD | 04/11/2021 | 09/18/2023 |  | 08/04/2023 |  |  | 04/11/2021 | 08/04/2023 |  |  | 02/28/2023 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022049823 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male patient on impella 5.5 support experienced a fluid leak from what appeared to be a small crack in the microfilter of the line. the physician decided to swap the pump for a new one. there was no harm to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the purge leak issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the purge leak issue was most likely sidearm filter damage.
35136 | 1220648-2023-02585 | 17461627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461627&pc=OZD | 06/22/2023 | 04/15/2025 |  | 08/04/2023 |  |  | 07/10/2023 | 08/04/2023 |  |  | 06/22/2024 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/10/2023 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced significant bleeding from the repositioning sheath. sheath mismatch from peel away to the repositioning sheath and there was still an issue with bleeding. it was only a slight ooze when patient was leaving the cath lab. staff had perclose placed and left in and fellow over night tightened perclose in ccu, but was still was bleeding significantly. staff applied manual pressure. representative stated there were unknown number of blood products given and a massive transfusion protocol was initiated. patient expired after 6.23 hours of support. death is unrelated to the impella. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | Medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the access site bleeding was patient condition related because the patient was coagulopathic in shock, and was not escalation candidate. this report is being filed as part of a retrospective review of historical records.
35137 | 1220648-2023-02587 | 17461728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461728&pc=OZD | 07/08/2023 | 01/19/2024 |  | 08/04/2023 |  |  | 07/08/2023 | 08/04/2023 |  |  | 05/31/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356154 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2023 |  |  |  |  | The user facility reported a 66-year-old male scheduled for coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a cerebral vascular accident with residual left sided weakness. the patient remained on impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the patient had a small, non-significant obstruction with residual left sided weakness. data logs were reviewed, and no motor current spike were observed relative to the event. the cause of the stroke issue was most likely patient condition related and unknown relation with insufficient clinical/log review.
35138 | 1220648-2023-02588 | 17461734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461734&pc=OZD | 07/09/2023 | 02/28/2024 |  | 08/04/2023 |  |  | 07/09/2023 | 08/04/2023 |  |  | 04/30/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333559 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/09/2023 |  |  |  |  | The user facility reported a patient (unknown age, gender, race) on impella 5.5 support experienced hemolysis. the medical team repositioned the device several times in attempt to resolve the issue but they were unsuccessful and the device had to be explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed. the device was not returned for investigation. data logs show no signs related to improper position or motor current spikes. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided.
35139 | 1220648-2023-02589 | 17461742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461742&pc=OZD | 07/07/2023 | 01/19/2024 |  | 08/04/2023 |  |  | 07/07/2023 | 08/04/2023 |  |  | 02/28/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324715 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported an 81-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support for upcoming percutaneous cardiac insertion. during explant, an angiograph demonstrated slow flow down leg. a balloon was dilated for three minutes. a clot demonstrated on angiography, second balloon inflation restored robust flow showing chronically occluded superficial femoral artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient is having peripheral vascular disease over right femoral artery, and the superficial femoral artery is chronically occluded leading to ischemia complication. the cause of the limb ischemia issue was due to patient condition related with patient having peripheral vascular disease over right femoral and angioplasty was performed to push the clot resolving complication.
35140 | 1220648-2023-02590 | 17461759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461759&pc=OZD | 07/07/2023 | 02/28/2024 |  | 08/04/2023 |  |  | 07/07/2023 | 08/04/2023 |  |  | 02/28/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334455 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 61-year-old female in cardiac arrest was implanted with an impella cp device for mechanical circulatory support. the impella was inserted into left tunnel artery using normal fashion for micro puncture and ultrasound. access gained to left ventricle and impeller inserted over .018 wire. percutaneous cardiac insertion (pci) of left anterior descending artery performed implanting one drug eluding stent. post pci ,it was identified that decreased flows and erratic wave forms may be due to a kinked catheter. physician decided to replace the impella cp while still in the cath lab. site access port accessed with over .035 wire, impella sheath was removed. new impella sheath was inserted to maintain hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported product damage has been completed since the original report was filed. pump was not returned for the investigation. data logs shows the low pump flow and suction alarms from the beginning of the case. as per the clinical details, doctor reported the low pump flow during the case run. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided.
35141 | 1220648-2023-02591 | 17461766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461766&pc=OZD | 07/07/2023 | 07/24/2024 |  | 08/04/2023 |  |  | 07/07/2023 | 08/04/2023 |  |  | 02/28/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341820 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 60-year-old male awaiting surgery on circumflex artery was implanted with an impella cp device for mechanical circulatory support. the patient inadvertently pulled their impella out when they suddenly went into ventricular tachycardia. as a result, bleeding was noted at the access site. femostop was applied to control the bleed and the patient was given one unit of blood. the physician stated that they would re-evaluate the patient to determine if the impella needed to be re-implanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. per the clinical details provided, the root cause of the access site bleeding issue was patient movement related. d6a/d6b added. h6 component code added 925 the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-02591: f6/f8-should have been left blank.
35142 | 1220648-2023-02592 | 17461770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461770&pc=OZD | 01/09/2022 | 01/19/2024 |  | 08/04/2023 |  |  | 01/09/2022 | 08/04/2023 |  |  | 07/31/2023 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022118273 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male patient on impella support experienced hematoma and oozing at the surgical site. the medical team administered pressure and 1 unit of rbc to the patient. the patient remained on impella cp support until explanted on (b)(6) 2022. | [5 empty] | This report is being filed as part of a retrospective review of historical records the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's access site bleeding has been completed. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation.
35143 | 1220648-2023-02593 | 17461782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461782&pc=OZD | 03/13/2023 | 08/04/2023 |  | 08/04/2023 |  |  | 03/13/2023 | 08/04/2023 |  |  | 10/31/2024 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287303 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 125 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 54-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. there was good signal distal to cp, patient's right leg mottled related to incorrect placement of 23fr sheath to right femoral artery (rfa). taken to operating room (or) for fem cutdown for 23fr sheath removal, thrombectomy right leg (rle), and rle fasciotomy; restoration of signal to rle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. (h3) the cause of the introducer issue was the use of a sheath since right femoral artery (rfa) sheath was mistakenly wired out and dilated to 23fr sheath. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35144 | 1220648-2023-02594 | 17461790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461790&pc=OZD | 02/21/2022 | 08/04/2023 |  | 08/04/2023 |  |  | 02/21/2022 | 08/04/2023 |  |  | 08/31/2023 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022140830 |  | 0550-0002 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 14 YR |  | Female |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 14-year-old female had arterial thrombus formation in the left lower leg and in the mesenteric vessels, requiring leg amputation and the purge filter near red plug is broken and leaking, requiring purge bypass during impella 5.5 support. the patient remained on impella 5.5 support until explanted on (b)(6) 2022. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. (h3) the root cause of the limb ischemia issue was biomaterial found in leg artery. the cause of the purge leak issue was not determined since product was not returned. instructions for use for the related event are as follows: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
35145 | 1220648-2023-02586 | 17461831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17461831&pc=OZD | 12/07/2021 | 01/12/2024 |  | 08/04/2023 |  |  | 12/07/2021 | 08/04/2023 |  |  |  |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONNECT |  |  | OZD |  |  | IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 2022132668 |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that their aic displayed a controller error during patient support. troubleshooting was attempted by the operator; however, the decision was ultimately made to swap out the console. support continued following the swap. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "alarm header displayed in the left column; window is color-coded red for critical alarms, yellow for serious alarms, white for advisory notifications, gray for resolved alarms.¿ "abnormal situation. prompt action needed. 3 beeps every 15 seconds alarm header on yellow background". "overlaid with a translucent yellow when the controller cannot determine catheter position, position is wrong, and placement monitoring is suspended or disabled". "when the impella catheter is operating in a low flow range, placement monitoring may be suspended and the flow rate in the lower left corner of the controller display screen will turn yellow to indicate that impella position is unknown". this report is being filed as part of a retrospective review of historical records. | The investigation into the reported aic - controller error (yellow alarm) has been completed since the original report was filed. imc logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. .real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. the reported issue is a known pattern failure caused by a temporary loss of optical data (placement, snr, contrast, lamp, gain). abiomed is aware of the issue and is working on appropriate corrective action. there¿s no need to send the console in for repairs if the condition self-corrects. this is a low probability event and is very difficult to reproduce. if needed, the clinical team can monitor patient hemodynamics and impella position with imaging. the device was not requested to be returned because the issue is self-correcting.
35146 | 1220648-2023-02595 | 17463305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17463305&pc=OZD | 07/07/2023 | 01/15/2024 |  | 08/04/2023 |  |  | 07/07/2023 | 08/04/2023 |  |  | 05/31/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351804 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 40-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the physician decided on escalation of care and the cp was removed and an impella 5.5 was inserted. the patient developed a hematoma at their impella access site. a wound vac was utilized to treat the condition and support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. as per clinical review, patient had hematoma at insertion site, later resolved with intervention of wound vac in place. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
35147 | 1220648-2023-02596 | 17463331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17463331&pc=OZD | 07/07/2023 | 01/23/2024 |  | 08/04/2023 |  |  | 07/07/2023 | 08/04/2023 |  |  | 02/28/2025 |  | 08/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335115 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 72-year-old female in non-st elevated myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when the patient was being weaned off the impella cp and device was completely removed, the patient's artery closed down. a balloon angioplasty was used to open artery but then patient did not stop bleeding. according to the md, multiple perclose advancements/failures from the day before and the placement of them in ¿her gritty artery¿ may have been the main cause of bleeding. multiple balloon inflations and manual pressure allowed for achievement of hemostasis. patient returned to unit on dobutamine drip. pulses remained dopplerable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, balloon angioplasty was performed after impella removal. the patient had diseased artery and bleeding was resolved with inflated balloon and manual pressure technique. the cause of the access site bleeding issue was most likely patient condition related to diseased vessel.
35148 | 17450038 | 17450038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17450038&pc=UNK | 06/27/2023 | 07/19/2023 | 07/19/2023 | 08/03/2023 | 07/19/2023 | 08/03/2023 |  | 08/03/2023 |  |  |  |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | UNK |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 416923 |  | 0048-0003 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No |  |  |  | User Facility report |  | Non-Healthcare Professional | [6 empty] | 84 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Displays Incorrect Message | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | Followup |  |  |  |  | Y | HOSPITAL |  |  | Placed impella catheter into left ventricle via rfa which gave an error code when device turned on. removed device per rep's direction and a new impella placed without issues. reported to rep 6/28/2023, no rga# issued, rep picked up. manufacturer response for abiomed impella, (brand not provided) (per site reporter), rep picked up and sent in no charge replacment. | Placed impella catheter into left ventricle via rfa which gave an error code when device turned on. removed device per rep's direction and a new impella placed without issues. reported to rep [redacted], no rga# issued, rep picked up. manufacturer response for abiomed impella, (brand not provided) (per site reporter). rep picked up and sent in no charge replacment.
35149 | 1220648-2023-02570 | 17450159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17450159&pc=OZD | 08/29/2022 | 10/16/2023 |  | 08/03/2023 |  |  | 08/29/2022 | 08/03/2023 |  |  | 05/31/2024 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217664 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 70-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support during bypass surgery. while trying to implant the impella 5.5, there was a left ventricle (lv) perforation. patches were sewn into the lv and venoarterial extracorporeal membrane oxygenation was cannulated. patient was given two units of blood products. the decision was subsequently made to place the impella 5.5 via direct approach and not use a wire/catheter to place. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical account, the md decided to place impella 5.5 directly without a wire. this technique is outlined in the impella 5.5 ifu, and abiomed support was not on sight. without further details on physician technique and patient condition, the cause of the ventricle perforation cannot be determined.
35150 | 1220648-2023-02571 | 17450746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17450746&pc=OZD | 06/04/2022 | 02/06/2024 |  | 08/03/2023 |  |  | 06/04/2022 | 08/03/2023 |  |  | 02/29/2024 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197493 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 53-year-old male patient had active bleeding and large hematoma to axillary insertion site progressively worsening during impella support. the patient was transfused 1 unit prbc as a result of the bleed and discussions with the physician were completed regarding additional treatments that could be undertaken for the site. support continued with the implanted pump following the transfusion. no additional intervention was required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿this report is being filed as part of a retrospective review of historical records.¿ | The investigation into the reported "access site bleeding - major" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. no further details are known. due to the lack of clinical details and no product returned, the cause of the access site bleeding cannot be determined. | Since the medical center did not return the impella device, no benchtop analysis was possible. due to the lack of clinical details and no product returned, the cause of the access site bleeding cannot be determined.
35151 | 1220648-2023-02572 | 17451391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17451391&pc=OZD | 07/06/2023 | 01/19/2024 |  | 08/03/2023 |  |  | 07/06/2023 | 08/03/2023 |  |  | 03/31/2024 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023209935 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported a 41-year-old male in cardiac arrest was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to possible ventricular assist device. the patient had access site bleeding. patient was on aggrastat. staff applied manual compression and the patient received blood platelets. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
35152 | 1220648-2023-02573 | 17451557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17451557&pc=OZD | 07/06/2023 | 09/26/2024 |  | 08/03/2023 |  |  | 07/06/2023 | 08/03/2023 |  |  | 05/19/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337975 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/06/2023 |  |  |  |  | The user facility reported a 75-year-old female in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on impella support, a ct scan was performed and it revealed a left side ischemic stroke. the physician noted the impella was not removed. do not believe it was the cause of the stroke and it is helping promote forward flow. pt was therapeutic on heparin drip with act 182 when stroke discovered. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02573 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
35153 | 1220648-2023-02574 | 17455624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455624&pc=OZD | 02/24/2023 | 08/31/2023 |  | 08/03/2023 |  |  |  | 08/03/2023 |  |  | 10/31/2024 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290773 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Alarm System | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported an 80-year-old female with worsening cardiogenic shock symptoms was implanted with an impella cp device for mechanical circulatory support. the procedure was tolerated well, however, the patient started to develop limb ischemia and the md decided to remove the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was patient condition related due to history of medical conditions. the patient had known coronary artery disease, pulmonary embolism, deep vein thrombosis, hypertension, diabetes, and cancer.
35154 | 1220648-2023-02576 | 17455625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455625&pc=OZD | 07/07/2023 | 01/23/2024 |  | 08/03/2023 |  |  |  | 08/03/2023 |  |  | 09/30/2024 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis and the physician made the decision to escalate the patient to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed. the device was not returned for benchtop investigation. data logs were returned, and no abnormalities were found. the root cause of the hemolysis issue was not determined since the product was not returned and insufficient clinical information provided. the relation to the pump is unknown.
35155 | 1220648-2023-02575 | 17455627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455627&pc=OZD | 07/06/2023 | 01/19/2024 |  | 08/03/2023 |  |  |  | 08/03/2023 |  |  | 01/31/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309041 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported a 64-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient on impella support in the cath lab woke up agitated and pulled the sheath away. the md took out the repositioning sheath and exchanged it with a new one. the patient however experienced heavy bleeding and hematoma at the access site and for that reason the medical team decided that it was better to explant impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into access site bleeding issue has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding issue was patient movement related.
35156 | 1220648-2023-02578 | 17455631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455631&pc=OZD | 07/07/2023 | 02/14/2024 |  | 08/03/2023 |  |  |  | 08/03/2023 |  |  | 03/31/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344888 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 63-year-old female in cardiac arrest was implanted with an impella cp device for mechanical circulatory support. during support, the device reported low flows. the pump was repositioned, but unable to obtain a flow greater than p5. it was further reported that there was oozing from the impella site, anticoagulation was stopped, and one unit of packed red blood cells was given. the patient remained on support for five days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major and the low pump flow issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the low pump flow issue and access site bleeding issue cannot be determined since the complaint device not returned for investigation. no data logs available for time the issues occurred.
35157 | 1220648-2023-02579 | 17455632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455632&pc=OZD | 07/04/2023 | 01/31/2024 |  | 08/03/2023 |  |  |  | 08/03/2023 |  |  | 01/31/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328602 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a female patient was implanted with an impella cp device for mechanical circulatory support. the user facility reported the impella access site developed a large hematoma. the hematoma was about 13cm in diameter. the operator immediately held pressure and tried to express the site, but this made little improvement. the physicians agreed that the impella would need to be removed. the patient was completely impella dependent and would not survive without the support. a second device was inserted into the left femoral artery while the first impella was removed from the right. while closing the right femoral access site, the perclose device malfunctioned. the device broke into two parts. the hydrophilic dilator was left inside the artery, and the handle was removed with force. the physicians used the peel away sheath to introduce a 6fr sheath, glide advantage wire and peripheral balloon to occlude the vessel proximal to the access site. bleeding was controlled, and the patient was taken to the operating room (or) for vessel repair and removal of the closer device left behind. patient was implanted with an impella cp device for mechanical circulatory support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding and vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the cause of the access site bleeding issue was not determined since product was not returned and with insufficient clinical information. the cause of the vascular damage issue was most likely use related due to perclose device malfunction by breaking into two parts during site closure. correction : section h6: the clinical code and the medical device problem code was inadvertently left out in the initial report which has been added to this report.
35158 | 1220648-2023-02580 | 17455637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455637&pc=OZD | 07/06/2023 | 11/08/2023 |  | 08/03/2023 |  |  |  | 08/03/2023 |  |  | 03/31/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported a 63-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had impella cp support for a few hours when the pump was seen to be causing limb ischemia. the pulses were lost in the limb. the limb ischemia also could have been due to cardiogenic shock and the vessel being clamped down. the team explanted the cp and placed a 5.5 pump via the axillary. the 5.5 was successfully placed and support proceeds. | [5 empty] | The investigation into the reversible limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the cause of the limb ischemia was patient condition related as the patient had a large calcific plaque. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35159 | 1220648-2023-02581 | 17455638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455638&pc=OZD | 07/08/2023 | 10/16/2023 |  | 08/03/2023 |  |  |  | 08/03/2023 |  |  | 05/31/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352392 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2023 |  |  |  |  | The user facility reported a 42-year-old male with non-ischemic cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the pump had low flows and the team believed this to be due to hemolysis and clot ingestion. the pump was explanted and replaced with a new impella 5.5 for the hemolysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
35160 | 1220648-2023-02582 | 17455643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455643&pc=OZD | 07/07/2023 | 01/12/2024 |  | 08/03/2023 |  |  | 07/07/2023 | 08/03/2023 |  |  | 01/31/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310004 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 68-year-old male on impella cp support. the patient was to be escalated to an impella 5.5. upon insertion of the axillary impella 5.5, the impella cp was removed. the impella pump stopped. the 5.5 pump stop prompted the team to remove and when the restart failed. the team instead placed an intra-aortic balloon pump. there was no harm or injury noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follow: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. the clinical evaluation and the data logs were reviewed and were sufficient to determine the root cause. the data logs revealed an alarm 115 pump triggered after a loss of purge resulting in air in purge alarm and unsuccessful de-air and open purge line. the cause of the pump stoppage was most likely due to clamping the purge line and the lack of purge led to a blood ingress pump stop.
35161 | 1220648-2023-02577 | 17455761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17455761&pc=OZD | 07/07/2023 | 01/19/2024 |  | 08/03/2023 |  |  | 07/07/2023 | 08/03/2023 |  |  | 01/31/2025 |  | 08/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024323211 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 57-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during support, the team was receiving suction alarms on p2. the customer was advised to check the position of the impella cp but the customer was unable to reposition the device. the impella cp was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
35162 | 1220648-2023-02557 | 17442620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17442620&pc=OZD | 11/08/2022 | 08/31/2023 |  | 08/02/2023 |  |  | 11/08/2022 | 08/02/2023 |  |  | 06/30/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238073 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device-Device Incompatibility | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported a 50-year-old male undergoing surgery was implanted with an impella cp device for mechanical circulatory support. the device had low flow during the procedure due to interaction with other device. procedure successfully completed, no harm done to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the low pump flow was interaction with other device and obstructed impella flow path due to improper technique use as mapping catheter became caught in impella pigtail. the doctor was able to free pigtail from mapping catheter to solve the low flow issue.
35163 | 1220648-2023-02558 | 17442650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17442650&pc=OZD | 07/07/2023 | 02/08/2024 |  | 08/02/2023 |  |  | 07/07/2023 | 08/02/2023 |  |  | 04/30/2025 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024335027 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 80-year-old male have multiple fluid leaks during impella support. one day following implant, it was reported that a fluid leak was seen from the purge cassette housing. the purge cassette was replaced to resolve the noted issue. however, three days later, it was documented that purge fluid was also leaking from the luer. the pump was replaced to resolve the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the reported purge leak- pump has been completed since the original report was filed. only purge cassette was returned for investigation. the device was tested and functioned without issue, no leaks were observed. the root cause of the purge leak-pump issue was not determined to be a filter leak from the clinical description and the image uploaded form the field. the cause of the purge leak was sidearm filter damage. section d9 updated.
35164 | 1220648-2023-02559 | 17442749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17442749&pc=OZD | 07/04/2023 | 01/19/2024 |  | 08/02/2023 |  |  | 07/04/2023 | 08/02/2023 |  |  | 03/31/2025 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024333433 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 86-year-old female patient was implanted with an impella cp device for mechanical circulatory support. during case, the patient experienced loss of pulsatility, loss of blood pressure, loss of impella flow due to huge tamponade. a pericardiocentesis was performed and patient was stabilized. a protamine was administered, the impella was extracted and the patient was transferred to icu. a tamponade probably created during insertion guidewire or pigtail catheter. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the ventricle perforation has been completed since the original report was filed. no product was returned. as per clinical, there was loss of pulsatility, blood pressure, and the impella pump flow during rotablator use. tamponade was created due to insertion guidewire or pigtail catheter and pericardiocentesis was done. data logs were not returned to confirm the timing of perforation. the cause of the ventricle perforation issue was most likely patient condition related and relation to the impella is unknown.
35165 | 1220648-2023-02560 | 17442895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17442895&pc=OZD | 07/04/2023 | 08/02/2024 |  | 08/02/2023 |  |  | 07/04/2023 | 08/02/2023 |  |  | 07/31/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257412 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Break | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 48-year-old male reported that following impella implant, difficulty was met when attempting to remove the guidewire. the component was noted to have gotten stuck and, upon eventual removal, part of the distal tip of the wire broke off and remained in the outlet cage. support continued with the implanted pump as the patient was dependent on the device. | It was further reported that the patient expired in the procedural area. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that the use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the product damage has been completed. the impella was returned for evaluation. upon visual inspection, there were no guidewire fragments found within the pump. additionally, no damage was found on the impeller or the cage. the guidewire was not returned for analysis. data log analysis confirmed no motor current spikes that would suggest a piece of the wire was within the pump upon activation. the cause of the guidewire damage was not determined because the guidewire was not returned and not enough clinical information was given. the instructions for use for the impella cp with smartassist system states the following: section: warnings & cautions: warnings ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ added b5 mso review, d9 device return date added- d6a/d6b.
35166 | 1220648-2023-02561 | 17443082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17443082&pc=OZD | 07/07/2023 | 01/19/2024 |  | 08/02/2023 |  |  | 07/07/2023 | 08/02/2023 |  |  | 05/31/2025 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352132 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported an 89-year-old female undergoing surgery was implanted with an impella cp device for mechanical circulatory support. during surgery for impella implantation, the pump was pulled back into the aorta and the md was unable to get it into the left ventricle. the patient had already been separated from the bypass so the decision was made by the medical team to leave impella in the aorta and remove it in the morning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the root cause of the positioning issue was most likely use related.
35167 | 1220648-2023-02562 | 17443142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17443142&pc=OZD | 07/03/2023 | 10/11/2023 |  | 08/02/2023 |  |  | 07/03/2023 | 08/02/2023 |  |  | 07/31/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255189 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 20-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a lack of blood flow to the right lower extremity on impella support. a distal reperfusion catheter was placed in the left superficial femoral artery (sfa) and bridged to right sfa. the patient's right groin was within normal limits and there was no bleeding nor hematoma. the angle of insertion was maintained with dressing. the patient remained on impella cp support until explanted on (b)(6) 2023. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient had no flow over right femoral leg due to diseased vessel and fem-fem bypass was done from right sfa to left sfa to restore flow. the cause of the limb ischemia issue was patient condition related with patient having diseased vessels in right femoral and fem-fem bypass resolved ischemia.
35168 | 1220648-2023-02563 | 17443307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17443307&pc=OZD | 07/04/2023 | 09/22/2024 |  | 08/02/2023 |  |  | 07/04/2023 | 08/02/2023 |  |  | 09/30/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023279063 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 66-year-old male had an impella cp inserted and placed to the left ventricle (lv) when the pump jumped in position. the jump caused the lv perforation when the pump advanced too fully to the wall. the team performed a pericardiocentesis and transfusion, but the patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the malfunctio, has been completed since the original report was filed. the product was returned and the anticontamination sleeve was found to be damaged and disconnected from the hub on the gwru side. the plastic of the anticontamination sleeve was deformed. the deformation is consistent with the use of excessive force, the timing of the anticontamination sleeve damage is unknown. the cause of the initial resistance leading to the need for excess force is unknown,the root cause of the ventricle perforation was most likely use related as evidence shows excessive force was applied to the anticontamination sleeve leading to the ventricle perforation. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02563 was provided in sections b5, d6a, d6b, and h1. note: corrected information provided in h6 is an addition to previous coding.
35169 | 1220648-2023-02564 | 17446116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17446116&pc=OZD | 12/12/2022 | 01/19/2024 |  | 08/02/2023 |  |  | 12/12/2022 | 08/02/2023 |  |  | 09/30/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266020 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 63-year-old male noted for high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. during support, the team was unable to increase p-level past p5 secondary to diastolic suction and a lack of sufficient preload in the left ventricle. the doctor was consulted and the right ventricle was noted to be enlarged. the pump experienced low flows and was now unable to go above p4 support levels. the decision was made to re-open the chest and place central veno-arterio-venous extracorporeal membrane oxygenation. while open, chest evacuated of >1150ml blood and multiple large clots. patient received multiple blood products. the patient remained on support for four additional days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the low pump flow and the thrombus has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based off the clinical info available, the patient had an enlarged right ventricle (rv) and high central venous pressure (cvp). thus, the root cause of the low pump flow was most likely due to patient condition, specifically right heart failure and lack of preload. the root cause of the thrombus evacuated from the chest is unknown due to lack of sufficient clinical information.
35170 | 1220648-2023-02566 | 17447006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17447006&pc=OZD | 12/16/2022 | 01/18/2024 |  | 08/02/2023 |  |  |  | 08/02/2023 |  |  | 08/31/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023276905 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 64-year-old male patient developed a hematoma at their axillary impella site during pump support. the hematoma was evacuated in the operating room and support continued with the implanted pump. | [5 empty] | (h3) the investigation into the reported "access site bleeding - major" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. a 64-year-old male patient developed a hematoma at their axillary impella site during pump support. the hematoma was evacuated in the operating room and support continued with the implanted pump. all other details are unknown. the root cause of the access site bleeding could not be determined as no clinical details are provided. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | G6: 8/14/2023.
35171 | 1220648-2023-02567 | 17447012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17447012&pc=OZD | 12/09/2022 | 09/05/2023 |  | 08/02/2023 |  |  |  | 08/02/2023 |  |  | 09/30/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259130 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported a 47-year-old male (unknown race) with non-ischemic covid cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma around their axillary impella site during pump support. the patient underwent a washout and a jackson pratt drain was placed to treat the hematoma. support continued with the implanted pump following the treatment. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the access site bleeding was not determined since the product was not returned and no other clinical details were provided. the failure mode will be monitored and trended.
35172 | 1220648-2023-02568 | 17447029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17447029&pc=OZD | 07/05/2023 | 01/23/2024 |  | 08/02/2023 |  |  |  | 08/02/2023 |  |  | 01/31/2025 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2023 |  |  |  |  | The user facility reported a 75 year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a severe hematoma which required vascular surgery to intervene and impella cp to be removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding has been completed. the device was not returned for benchtop investigation. the root cause of the access site bleeding issue was patient's high activated clotting time (act) values.
35173 | 1220648-2023-02569 | 17447040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17447040&pc=OZD | 05/30/2023 | 09/19/2025 |  | 08/02/2023 |  |  | 05/30/2023 | 08/02/2023 |  |  | 11/30/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023280932 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Degraded | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2023 |  |  |  |  | The user facility reported a 47-year-old female in myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump had to be replaced after over 120 days of use due to damage sustained to the pump (exposed wiring near the red plug). the pump was removed and replaced with no harm to the patient. | Additional information indicated that, during patient support, the pump experienced placement signal issues. |  |  |  |  | The investigation into the optical signal issue and product damage has been completed. the returned product was inspected for optical signal issue and product damage. optical bench test, light test was done, and both passed with 5s parameters verified on console to be in range and light signal was observed at the receiving end of the sensor. visual inspection was done, and biomaterial ingestion was observed on the outflow cage of the pump and over the optical sensor. the cause of the sterile sleeve damage issue was use related with also catheter curling or twisting observed near the red plug and splinted during use. product failure has been observed beyond recommended design life period per design trace matrix 0550-9002. the failure modes will be monitored and trended. | B1 updated with product problem. b5 updated with additional information. d6a/d6b updated with additional information. d9 device return was updated. f6 and f8 should have been left blank. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-02569.
35174 | 1220648-2023-02556 | 17447084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17447084&pc=OZD | 07/05/2023 | 01/19/2024 |  | 08/02/2023 |  |  |  | 08/02/2023 |  |  | 01/31/2025 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316088 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2023 |  |  |  |  | The user facility reported 67-year-old female in st elevation myocardial infarction with an impella cp device for mechanical circulatory support. the patient developed sepsis during the course of impella support. as the pump was suspected as a source of the infection, the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the stroke/cva has been completed since the original report was filed. no product was returned. as per clinical, sepsis was suspected with elevated white blood cells (wbc), cardiac output (co) levels and systemic vascular resistance (svr) at 300. the cause of the infection/sepsis issue was most likely patient condition related and unknown relation to the impella.
35175 | 1220648-2023-02565 | 17447282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17447282&pc=OZD | 12/02/2022 | 09/19/2024 |  | 08/02/2023 |  |  | 12/02/2022 | 08/02/2023 |  |  | 02/29/2024 |  | 08/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023194865 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that after a ct scan the 72-year-old female patient had significant mca clot/stoke. patient was on impella support for 72.45 hours before the patient expired. the death has been determined to unrelated to the impella. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | ¿this report is being filed as part of a retrospective review of historical records.¿ the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported "stroke/cva" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the 72-year-old female patient had significant mca clot/stoke. with no product, data, or clinical details provided, the cause of the stroke could not be determined. correction :section d4: the udi was submitted incorrect inadvertently in the initial report which has been corrected to this report. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02565 was provided.
35176 | 1220648-2023-02554 | 17430566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17430566&pc=OZD | 07/02/2023 | 09/15/2023 |  | 07/31/2023 |  |  | 07/02/2023 | 07/31/2023 |  |  | 04/30/2025 |  | 07/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024346833 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in the impella leg during pump support. it was noted that the patient did suffer from peripheral arterial occlusive disease. due to the noted condition, the pump was explanted and the groin was surgically treated. | [5 empty] | The investigation into the limb ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the ischemia issue was patient condition related to the diseased/ small vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35177 | 1220648-2023-02555 | 17430573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17430573&pc=OZD | 06/07/2023 | 09/12/2023 |  | 07/31/2023 |  |  | 06/07/2023 | 07/31/2023 |  |  | 02/28/2025 |  | 07/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024325876 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during support, the device had a low purge pressure with purge filter leak. a purge system bypass was performed to resolve the issue. there was no reported harm to the patient. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the purge leak issue was a leak from the sidearm, but the exact location of the leak was unknown. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿.
35178 | 1220648-2023-02552 | 17414354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17414354&pc=OZD | 06/11/2021 | 07/28/2023 |  | 07/28/2023 |  |  | 06/11/2021 | 07/28/2023 |  |  |  |  | 07/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Literature | Health Professional | User facility | Company representation |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/11/2021 |  |  |  |  | A publication found by abiomed japan literature review on (b)(6) 2023 reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. "when the patient stood up this evening, the purge line from the red port to the purge filter broke. the field service representative reached out to the team and they suggested the 20 gauge access needle method and walked them through the process. the purge flows and pressures normalized and the pump returned to normal functions. there were no issues at that time. the team is aware that the patient had been on support for over 120 days. it was reported there was a successful extended use of impella5.5 as a bridge to heart transplantation". | [5 empty] | The investigation into the reported purge leak has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the purge leak was due to damage to the filter to tubing joint specifically at the filter neck or tubing to filter neck bond. the tubing to filter neck bond damage occurred after 114 days which is beyond 60 days impella 5.5 indication for design life. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35179 | 1220648-2023-02553 | 17415135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17415135&pc=OZD | 07/03/2023 | 09/20/2023 |  | 07/28/2023 |  |  | 07/03/2023 | 07/28/2023 |  |  | 12/31/2024 |  | 07/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302771 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/03/2023 |  |  |  |  | The user facility reported a 75-year-old female undergoing surgery was implanted with an impella cp device for mechanical circulatory support. the md decided to remove impella from the right femoral artery because the patient had ischemic leg due to small arteries. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. as per clinical details, impella leg was ischemic due to patient having small arteries on the right femoral site. the root cause is native patient condition. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35180 | 1220648-2023-02538 | 17406864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17406864&pc=OZD | 06/30/2023 | 09/20/2024 |  | 07/27/2023 |  |  | 06/30/2023 | 07/27/2023 |  |  | 02/28/2025 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321261 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 76-year-old female patient who required impella cp support suffered from a perforation of the femoral and access site bleeding, requiring 2 units of blood. the patient subsequently expired. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the access site bleeding and the vascular damage issues was not determined since product was not returned and there was insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
35181 | 1220648-2023-02539 | 17406935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17406935&pc=OZD | 06/30/2023 | 01/11/2024 |  | 07/27/2023 |  |  | 06/30/2023 | 07/27/2023 |  |  | 03/12/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315910 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support; however, the impella 5.5 reported a pump stop due to clot ingestion leading to pump explant. there was no reported patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. the device was not returned for evaluation. the data logs were reviewed and showed the pump stop within four minutes with several motor current spikes and 13-impella stop alarms. this indicates possible biomaterial interaction with the pump impeller. additionally, upon clinical review, the physician reported a presence of a thrombus. the root cause of the pump stop issue was most likely biomaterial in the pump. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35182 | 1220648-2023-02540 | 17406947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17406947&pc=OZD | 04/27/2023 | 08/11/2023 |  | 07/27/2023 |  |  | 05/30/2023 | 07/27/2023 |  |  | 11/30/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314022 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2023 |  |  |  |  | The user facility reported a 27-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, there were alarms for pump in aorta. the signal issue did not cause any harm or injury. the pump remained on for support for 2 days when successfully explanted. | [5 empty] | The investigation into the reported device malfunction has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report and data logs were evaluated. based on clinical description & data, the root cause of positioning issue was most likely due to use issue. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35183 | 1220648-2023-02541 | 17406976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17406976&pc=PYX | 06/30/2023 | 09/20/2024 |  | 07/27/2023 |  |  | 06/30/2023 | 07/27/2023 |  |  | 04/30/2025 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337974 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  |  |  | Death | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a female patient had a bipella support of impella cp- impella rp planned for a patient in the er with vtach and a lucas device supporting the heart function. the impella cp was placed but, the impella rp was unable to be successfully delivered. the team troubleshot with wire exchanges but when advancing the pump the sheath split and catheter became bent and unable to advance. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the cause of the delivery issue was use related since excessive force were applied, which resulted into introducer damage and catheter kink. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella rp with smart assist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
35184 | 1220648-2023-02542 | 17407505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17407505&pc=OZD | 07/03/2023 | 07/27/2023 |  | 07/27/2023 |  |  | 07/03/2023 | 07/27/2023 |  |  | 05/31/2025 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352469 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2023 |  |  |  |  | The user facility reported an 80-year-old male undergoing high risk coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a "very bruised" heart from recent acute myocardial infarction. the 5.5 pump perforated the left ventricle (lv) wall. the patient was surgically repaired with a patch and then the chest was closed. the lv perforation was repaired and the patient had the 5.5 support ongoing for 94 hours when weaned and explanted. | [5 empty] | The investigation into the reported ventricle perforation has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. no product was returned for investigation. the cause of the ventricle perforation was use related and related to pushing the pump too far into the ventricle. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35185 | 1220648-2023-02543 | 17410935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410935&pc=OZD | 03/06/2023 | 08/30/2023 |  | 07/27/2023 |  |  | 03/06/2023 | 07/27/2023 |  |  | 11/30/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024290569 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Migration or Expulsion of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during a switch of the automated impella controller, the impella was dislocated into aorta and was not able to be repositioned. impella cp removed, patient stable with va-ecmo. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the positioning issue was use related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35186 | 1220648-2023-02544 | 17410937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410937&pc=OZD | 06/29/2023 | 09/29/2023 |  | 07/27/2023 |  |  | 06/29/2023 | 07/27/2023 |  |  | 01/31/2025 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306294 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68 year-old female patient was taken to the cath lab with symptoms of atrial fibrillation and rapid ventricular rate and ventricular tachycardia. the medical team decided to implant impella for support, however, soon after the procedure the patient experienced a stroke. the pump was explanted two days after implantation and a day after that the patient was transitioned to comfort measures only. | [5 empty] | Since the original report was filed, the investigation has closed. the pump was not returned for benchtop analysis, only the data logs and clinical report were shared. the data logs revealed no indication of biomaterial. the root cause of the cva was determined to be patient condition. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35187 | 1220648-2023-02545 | 17410946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410946&pc=OZD | 06/29/2023 | 01/12/2024 |  | 07/27/2023 |  |  | 06/29/2023 | 07/27/2023 |  |  | 03/31/2025 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024335857 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2023 |  |  |  |  | The user facility reported person (unknown age, race, gender) was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site bleed with hematoma. the patient was give 2 units of red blood cells to treat the condition. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. per the provided clinical information, patient had hematoma and bleeding on impella site and ptt was 100 seconds. bleeding improved and resolved with reduction of heparin. the cause of the access site bleeding issue was most likely due to anticoagulation management related with bleeding resolved after reducing heparin per clinical. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35188 | 1220648-2023-02546 | 17410960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410960&pc=OZD | 07/02/2023 | 01/22/2024 |  | 07/27/2023 |  |  | 07/02/2023 | 07/27/2023 |  |  | 03/31/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209371 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2023 |  |  |  |  | The user facility reported a 75-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced persistent bleeding and a loss in flow to the impella leg during patient support. upon recognizing the loss in doppler signal, a reperfusion cannula was placed from the left common femoral artery to the right superficial femoral artery. the encountered bleeding was initially countered with manual pressure, however, the decision was eventually made to surgically remove the pump and repair the site. | [5 empty] | The investigation into the patient's limb ischemia and access site bleeding has been completed. per the given clinical details, the root cause of the limb ischemia issue was patient condition related to diseased/ small vessels. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35189 | 1220648-2023-02547 | 17410977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410977&pc=OZD | 07/03/2023 | 09/11/2023 |  | 07/27/2023 |  |  | 07/03/2023 | 07/27/2023 |  |  | 12/31/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315412 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/03/2023 |  |  |  |  | The user facility reported a 52-year-old male with heart failure was implanted with an impella 5.5 device for mechanical circulatory support. on day 26 of support, the pump stopped after they where getting patient from commode to chair. they tried to restart and kept getting controller failure alarm. the field service representative had them try new automated impella controller with no resolution. patient had to be taken to or for unplanned impella 5.5 pump exchange. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the medical center returned the impella device and data logs for analysis. the root cause of the pump stop issue was an open electrical line with all three-motor phase line found to be open in red handle. the data logs also noted the abrupt occurrence of the motor current spike and pump stop. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35190 | 1220648-2023-02548 | 17410979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410979&pc=OZD | 07/03/2023 | 09/01/2023 |  | 07/27/2023 |  |  | 07/03/2023 | 07/27/2023 |  |  | 08/31/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267019 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/03/2023 |  |  |  |  | The user facility reported post case, upon moving patient over, team noticed blood leaking out of the red impella plug. the leak was coming from the gray gasket that was holding the purge tubing in place. drips 1-2 drops of blood every 3-4 seconds. patient was given 2 units of blood. | [5 empty] | The investigation into the product damage (hole in catheter) has been completed since the original report was filed. the medical center did retun for investigation the impella pump. two holes were found which allowed for the blood leak at the red plug. the root cause of the product damage issue was use related. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35191 | 1220648-2023-02549 | 17410983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410983&pc=OZD | 06/30/2023 | 09/20/2024 |  | 07/27/2023 |  |  | 06/30/2023 | 07/27/2023 |  |  | 12/31/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024301289 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information noted that the patient survived the explant but subsequently expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 33-year-old female patient was on 5.5 support for around 3 weeks. during support, a spike was noted in the motor current and the placement signals became elevated. over the course of the next couple hours, the impella flows dropped to .1-.5 at p5. the team decided the most likely cause was clot intake. patient was taken to operating room (or) for new 5.5 placement. patient was stable. |  |  |  |  | The investigation into the thrombus has been completed since the original report was filed. the product was returned. data logs were reviewed as well. logs showed about 548 hours into the case, the motor current spiked & the flow dropped immediately to 0.1 l/min that triggered suction alarms. two hours later, the pump was stopped manually & disconnected from the console. pump analysis on benchtop showed biomaterial inside the pump housing & a big clot on the catheter. no other issues were found on the rest of the pump. based on clinical description, data & product analysis, the root cause of low flow was due to biomaterial ingestion. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
35192 | 1220648-2023-02551 | 17410990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410990&pc=OZD | 06/14/2022 | 01/12/2024 |  | 07/27/2023 |  |  | 06/14/2022 | 07/27/2023 |  |  | 01/31/2024 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183694 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2022 |  |  |  |  | The user facility reported a 72-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support prior to mitral valve surgery. the patient developed access site bleeding with a hematoma. the team performed additional surgery to stop the bleeding. the patient was given two units of packed red blood cells and sent to unit. | [5 empty] | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. no product was returned. as per clinical, due to severe tortuosity in lfa it was unable to advance sheath without kinking and later the pigtail catheter into sheath. impella 5.5 inserted via right axillary using surgical method. bleeding was reported from anastomosis site after gwru insertion in cath lab with best practices being followed. access site bleeding- the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35193 | 1220648-2023-02550 | 17410991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17410991&pc=OZD | 06/30/2023 | 01/18/2024 |  | 07/27/2023 |  |  | 06/30/2023 | 07/27/2023 |  |  | 05/31/2025 |  | 07/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351910 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2023 |  |  |  |  | The user facility reported a 58-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma around their axillary impella site prompting an evacuation. the bleeding was noted to be under control following the intervention and patient support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. no product was returned. as per clinical, patient was bleeding at axillary access site after impella insertion and manual pressure was held to control bleeding. the cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided.
35194 | 1220648-2023-02531 | 17396363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17396363&pc=OZD | 05/25/2023 | 01/19/2024 |  | 07/26/2023 |  |  | 06/12/2023 | 07/26/2023 |  |  |  |  | 07/26/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1338086 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/12/2023 |  |  |  |  | The us complainant had an impella cp that was caught under the papillary and was being repositioned but, the pump came across the valve and was unable be re-deliver. the cp was explanted and replaced without any harm or injury. during investigation of this incident, the team analyzed the automated impella controller and found the purge disc retainer to have visible damage. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the console purge disc/flag issues has been completed. the console was returned for evaluation finding visible damage to the console purge assembly. the purge disc retainer was completely broken off confirming the reported failure mode. data logs were reviewed that revealed case wizard was not completed due to the inability to detect purge fluid. therefore the root cause of the reported purge disc/flag issues was due to a broken purge disc retainer and purge flag.
35195 | 1220648-2023-02532 | 17396396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17396396&pc=OZD | 06/27/2023 | 09/26/2023 |  | 07/26/2023 |  |  | 06/27/2023 | 07/26/2023 |  |  | 09/30/2024 |  | 07/26/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274456 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old female patient had suffered a brief cardiac arrest and impella cp was placed in the right femoral artery. one stent was placed in marginal branch of left circumflex artery but the team was not able to adequately revascularize. the following morning, there was no pulse in the distal extremity despite the high vasopressor doses. the medical team decided for escalation of care by using a impella 5.5, however, due to the difficulty with the impella 5.5 delivery a new impella cp was placed via axillary for support. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. as per clinical, fem to fem bypass was done to recover flow on right femoral leg and ischemia resolved with pump explant. the root cause of the limb ischemia issue was patient condition. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events, acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35196 | 1220648-2023-02533 | 17396411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17396411&pc=OZD | 06/29/2023 | 01/11/2024 |  | 07/26/2023 |  |  | 06/29/2023 | 07/26/2023 |  |  | 02/28/2025 |  | 07/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338776 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old female patient undergoing ep/ablation was implanted with an impella cp device for mechanical circulatory support. the patient had access site bleeding at the groin occur and the patient required three units of blood. | [5 empty] | The investigation for the reported access site bleed has been completed. the clinical report was evaluated and provides sufficient information to determine the root cause. it was revealed that the groin angle was altered during a patient transfer. therefore the root cause of the reported access site bleeding issue was user-related as the groin angle was altered. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events, acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35197 | 1220648-2023-02534 | 17399014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17399014&pc=OZD | 07/07/2023 | 04/29/2024 |  | 07/26/2023 |  |  | 07/02/2023 | 07/26/2023 |  |  | 12/31/2024 |  | 07/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306131 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | Pressure Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2023 |  |  |  |  | The user facility reported a 29-year-old male in cardiogenic shock and deteriorating was implanted with an impella cp device for mechanical circulatory support. four days later, the patient the patient had a pump positioning issue which required the medical team to make some adjustments. the following day, the patient was bleeding from the access site which prompted the medical team to give three units of blood. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported high purge pressure, positioning issues, and access site bleed has been completed. the pump was returned for evaluation and was visually inspected. there were two points on the purge tubing that appeared to be clamped or twisted. the purge tubing was cut around 1cm before the blockage areas and de-air leak testing was performed for approximately 15 minutes. the purge cassette ran well as there were no other abnormalities found. the root cause of the high purge pressure is most likely due to a kinked purge line as the tubing appears to have been clamped causing purge fluid to not be able to pass through. the root cause of the access site bleeding is most likely due to use as when the sutures were fixed and replaced, the bleeding resolved. the root cause of the positioning issues is most likely due to use as the pump was repositioned the day before and then it came out of position and was already shallow from moving it so they could not pull it back further. added- d9 device return date, h6 impact code 4628, h6 med dev problem code 3009, 3012 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
35198 | 1220648-2023-02535 | 17399058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17399058&pc=OZD | 07/02/2023 | 01/15/2024 |  | 07/26/2023 |  |  | 07/02/2023 | 07/26/2023 |  |  | 04/30/2025 |  | 07/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336740 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2023 |  |  |  |  | The user facility reported a 66-year-old male awaiting a lvad implant was implanted with an impella 5.5 device for mechanical circulatory support. during insertion of the impella 5.5, there was significant difficulty advancing the impella from the innominate into the ascending aorta. the physician decided to remove the impella and reattempt with rotaglide lubrication. during removal of the pump, the impella motor and cannula became lodged in the graft and the subclavian artery. in the process of attempting removal, the catheter became disconnected from the motor and cannula. the physician removed the introducer from the graft and was able to retrieve the impella motor and cannula from the graft and subclavian artery. the physician replaced the introducer and attempted to regain access to the lv. unfortunately unable to advance the wire and catheter from the innominate into the ascending aorta. due to concerns with possible harm to the patient, the surgeon and team decided to stop further attempts at impella implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the removal issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, there was significant difficulty advancing the impella from the innominate into the ascending aorta due to tortuous vessels and planned to remove pump to reinsert using rotaglide lubrication. during the removal of the pump, the cannula detached from the catheter in the peripheral vessels with excessive force being applied. the cause of the cannula detachment issue was most likely use related with excessive force applied during removal leading the motor housing to be dislodged around the graft. the cause of the delivery issue was due to patient condition related with patient having tortuous vessels.
35199 | 1220648-2023-02536 | 17399128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17399128&pc=OZD | 07/02/2023 | 11/03/2023 |  | 07/26/2023 |  |  | 07/02/2023 | 07/26/2023 |  |  | 04/30/2025 |  | 07/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349536 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2023 |  |  |  |  | The user facility reported a 61-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed a clot at the access site. the clot was aspirated from the peel-away sheath. in addition, the patient developed limb ischemia with no pulses in the leg. the patient was given blood products and the impella cp was explanted. | [5 empty] | The investigation into the access site bleeding and the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the cause of limb ischemia was use related as the doctor left the peel away sheath in. the cause of the access site bleeding cannot be determined since complaint device not returned for investigation and no sufficient data provide to indicate the cause. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35200 | 1220648-2023-02537 | 17399186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17399186&pc=OZD | 07/01/2023 | 09/12/2023 |  | 07/26/2023 |  |  | 07/01/2023 | 07/26/2023 |  |  | 08/31/2024 |  | 07/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268064 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2023 |  |  |  |  | The user facility reported a 79-year-old male who had suffered a cardia arrest was implanted with an impella cp device for mechanical circulatory support. five days after the device was implanted, the patient was shown to have hemolysis. the prognosis for the patient was not good and subsequently the patient expired. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis issue was not determined since the product was not returned for investigation and insufficient clinical information provided. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35201 | 1220648-2023-02526 | 17390225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17390225&pc=OZD | 06/28/2023 | 08/28/2023 |  | 07/25/2023 |  |  | 06/28/2023 | 07/25/2023 |  |  | 04/30/2025 |  | 07/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024336554 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2023 |  |  |  |  | The user facility reported a 68-year-old male in myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the device stopped and the team was unable to restart. the device was exchanged for a new impella 5.5. the patient remained on support for two days. there was no reported harm to the patient. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the medical center has returned for analysis the impella pump and the benchtop analysis revealed that the root cause of the pump stop issue was a significant amount of ingested biomaterial observed around the impeller from halfway inside the cannula through it and also in the inflow cage. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35202 | 1220648-2023-02527 | 17392970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17392970&pc=OZD | 06/28/2023 | 09/11/2023 |  | 07/25/2023 |  |  | 06/28/2023 | 07/25/2023 |  |  | 04/30/2025 |  | 07/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338722 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2023 |  |  |  |  | The user facility reported a 44-year-old male awaiting heart and kidney transplant was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump had a fluid leak from a cracked purge sidearm. as a result of the noted issue, the pump was removed and replaced with another impella 5.5 pump for ongoing support. there was no patient harm. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the purge leak was most likely sidearm yellow luer damage. the characteristic of the damage was unknown since the product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
35203 | 1220648-2023-02528 | 17393056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17393056&pc=OZD | 06/28/2023 | 09/19/2024 |  | 07/25/2023 |  |  | 06/28/2023 | 07/25/2023 |  |  | 03/31/2025 |  | 07/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331071 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/28/2023 |  |  |  |  | The user facility reported a 69-year-old male with heart failure was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma coinciding with impella support. manual pressure and femostop were utilized to counteract the condition, however a hematoma then began to develop around the femostop. the hematoma grew in size and a scrotal edema was identified. the patient was subsequently provided four units packed red blood cells, 4 units of platelets, and 2 units of plasma as a result of the issue. due to the patient's overall condition, the decision was made to withdraw care. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that that post procedure, the patient's activated coagulation time (act) was prolonged at 351. therefore, the root cause of the access site bleed was most likely due the patient¿s act values. the cause was anticoagulation management issue. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
35204 | 1220648-2023-02529 | 17393184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17393184&pc=OZD | 06/28/2022 | 08/15/2023 |  | 07/25/2023 |  |  | 06/28/2022 | 07/25/2023 |  |  | 03/31/2024 |  | 07/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203484 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 62-year-old female undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. the patient was very stable hemodynamically and weaned off/explanted well. patient experienced an embolic stroke and sent to neuro interventional radiololgy (ir) for thrombectomy after explant. the patient had extremely occluded/stenosed carotids and neuro was able to successfully remove the thrombus in ir. based on the findings, the team believes the stroke was unrelated to impella and was due to existing carotid occlusions. patient neurologic status improving. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the issue was most likely patient condition related since the thrombus was observed in carotids and during the support, the patient was hemodynamically stable during the surgery. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
35205 | 1220648-2023-02530 | 17393237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17393237&pc=OZD | 06/17/2023 | 09/15/2023 |  | 07/25/2023 |  |  | 06/30/2023 | 07/25/2023 |  |  | 02/28/2025 |  | 07/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024326735 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2023 |  |  |  |  | The user facility reported a person (unknown age, race, gender) noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the user facility reported positioning issues and stated the device moved to the aortic root. they device was explanted due to the patient's hemodynamic stable situation. no harm was done to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely due to use issue. the failure mode will be monitored and trended.
35206 | 1220648-2023-02517 | 17378325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17378325&pc=OZD | 03/11/2023 | 09/05/2023 |  | 07/24/2023 |  |  | 03/11/2023 | 07/24/2023 |  |  | 11/30/2024 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301322 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported a 50-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during support, the patient showed no pedal pulse on both right and left limb. there was no additional harm to the patient as a result of this incident and the pump implant was completed successfully. the patient remained on support for 11 days. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of limb ischemia was most likely patient condition with both legs being ischemic and unrelated to impella. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
35207 | 1220648-2023-02519 | 17379230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17379230&pc=OZD | 06/27/2023 | 09/18/2023 |  | 07/24/2023 |  |  | 06/27/2023 | 07/24/2023 |  |  | 12/31/2024 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315408 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported a 68 year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the support was ongoing for the patient awaiting possible transplant. during the support there was a cerebral vascular accident and the patient was unable to move the left side. the patient was having seizures and having ct imaging. the belief/presumption was the embolic stroke occurred. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. the product was not returned for review. the cause of the stroke was not determined due to insufficient clinical details and unknown relation to impella.
35208 | 1220648-2023-02518 | 17379840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17379840&pc=OZD | 07/04/2022 | 07/24/2023 |  | 07/24/2023 |  |  | 06/26/2023 | 07/24/2023 |  |  | 02/29/2024 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2023183284 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2023 |  |  |  |  | The user facility reported a 50-year-old male in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during insertion of the 5.5 in 2022 , the team found the 23fr sheath had a crack. the team observed prior to insertion and this prompted the team to instead place a new sheath. the team, a year later in 6/2023, found that there was a publication alleging valve damage attributed to use of impella. the publication notes that moderate aortic regurgitation was seen on transesophageal echocardiography and the team made the choice to replace the valve on the same surgical date as the lvad placement. | [5 empty] | The investigation into the reported heart valve damage has been completed. the medical center did not return any impella products for benchtop analysis. the root cause for heart valve damage remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35209 | 1220648-2023-02520 | 17379882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17379882&pc=OZD | 03/26/2023 | 07/24/2023 |  | 07/24/2023 |  |  | 03/26/2023 | 07/24/2023 |  |  | 02/07/2023 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297297 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/26/2023 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was an alarm stating motor current high followed by pump stop. failure of device occurred within four hours of implant. the team later discovered a left ventricle thrombus that likely was ingested and caused the pump failure. the device emergently explanted and the patient was converted to intra-aortic balloon pump. | [5 empty] | The investigation into the reported pump stop has been completed. the medical center returned the impella products for benchtop analysis. the root cause of the pump stop was due to a build up of biomaterial in the purge gap. the failure modes will be monitored and trended. instructions for use for the related event: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35210 | 1220648-2023-02521 | 17379936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17379936&pc=OZD | 05/01/2023 | 08/21/2023 |  | 07/24/2023 |  |  | 05/01/2023 | 07/24/2023 |  |  | 05/31/2024 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222070 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2023 |  |  |  |  | The user facility a 58-year-old male was implanted with an impella 5.5 device for circulatory support. the device was positioned directly towards the mitral valve apparatus with the inability to reposition away from this area. no harm or premature explant was endured, however, flows were noted as suboptimal. | [5 empty] | The investigation into the low pump flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the low pump flow was improper positioning of the device. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35211 | 1220648-2023-02522 | 17379991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17379991&pc=OZD | 06/27/2023 | 11/08/2023 |  | 07/24/2023 |  |  | 06/27/2023 | 07/24/2023 |  |  | 10/31/2024 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292341 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported a 62-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. there was an observation made of a hematoma at the impella site. additionally, there was a concern for hemolysis. the team treated the patient with rbc infusion for the bleed and hemolyzed labs. the patient did expire on support, but no allegation made of the impella causing the patient outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the access site bleeding and the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs showed several suction alarms were triggered. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the access site bleeding issue was patient condition related since bleeding was from multiple sites.
35212 | 1220648-2023-02523 | 17386537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17386537&pc=OZD | 06/27/2023 | 01/26/2024 |  | 07/24/2023 |  |  | 06/27/2023 | 07/24/2023 |  |  | 12/31/2024 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024285876 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported a 66-year-old in myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a delivery issue involving an impella 5.5 pump. it was noted that the device was unable to advance from the axillary artery into the innominate. during this complication, the graft anastomosis tore and the device had to be removed to repair the graft. due to the complication, the implant was aborted and an impella cp was placed via axillary access for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into delivery issues and vascular damage has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the delivery issue was patient condition related to small vessels. the root cause of the vascular damage - peripheral issue was patient condition related to small vessels.
35213 | 1220648-2023-02524 | 17386543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17386543&pc=OZD | 06/26/2023 | 09/20/2024 |  | 07/24/2023 |  |  | 06/26/2023 | 07/24/2023 |  |  | 07/31/2024 |  | 07/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 50-year-old male patient suffering from acute myocardial infarction and cardiogenic shock required impella placement. roughly five minutes after implant, it was documented that the impella cp pump experienced a spike in placement signal followed by a drop in flows. multiple attempts were made to reposition the pump; however, the device still produced a flat motor current and placement signal. the decision was made to continue with support despite the resulting position and suction alarms. the patient decompensated as a result of their overall condition and later passed away. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The investigation into the low pump flow issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, low pump flow were observed with continuous suction alarm until p1 level and clot was found to present on the inflow and outflow. data logs were reviewed, and motor current spike were found relative to suction alarms causing pump flow issue. the cause of the low pump flow issue was most likely due to biomaterial based on the clinical information and data log review.
35214 | 1220648-2023-02506 | 17370575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17370575&pc=OZD | 06/22/2023 | 09/19/2024 |  | 07/21/2023 |  |  | 06/22/2023 | 07/21/2023 |  |  | 03/31/2025 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342632 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/22/2023 |  |  |  |  | The user facility reported an 87-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding and hematoma at the impella access site which required blood products and device removal. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the major access site bleeding issue was not determined since product was not returned.
35215 | 1220648-2023-02507 | 17372038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372038&pc=OZD | 06/21/2023 | 09/18/2023 |  | 07/21/2023 |  |  | 06/21/2023 | 07/21/2023 |  |  | 04/30/2025 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336777 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2023 |  |  |  |  | The user facility reported a 66-year-old male in non-st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the patient originally had an iabp for support but was escalated to the 5.5 pump. on day three of the support, there was a concern noted of infection/sepsis. the team consulted with infectious diseases (id) for care. the id team began new antibiotics and the pump remained on for support. to date the pump has been supporting the patient for 33 days. the 5.5 has been used well beyond indication. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the reported 'infection/sepsis" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the infection/sepsis and it is unknown relation to impella.
35216 | 1220648-2023-02508 | 17372370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372370&pc=KFM | 02/27/2023 | 01/23/2024 |  | 07/21/2023 |  |  | 02/27/2023 | 07/21/2023 |  |  |  |  | 07/21/2023 |  | IMPELLA CONTROLLER WITH FLOW CONTROL |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | KFM |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1692213 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Failure to Shut Off | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/27/2023 |  |  |  |  | The abiomed field service representative reported that after turning on the automated impella controller console, the system self-check failed and did not allow the operator to turn the console off. the console was changed immediately. there was no patient involvement. | [5 empty] | The investigation into the aic - system self-check error has been completed since the original report was filed. product was returned for evaluation. console (b)(6) was tested, but the issue of failed boot-up was not reproduced. instances of pbm not withdrawing power from impellatronic after aic shutdown were observed, indicating pbm failure. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. follow up with the field service representative has not been completed. upon investigation closure, a supplemental mdr will be filed.
35217 | 1220648-2023-02509 | 17372378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372378&pc=PYX | 06/25/2023 | 01/18/2024 |  | 07/21/2023 |  |  | 06/25/2023 | 07/21/2023 |  |  | 05/31/2025 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024345509 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Female |  |  |  | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Obstruction/Occlusion | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2023 |  |  |  |  | The user facility reported female (unknown age and race) with an orthotopic heart transplant on (b)(6), 2023 was implanted with an impella rp device for mechanical circulatory support. the impella rp was causing a left ventricular outflow tract obstruction when the flow rate was increased. the pump was subsequently explanted as a result of the noted issue and the patient was placed on venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the reported ventricle obstruction/cardiovascular insufficiency has been completed since the original report was filed. no product was returned. as per clinical, the pump was having flow issues with left ventricular outflow tract obstruction noted causing the patient to become hemodynamically unstable. data log was reviewed and low flows at start were observed at p-8 to be less than 3.0 l/min. placement signal trend was observed due to likely poor patient condition with suction alarms. the cause of the low pump flow issue was most likely patient condition related per clinical and log review.
35218 | 1220648-2023-02510 | 17372385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372385&pc=OZD | 07/19/2023 | 11/14/2025 |  | 07/21/2023 |  |  | 07/19/2023 | 07/21/2023 |  |  |  |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251859 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/19/2023 |  |  |  |  | The user facility reported that a 67-year-old male with a history of heart failure required to have heart surgery. the procedure required the assistance of an impella 5.5 pump which had a high purge pressure/system blocked alarm occur. in addition to that, the user also encountered a "purge disk not detected" alarm following the cassette change and an automated impella controller "unable to identify the placement of the purge disk" after the cassette change. despite these issues the doctor was able to complete the procedure successfully. there was no harm to the patient reported because of these incidents. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿. | The investigation has been completed. console logs from the reported day of event are consistent with complaint and show that during purge change on (b)(6) 2023, purge pressure disk could no longer be detected (as evidenced by alarms #402). multiple attempts were made to re-seat the purge cassette (and presumably purge disk as well), but issue remained unresolved. prior to disk detection issues, there were multiple purge pressure high (#408) and purge system blocked (#409) alarms, that were investigated under (b)(4) and considered unrelated to purge disk detection issues. inspection of the console performed in (b)(6) revealed missing/broken purge disk detect flag. b7 other relevant history, including preexisting medical conditions updated. d2 common device name updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated. g3/g4 pma/510(k)number updated.
35219 | 1220648-2023-02511 | 17372389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372389&pc=OZD | 06/24/2023 | 01/19/2024 |  | 07/21/2023 |  |  | 06/24/2023 | 07/21/2023 |  |  | 01/31/2025 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334459 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2023 |  |  |  |  | The user facility reported a 64-year-old male undergoing left heart catheterization was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during impella support. p levels were reduced and volume was provided to initially counteract the condition, however, the decision was ultimately made to wean the patient and explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the patient's hemolysis has been completed. data logs were returned and no discrepancies found. no product was returned for investigation. hemolysis was reported on the first day of support to the clinical information. the motor current spike observed was after multiple hours of hemolysis observed. as per the clinical information provided, the root cause of the hemolysis issue was most likely patient condition related to low volume.
35220 | 1220648-2023-02512 | 17372403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372403&pc=OZD | 06/22/2023 | 01/19/2024 |  | 07/21/2023 |  |  | 06/22/2023 | 07/21/2023 |  |  | 05/31/2025 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024346447 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2023 |  |  |  |  | The user facility reported a 48-year-old male (race unknown) was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. the patient developed a large hematoma at the axillary site. a washout had to be performed to manage the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed was completed. since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, impella 5.5 inserted via right axillary surgical access and hematoma at axillary site was found getting larger. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information.
35221 | 1220648-2023-02513 | 17372425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372425&pc=OZD | 06/26/2023 | 02/07/2024 |  | 07/21/2023 |  |  | 06/26/2023 | 07/21/2023 |  |  | 02/28/2025 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325588 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis. the pump was repositioned and volume levels were adjusted accordingly to initially counteract the hemolysis. the patient was then provided 4000 units of haptoglobin to treat the condition and the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis that occurred during device support has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the hemolysis was patient condition.
35222 | 1220648-2023-02514 | 17372454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372454&pc=OZD | 06/25/2023 | 01/23/2024 |  | 07/21/2023 |  |  | 06/25/2023 | 07/21/2023 |  |  | 11/30/2024 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023283643 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2023 |  |  |  |  | The user facility reported a 72-year-old male with decompensated heart failure was implanted with an impella cp device for mechanical circulatory support prior to coronary artery bypass grafting procedure. during support, the patient's ct scan showed a small cerebellar infarction. the source is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. logs were returned and investigated. no abnormalities were found. per the given clinical information, the root cause of the stroke issue was most likely patient condition related. the failure mode will be monitored and trended.
35223 | 1220648-2023-02515 | 17372466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17372466&pc=OZD | 06/24/2023 | 09/18/2023 |  | 07/21/2023 |  |  | 06/24/2023 | 07/21/2023 |  |  | 03/31/2025 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024346128 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2023 |  |  |  |  | The user facility reported a (b)(6) female (race unknown) in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the access site which worsened with mobilization of the patient. the patient was stable after transfusion of blood units. the device was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the major access site bleeding issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the access site bleeding was most likely due to the user issue of patient movement.
35224 | 1220648-2023-02516 | 17373604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17373604&pc=OZD | 06/26/2023 | 07/21/2023 |  | 07/21/2023 |  |  | 06/26/2023 | 07/21/2023 |  |  | 12/31/2024 |  | 07/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023284563 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 69-year-old female was implanted impella 5.5 was placed for mechanical circulatory support. the impella needed to be torqued to obtain correct orientation. multiple unsuccessful attempts were made. after further discussion, decision made to removed the impella and place iabp. at this time, it was discovered that the left ventricle (lv) perforated. | [5 empty] | The investigation into the reported ventricle perforation has been completed. the impella 5.5 was placed but it needed to be torqued to obtain correct orientation, however it was unsuccessful. a decision was made to remove the impella and place iabp. at the time of replacement, it was discovered that the lv perforated. the physician was aware of the patients small ventricle and understood the risks of the 5.5 placement, and does not blame the impella for the lv perforation. the root cause of the ventricle perforation was a use issue of trying to reposition the 5.5 in the patients small left ventricle. the failure mode will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. furthermore, dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.
35225 | 1220648-2023-02500 | 17360760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17360760&pc=OZD | 05/15/2023 | 09/11/2023 |  | 07/20/2023 |  |  | 05/15/2023 | 07/20/2023 |  |  | 12/31/2024 |  | 07/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325650 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2023 |  |  |  |  | The complainant reported a guidewire tip ¿unfurled¿ upon removal from the device, after impella insertion. the team proceeded with impella support using the inserted pump without any harm to the patient. the guidewire malfunction could have caused harm or injury. | [5 empty] | The impella device was received from the customer and the evaluation of the device has not been completed. upon investigation closure, a supplemental mdr will be filed. | The investigation into the guidewire damage has been completed since the original report was filed. the device was returned and investigated. the benchtop analysis and investigation determined that the root cause of the guidewire damage was use related, or improper technique. the failure mode will be monitored and trended.
35226 | 1220648-2023-02491 | 17360788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17360788&pc=OZD | 06/20/2023 | 08/01/2023 |  | 07/20/2023 |  |  | 08/01/2023 | 07/20/2023 |  |  | 12/31/2024 |  | 07/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024322759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2023 |  |  |  |  | The us complainant had a pump support ongoing for a protect iv patient who had an access site bleed. the hematoma was termed to have a definite relationship. the hematoma had developed when during ecmo removal the pump shifted to the aorta and the impella also was replaced. the bleed prompted the team to infuse one unit of red blood cells to the patient. additionally another adverse event of acute heart failure was attributed to the impella per the primary investigator (pi). the pi states upon the certificate of death, the heart failure was due to the cad and diabetes type ii. | [5 empty] | Upon review of the originally reported complaint, an incorrect serial number was provided by the study team. we have since been able to identify the correct case and serial number that pertains to this patient support, thus, this information has been reflected in sections a and d of this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35227 | 1220648-2023-02501 | 17365254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17365254&pc=OZD | 10/13/2022 | 08/24/2023 |  | 07/20/2023 |  |  | 10/13/2022 | 07/20/2023 |  |  | 06/30/2024 |  | 07/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247438 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported a 51-year-old male noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a small hematoma which required a washout to be removed. the patient successfully bridged to heart transplant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information and the data log provided, the root cause of the access site bleeding issue was not determined since product was not returned and enough clinical information was not given. the failure mode will be monitored and trended.
35228 | 1220648-2023-02502 | 17365283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17365283&pc=OZD | 06/26/2023 | 10/14/2024 |  | 07/20/2023 |  |  | 06/26/2023 | 07/20/2023 |  |  | 02/28/2025 |  | 07/20/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334117 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male patient was transferred from osh with peel away sheath in right femoral. there were no groin issues until patient sat up and cracked the peel away. after that there was continuous bleeding from the peel away that did not slow down as act came down. the two possibilities that were discussed were taking the peel away out and advancing the repo sheath, or removing the whole impella. the patient's ef is 5-10% so the consensus was to remove the peel away and advance the repo sheath. the surgeon who placed ecmo decided to remove everything. once the impella came out the patient's map decreased to 35 requiring an epi drip and fluid resuscitation. | Additional information was reported indicating due to the tuohy not being locked and the blue butterfly not being sutured that the impella moved forward, and the patient had a sustained run of ventricular tachycardia. an echocardiogram confirmed the impella placement then the tuohy was locked in place and the sheath sutured. |  |  |  |  | The investigation into the reported access bleed related to the cracked sheath has been completed. the medical facility did not return any impella products as they were discarded. therefore, the clinical report was evaluated. the evaluation revealed that the most likely root cause of the issue was the peel-away introducer sheath split along the scorelines due to patient sitting up and cracking the sheath. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr was provided in sections b5, d6 and h6.
35229 | 1220648-2023-02503 | 17365416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17365416&pc=OZD | 11/21/2022 | 08/24/2023 |  | 07/20/2023 |  |  | 11/21/2022 | 07/20/2023 |  |  | 06/30/2024 |  | 07/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234682 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella 5.5 device for mechanical circulatory support awaiting orthotopic heart transplant. the patient was in a stable condition had hemolysis and hematuria. the doctor discussed echo and preload as things to consider to resolve. after 8 days on impella support, the patient had to be intubated, pupils were 4 and fixed and the patient had no response to pain. the patient expired. there was no reported product malfunction or allegation toward impella to related it to the patient expiry. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information and the data log provided, the root cause of the hemolysis issue was not determined since product was not returned and inadequate clincal data was shared. the failure mode will be monitored and trended.
35230 | 1220648-2023-02504 | 17368441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17368441&pc=OZD | 12/12/2022 | 08/15/2023 |  | 07/20/2023 |  |  | 12/12/2022 | 07/20/2023 |  |  | 01/31/2024 |  | 07/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183809 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 60-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced lessened flow, mottling and cold extremity of the left leg/foot. because of this, the decision was made to place a new impella. unfortunately due to the patient's anatomy, the impella 5.5 did not cross despite multiple attempts. the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was most likely patient condition related as ischemia resolved after implant.
35231 | 1220648-2023-02505 | 17368455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17368455&pc=OZD | 06/12/2023 | 02/07/2024 |  | 07/20/2023 |  |  | 06/12/2023 | 07/20/2023 |  |  | 04/30/2025 |  | 07/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330361 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2023 |  |  |  |  | The user facility reported that a 67-year-old male with a history of heart failure required to have heart surgery. the procedure required the assistance of an impella 5.5 pump which had a high purge pressure/system blocked alarm occur. in addition to that, the user also encountered a "purge disk not detected" alarm following the cassette change and an automated impella controller "unable to identify the placement of the purge disk" after the cassette change. despite these issues the doctor was able to complete the procedure successfully. there was no harm to the patient reported because of these incidents. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿. | The investigation into the reported high purge pressure has been completed since the original report was filed. high purge pressure occurred on june 12 and june 27th, both instances resolved by using tpa in the purge. the cause most likely was biomaterial issue. no complaint product return. data log was reviewed and high pp was confirmed and was resolved when tpa was added. long purge change preceded the second hpp event on the 27th.
35232 | 1220648-2023-02497 | 17350648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17350648&pc=OZD | 06/22/2023 | 09/12/2023 |  | 07/19/2023 |  |  | 06/22/2023 | 07/19/2023 |  |  | 03/31/2025 |  | 07/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346538 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2023 |  |  |  |  | The user facility reported a 66-year-old female with heart failure was implanted with an impella cp device for mechanical circulatory support. after the impella pump was explanted, the patient experience reduced run-off in the leg. vascular surgery performed a successful thrombectomy procedure. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was biomaterial since issue was resolved with thrombectomy. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35233 | 1220648-2023-02498 | 17352559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17352559&pc=OZD | 06/22/2023 | 01/24/2024 |  | 07/19/2023 |  |  | 06/22/2023 | 07/19/2023 |  |  | 03/31/2025 |  | 07/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024326321 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2023 |  |  |  |  | The user facility reported a (b)(6) male (race unknown) with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient could only move left arm and leg. the patient remained on support until explant 17 days after event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed. the device was not returned for investigation. as per the clinical details provided, the root cause of the reported stroke was not determined since product was not returned, logs were not returned, and insufficient clinical information was provided.
35234 | 1220648-2023-02499 | 17353705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17353705&pc=OZD | 06/24/2023 | 09/26/2024 |  | 07/19/2023 |  |  | 06/24/2023 | 07/19/2023 |  |  | 02/28/2025 |  | 07/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342047 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Unspecified Heart Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/24/2023 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella had had an immediate loss of flow to 0.1 occur. the patient coded, return of spontaneous circulation was not achieved and the patient subsequently expired. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. product was returned in damaged condition as the patient cable was cut off and missing. visual inspection found no issues on the pump. the pump run test was unable to be performed due to missing patient cable. as the patient coded and the low flow occurred afterward, the root cause of low flow was most likely due to patient condition. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in h1. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35235 | 1220648-2023-02487 | 17342411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17342411&pc=OZD | 06/20/2023 | 09/18/2023 |  | 07/18/2023 |  |  | 06/20/2023 | 07/18/2023 |  |  | 02/28/2025 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337403 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2023 |  |  |  |  | The user facility reported an 81-year-old female noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. there was a blood ooze from the access site, and the team treated the site with pressure. this did not fully remedy the issue and as such they evaluated and diagnosed a retroperitoneal hematoma. the bleeding and hematoma were treated by 3 units of red blood cells and pump explant. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. product was not returned for review. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35236 | 1220648-2023-02488 | 17342923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17342923&pc=OZD | 03/11/2023 | 10/14/2024 |  | 07/18/2023 |  |  | 03/11/2023 | 07/18/2023 |  |  | 07/31/2024 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250079 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/11/2023 |  |  |  |  | The complainant reported that the doctor was unable to advance repositioning sheath into secure impella properly. upon arrival, the peel away sheath was found to be left in place with repositioning sheath partially inserted into the peel away sheath. site was dressed with multiple transparent dressings, blue suture hub sutured to patient¿s mid-thigh. moderate amount of bleeding noted at the insertion site. | The user facility reported a 45-year-old male with heart failure was implanted with an impella cp device for mechanical circulatory support. the patient developed leg ischemia and was treated with heparin. the patient was escalated to an impella 5.5 device. |  |  |  |  | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was patient condition related with peel away sheath left in place to maintain hemostasis of cfa due to patient habitus. | Additional information for the initial mfr was provided in sections b1, b5, h6, h10 as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35237 | 1220648-2023-02489 | 17342991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17342991&pc=OZD | 06/20/2023 | 05/23/2025 |  | 07/18/2023 |  |  | 06/20/2023 | 07/18/2023 |  |  | 02/28/2025 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342266 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 06/20/2023 |  |  |  |  | Additional information was reported that the patient partial thromboplastin time had been high. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported an 82-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp limb had signs of ischemia. on the 2nd day of support, the team performed a fasciotomy. |  |  |  |  | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Corrected information for the initial mfr 1220648-2023-02489 was provided. | Per the provided clinical information, the root cause of the limb ischemia issue was patient condition related. the cause of the thrombocytopenia was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35238 | 1220648-2023-02490 | 17344518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17344518&pc=OZD | 06/21/2023 | 12/07/2023 |  | 07/18/2023 |  |  | 06/21/2023 | 07/18/2023 |  |  | 03/31/2025 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327054 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2023 |  |  |  |  | The user facility reported a 72-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support post procedure. on day three of the support the oozing from the site prompted intervention. the team placed a jp drain and were giving one unit of blood/every 12 hours. the team has not explanted the pump but, anticoagulation is being monitored. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's hemorrhage/bleeding has been completed. the cause of the access site bleeding cannot be determined since the complaint device not returned for investigation and insufficient clinical details were provided.
35239 | 1220648-2023-02492 | 17347556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17347556&pc=OZD | 06/21/2023 | 07/19/2024 |  | 07/18/2023 |  |  | 06/21/2023 | 07/18/2023 |  |  | 03/31/2024 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211667 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2023 |  |  |  |  | The user facility reported a 59-year-old female in st elevated myocardial infarction was to be implanted with an impella cp device for mechanical circulatory support. the insertion was aborted due to patient anatomy of extensive peripheral artery disease in iliac arteries and descending aorta. during removal of the device, dissection of iliac was noted. the patient was brought to the or for vascular repair and potential axillary insertion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for vascular damage has been completed. the impella device was not returned for evaluation. product was not returned for investigation. as per clinical details, patient had extensive peripheral artery disease (pad) in the iliacs and descending aorta. impella couldn't pass the iliac lesion, was removed, and a dissection of the iliac was noted. the patient was taken to the operating room for vascular repair and potential axillary impella insertion, but the procedure was aborted due to the patient's anatomy. the root cause of the vascular damage issue was most likely patient condition related to small/diseased vessels. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank in the initial report.
35240 | 1220648-2023-02493 | 17347814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17347814&pc=OZD | 06/20/2023 | 02/06/2024 |  | 07/18/2023 |  |  | 06/20/2023 | 07/18/2023 |  |  | 02/28/2025 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024321945 |  | 0048-0024 JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2023 |  |  |  |  | The user facility reported a 48-year-old male noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient experienced an access site bleed coinciding with impella support. the bleed was surgically treated, and support continued with the implanted pump. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the cause of the access site bleeding- major issue cannot be determined since the complaint device not returned for investigation and no sufficient data provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35241 | 1220648-2023-02494 | 17347815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17347815&pc=OZD | 06/20/2023 | 09/01/2023 |  | 07/18/2023 |  |  | 06/20/2023 | 07/18/2023 |  |  | 12/31/2024 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317959 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2023 |  |  |  |  | The user facility reported a 42-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in their impella leg during pump support. an external fem-fem bypass was completed to return flow to the leg. impella support continued following the noted intervention. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was patient condition related since fem-fem bypass resolved the issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35242 | 1220648-2023-02495 | 17347816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17347816&pc=OZD | 06/21/2023 | 08/11/2023 |  | 07/18/2023 |  |  | 06/21/2023 | 07/18/2023 |  |  | 12/31/2024 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318689 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2023 |  |  |  |  | The user facility reported a 72-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed limb ischemia and an external fem-fem bypass was performed. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35243 | 1220648-2023-02496 | 17347832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17347832&pc=OZD | 03/11/2023 | 09/26/2023 |  | 07/18/2023 |  |  | 03/11/2023 | 07/18/2023 |  |  | 09/30/2024 |  | 07/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278990 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2023 |  |  |  |  | The user facility reported persistent suction issues involving an impella cp pump during patient support. it was stated that the user was unable to raise support levels above p1 without triggering a suction alarm. due to the ongoing issue and the need for continued support, the decision was made to explant the pump and escalate the patient to impella 5.5 support. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. per the clinical information provided, the root cause of the low pump flow issue was not determined since product was not returned and data logs were not returned. the failure mode will be monitored and trended.
35244 | 1220648-2023-02483 | 17327192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17327192&pc=OZD | 04/09/2023 | 09/26/2023 |  | 07/14/2023 |  |  | 04/09/2023 | 07/14/2023 |  |  | 06/30/2024 |  | 07/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247288 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2023 |  |  |  |  | The user facility reported a 69-year-old male undergoing high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during implant using single access, blood was leaking through the seal around the guide. the 14f catheter was occlusive so internal bypass was placed. the physician removed single access, then advanced repositioning sheath and accessed the contralateral side for guide. an internal bypass replaced at the conclusion of the procedure. the pump was explanted a day later after implant. patient stable. | [5 empty] | The investigation into the reversible limb ischemia and the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. per the provided clinical information, the root cause of the limb ischemia issue was patient condition related to occluded catheter and low volume of blood. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical information provided. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35245 | 1220648-2023-02484 | 17328560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17328560&pc=OZD | 03/07/2023 | 08/11/2023 |  | 07/14/2023 |  |  | 03/07/2023 | 07/14/2023 |  |  | 09/30/2024 |  | 07/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2023 |  |  |  |  | The user facility reported a 63-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support prior to surgery. the physician had concern for distal limb ischemia and impella cp was explanted. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia related to the patient's small/diseased vessels. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35246 | 1220648-2023-02486 | 17328656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17328656&pc=OZD | 03/10/2023 | 08/04/2023 |  | 07/14/2023 |  |  | 03/10/2023 | 07/14/2023 |  |  | 11/30/2024 |  | 07/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2023 |  |  |  |  | The user facility reported a 55-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in their left leg during impella support. the leg was initially noted to be mottled with no pulse and it was advised to the staff that the patient be escalated to an impella 5.5 for increased support. no immediate action was taken, and two days later, the foot was noticeably colder and mottled. a ct confirmed the ischemic leg coinciding with an ischemic gut due to multiple organ failure (mof). a fem-fem bypass was attempted however flow could not be re-established to the leg and the patient was taken to the operating room for limb amputation. a concurrent laparotomy revealed that the patient's gut was not compatible with life and the decision was made to withdraw care. the patient passed away from mof upon removal of support. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the cause of the limb ischemia issue was most likely due to patient condition related based on the clinical information provided that patient had ischemic gut. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35247 | 1220648-2023-02485 | 17329056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17329056&pc=OZD | 03/10/2023 | 08/11/2023 |  | 07/14/2023 |  |  | 03/10/2023 | 07/14/2023 |  |  | 11/30/2024 |  | 07/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303088 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2023 |  |  |  |  | The user facility reported a 70-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had bleeding from the impella site which required packed red blood cells but patient was stable. the pump was explanted three days after implant. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia since the patient was bleeding from multiple sites. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35248 | 1220648-2023-02471 | 17317491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17317491&pc=OZD | 06/16/2023 | 10/13/2023 |  | 07/13/2023 |  |  | 06/16/2023 | 07/13/2023 |  |  | 04/30/2025 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332843 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 33 YR |  | Female |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup |  |  |  | 06/16/2023 |  |  |  |  | The user facility reported a 33-year-old female (race unknown) was implanted with an impella 5.5 device for mechanical circulatory support. upon removal of impella in or, the device was stuck at innominate and axillary bifurcation. decision made by the doctor to do sternotomy and place central va ecmo. sternotomy was performed to help facilitate the removal of the impella, but to also place va ecmo. impella was then removed by strict tension and wiggling of catheter out through axillary graft. | [5 empty] | The investigation into the reported "low pump flow and removal issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the data logs were returned for a period of ten days and not for the entire case. the low pump flow issue occurred on 16th day of the case and data logs were provided for that duration but no abnormalities were found. there is no evidence in the logs pointing towards low pump flow hence, there was no issue found during the case. the device functioned/operated to specification. the root cause of the removal issue was use related to excessive force. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35249 | 1220648-2023-02472 | 17322357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322357&pc=OZD | 04/07/2023 | 02/02/2024 |  | 07/13/2023 |  |  | 04/07/2023 | 07/13/2023 |  |  | 11/30/2024 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Renal Impairment | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2023 |  |  |  |  | The user facility reported a 47-year-old male undergoing high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient had lab values indicating loss of platelets and loss of renal function. the patient had to receive an unknown amount of units of blood and the pump had to be explanted. | [5 empty] | The investigation for the reported thrombocytopenia and access site bleeding has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. clinical details states that the activated coagulation time (act) was noted at 240 seconds initials and then found later to be 407 seconds with complications. the cause of the thrombocytopenia issue was most likely patient condition related and unknown relation to impella due to insufficient clinical details. the cause of the access site bleeding issue was most likely anticoagulation management related due to high patient act values per clinical review. section h6 : the clinical codes in section h6, was inadvertently left out in the initial report which has been added to this report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35250 | 1220648-2023-02473 | 17322359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322359&pc=OZD | 06/16/2023 | 08/23/2023 |  | 07/13/2023 |  |  | 06/16/2023 | 07/13/2023 |  |  | 04/30/2025 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024334914 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2023 |  |  |  |  | The user facility reported a male of unknown age underwent mitral valve replacement. the patient was implanted with an impella 5.5 for circulatory support post-procedure. during support, the pump stopped. attempts to restart the device were unsuccessful. when the pump was removed, a pledgeted suture was noted in the outflow cage and thought to have been a cause of the pump stop. an intra-aortic balloon pump was placed and the patient was returned to the icu with stable hemodynamics. | [5 empty] | (h3) the investigation into the reported 'pump stop" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed mitral valve repair (mvr) was performed under surgical mode. when the aortic clamp was released and the support level was raised, the motor current increased and the pump stopped. an attempt was made to restart, but it did not start and was removed. when the pump was removed, there was foreign material caught in the outflow cage, which was determined to be the cause of the pump stoppage. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35251 | 1220648-2023-02474 | 17322360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322360&pc=OZD | 06/17/2023 | 09/19/2024 |  | 07/13/2023 |  |  | 06/17/2023 | 07/13/2023 |  |  | 03/31/2025 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343810 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/17/2023 |  |  |  |  | Additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 65-year-old male was implanted with an impella cp device for mechanical circulatory support pre-procedure. the patient started to bleed from the access site. the patient required 12 units of blood and the condition was stabilized. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35252 | 1220648-2023-02476 | 17322362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322362&pc=OZD | 06/19/2023 | 02/07/2024 |  | 07/13/2023 |  |  | 06/19/2023 | 07/13/2023 |  |  | 03/31/2025 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343775 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2023 |  |  |  |  | The user facility reported a 79-year-old female (race unknown) was implanted with an impella cp device for mechanical circulatory support pre mitral clip procedure. the physician left the peel-away sheath in, patient had continued to bleed warranting two units of packed red blood cells. patient was weaned and explanted and patient recovered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding issue has been completed. the device was not returned for investigation. the cause of the access site bleeding was most likely use related because the physician left the peel away sheath in place.
35253 | 1220648-2023-02475 | 17322363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322363&pc=OZD | 06/19/2023 | 09/19/2024 |  | 07/13/2023 |  |  | 06/19/2023 | 07/13/2023 |  |  | 04/30/2025 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024336557 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/19/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 47-year-old male (race unknown) was implanted with an impella 5.5 device for mechanical circulatory support. the pump stopped during support. several unsuccessful restart attempts were made. changing the automated impella controller also did not resolve the issue. the plan is to change impella 5.5. |  |  |  |  | The investigation into the pump stop has been completed since the original report was filed. the medical center did return the impella pump and purge cassette. based on the clinical information provided, the root cause of the pump stop issue was an open electrical line found with all 3 motor phase cables disconnected in the pump red plug, due to excessive force which is a use error. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35254 | 1220648-2023-02477 | 17322365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322365&pc=OZD | 06/18/2023 | 02/05/2024 |  | 07/13/2023 |  |  | 06/18/2023 | 07/13/2023 |  |  | 05/31/2023 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342482 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2023 |  |  |  |  | The user facility reported a 53-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed hematoma at the incision site, limb ischemia and blood clots. the md decided to surgically remove the pump, repair the vessels and remove the blood clots. flow was restored in the patient's leg and the patient was taken to the room for recovery. | [5 empty] | The investigation into the reported access site bleed, thrombus, and limb ischemia has been completed. the device was not returned for evaluation. however, clinical details was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely due to anticoagulation management. the root cause of the limb ischemia issue was most likely due to the patients condition as a clot was found in right iliac artery and from somewhere distal to the right femoral artery. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35255 | 1220648-2023-02478 | 17322366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322366&pc=OZD | 06/19/2023 | 01/19/2024 |  | 07/13/2023 |  |  | 06/19/2023 | 07/13/2023 |  |  | 04/30/2025 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024347334 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2023 |  |  |  |  | The user facility reported a patient (unknown age, gender, race) noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient had a right brachiofacial paresis on the ward. the exact onset of the symptom cannot be determined. no information available yet if the stroke has been ischemic or hemorrhagic. | [5 empty] | The investigation into the stroke/cva issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the stroke issue was patient condition related since the stroke occurred a day after impella support and therefore is unrelated to impella. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35256 | 1220648-2023-02479 | 17322370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322370&pc=OZD | 04/11/2023 | 08/30/2023 |  | 07/13/2023 |  |  | 04/11/2023 | 07/13/2023 |  |  | 05/31/2024 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2023 |  |  |  |  | The user facility reported a 79-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced loss of distal pulses in the impella inserted extremity. therefore, the pump was removed. | [5 empty] | The investigation into the reversible limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the ischemia issue was patient condition related to the diseased/ small vessels. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35257 | 1220648-2023-02480 | 17322394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322394&pc=OZD | 03/07/2023 | 02/02/2024 |  | 07/13/2023 |  |  | 03/07/2023 | 07/13/2023 |  |  | 11/30/2024 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302966 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2023 |  |  |  |  | The user facility reported a 78-year-old male in st elevated myocardial infarction / cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding from their sites during impella support. the patient was provided two units of blood to counteract the blood loss and support continued with the implanted pump. | [5 empty] | The investigation into the access site bleeding has been completed. the device was not returned for investigation. as per the given clinical information, the root cause of the access site bleeding was determined to be patient condition as the patient was bleeding from multiple sites. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35258 | 1220648-2023-02481 | 17322408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322408&pc=OZD | 06/15/2023 | 01/19/2024 |  | 07/13/2023 |  |  | 06/15/2023 | 07/13/2023 |  |  | 02/28/2025 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321265 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2023 |  |  |  |  | The user facility reported a 62-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had decreased perfusion to the impella limb during pump support. due to the noted issue, a distal perfusion sheath was placed on the impella access side. the condition continued to worsen despite the perfusion sheath and the decision was made to explant the pump. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the root cause of the ischemia issue was patient condition related to a clot in the distal to the right femoral artery which was the artery used for pump access. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35259 | 1220648-2023-02482 | 17322419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17322419&pc=OZD | 06/18/2023 | 09/05/2023 |  | 07/13/2023 |  |  | 06/18/2023 | 07/13/2023 |  |  | 12/31/2024 |  | 07/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324671 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2023 |  |  |  |  | The user facility reported a 76-year-old male (race unknown) in st elevated myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was accidentally removed during 14 fr sheath peel-away removal. stent placed to left common femoral artery and patient was transfused one unit of packed red blood cells. an intra-aortic balloon pump was placed. | [5 empty] | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the vascular damage issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35260 | 1220648-2023-02469 | 17308862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17308862&pc=OZD | 03/12/2023 | 09/20/2024 |  | 07/12/2023 |  |  | 03/12/2023 | 07/12/2023 |  |  | 09/30/2024 |  | 07/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/12/2023 |  |  |  |  | The user facility reported a 50-year-old male with multi vessel disease was implanted with an impella cp device for mechanical circulatory support. the patient had strokes while on impella support. the impella pump was explanted two days after implant and the patient expired. the patient expired after 53.93 hours of support. the field representative was not able to determine whether the strokes were due to hemorrhagic or embolic/ischemic conditions. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the stroke/cva issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient had a large ventricular shift and multiple strokes in the brain. data logs were returned, and no motor current spike was found relative to ischemic stroke. the cause of the ischemic stroke issue was most likely patient condition related. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02469 was provided in sections b2, b5, d6a, d6b, h1, and h6.
35261 | 1220648-2023-02470 | 17310547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17310547&pc=OZD | 03/08/2023 | 08/24/2023 |  | 07/12/2023 |  |  | 03/08/2023 | 07/12/2023 |  |  | 10/31/2024 |  | 07/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023269047 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2023 |  |  |  |  | The user facility reported a 75-year-old female with intra-aortic balloon pump was deteriorating and was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient oozing from the access site. the patient required 2 units of blood as a result. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the root cause of the access site bleeding was not determined since product was not returned. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35262 | 1220648-2023-02457 | 17297785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17297785&pc=OZD | 03/06/2023 | 01/24/2024 |  | 07/11/2023 |  |  | 03/06/2023 | 07/11/2023 |  |  | 12/31/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303378 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 61-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the site of the impella cp had an access site ooze that was treated by multiple units of blood/blood products. the repo-sheath bleed was due to patient movement in bed. it was noted the patient would not lay still in the bed. the impella cp was explanted | [5 empty] | The investigation into the access site bleeding has been completed. the device was not returned for investigation. as per the clinical details provided, the root cause of the access site bleeding issue was most likely patient movement. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35263 | 1220648-2023-02458 | 17298455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17298455&pc=OZD | 03/06/2023 | 07/11/2023 |  | 07/11/2023 |  |  | 03/06/2023 | 07/11/2023 |  |  | 12/31/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302740 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 79-year-old female in cardiac arrest was implanted with an impella cp device for mechanical circulatory support. during support, the patients' foot had gone cold, became mottled, and they were unable to find a pulse. rn communicated that md was planning on removing the impella support despite patient requiring max impella flows, multiple vasoactive infusions. us/doppler of left lower extremity performed, per report, there was evidence of flow, but significant peripheral artery disease, with no clot present. impella removal and balloon pump placement for continued mcs support. device removed in or, pt. returned to icu. care withdrawn, patient expired. | [5 empty] | The investigation into the reported limb ischemia has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35264 | 1220648-2023-02459 | 17299429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17299429&pc=OZD | 03/13/2023 | 08/09/2023 |  | 07/11/2023 |  |  | 03/13/2023 | 07/11/2023 |  |  | 06/30/2023 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095651 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2023 |  |  |  |  | The user facility reported an 84-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient lost distal pulses in right femoral artery (impella side). an external fem/fem bypass was initiated. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35265 | 1220648-2023-02460 | 17300635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17300635&pc=OZD | 05/09/2023 | 09/05/2023 |  | 07/11/2023 |  |  | 05/09/2023 | 07/11/2023 |  |  | 10/31/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279419 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2023 |  |  |  |  | The user facility reported a (b)(6) female who coded during diagnostic cath was implanted with an impella cp device for mechanical circulatory support. upon insertion of the impella cp into the 14fr sheath, the pump would not advance upon exiting the sheath. after removing the impella, the angiogram demonstrated dissection of the right femoral artery (rfa). the 14fr short sheath was exchanged for a 14fr long sheath and impella was reinserted and implanted without problems. rfa site was sutured and dressed using best practices and the patient was sent to icu. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the vascular damage issue was not determined since the product was not returned and with insufficient clinical information. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35266 | 1220648-2023-02461 | 17300765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17300765&pc=OZD | 03/05/2023 | 09/06/2023 |  | 07/11/2023 |  |  | 03/05/2023 | 07/11/2023 |  |  | 09/30/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279525 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2023 |  |  |  |  | The user facility reported a 60-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after the device implant, the patient showed to have hemolysis and decision was made to explant impella. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the root cause of the hemolysis was not determined since product was not returned and insufficient clinical information provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
35267 | 1220648-2023-02462 | 17303268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17303268&pc=OZD | 03/09/2023 | 10/02/2023 |  | 07/11/2023 |  |  | 03/09/2023 | 07/11/2023 |  |  | 06/30/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234699 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2023 |  |  |  |  | The user facility reported a 74-year-old male requiring support post aortic valve replacement was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed hemolysis during the course of impella support. due to the noted condition, the patient was provided 2 units packed red blood cells, 1 unit of platelets, and 2 units of fresh frozen plasma. the pump was explanted later that day after the family was made aware of the patient's overall condition. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible and based on the clinical information, the root cause of the hemolysis was not determined. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
35268 | 1220648-2023-02463 | 17303273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17303273&pc=OZD | 03/13/2023 | 09/29/2023 |  | 07/11/2023 |  |  | 03/13/2023 | 07/11/2023 |  |  | 04/30/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218624 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2023 |  |  |  |  | The user facility reported a 24-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during support. staff chose to escalate the patient to the impella 5.5. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, it was found the pump to be flipped upside down in left ventricle and with the plasma free hgb levels elevated complicating the repositioning. data logs were returned and suction, impella position unknown alarms were observed. the cause of the hemolysis issue was most likely improper positioning of the device based on the clinical information and data log review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35269 | 1220648-2023-02464 | 17303280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17303280&pc=OZD | 03/06/2023 | 01/22/2024 |  | 07/11/2023 |  |  | 03/06/2023 | 07/11/2023 |  |  | 10/31/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283411 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 65-year-old male (race unknown) undergoing surgery to left descending artery was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia coinciding with impella support. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia issue has been completed. the device was not returned for investigation. as per clinical information provided, the root cause of the limb ischemia issue was patient condition related to small/ diseased vessels.
35270 | 1220648-2023-02465 | 17303291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17303291&pc=OZD | 06/15/2023 | 09/20/2023 |  | 07/11/2023 |  |  | 06/15/2023 | 07/11/2023 |  |  | 04/30/2025 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332941 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2023 |  |  |  |  | The user facility reported a 61-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient¿s urine was darkening and both ldh and creatine levels increased. the physician identified the ingestion of a clot by the impella as the cause of the sudden hemolysis. this resulted in limiting the impella¿s ability to run above p-4 without resulting in hematuria. the decision was made to exchange the impella at the same surgical site. the hemolysis was resolved by the removal of the initial impella and replacement with a new impella. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis issue cannot be determined since the complaint 5.5 pump was not returned for investigation and lack of clinical data. data logs reviewed. no motor current or other abnormality observed that indicated the cause of the issue or ingested biomaterial. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35271 | 1220648-2023-02466 | 17303306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17303306&pc=OZD | 03/07/2023 | 09/01/2023 |  | 07/11/2023 |  |  | 03/07/2023 | 07/11/2023 |  |  | 10/31/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294016 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2023 |  |  |  |  | The user facility reported a 68-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis and the device had to be explanted the following day. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis issue was not determined since product was not returned and the relation to impella is unknown based on the clinical information provided. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35272 | 1220648-2023-02467 | 17303309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17303309&pc=OZD | 03/09/2023 | 09/26/2024 |  | 07/11/2023 |  |  | 03/09/2023 | 07/11/2023 |  |  | 10/31/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023276844 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/09/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing high risk coronary artery bypass grafting and valve surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient had trouble coming off the impella support pump. the patient was placed back on the pump to repair a hematoma and bleeding, however, the md was unable to repair multiple bleeding areas. the patient expired in the operating room. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections h1. | The investigation into the access site bleeding ¿ major has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the root cause of the access site bleeding issue was most likely patient condition related and unrelated to impella due to multiple bleeding areas.
35273 | 1220648-2023-02468 | 17303314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17303314&pc=OZD | 03/06/2023 | 08/11/2023 |  | 07/11/2023 |  |  | 03/06/2023 | 07/11/2023 |  |  | 08/31/2024 |  | 07/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265389 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 62-year-old male noted as high risk surgery was to be implanted with an impella 5.5 prior to surgery. delivery of the impella 5.5 was unsuccessful due to the anatomy of the patient. the team implanted an impella cp. once transferred to the floor, the patient lost flow down the leg. the device had to be explanted and manual pressure was used to close the access site. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia is due to the patient's condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35274 | 1220648-2023-02427 | 17286671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17286671&pc=OZD | 06/10/2023 | 02/12/2024 |  | 07/10/2023 |  |  | 06/10/2023 | 07/10/2023 |  |  | 12/31/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297564 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2023 |  |  |  |  | The user facility reported 56-year-old male with multi vessel disease including severe left main artery disease was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding issues coinciding with impella support. the patient was considered hemodynamically unstable and was resuscitated with packed red blood cells, albumin, and fluids. the onsite physician stated that they believed the impella was damaging the platelets, resulting in the ongoing bleeding. the pump was subsequently removed as a result of the noted issue | [5 empty] | The investigation into the access site bleeding has been completed. the device was not returned for investigation. data logs shows several suction alarms during case. the cause of the access site bleeding issue was most likely patient condition related and determined to be unrelated to impella since bleeding was reported from non-impella site. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35275 | 1220648-2023-02434 | 17286989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17286989&pc=OZD | 06/11/2023 | 07/10/2023 |  | 07/10/2023 |  |  | 06/11/2023 | 07/10/2023 |  |  | 12/31/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1693613 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/11/2023 |  |  |  |  | The user facility reported a 79-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. at conclusion of case, impella weaned and explanted. the 14f sheath was removed, but perclose unsuccessful in obtaining hemostasis and patient developed hematoma. manual pressure and a balloon tamponade insertion was performed to achieve homeostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35276 | 1220648-2023-02435 | 17288080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17288080&pc=OZD | 06/11/2023 | 11/08/2023 |  | 07/10/2023 |  |  | 06/11/2023 | 07/10/2023 |  |  | 02/28/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338478 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing advanced cardiac support was implanted with an impella cp device for mechanical circulatory support. during support, the peel-away sheath was left in the femoral artery. the sheath was tipped and allowed for blood loss to occur. the team infused two units of red blood cells. of note the cp was placed after the team had placed ecmo and bypass sheath for limb perfusion. the decision was made to explant the impella cp, exchange to a new peel-away sheath and implant a new impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the cause of the issue was use related since the introducer was left in the patient and the repo unit sheath was inserted into the introducer valve.
35277 | 1220648-2023-02436 | 17289357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17289357&pc=OZD | 06/14/2023 | 09/26/2023 |  | 07/10/2023 |  |  | 06/14/2023 | 07/10/2023 |  |  | 04/30/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332424 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported a 62-year-old female was to be upgraded from intra-aortic balloon pump (iabp) to an impella 5.5 device for circulatory support. the support was ongoing for two days when the team made choice to explant the pump due to hemolysis concerns. there was thought that the mitra-clip could have been the cause of the hemolysis. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. data logs do not indicate any abnormal events. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35278 | 1220648-2023-02437 | 17289368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17289368&pc=OZD | 06/14/2023 | 01/23/2024 |  | 07/10/2023 |  |  | 06/14/2023 | 07/10/2023 |  |  | 12/31/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298254 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported a 61-year-old male had impella cp support ongoing when the pump had suction and low flows. the team was unable to remedy the issues, so explanted the cp and delivered a new 5.5 via the axillary artery access. upon closure of the femoral access of the impella cp, the team observed clot to have cascaded down to the superficial femoral artery and popliteal. intervention was performed and thrombolytics were considered but not used due to fresh post-op nature. | [5 empty] | The investigation into the low pump flow and the thrombus has been completed. the device was not returned for benchtop evaluation and investigation. however, data logs were returned and motor current spike corresponding to suction was found. the root cause of the low pump flow issue was most likely biomaterial. the root cause of the thrombus issue was patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35279 | 1220648-2023-02438 | 17290927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17290927&pc=OZD | 06/14/2023 | 03/10/2025 |  | 07/10/2023 |  |  | 06/14/2023 | 07/10/2023 |  |  | 03/31/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024320161 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported a 62-year-old male in end stage heart failure was implanted with an impella 5.5 device for mechanical circulatory support. the patient had positive blood cultures. discussion was to plan lvad when cultures cleared, but they were concerned the 5.5 is the source. impella was turned off and patient was no longer lvad candidate due to head bleed. patient deteriorated after the impella explant and the patient expired when care was withdrawn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | An infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician physical examination findings results of special investigations : laboratory test results & imaging studies microbiological culture and sensitivity results response to previous treatments and therapies any relevant epidemiological information or exposure history concomitant devices in use preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. this report is being filed as part of a retrospective review of historical records.
35280 | 1220648-2023-02439 | 17292618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17292618&pc=OZD | 02/24/2023 | 09/14/2024 |  | 07/10/2023 |  |  | 02/24/2023 | 07/10/2023 |  |  | 11/30/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023279022 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/24/2023 |  |  |  |  | The user facility reported a 69-year-old was escalated to an impella 5.5 device from an impella cp for mechanical circulatory support. the physician noted the tortuous subclavian artery made implant challenging. the patient became very hypotensive requiring pressers and packed red blood cells. the patient experienced an access site bleed. the impella 5.5 was explanted. | It was noted care was withdrawn and the patient expired while on impella cp support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the access site bleeding has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible and based on the clinical information, the root cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02439 was provided.
35281 | 1220648-2023-02440 | 17292777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17292777&pc=OZD | 03/31/2023 | 02/07/2024 |  | 07/10/2023 |  |  | 03/31/2023 | 07/10/2023 |  |  | 09/30/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281615 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2023 |  |  |  |  | The user facility reported an 80-year-old male in non-st elevated myocardial infarction and ventricular tachycardia was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing from the impella access site. blood products were given to the patient. the patient was weaned from the impella cp with heart improvement two days after implant. | [5 empty] | The investigation for the reported access site bleed has been completed. the product was not returned for investigation. the cause of the access site bleeding issue was not determined since the product was not returned and sufficient clinical information was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35282 | 1220648-2023-02441 | 17292857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17292857&pc=OZD | 03/06/2023 | 09/29/2023 |  | 07/10/2023 |  |  | 03/06/2023 | 07/10/2023 |  |  | 06/30/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238258 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 61-year-old female in supraventricular tachycardia was implanted with an impella cp device for mechanical circulatory support. the impella was removed due to ventricular tachycardia storm. the patient survived on continued ecmo support. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the ventricular tachycardia issue was not determined due to insufficient clinical information and based on data log review. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35283 | 1220648-2023-02442 | 17292908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17292908&pc=OZD | 05/18/2023 | 01/19/2024 |  | 07/10/2023 |  |  | 05/18/2023 | 07/10/2023 |  |  | 06/30/2023 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022106026 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2023 |  |  |  |  | The user facility reported, a 66-year-old male noted as high risk percutaneous cardiac insertion (hrpci). was implanted with an impella cp device for mechanical circulatory support. following the insertion of the device, the patient sat up and bent both legs while, experiencing ventricular tachycardia. and subsequently, caused a perforation in the left femoral artery. the perforation was discovered, when the pump was explanted following the scheduled pci. vascular surgery was consulted to repair the noted perforation. | [5 empty] | The impella device was discarded by the customer. and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "be careful not to pull on the impella catheter, when transferring a patient from one bed to another". ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs were returned and investigated. the root cause of the vascular damage issue was use related since vascular damage was noted during patient movement. the failure mode will be monitored and trended.
35284 | 1220648-2023-02443 | 17293098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293098&pc=OZD | 03/07/2023 | 01/19/2024 |  | 07/10/2023 |  |  | 03/07/2023 | 07/10/2023 |  |  | 09/30/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281381 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2023 |  |  |  |  | The user facility reported a 31-year-old female with heart failure was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing from their access site during impella support. the patient was provided two units of blood to counteract the ooze. patient support continued with the implanted pump throughout the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided.
35285 | 1220648-2023-02444 | 17293155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293155&pc=OZD | 06/11/2023 | 02/02/2024 |  | 07/10/2023 |  |  | 06/11/2023 | 07/10/2023 |  |  | 04/30/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332328 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2023 |  |  |  |  | The user facility reported a 75-year-old male with decompensated heart failure was implanted with an impella cp device for mechanical circulatory support. fluid leaked from the red plug of the impella 5.5 assembly during patient support. the pump was explanted as a result of the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. clinical details were reviewed but were limited. therefore, the root cause of the purge leak was unable to be determined because the product was not returned.
35286 | 1220648-2023-02445 | 17293333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293333&pc=OZD | 06/11/2023 | 02/12/2024 |  | 07/10/2023 |  |  | 06/11/2023 | 07/10/2023 |  |  | 07/31/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253366 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2023 |  |  |  |  | The user facility reported a 24-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site bleed coinciding with impella support. while it was originally believed that the blood leak was coming from a compromised repositioning sheath, it was later discovered that the forward tension sutures were loose and not providing the needed tension at the access site. after replacing the forward sutures and placing a mattress suture, the bleeding at the site stopped. in addition to the sutures, the patient was provided two units of packed red blood cells to counteract the blood loss. the patient was considered to be stable following the event. | [5 empty] | The investigation into the access site bleeding has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was due to a use issue as the forward sutures were loose and when they were fixed, the bleeding subsided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35287 | 1220648-2023-02446 | 17293511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293511&pc=OZD | 06/12/2023 | 01/19/2024 |  | 07/10/2023 |  |  | 06/12/2023 | 07/10/2023 |  |  | 03/31/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327578 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2023 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump stopped functioning during patient support. the automated impella controller was swapped in an attempt to troubleshoot the issue, however, the pump still failed to restart. the pump was subsequently removed as a result of the failure. there was no patient harm. | [5 empty] | The investigation into the pump stop has been completed. the impella pump was not returned for investigation. data logs shows the motor current high and impella stop 13 alarm at the end of the case. data logs were not provided for backup/second aic. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup. unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop¿. ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency¿.
35288 | 1220648-2023-02447 | 17293631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293631&pc=OZD | 06/12/2023 | 01/19/2024 |  | 07/10/2023 |  |  | 06/12/2023 | 07/10/2023 |  |  | 04/30/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336741 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2023 |  |  |  |  | The user facility reported a 76-year-old male in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient was suspected to have experienced an embolic stroke during impella support. the patient suffered from expressive aphasia, however, intervention was not required and support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, ct scan confirmed patient had embolic stroke and it was unknown if it was related to any clot or impella. data logs were reviewed, and no motor current spike was observed relative to the event. the cause of the stroke issue was most likely patient condition related and unknown relation with insufficient clinical/log review. the failure mode will be monitored and trended.
35289 | 1220648-2023-02448 | 17293648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293648&pc=OZD | 06/12/2023 | 10/19/2024 |  | 07/10/2023 |  |  | 06/12/2023 | 07/10/2023 |  |  | 12/31/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325708 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/12/2023 |  |  |  |  | The user facility reported an 81-year-old female scheduled for percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pigtail of the implanted impella cp pump perforated the patient's left ventricle during support. the patient was taken to the operating room where the pump was pulled back into the aorta and the perforation was surgically repaired. however, following surgery, the pigtail was unable to re-cross the aortic valve and the pump was subsequently replaced. | Additional information was received that reported a second impella cp was successfully inserted and started under transesophageal echocardiogram (tee) guidance. the patient was weaned from bypass. due to severe right ventricular failure and unable to maintain flows from the pump, the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the ventricle perforation and the positioning issue has been completed. no product was returned for evaluation. data logs were reviewed which showed several suction alarm and low pump flow, with two impella position unknown alarms followed by a disabled placement signal monitoring alarm. according to the clinical review, the patient had an lv perforation with the impella pigtail, requiring pericardiocentesis and emergent surgery after cpr. the pump was pulled back into the aorta, and the perforation was surgically repaired. unable to recross the aortic valve with the pigtail manually when coming off the pump. the cause of the ventricle perforation issue was most likely improper positioning due to cpr performed during use of impella. the cause of the positioning issue was use related since doctor was unable to cross aortic valve after pump pulled back into aorta. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 med dev prob code 1158 added- d6a/d6b f6/f8 should have been left blank in the initial report. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. additional information for the initial mfr 1220648-2023-02448 was provided in sections b2, b5, g3, h1, and h6.
35290 | 1220648-2023-02449 | 17293692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293692&pc=OZD | 06/12/2023 | 02/28/2024 |  | 07/10/2023 |  |  | 06/12/2023 | 07/10/2023 |  |  | 04/30/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333518 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2023 |  |  |  |  | The user facility reported a 51-year-old male with non-ischemic cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at the impella access site due to trauma to the venous area during initial insertion. the following day, the hematoma was drained of roughly a liter of fluid. the patient was also provided two units of packed red blood cells to counteract the condition. two days after the drain, the hematoma was evacuated and the condition stabilized. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ vessel damage: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the reported hematoma has been completed since the original report was filed. pump was not returned for investigation. as per the clinical details, hematoma was noted during impella insertion. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided.
35291 | 1220648-2023-02451 | 17293786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293786&pc=OZD | 06/14/2023 | 10/06/2023 |  | 07/10/2023 |  |  | 06/14/2023 | 07/10/2023 |  |  | 09/30/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281138 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported a 75-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the impella limb became ischemic. the team placed 5fr retrograde sheath but the patient was suffering from acidosis. the team made choice to infuse blood products and place continuous renal replacement therapy for the acidosis. the patient was a full code and the patient expired despite efforts. | [5 empty] | The investigation into the limb-ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the root cause of the limb-ischemia issue was patient condition related since patient was given multiple vasopressors. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35292 | 1220648-2023-02452 | 17293796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17293796&pc=OZD | 06/14/2023 | 08/11/2023 |  | 07/10/2023 |  |  | 06/14/2023 | 07/10/2023 |  |  | 08/31/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265814 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported a 81-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during impella support. repositioning was attempted, however the issue remained throughout the remainder of support. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the data logs and clinical report was evaluated. data review: data showed pump started with low flow that triggered auto suction response & suction alarms. about 11 hours into support, pump was in aorta area with alarms impella position in aorta & impella position unknown. 23 hours later, pump was in proper position & ran without issue to the rest of the case. the root cause of hemolysis was most likely due to pump was not in the proper position. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35293 | 1220648-2023-02450 | 17294232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17294232&pc=OZD | 03/06/2023 | 08/11/2023 |  | 07/10/2023 |  |  | 03/06/2023 | 07/10/2023 |  |  | 11/30/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299123 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 48-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had red urine and the team was hesitant to give volume due to pulmonary status. the staff made the decision to remove the impella due to hemolysis. the patient remained off drips and was hemodynamically stable. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible and based on the clinical information, the root cause of the hemolysis was not determined. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35294 | 1220648-2023-02453 | 17294535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17294535&pc=OZD | 06/14/2023 | 10/13/2025 |  | 07/10/2023 |  |  | 06/14/2023 | 07/10/2023 |  |  | 02/28/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338553 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 06/14/2023 |  |  |  |  | The user facility reported a 63-year-old female in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. immediately following pump removal, the patient lost pulse in the impella leg due to a clot. a penumbra was utilized to remove as much clot as possible before the patient was sent to the operating room for a surgical cutdown. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. additionally, it was noted that there was significant oozing near impella site. a stitch was placed around the large impella sheath to help with bleeding. | Additionally, there was significant oozing near impella site for which a stitch was placed around the large impella sheath to help with bleeding. |  |  |  | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was biomaterial found in impella sided limb, after the pump was removed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issue of access site bleeding is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of the access site bleeding - major could not be determined as insufficient clinical details were provided. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2023-02453. d6a and d6b revised from the initial manufacturer device report 1220648-2023-02453 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-02453 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-02453 in accordance with updated procedures. health effect - clinical code 1888 and health effect - impact code 4641 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-02453. h6 investigation conclusions 4315 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-02453. h10 investigation results for access site bleeding was inadvertently omitted on the initial manufacturer device report number 1220648-2023-02453-1.
35295 | 1220648-2023-02423 | 17294583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17294583&pc=OZD | 08/04/2022 | 02/12/2024 |  | 07/10/2023 |  |  | 06/08/2023 | 07/10/2023 |  |  |  |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2023 |  |  |  |  | A publication found via abiomed japan reported that an implanted impella 5.0 pump required unspecified revision during patient support due to a "broken pin in the manifold" roughly two days after implant. patient support continued following the event. there was no patient harm. european society of cardiology, european heart journal, case report (2023) 7, 1-4: the publication is entitled: "tissue plasminogen activator for axillary impella5.0 with heparin-induced thrombocytopenia as a treatment of choice for acute impella thrombosis: a case report" | [5 empty] | The investigation into the product damage has been completed. the device was not returned for investigation. as per clinical review, broken pin in the manifold was seen and structural damage of the outflow cage noted (blood outlet component of the cannula outlet). the 5.0 device was explanted and revised to new pump later. no relevant case information or product serial number was reported to review exact clinical information and data logs. the cause of the pump stop issue was most likely damaged outflow cage per publication information. h6: impact code updated from 2199 to 4627. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35296 | 1220648-2023-02454 | 17295160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17295160&pc=OZD | 06/13/2023 | 01/22/2024 |  | 07/10/2023 |  |  | 06/13/2023 | 07/10/2023 |  |  | 07/31/2024 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258015 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/13/2023 |  |  |  |  | Abiomed representative contacted by hospital administrator stating that a legal investigation has been started due to patient stating a piece of the impella cp device was left in his leg. no difficulties during pump explantation reported from hospital. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The medical center did not return the impella device. no benchtop analysis was possible. the root cause of the removal issue was not determined since product was not returned and clinical details were not sufficient.
35297 | 1220648-2023-02455 | 17295171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17295171&pc=OZD | 06/14/2023 | 08/30/2023 |  | 07/10/2023 |  |  | 06/14/2023 | 07/10/2023 |  |  | 03/31/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343002 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported a 77-year-old female undergoing heart surgery was implanted with an impella cp device for mechanical circulatory support. during support, the patient's foot was noted to be turning white with no evidence of perfusion to leg. sheath manipulations by md performed without resolve. decision made to remove impella and replace with iabp due to occlusive perfusion to leg despite distal perfusion sheath (noted to have a clot distal to sheath preventing flow ). patient tolerated well coming off impella support as it allowed her heart to off load for one hour. iabp implanted and patient was sent to icu. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was due to presence of clot distal to the sheath preventing the flow down the leg. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35298 | 1220648-2023-02456 | 17295172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17295172&pc=OZD | 04/12/2023 | 01/18/2024 |  | 07/10/2023 |  |  | 04/12/2023 | 07/10/2023 |  |  | 01/31/2025 |  | 07/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024309032 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2023 |  |  |  |  | The user facility reported a 59-year-old male (race unknown) with exacerbated heart failure was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma, which required a washout to be removed. the patient continued with impella 5.5 support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident. | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
35299 | 17277333 | 17277333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17277333&pc=OZD | 05/20/2023 | 06/20/2023 | 06/20/2023 | 07/07/2023 | 06/20/2023 | 07/07/2023 |  | 07/07/2023 |  |  |  |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1711420 |  |  |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 22630 DA |  | Male |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella purge cassette routine change. when the new cassette was inserted, it immediately cracked and started leaking purge solution.
35300 | 1220648-2023-02416 | 17278138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278138&pc=OZD | 06/09/2023 | 09/19/2024 |  | 07/07/2023 |  |  | 06/09/2023 | 07/07/2023 |  |  | 10/31/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290659 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/09/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing left heart catheterization was implanted with an impella cp device for mechanical circulatory support. it was noted that the catheter had receded across the aortic valve during the course of support. upon attempting to re-cross the valve, the catheter prolapsed on itself and could not be advanced. the pump was pulled back into the descending aorta, however, attempts to straighten the catheter were unsuccessful. the vascular surgeon was advised that in order to continue using the pump, it would need to be straightened out using a snare. due to the patient's existing overall condition, the decision was made to withdraw impella support. the pump was left in the descending aorta until the patient passed away. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02416 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the positioning issues has been completed. the device was not returned for benchtop investigation. the root cause of the positioning issue was use related as the pump was pulled back across the aortic valve and was kinked when attempting to recross.
35301 | 1220648-2023-02419 | 17278192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278192&pc=OZD | 02/27/2023 | 07/13/2023 |  | 07/07/2023 |  |  | 02/27/2023 | 07/07/2023 |  |  | 11/30/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298000 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2023 |  |  |  |  | The user facility reported an 82-year-old female undergoing left main coronary artery surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in the impella leg during support. it was documented that the patient had notable stenosis and tortuosity in both femoral and iliac systems prior to impella implant. due to the noted condition, the decision was made to explant the pump after the patient's scheduled left main intervention was completed. | [5 empty] | The impella device has not been returned by the customer and therefore, evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the cause of the limb ischemia issue was patient condition due to notable stenosis and tortuosity in both femoral and iliac arteries. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35302 | 1220648-2023-02420 | 17278209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278209&pc=PYX | 03/02/2023 | 09/11/2023 |  | 07/07/2023 |  |  | 03/02/2023 | 07/07/2023 |  |  | 07/31/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023234606 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2023 |  |  |  |  | The user facility reported an 84-year-old female in st elevated myocardial infarction was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis during impella support. the pump was subsequently removed as a result of the noted condition. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical data was returned. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35303 | 1220648-2023-02421 | 17278245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278245&pc=OZD | 02/28/2023 | 01/18/2024 |  | 07/07/2023 |  |  | 02/28/2023 | 07/07/2023 |  |  | 09/30/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273649 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2023 |  |  |  |  | The user facility reported a 55-year-old female in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had poor perfusion down the impella leg. a fem-fem bypass was placed to resolve the noted condition and support continued with the implanted pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related to diseased/small vessels.
35304 | 1220648-2023-02422 | 17278269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278269&pc=OZD | 06/08/2023 | 10/23/2023 |  | 07/07/2023 |  |  | 06/08/2023 | 07/07/2023 |  |  | 12/31/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291943 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/08/2023 |  |  |  |  | During priming of the impella 5.5 pump, it was noted that fluid was leaking from the yellow luer at the end of the purge cassette. the connection were tightened to resolve the corresponding alarm; however, fluid was still seen leaking from the yellow luer. the purge cassette was subsequently replaced to troubleshoot the issue prior to attempted delivery. there was no patient harm resulting from the leak. | The user facility reported a 61-year-old male in ventricular fibrillation was implanted with an impella cp device for mechanical circulatory support. the patient was to be escalated to an impella 5.5 device and the physician experienced difficulties in the delivery of the device. after the first suture line, the artery was tearing with every suture they put in. due to the fragile nature of the artery, multiple repair stitches had be used to control the bleeding and nu-knit had to be placed as well. the physician experienced difficulty advancing a standard .035 ¿j¿wire. a stiff angled glidewire and jr 4 catheter had to be used to get access into ascending aorta. with a lot of difficulty, the pigtail finally crossed into the lv. ¿j¿ wire removed and .018 wire advanced into the lv. doctor then attempted to advance the 5.5. however despite manipulation, they could not get the device to advance out of the graft into the artery. doctor then advanced a crosslead .089 wire as a buddy wire via the axillary kit and despite having the buddy wire, the pump would not advance into the artery. the decision made to abort procedure and leave patient on impella cp. all impella 5.5 equipment was removed and right axillary incision closed. |  |  |  |  | Please see manufacturing report number 1220648-2023-02422, follow up 1 for the results of the investigation pertaining to the reported leak. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the device, including the pump and both pump cassettes, was sent back for investigation. both purge cassettes were tested, and no leaks were observed from the yellow luer. it was determined that the root cause of the delivery issue was most likely patient condition related. and the cause of the purge leak issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35305 | 1220648-2023-02424 | 17278352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278352&pc=OZD | 03/03/2023 | 08/11/2023 |  | 07/07/2023 |  |  | 03/03/2023 | 07/07/2023 |  |  | 10/31/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288200 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | The user facility reported a 23-year-old female with decompensated congestive heart failure was implanted with an impella cp device for mechanical circulatory support. during support, the patient had no pulse in the left foot. the pump was explanted. patient is awaiting heart transplant. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia is due to the patient's condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35306 | 1220648-2023-02425 | 17278386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278386&pc=OZD | 02/22/2023 | 08/11/2023 |  | 07/07/2023 |  |  | 02/22/2023 | 07/07/2023 |  |  | 09/30/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282543 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2023 |  |  |  |  | The user facility reported an 83-year-old female undergoing transcatheter aortic valve replacement was implanted with an impella cp device for mechanical circulatory support after the patient coded. at procedure end, runoff revealed no distal perfusion to left lower extremity. impella options were considered due to device being rapidly inserted through 14f edwards tavr sheath. options considered were to leave as is with fem-fem bypass, remove device & exchange impella sheath for device reinsertion, or alternative access site & device placement. due to patients hemodynamic improvement at mean arterial pressures remained 80-85¿s. decision was made to explant device. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was patient condition related due to pad/pvd vessel conditions. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35307 | 1220648-2023-02417 | 17278438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278438&pc=OZD | 02/28/2023 | 08/11/2023 |  | 07/07/2023 |  |  | 02/28/2023 | 07/07/2023 |  |  | 10/31/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023268082 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2023 |  |  |  |  | The user facility reported a 62-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. due to lactate rising and additional pressor requirements, the patient was escalated to an impella 5.5 device. the developed access site bleeding at the auxilliary cutdown site. three units of blood were given to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the access site bleeding was not determined. the failure mode will be monitored and trended.
35308 | 1220648-2023-02418 | 17278442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278442&pc=OZD | 06/09/2023 | 01/15/2024 |  | 07/07/2023 |  |  | 06/09/2023 | 07/07/2023 |  |  | 02/28/2025 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342031 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2023 |  |  |  |  | The user facility reported a 49-year-old male in st undergoing heart surgery was implanted with an impella cp device for mechanical circulatory support. following advancement of the repositioning sheath, a subdural hematoma with pulsatile blood formed. manual pressure was applied around the sheath, however, the bleeding persisted. the pump was subsequently explanted and the groin was surgically closed to resolve the bleed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that post procedure, the patient's activated coagulation time (act) was prolonged at greater than 400. therefore, the root cause of the access site bleeding issue was high patient act values during pump implant.
35309 | 1220648-2023-02426 | 17278788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17278788&pc=OZD | 02/26/2023 | 08/09/2023 |  | 07/07/2023 |  |  | 02/26/2023 | 07/07/2023 |  |  | 09/30/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2023 |  |  |  |  | The user facility reported a 59-year-old male undergoing left heart surgery was implanted with an impella cp device for mechanical circulatory support. during support, the patient bent the leg and bleeding from the groin followed. decision made by attending physician to replace impella with a new impella and leave peel away sheath in to achieve hemostasis. patient stable. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the access site bleeding was patient movement due to improper technique that led to bleeding from the insertion site. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35310 | 1220648-2023-02428 | 17283294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17283294&pc=OZD | 06/09/2023 | 07/13/2023 |  | 07/07/2023 |  |  | 06/09/2023 | 07/07/2023 |  |  | 05/31/2025 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024340870 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2023 |  |  |  |  | The complainant reported a 60 year old female patient presenting for an elective procedure using the impella device for hemodynamic support. during support, the patient required a trip to the operating room for a surgical washout. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis. follow up with the field representative revealed the reported patient bleeding was unrelated to the impella device. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35311 | 1220648-2023-02429 | 17283486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17283486&pc=OZD | 04/05/2023 | 09/11/2023 |  | 07/07/2023 |  |  | 04/05/2023 | 07/07/2023 |  |  | 11/30/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315939 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2023 |  |  |  |  | The user facility reported a 61-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. there were reported hemolyzed labs during support, prompting suspicion of hemolysis. the pump was removed as a result. secondary to hemolysis, the pump was removed after impella in aorta alarms occurred as a result of the pump falling out of the ventricle. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information not provided. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35312 | 1220648-2023-02430 | 17283492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17283492&pc=OZD | 02/27/2023 | 01/18/2024 |  | 07/07/2023 |  |  | 02/27/2023 | 07/07/2023 |  |  | 10/31/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286648 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2023 |  |  |  |  | The user facility reported a 69-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed in the left subclavian due to severe bilateral iliac disease. peel-away sheath removed at the end of the case and subclavian angiogram performed revealed perforation. a covered stent was placed to left subclavian and bleeding controlled. sent to unit post procedure. patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure. | The investigation into the vascular damage- peripheral vessel has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the vascular damage to the peripheral vessel issue was patient condition related to the patient¿s poor vasculature which was a result of chemo/radiation.
35313 | 1220648-2023-02431 | 17283493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17283493&pc=OZD | 02/28/2023 | 08/09/2023 |  | 07/07/2023 |  |  | 02/28/2023 | 07/07/2023 |  |  | 09/30/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281446 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2023 |  |  |  |  | The user facility reported a 49-year-old male in myocardial infarction was implanted with an impella cp device for mechanical circulatory support. upon implantation of the device, it was noticed that the patient did not have distal flow. the patient condition did not improve and the device had to be explanted. the patient ultimately expired, however the patient's death is not attributed to impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the limb ischemia was not determined. the failure mode will be monitored and trended.
35314 | 1220648-2023-02432 | 17283497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17283497&pc=OZD | 05/17/2023 | 01/19/2024 |  | 07/07/2023 |  |  | 05/17/2023 | 07/07/2023 |  |  | 01/31/2025 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312468 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2023 |  |  |  |  | The user facility reported a 67-year-old male had an impella cp explanted and subsequently developed bradycardia and ventricular fibrillation/ventricular tachycardia arrest. the patient was implanted with a second impella cp device for mechanical circulatory support. the patient developed a hematoma and access site bleeding occur. the patient was given two units of blood and the pump was explanted six days after the implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement since it was reported after the patient transferred to intensive care unit.
35315 | 1220648-2023-02433 | 17283500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17283500&pc=OZD | 02/23/2022 | 08/09/2023 |  | 07/07/2023 |  |  | 02/23/2022 | 07/07/2023 |  |  | 08/31/2024 |  | 07/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279395 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2022 |  |  |  |  | The user facility reported a 65-year-old was implanted with an impella cp device for mechanical circulatory support. wcmo was added to support and signs and symptoms of leg ischemia were present at the impella cp accessed leg. the decision was made by the medical team to take the patient to the or and explant impella cp, escalate to impella 5.5 and remain on ecpella configuration, post procedure. impella cp was removed, however, the medical team was unable to insert impella 5.5 as it would not fit. the graft was left long, and the access site was disinfected and dressed accordingly. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was patient condition related to small vessels.
35316 | 1220648-2023-02404 | 17275016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17275016&pc=OZD | 06/03/2023 | 08/02/2024 |  | 07/06/2023 |  |  | 06/07/2023 | 07/06/2023 |  |  | 04/30/2025 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024341024 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported a male (age unknown) post cardiotomy cardiogenic shock was implanted with an impella 5.5 device for circulatory support. the patient experienced hemolysis during support. the position of the pump was adjusted, however that showed no improvement. dialysis was performed as supportive therapy however it was not sufficient, and the patient expired. the physician noted there was no association between the impella and death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. per medical safety officer review use of the device cannot conclusively be associated with the outcome. d6a/d6b added. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-02404. f6/f8-should have been left blank.
35317 | 1220648-2023-02405 | 17275135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17275135&pc=OZD | 04/03/2023 | 08/31/2023 |  | 07/06/2023 |  |  | 04/03/2023 | 07/06/2023 |  |  | 01/31/2025 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024313604 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2023 |  |  |  |  | The user facility reported a 63-year-old female who was decompensating prior to a splenectomy procedure was implanted with an impella cp device for mechanical circulatory support. when inserting the pump, the catheter was advanced in the left ventricle (lv) and the lv collapsed. the md was unable to remedy and determined to explant that impella and insert a new one due to the patient's deteriorating condition. the new impella was inserted successfully, however, the patient experienced bilateral ischemia. the pump was explanted four days after implant due to the patient's multiple comorbidities. the patient expired. additional information from the clinical consultant indicated that the patient death was not related to impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the limb ischemia was not determined since the product was not returned and sufficient clinical information was not provided.
35318 | 1220648-2023-02406 | 17275151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17275151&pc=OZD | 03/06/2023 | 08/11/2023 |  | 07/06/2023 |  |  | 03/06/2023 | 07/06/2023 |  |  | 09/30/2024 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282051 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported a 66-year-old male in ventricular tachycardia/ ventricular fibrillation arrest was implanted with an impella cp device for mechanical circulatory support. during support, the impella access site was oozing. blood was given to the patient. the patient was successfully weaned from the device four days after implant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the access site bleeding issue was not determined since product was not returned and it is unrelated to impella based on the provided clinical information. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35319 | 1220648-2023-02407 | 17276311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276311&pc=OZD | 06/07/2023 | 08/24/2023 |  | 07/06/2023 |  |  | 06/07/2023 | 07/06/2023 |  |  | 01/31/2025 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331552 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported an 80-year-old female undergoing high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient became profoundly hypotensive and, upon inspection of the right coronary artery, it was determined that it had no flow and required an additional stent. a cardiac ultrasound found a large effusion requiring pericardiocentesis. the physician also determined that the left ventricle (lv) was perforated by the impella catheter. the physician repaired the lv and then removed the impella cp. the patient was transferred to icu in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the ventricle perforation was most likely improper positioning due to cpr performed during use of impella.
35320 | 1220648-2023-02408 | 17276313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276313&pc=OZD | 06/07/2023 | 07/25/2024 |  | 07/06/2023 |  |  | 06/07/2023 | 07/06/2023 |  |  | 02/28/2025 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341064 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported a 77-year-old male undergoing surgery on the left main coronary artery was implanted with an impella cp device for mechanical circulatory support. after implant, the impella pump perforated the patient's left ventricle. the pump was subsequently removed and the patient's lv was surgically repaired to treat the perforation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for ventricle perforation has been completed. according to the clinical, shortly after beginning impella cp support the patient developed an effusion. a pericardial drain was placed and the patient was sent to the or. upon further inspection it was discovered that the impella had perforated the lv. the physician noted that the patient had fragile tissue and confirmed that the insertion was never too deep. no product was returned for a visual inspection. data log analysis was not required for this complaint. the most likely cause of the ventricular perforation was patient condition related. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
35321 | 1220648-2023-02409 | 17276354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276354&pc=OZD | 05/18/2023 | 01/22/2024 |  | 07/06/2023 |  |  | 06/08/2023 | 07/06/2023 |  |  | 11/30/2024 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308870 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2023 |  |  |  |  | The user facility reported an 85-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the team found the repo-sheath to be occlusive to the vasculature. the team explanted the pump and there was no lasting harm or injury. vessel hemostasis was achieved with perclose and angioseal. | [5 empty] | The investigation into the limb ischemia issue has been completed. the device was not returned for investigation. per the provided clinical details, the root cause of the limb ischemia issue was patient condition related to small/ diseased vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35322 | 1220648-2023-02410 | 17276372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276372&pc=OZD | 06/08/2023 | 08/02/2024 |  | 07/06/2023 |  |  | 06/08/2023 | 07/06/2023 |  |  | 04/30/2024 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220646 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2023 |  |  |  |  | The user facility reported a 79-year male with severe multi-vessel disease was implanted with an impella cp device for mechanical circulatory support. the patient was escalated to an impella 5.5 for continued support while the impella cp was still in position. after explant of the impella cp, a perforation of the left common femoral artery was visualized. the physician performed a balloon tamponade and inflation was held for ten minutes. the patient received fluid and blood products. the procedure was successful and the impella 5.5 was implanted. the patient was sent to icu with the support of an impella 5.5. | It was further reported that the patient expired sometime after the pump was explanted. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that the use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the vascular damage has been completed. the product was not returned for investigation. the cp pump was explanted, and the patient was escalated to 5.5 pump. the doctor reassessed the cfa and performed a femoral angiogram after the blood pressure began to plummet. a perforation was found in the left common femoral artery, and hemostasis was successfully achieved after balloon tamponade. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ added- b5 mso review, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
35323 | 1220648-2023-02411 | 17276376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276376&pc=OZD | 06/08/2023 | 09/26/2023 |  | 07/06/2023 |  |  | 06/08/2023 | 07/06/2023 |  |  | 01/31/2025 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310016 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2023 |  |  |  |  | The user facility reported a 64-year-old male undergoing a left heart catheterization was implanted with an impella 5.5 device for mechanical circulatory support. the patient was seen to have mitral regurgitation and tricuspid regurgitation in addition to cardiomyopathy. during the support, the team observed microbubbles under transesophageal echocardiography. there was no issue or malfunction observed that may have caused air embolism. there has been no harm or cva to date. | [5 empty] | Since the product was not returned for review, no benchtop analysis was possible. per clinical information, the micro bubbles were dissipated after couple minutes when pump started. the root cause of the embolism was unable to be determined as no product was returned for review. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35324 | 1220648-2023-02412 | 17276379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276379&pc=OZD | 06/08/2023 | 10/16/2023 |  | 07/06/2023 |  |  | 06/08/2023 | 07/06/2023 |  |  | 01/31/2025 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306214 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2023 |  |  |  |  | The user facility reported a 60-year-old male with chronic heart failure was implanted with an impella cp device for mechanical circulatory support. the patient was transferred with a peel-away sheath retained in the vessel. the limb became ischemic. pulses were lost and the limb was mottled. the team made choice to explant the pump from the femoral and place an impella 5.5 pump via axillary. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the given clinical information, the root cause of the ischemia issue was use related to improper technique of suturing the introducer to the patient and retention of the introducer.
35325 | 1220648-2023-02413 | 17276395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276395&pc=OZD | 03/03/2023 | 08/09/2023 |  | 07/06/2023 |  |  | 03/03/2023 | 07/06/2023 |  |  | 10/31/2024 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290549 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | The user facility reported a 46-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the device was placed electively prior to coronary artery bypass grafting. the patient lost pulses on the impella-accessed side. the leg was cold and ischemic. the impella cp was explanted. heparin was administered to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was most likely patient condition related, as reported absence of limb pulses after removal of impella.
35326 | 1220648-2023-02414 | 17276396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276396&pc=OZD | 02/26/2023 | 08/11/2023 |  | 07/06/2023 |  |  | 02/26/2023 | 07/06/2023 |  |  | 05/31/2024 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231131 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2023 |  |  |  |  | The user facility reported a 73-year-old female undergoing percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed limb ischemia while on support. the decision was made to removed the impella cp device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible and based on the clinical information, the root cause of the issue was not determined. the failure mode will be monitored and trended.
35327 | 1220648-2023-02415 | 17276399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17276399&pc=OZD | 03/03/2023 | 08/09/2023 |  | 07/06/2023 |  |  | 03/03/2023 | 07/06/2023 |  |  | 12/31/2024 |  | 07/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300115 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | The user facility reported a 76-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the limb was seen to be ischemic. the impella was in the limb via a 30cm cook sheath when both the short and long 14fr sheaths failed to deploy beyond iliac bifurcation. the cook sheath and cp pump were explanted due to the ischemia. the limb was thrombosed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was most likely patient condition related to thrombosis.
35328 | 1220648-2023-02392 | 17255316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17255316&pc=OZD | 06/07/2023 | 09/15/2023 |  | 07/05/2023 |  |  | 06/07/2023 | 07/05/2023 |  |  | 12/31/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300843 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported a 62-year-old female experienced a left anterior descending artery perforation during a balloon inflation. the patient was implanted with an impella cp device for mechanical circulatory support. the patient's limb became mottled and the pulses were diminished as shown on doppler. the impella was explanted prematurely. the device was not replaced as the patient's condition was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the limb ischemia issue was not determined since the product was not returned.
35329 | 1220648-2023-02394 | 17259669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17259669&pc=OZD | 02/24/2023 | 08/30/2023 |  | 07/05/2023 |  |  | 02/24/2023 | 07/05/2023 |  |  | 12/31/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301938 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2023 |  |  |  |  | The user facility reported a 75-year-old male in cardiogenic and septic shock was implanted with an impella cp device for mechanical circulatory support. a hematoma developed at the patient's lateral circumflex femoral artery post implant. .manual pressure was applied to the hematoma to decompress. multiple attempts were made but the hematoma continued to get larger. the patient started to become hemodynamically unstable requiring more pressors/volume/blood products. the pump was emergently removed. patient received more than 2 blood products. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the vascular damage was unable to be determine as no product and insufficient clinical details were provided. the failure mode will be monitored and trended.
35330 | 1220648-2023-02393 | 17259757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17259757&pc=OZD | 02/22/2023 | 08/11/2023 |  | 07/05/2023 |  |  | 02/22/2023 | 07/05/2023 |  |  | 11/28/2022 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272591 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2023 |  |  |  |  | The user facility reported a 77-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support after the patient became hypotensive. the physician immediately put in 14fr sheath, wired lv with 0.018 and backloaded the impella cp, and then patient became stabilized. aborted implantation of the impella cp due to distal limb ischemia. the patient was transferred to another facility. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was patient condition related to small/diseased vessels.
35331 | 1220648-2023-02398 | 17261771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17261771&pc=OZD | 04/03/2023 | 09/28/2023 |  | 07/05/2023 |  |  | 04/03/2023 | 07/05/2023 |  |  | 10/31/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023268913 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2023 |  |  |  |  | The user facility reported that while a patient was on impella cp support, suction alarms and no flow events occurred. after repositioning failed, the impella cp was removed. the physician made the decision to bridge to veno-arterio-venous ecmo due to repetitive ventricular tachycardia. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the vf/vt/af/at has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. per the provided clinical information, the root cause of the low pump flow issue was improper positioning of the device based on provided clinical information; however, the root cause of the ventricular fibrillation, ventricular tachycardia, atrial fibrillation, atrial tachycardia issues was not determined since product was not returned and sufficient clinical information was not provided.
35332 | 1220648-2023-02397 | 17262255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17262255&pc=OZD | 02/25/2023 | 08/24/2023 |  | 07/05/2023 |  |  | 02/25/2023 | 07/05/2023 |  |  | 06/30/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248746 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2023 |  |  |  |  | The user facility reported a 63-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the physician explanted the pump due to concerns of the superficial femoral artery being occluded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the reversible ischemia was patient condition related to the patient vessels.
35333 | 1220648-2023-02399 | 17262763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17262763&pc=OZD | 06/05/2023 | 02/16/2024 |  | 07/05/2023 |  |  | 06/05/2023 | 07/05/2023 |  |  | 02/28/2025 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024341041 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2023 |  |  |  |  | The complainant reported a 51-year-old male patient suffering from acute myocardial infarction and cardiogenic shock was presented for impella placement. following impella implant, suction was immediately encountered from levels p4 down to p1. an echo was performed and a free-floating clot was seen extending from the left atrium and pulmonary veins into the left ventricle. the pump was explanted and surgery was performed to remove the clot. the patient survived the procedure. | [5 empty] | The impella device has not been returned by the customer and therefore, evaluation of the device was not possible. upon completion of the investigation, a supplemental report will be filed. | The investigation into the thrombus and the low pump flow issue has been completed. the device was not returned for investigation. per the clinical information provided, doctor found floating biomaterial in left ventricle. the cause of the low pump flow issue was most likely biomaterial found in lv and biomaterial interaction with pump was noted during data log review.
35334 | 1220648-2023-02400 | 17263018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17263018&pc=OZD | 02/22/2023 | 08/09/2023 |  | 07/05/2023 |  |  | 02/22/2023 | 07/05/2023 |  |  | 08/31/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259652 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2023 |  |  |  |  | The user facility reported a 65-year-old female with heart failure was implanted with an impella cp device for mechanical circulatory support. the patient experienced signs of ischemia while on support. the physician planned to escalate the patient to an impella 5.5 device, but was unable due to the patient's small axillary arteries. the patient was implanted with a new impella cp device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the limb ischemia was most likely patient condition related.
35335 | 1220648-2023-02402 | 17263021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17263021&pc=OZD | 06/06/2023 | 01/18/2024 |  | 07/05/2023 |  |  | 06/06/2023 | 07/05/2023 |  |  | 03/31/2025 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327430 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2023 |  |  |  |  | The user facility reported a 74-year-old female undergoing repair of the mitral valve was implanted with an impella cp device for mechanical circulatory support. the patient endured distal limb ischemia on (b)(6) 2023 that prompted an external perfusion catheter to restore blood flow. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. product was not returned for review. based on clinical description, the root cause of limb ischemia were most likely due to patient condition.
35336 | 1220648-2023-02403 | 17263022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17263022&pc=OZD | 02/24/2023 | 08/25/2023 |  | 07/05/2023 |  |  | 02/24/2023 | 07/05/2023 |  |  | 07/31/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2023 |  |  |  |  | The user facility reported a 63-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and limb ischemia while on support. the patient was escalated to an impella 5.5 and conditions improved. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis and the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the hemolysis issue was most likely improper positioning of the device with impella in aorta alarms found per data log review, and the root cause of the limb ischemia issue was not determined based on the insufficient information from clinical review.
35337 | 1220648-2023-02401 | 17263048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17263048&pc=OZD | 06/06/2023 | 07/05/2023 |  | 07/05/2023 |  |  | 06/06/2023 | 07/05/2023 |  |  | 09/30/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273952 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2023 |  |  |  |  | The user facility reported a 59-year-old female with a proximal left main lesion was to undergo surgery for coronary artery bypass grafting. the patient was implanted with an impella cp device for mechanical circulatory support. the patient developed pericardial effusion during support. the etiology of the effusion was unknown and a drain was required as treatment. the pump was ultimately removed after the patient regained hemodynamic stability, however, a "small" pericardial effusion remained. | [5 empty] | The impella device was not received from the customer and, therefore, an evaluation of the device was not possible. the data logs and clinical report were evaluated. there is no evidence in the data logs that the pump was deep into the ventricle. the clinical account also mentions the relationship to impella is unknown. with this evidence, the root cause of the ventricle perforation leading to the pericardial effusion could not be determined. any relationship to impella could not be confirmed. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35338 | 1220648-2023-02396 | 17263552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17263552&pc=OZD | 03/02/2023 | 08/09/2023 |  | 07/05/2023 |  |  | 03/02/2023 | 07/05/2023 |  |  | 12/31/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301636 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2023 |  |  |  |  | The user facility reported a 59-year-old female in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while under support, the pulses were lost in the leg. the team explanted via a femoral cut down, treated the the limb by fogarty embolectomy, and the pulse was restored. the cp was escalated to the 5.5 pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was verified due to the presence of thrombus in the patient vessel since the embolectomy resolved the limb ischemia. the failure mode will be monitored and trended.
35339 | 1220648-2023-02395 | 17263602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17263602&pc=OZD | 02/24/2023 | 08/11/2023 |  | 07/05/2023 |  |  | 02/24/2023 | 07/05/2023 |  |  | 11/30/2024 |  | 07/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2023 |  |  |  |  | The user facility reported a 71-year-old male with multiple comorbidities and undergoing percutaneous coronary intervention was implanted with an impella cp for circulatory support. the impella-accessed limb became cold and mottled. the team placed fem-fem bypass to restore limb perfusion. after two days of the cp support, the family made choice to withdraw care and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the reported limb ischemia has been completed since the original eport was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the most likely small vessel patient condition was the cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35340 | 17236669 | 17236669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17236669&pc=OZD | 06/16/2023 | 06/27/2023 | 06/27/2023 | 06/30/2023 | 06/27/2023 | 06/30/2023 |  | 06/30/2023 |  |  |  |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMART ASSIST |  |  | OZD |  |  | 0550-0008 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 24090 DA |  | Male | 82 KG |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5.5 smartassist device placed. nine days later, the device started to alarm "impella stopped" during patient ambulation. the patient was returned to his room and the impella was placed to p8 by the critical care physicians. the alarm continued with any patient movement. the device was removed later that day during heart transplant procedure. the explanted device is sequestered at the facility.
35341 | 1220648-2023-02381 | 17238017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17238017&pc=OZD | 04/12/2023 | 09/01/2023 |  | 06/30/2023 |  |  | 04/12/2023 | 06/30/2023 |  |  | 03/31/2025 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024320245 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2023 |  |  |  |  | The user facility reported a 74-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was to be escalated to an impella 5.5 device. during implant of the 5.5, the patient experienced a dissection of the ascending aorta. there was a noted change in transesophageal echocardiography. the implant was aborted, patient was emergently placed on bypass. a coronary artery bypass grafting and endarterectomy was performed. | [5 empty] | The investigation into the delivery issues and the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the delivery issue was due to patient condition related with patient having stenosis at innominate artery and the root cause of the vascular damage issue was most likely use related due to j-wire advancement through stenosis causing dissection. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35342 | 1220648-2023-02382 | 17238707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17238707&pc=OZD | 06/04/2023 | 01/23/2024 |  | 06/30/2023 |  |  | 06/04/2023 | 06/30/2023 |  |  | 01/31/2025 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330536 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2023 |  |  |  |  | The user facility reported a 80-year-old female in ventricular arrest was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site bleed and a hematoma during impella support. manual pressure was applied and it was noted that the blue hub was inserted through the skin. two units of blood were subsequently transfused and femostop was utilized to treat the bleed. the pump was explanted and the patient was supported via an intra-aortic balloon pump. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the cause of the access site bleeding issue was most likely use related since blue suture was found deep into skin after patient transfer to cvicu. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35343 | 1220648-2023-02384 | 17240087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17240087&pc=OZD | 06/05/2023 | 11/08/2023 |  | 06/30/2023 |  |  | 06/05/2023 | 06/30/2023 |  |  | 12/31/2024 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325648 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported an 84-year-old male undergoing coronary artery bypass grafting had an impella cp device implanted for circulatory support. the team peeled the sheath away and inserted the repo-sheath. the team then moved the limb and a hematoma and significant access site bleed occurred. the team gave blood x2 units and surgically cut down and made a purse string suture to achieve hemostasis. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on clinical description, the root cause of access site bleeding was most likely due to a lack of vessel recoil. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35344 | 1220648-2023-02383 | 17240533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17240533&pc=OZD | 04/13/2023 | 07/15/2024 |  | 06/30/2023 |  |  | 04/13/2023 | 06/30/2023 |  |  | 12/31/2024 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316425 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Heart Failure/Congestive Heart Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2023 |  |  |  |  | It was further reported that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. | The user facility reported a 68-year-old male undergoing ventricular tachycardia ablation was implanted with an impella cp device for mechanical circulatory support. during the procedure, the patient went into pulseless electrical activity (pea) arrest requiring multiple shock attempts along with compressions. after return of spontaneous circulation (rosc) was achieved, the patient began to decompensate rapidly requiring mass blood transfusion along with multiple inotropes and pressors. the patient received multiple blood products, medications and five shocks. the physician and the family decided to withdraw care and the patient expired. |  |  |  |  | B.5 additional information was received and abiomed's medical safety officer review was completed. the investigation has been completed. the impella has not been returned for evaluation. the root cause of the bleeding could not be determined. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from manufacturer device report 1220648-2023-02383. d.6a and d.6b were revised in accordance with updated procedures since manufacturer device report 1220648-2023-02383 was submitted. f.6 and f.8 dates were reported on manufacturer device report 1220648-2023-02383 and should not have been. h.6 code 4115 was reported incorrectly on the previously submitted report. a new code has been added to type of investigation codes. h.6 added codes 925 1888 and 4647 as they were inadvertently left off manufacturer device report 1220648-2023-02383. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report 1220648-2023-02383 in accordance with updated procedures. additional manufacturer narrative on the manufacturer device report 1220648-2023-02383 stated that the device was discarded but the device was not returned. type of investigation codes has been updated to reflect not returned. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35345 | 1220648-2023-02385 | 17242836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17242836&pc=OZD | 04/17/2023 | 09/20/2024 |  | 06/30/2023 |  |  | 04/17/2023 | 06/30/2023 |  |  | 12/31/2024 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/17/2023 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 61-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was found to have multiple strokes (temporal and middle cerebral artery (mca). patient was responding to commands. pump was explanted 10 days after implant and patient ultimately expired from withdrawn care. the clinical representative could not determine what caused the strokes. |  |  |  |  | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02385 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury."
35346 | 1220648-2023-02386 | 17242862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17242862&pc=OZD | 04/07/2023 | 09/28/2023 |  | 06/30/2023 |  |  | 04/07/2023 | 06/30/2023 |  |  | 12/31/2024 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317966 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2023 |  |  |  |  | The user facility reported a 58-year-old male undergoing percutaneous cardiac insertion for ventricular tachycardia (vt) arrest was implanted with an impella cp device for mechanical circulatory support. while on impella support, the vt did not resolve and became even worse. multiple attempts to reposition the impella catheter were made and the impella was ultimately removed. the patient was no longer experiencing vt and the hemodynamics improved. the groin was closed without complication, however, the patient was diagnosed with clostridioides difficile infection in the groin. | [5 empty] | The investigation into the infection/sepsis issue and the vf/vt issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical details and data logs, the root cause of the infection/sepsis issue was most likely patient condition related and the root cause of the vf/vt issue was most likely patient condition related since patient had pre-existing vt. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35347 | 1220648-2023-02387 | 17242938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17242938&pc=OZD | 06/05/2023 | 08/02/2023 |  | 06/30/2023 |  |  | 06/05/2023 | 06/30/2023 |  |  | 03/31/2025 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327060 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported a 61-year-old male had an impella 5.5 device implanted for support during coronary artery bypass grafting procedure. during support, the pump stopped. the pump was removed and graft was closed. an impella cp was implanted without issue. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the pump stop was ingested biomaterial. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35348 | 1220648-2023-02388 | 17243056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17243056&pc=OZD | 03/22/2023 | 09/06/2023 |  | 06/30/2023 |  |  | 03/22/2023 | 06/30/2023 |  |  | 12/31/2024 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292127 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2023 |  |  |  |  | The user facility reported a 65-year-old male in the hospital for a left anterior descending artery work-up was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a transient ischemic attack/stroke which required the patient to be intubated. the family of the patient decided to withdraw care and the patient expired. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. data logs were returned and investigated. the root cause of the issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35349 | 1220648-2023-02389 | 17243061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17243061&pc=OZD | 04/12/2023 | 01/22/2024 |  | 06/30/2023 |  |  | 04/12/2023 | 06/30/2023 |  |  | 08/31/2024 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed an ooze at the extracorporeal membrane oxygenation (ecmo) and the impella site. a figure eight suture was placed at the impella site and a hematoma was pushed out by surgeons on the impella site and the ecmo site. however, patient developed significant bleeding and pump was removed to achieve hemostasis. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient had oozing at ecmo and impella sites. patient act was high at 347 during bleeding issues and before pump explant. the cause of the access site bleeding issue was anticoagulation management due to high patient act values per clinical review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35350 | 1220648-2023-02390 | 17243079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17243079&pc=OZD | 04/08/2023 | 08/09/2023 |  | 06/30/2023 |  |  | 04/08/2023 | 06/30/2023 |  |  | 09/30/2024 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272964 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2023 |  |  |  |  | The user facility reported an 83-year-old male undergoing left anterior descending artery procedure was implanted with an impella cp device for mechanical circulatory support. readjustment was needed on the pump and hemolysis lab values were noticed. the md repositioned the pump and after the procedure was completed, it was noticed that the patient was oozing heavily from the access site. the oozing could not be stopped and decision was made to explant the impella. the patient was put on continuous renal replacement therapy and was shown to be hemodynamically stable. | [5 empty] | The investigation into the reported hemolysis and access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report and data logs were evaluated. data logs shows the suction events during case. the cause of the hemolysis issue was improper positioning of the device since pump repositioning resolved the hemolysis based on clinical information. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not received. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35351 | 1220648-2023-02391 | 17243383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17243383&pc=OZD | 06/05/2023 | 03/01/2024 |  | 06/30/2023 |  |  | 06/05/2023 | 06/30/2023 |  |  | 02/28/2025 |  | 06/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325035 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported a 57-year-old female in respiratory arrest was implanted with an impella cp device for mechanical circulatory support. the patient experienced an ischemic stroke coinciding with impella support. patient support with the implanted pump continued following the event. | [5 empty] | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35352 | 1220648-2023-02379 | 17226960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17226960&pc=OZD | 02/24/2023 | 08/09/2023 |  | 06/29/2023 |  |  | 02/24/2023 | 06/29/2023 |  |  | 12/31/2024 |  | 06/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293905 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2023 |  |  |  |  | The user facility reported a 75-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after three days, the patient was escalated to an impella 5.5 device. the 5.5 was placed and support continued, however, with signs of hemolysis in this patient with impella and extracorporeal membrane oxygenation support. the patient had continuous renal replacement therapy and blood products were infused. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the hemolysis was not determined. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35353 | 1220648-2023-02380 | 17231120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17231120&pc=OZD | 05/28/2023 | 07/12/2023 |  | 06/29/2023 |  |  | 05/30/2023 | 06/29/2023 |  |  | 01/31/2025 |  | 06/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338688 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2023 |  |  |  |  | The user facility reported a 63-year-old male with persistent ventricular tachycardia was implanted with an impella cp device for mechanical circulatory support. the patient developed an embolic stroke approximately three hours post impella implant. the patient exhibited symptoms of facial drooping, aphasia, and inability to move the right arm. tissue plasminogen activator was administered with some improvement in symptoms, however, the patient continued to have right arm weakness. it's believed that a clot had existed in the patient's left ventricle that was missed via echo imaging prior to procedure and had become dislodged post procedure when the patient had returned to their room. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the stroke/cva was most likely patient condition related. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident / stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35354 | 1220648-2023-02367 | 17215867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17215867&pc=OZD | 04/15/2023 | 10/19/2023 |  | 06/28/2023 |  |  | 04/15/2023 | 06/28/2023 |  |  | 10/31/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294480 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2023 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was having issues lying still and positioning issues occurred. the team attempted to reposition impella which resulted in episode of ventricular tachycardia (vt) that did not require intervention. then a foley catheter was placed as patient was not able to urinate. the urine was bright red, so decision to remove impella . manual pressure x 50 minutes was applied and successful hemostasis achieved. | [5 empty] | The investigation into the hemolysis, the positioning issues and the vf/vt/af/at issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, there were concerns with patient lying still during support causing positioning issue and attempt of repositioning leading to multiple vt events. hemolysis was reported with red urine. no abnormalities were found relative to hemolysis and vt in the returned data logs. based on log review and due to insufficient clinical information, the cause of the hemolysis issue was not determined since product was not returned. the cause of the positioning issue was due to patient movement as there was concerns with patient lying still during support. the cause of the ventricular tachycardia issue was most likely patient condition related based on the clinical information and log review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35355 | 17215949 | 17215949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17215949&pc=OZD | 06/16/2023 | 06/22/2023 | 06/22/2023 | 06/28/2023 | 06/22/2023 | 06/28/2023 |  | 06/28/2023 |  |  |  |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMART ASSIST |  |  | OZD |  |  | 0550-0008 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 12410 DA |  | Female | 62 KG |  | Black Or African American | Injury | Yes | Yes | Protective Measures Problem | Therapeutic or Diagnostic Output Failure | [5 empty] | Nausea | Alteration in Body Temperature | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The patient's impella 5.5 device stopped working. echocardiogram showed device in the ventricle, but the motor current was flat. suction alarms and decreased flows continued. the patient began feeling cold and clammy, nauseous. inotropes were increased. following multiple attempts to troubleshoot by physicians from critical care and transplant cardiology, the patient taken emergently to the or (operating room) to have the impella 5.5 removed. patient switched to va-ecmo (veno-arterial extracorporeal membrane oxygenation) and listed for urgent heart transplant.
35356 | 1220648-2023-02368 | 17215971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17215971&pc=OZD | 02/20/2023 | 08/09/2023 |  | 06/28/2023 |  |  | 02/20/2023 | 06/28/2023 |  |  | 10/31/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266202 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2023 |  |  |  |  | The user facility reported a 73-year-old female in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical support. the patient developed a large chest wall hematoma. multiple blood products were given to the patient and a thoracentesis procedure was performed. | [5 empty] | The investigation into the patient injury has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the patient injury was not determined. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35357 | 1220648-2023-02369 | 17219254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17219254&pc=OZD | 04/15/2023 | 11/08/2023 |  | 06/28/2023 |  |  | 04/15/2023 | 06/28/2023 |  |  | 07/31/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP APAC |  |  | ABIOMED, INC. |  |  | US |  | 2023257868 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2023 |  |  |  |  | The user facility reported a 71-year-old female undergoing a left main artery and left anterior descending artery procedure was implanted with an impella cp device for mechanical circulatory support. at the time of peel-away sheath removal, there was blood loss. the patient's blood pressure dropped and inotropes were given. the field representative did note that "two pints" of blood were delivered back to the patient. additionally, the team had a successful closure of the artery with perclose and angioseal. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35358 | 1220648-2023-02370 | 17220070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17220070&pc=OZD | 04/10/2023 | 09/01/2023 |  | 06/28/2023 |  |  | 04/10/2023 | 06/28/2023 |  |  | 12/31/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024300098 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported a 56-year male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a stroke coinciding with the impella support. the patient was said to have recovered from the stroke and neuro symptoms after roughly five days. support continued with the implanted pump following the event. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. data logs show no signs of motor current spikes or biomaterial ingestion. the root cause of the issue was most likely patient condition related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35359 | 1220648-2023-02371 | 17220232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17220232&pc=OZD | 05/31/2023 | 08/22/2023 |  | 06/28/2023 |  |  | 05/31/2023 | 06/28/2023 |  |  | 12/31/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2023 |  |  |  |  | The user facility reported a (b)(6) male with 99% diffused disease in the left anterior descending artery was implanted with an impella cp for circulatory support. there was significant bleeding on peel-away sheath hemostatic valve. two units of blood product given. low pump flow and suction alarms occurred. imaging showed that impella was under a papillary muscle. during repositioning attempts, impella cp dislodged out of ventricle. the impella cp was removed. there appeared to be clot within the impella outlet cage. a new impella cp was successfully placed. | [5 empty] | The investigation into the low pump flow and the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the low pump flow issue was biomaterial, and the root cause of the access site bleeding was most likely anticoagulation management related as patient had high act values. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident.
35360 | 1220648-2023-02372 | 17220328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17220328&pc=OZD | 05/30/2023 | 01/25/2024 |  | 06/28/2023 |  |  | 05/31/2023 | 06/28/2023 |  |  | 12/31/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325179 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2023 |  |  |  |  | The user facility reported a 57-year-old female began to decompensate after a mitral valve/left anterior descending artery stenting procedure. the decision was made to place an impella cp for circulatory support. the patient experienced bleeding from an ecmo site and impella access site during impella support. the impella sheath was tented by the staff in order to dress the site and two units of blood were provided. the dressing was later removed and the device was placed in the correct angle match position to achieve hemostasis. | [5 empty] | The investigation into the access site bleeding issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the cause of the access site bleeding issue was use related due to improper angle matching. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35361 | 1220648-2023-02373 | 17220525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17220525&pc=OZD | 06/01/2023 | 07/25/2024 |  | 06/28/2023 |  |  | 06/01/2023 | 06/28/2023 |  |  | 12/31/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300838 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2023 |  |  |  |  | The user facility reported a 74-year-old female developed limb ischemia during impella cp support. the pump was explanted as a result of the condition. it was additionally noted that prior to explant, the placement signal monitoring had to be disabled due to a consistently low diastolic value despite attempts to reposition. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 added-d6a/d6b. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35362 | 1220648-2023-02374 | 17220574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17220574&pc=OZD | 04/12/2023 | 09/18/2023 |  | 06/28/2023 |  |  | 04/12/2023 | 06/28/2023 |  |  | 09/30/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Coagulation Disorder | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2023 |  |  |  |  | The user facility reported a 59-year-old female with a congenital heart defect was implanted with an impella cp device for mechanical circulatory support during a right heart catheterization procedure. the patient experienced disseminated intravascular coagulation and the impella cp was explanted six days after implant. the patient¿s vitals were stable, and no complications were noted with explant. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. data logs were provided and investigated but no abnormalities were found. the root cause of the hemolysis issue was not determined since the product was not returned and insufficient clinical information provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35363 | 1220648-2023-02375 | 17220814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17220814&pc=PYX | 06/03/2023 | 06/28/2023 |  | 06/28/2023 |  |  | 06/05/2023 | 06/28/2023 |  |  | 09/30/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023270489 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported a 60-year-old male with a pulmonary embolism and right heart failure was implanted with an impella rp device for mechanical circulatory support. the patient experienced oozing from their impella access site following impella implant. manual pressure and femostop were initially utilized in an attempt to achieve hemostasis. the patient then developed hemolysis, however, despite recommendations to reposition the pump, the medical director opted to treat the dropping hemoglobin levels via blood transfusions. two units of blood were subsequently provided to treat the hemolysis. the access site ooze continued to be present throughout support and three units of blood were given to counteract the condition. following the interventions, the decision was eventually made to withdraw care and explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
35364 | 1220648-2023-02376 | 17221174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17221174&pc=OZD | 04/15/2023 | 01/22/2024 |  | 06/28/2023 |  |  | 04/15/2023 | 06/28/2023 |  |  | 09/30/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275620 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2023 |  |  |  |  | The user facility reported a 69-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis. the patient was given one liter of fluid and the impella device was explanted 24 hours after implant to protect the patient's kidneys. | [5 empty] | The investigation into the patient's hemolysis has been completed. no product was returned. as per the clinical review, the pump pulled back 2 centimeters as it was just a fraction deep and central venous pressure (cvp) was found to be 6 mmhg. additional fluid with turning the pump p-level down to p-3 was done. urine color got some better later after additional fluid. data logs were returned and no positioning issues. motor current spike relative to hemolysis event was observed. the impella p-level was down to p-3 during management of blood volume and troubleshooting to resolve hemolysis. the cause of the hemolysis issue was most likely patient condition related with cvp being less and additional fluid given to manage low blood volume and p-level reduced to p-3 per clinical and log review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35365 | 1220648-2023-02377 | 17221336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17221336&pc=OZD | 04/14/2023 | 09/19/2024 |  | 06/28/2023 |  |  | 04/14/2023 | 06/28/2023 |  |  | 09/30/2024 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280270 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/14/2023 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a (b)(6) female in hospital for electrophysiology ablation was implanted with an impella cp device for mechanical circulatory support after coding in the catheterization lab. the patient lost pulses in the lower extremities following the device implant. vascular surgery attempted to treat the ischemia, however they were unsuccessful in placing fem-fem bypass reperfusion catheters. the patient passed away from their pre-existing disposition prior to further intervention taking place. |  |  |  |  | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. the root cause of the limb ischemia was most likely patient condition related; based on provided clinical details that the patient was ischemic in both legs. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02377 was provided. note: corrected information provided in h6 is an addition to previous coding. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35366 | 1220648-2023-02378 | 17221337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17221337&pc=OZD | 06/03/2023 | 09/15/2023 |  | 06/28/2023 |  |  | 06/03/2023 | 06/28/2023 |  |  | 04/30/2025 |  | 06/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333065 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2023 |  |  |  |  | The user facility reported a 47-year-old male (race unknown) with heart failure was admitted for more aggressive transplant evaluation. the patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma around the axillary access site during the support. the patient was taken to the operating room and the hematoma was evacuated. support continued with the implanted pump following the intervention. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding issue was unable to be determined due to insufficient clinical information and with product not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35367 | 1220648-2023-02363 | 17210408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17210408&pc=OZD | 06/02/2023 | 11/08/2023 |  | 06/27/2023 |  |  | 06/02/2023 | 06/27/2023 |  |  | 02/28/2025 |  | 06/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334089 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2023 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction (ami) and cardiogenic shock (cgs) was implanted with an impella cp device for mechanical circulatory support via the right femoral artery. the patient additionally had ecmo support via the limb. the team explanted the impella cp and found thrombosis to the limb as well as concerns for compartment syndrome and had a right lower extremity fasciotomy performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the original report was filed the investigation has concluded. no impella product was returned for any benchtop analysis. only the clinical report was evaluated. the root cause of the bleed and was patient movement and the cause of the limb ischemia was patient condition.
35368 | 1220648-2023-02364 | 17210415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17210415&pc=OZD | 06/02/2023 | 08/11/2023 |  | 06/27/2023 |  |  | 06/02/2023 | 06/27/2023 |  |  | 02/28/2025 |  | 06/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311106 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient was on support for months when a diagnosis of blood infection was made. the infection site/etiology was not known. the team proactively explanted the 5.5 and placed a new access for concerns of infection. | [5 empty] | The investigation into the infection/sepsis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the infection/sepsis issue was most likely due to patient condition based on clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
35369 | 1220648-2023-02365 | 17213751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17213751&pc=OZD | 02/17/2023 | 08/04/2023 |  | 06/27/2023 |  |  | 02/17/2023 | 06/27/2023 |  |  | 10/31/2024 |  | 06/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023275846 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2023 |  |  |  |  | The user facility report a 67-year-old female with cardiomyopathy was being transitioned from an impella cp to an impella 5.5. device. after impella 5.5 was implanted, the patient endured access site bleeding which required one unit of red blood cell, two units of platelets, and k-centra before being transferred to the icu. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the access site bleeding was not determined. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35370 | 1220648-2023-02366 | 17213790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17213790&pc=OZD | 05/28/2023 | 08/11/2023 |  | 06/27/2023 |  |  | 05/28/2023 | 06/27/2023 |  |  | 12/31/2023 |  | 06/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179241 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2023 |  |  |  |  | The user facility reported that a 71-year-old female with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for circulatory support. the patient developed ischemia in the impella leg following implant. during delivery, it was noted that the patient's anatomy appeared too small for the impella, however the medical director decided to move forward with the implant despite the field engineer's recommendation to the contrary. a fem-fem bypass was subsequently performed as a result of the loss of flow, but the flow still remained absent. the patient was transferred for ongoing support and care. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was patient condition related to the small vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35371 | 17197831 | 17197831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17197831&pc=OZD | 05/16/2023 | 06/13/2023 | 06/13/2023 | 06/26/2023 | 06/13/2023 | 06/26/2023 |  | 06/26/2023 |  |  |  |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 28835 DA |  | Male |  |  |  | Malfunction | No | Yes | Break | Therapeutic or Diagnostic Output Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | An impella cp catheter device failed in the cath lab today. this had no patient contact. the impella was being prepped and the physician was unable to extend the catheter shaft for insertion, then the sterile sleeve around the guidewire re-access sheath ripped. the clinical rep was present during this occurrence.
35372 | 1220648-2023-02354 | 17199931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17199931&pc=OZD | 03/30/2023 | 09/20/2024 |  | 06/26/2023 |  |  | 03/30/2023 | 06/26/2023 |  |  | 11/30/2024 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312038 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/30/2023 |  |  |  |  | The user facility reported a 59-year-old male in post cardiotomy cardiogenic shock and noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the physician found it extremely challenging to obtain access. there was at least once source of extravasation and a large hematoma. pressure was applied to the hematoma. the patient continued in nearly constant ventricular tachycardia. access was obtained, a stent graft was placed, and the impella was implanted. the patient received two units of packed red blood cells and fresh frozen plasma. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. pump was not returned for investigation. the patient had very constricted vessels near access point. per the given clinical information, the root cause of the access site bleeding issue was most likely use related to improper technique. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02354 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
35373 | 1220648-2023-02355 | 17201584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17201584&pc=OZD | 06/01/2023 | 11/08/2023 |  | 06/26/2023 |  |  | 06/01/2023 | 06/26/2023 |  |  | 10/31/2024 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285762 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2023 |  |  |  |  | The user facility reported a 67-year-old male had an impella cp inserted for support of the patient undergoing percutaneous coronary intervention (pci) for a coded and ventricular fibrillation arrest. the patient had the pci performed and was then sent to the icu for continued monitoring. the impella cp site developed a hematoma and the patient was sent for exploration and vessel repair in the or. after volume replacement and 8hrs of support, impella weaned and explanted successfully. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to lack of vessel recoil. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35374 | 1220648-2023-02356 | 17201746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17201746&pc=OZD | 04/10/2023 | 11/08/2023 |  | 06/26/2023 |  |  | 04/10/2023 | 06/26/2023 |  |  | 12/31/2024 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303745 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported a 69-year-old female with 100% proximal left anterior artery and mitral valve disease was implanted with an impella cp device for mechanical circulatory support. the patient had a femoral-placed impella cp when the team allowed the patient to become confused and move. the movement caused an access site bleed. the bleed was at the repositioning sheath backing out and oozing. the team gave four units of packed red blood cells and made the choice to have surgery intervene and explant the device. a new impella 5.5 pump was implanted via axillary. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical review, the patient woke up confused and tried to get out of bed resulting in improper position in aorta with the groin continued to ooze excessively. later, the physician pulled and repositioned pump to maintain angle, tied multiple sutures around the surrounding skin and bleeding resolved. the cause of the access site bleeding issue was most likely patient movement due to patient waking up and moving out of the bed leading to improper position per clinical review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35375 | 1220648-2023-02357 | 17205894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17205894&pc=OZD | 05/27/2023 | 01/19/2024 |  | 06/26/2023 |  |  | 05/27/2023 | 06/26/2023 |  |  | 06/30/2024 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248463 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2023 |  |  |  |  | The user facility reported a 64-year-old male was implanted with an impella 5.5 for pre-operation stabilization. the patient developed an embolic cardiovascular accident approximately one day into support. care was withdrawn due to the severity of the injury and no further surgical options being available. the patient expired after being withdrawn from care. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs were returned and investigated. no discrepancies or motor current spikes were found during the data log review. no imaging was returned as well. the cause of the ischemic stroke was most likely patient condition related and the relation to impella is unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35376 | 1220648-2023-02358 | 17205899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17205899&pc=OZD | 03/27/2023 | 01/18/2024 |  | 06/26/2023 |  |  | 03/27/2023 | 06/26/2023 |  |  | 10/31/2024 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285504 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/27/2023 |  |  |  |  | The user facility reported a 59-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed leg ischemia and the md decided to pull the impella pump out. the patient remains on ecmo. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the limb ischemia was not determined. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible and based on the clinical information, the root cause of the limb ischemia was not determined since product was not returned.
35377 | 1220648-2023-02359 | 17205904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17205904&pc=OZD | 04/06/2023 | 08/02/2023 |  | 06/26/2023 |  |  | 04/06/2023 | 06/26/2023 |  |  | 03/21/2023 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314636 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2023 |  |  |  |  | The user facility reported female (age unknown) awaiting coronary artery bypass grafting was implanted with an impella cp device for mechanical circulatory support. during support, the patient was taken to the lab to repair part of the coronaries that were diseased. the patient had no flow on the coronaries on the same side of impella, the opposite side was severely diseased and had developed limb ischemia. the md decided to explant impella. upon the impella being explanted, the patient was stable and vitals had improved. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the limb ischemia issue was most likely due to patient condition related with patient having diseased iliac artery. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35378 | 1220648-2023-02360 | 17205909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17205909&pc=OZD | 05/30/2023 | 09/19/2024 |  | 06/26/2023 |  |  | 05/30/2023 | 06/26/2023 |  |  | 12/31/2024 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024304668 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/30/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing ventricular tachycardia ablation required the impella 5.5 for circulatory support. a ct scan showed both new and old cerebrovascular accidents. it is undetermined if this was related to a thrombus or the impella. patient had been critically ill on impella 5.5 and ecmo support for weeks and not improving. pt was made dnr and family eventually planned to withdraw care. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the stroke/cerebral vascular accident (cva) has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, the patient computed tomography (ct) scan confirmed the patient had an embolic stroke and it was unknown if it was related to thrombus or the impella. data logs were reviewed, and no motor current spike was observed relative to the event. the cause of the stroke issue was most likely patient condition related and unknown relation with insufficient clinical/log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02360 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
35379 | 1220648-2023-02361 | 17205911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17205911&pc=OZD | 06/01/2023 | 09/26/2023 |  | 06/26/2023 |  |  | 06/01/2023 | 06/26/2023 |  |  | 11/30/2024 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295754 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 9 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2023 |  |  |  |  | The user facility in japan reported a 9-year-old female post cardiotomy cardiogenic shock was implanted with an impella cp for circulatory support. the patient developed limb ischemia distal to the impella insertion site. in response, the device was removed prematurely. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the limb ischemia cannot be determined since the complaint device not returned and lack of clinical details. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35380 | 1220648-2023-02362 | 17205916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17205916&pc=OZD | 06/01/2023 | 10/14/2024 |  | 06/26/2023 |  |  | 06/01/2023 | 06/26/2023 |  |  | 01/31/2025 |  | 06/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315265 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/01/2023 |  |  |  |  | The user facility reported a 51-year-old female with myocardial infarction/cardiogenic shock was implanted with an impella cp device for circulatory support. while the patient was in the icu, an ongoing high pressure alarm occurred. the purge tubing was inspected for kinks, but no kinks were observed. the purge solution was changed, however the alarms persisted. the patient was already in critical condition when the impella pump was implanted and there was no improvement of the patient¿s state. additionally, the patient¿s condition worsened and developed leg ischemia at the leg incision site. the impella was removed and the patient was placed on va ecmo. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that an ongoing high-pressure occurred during impella support. the purge tubing was inspected for kinks, but no kinks were observed. the purge solution was changed, however the alarms persisted. the patient was already in critical condition when the impella pump was implanted and there was no improvement of the patient¿s state. moreover, impella cp moved into the aorta, was not able to be readvanced into left ventricle (lv). patient had already been determined the need to escalate to venoarterial extracorporeal membrane oxygenation (va-ecmo). impella was explanted and patient was put on va-ecmo. |  |  |  |  | The investigation into the high purge pressure, the positioning issues and the limb ischemia has been completed since the original report was filed. the product was returned for investigation. the returned complaint impella cp pump was visually inspected. the cannula kinked near the outlet cage was observed. the returned complaint pump ran at the investigation lab to reproduce the alarms. no high purge pressure (hpp) alarm was reproduced, and the pump flow was with specific flow rate with p-levels. no roughness surface or other abnormality was observed in the pump. the cause of the hpp issue could not be determined since the issue could not be reproduced in the investigation lab and no data logs for review were returned. the cause of limb ischemia- reversible was use related issue due to the peel away left in place. the cause of the positioning issue was use related since the pump should not come out in the first place. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Returned complaint cp pump sn (b)(6) visually inspected. cannula kink near outlet cage observed. returned complaint pump run at the investigation lab to reproduce the alarms. no hpp alarm reproduced and the pump flow were with specific flow rate with p-levels. no roughness surface or other abnormality observed in the pump. the cause of the hpp issue cannot be determined since the issue cannot reproduced in the investigation lab and no data logs for review returned. the cause of limb ischemia- reversible was use related issue due to the peel away left in place. no other interventions were pursued to resolve the ischemic leg as patient was cannulated for va ecmo. the cause of the positioning issue was use related since the pump shouldn't come out in the first place. this report is being filed as part of a retrospective review of historical records.
35381 | 1220648-2023-02344 | 17191112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191112&pc=OZD | 05/29/2023 | 08/24/2023 |  | 06/23/2023 |  |  | 05/29/2023 | 06/23/2023 |  |  | 08/31/2024 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2023257048 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2023 |  |  |  |  | The user facility reported a 56-year-old male post cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was a lower than expected pump flow with evidence of hemolysis (dialysate effluent and elevated ldh labs). as a result, the pump was removed and replaced with another impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the low pump flow and the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the low pump flow and the hemolysis was most likely biomaterial.
35382 | 1220648-2023-02345 | 17191472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191472&pc=OZD | 05/27/2023 | 11/08/2023 |  | 06/23/2023 |  |  | 05/27/2023 | 06/23/2023 |  |  | 03/31/2025 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024332358 |  | 0048-0014-EU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2023 |  |  |  |  | The user facility reported a 54-year-old female in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the positioning during the first night of support, the patient started to bleed from the access site. the hospital decided to remove heparin, re-assess the anastomosis and finally to wean and explant the pump. successfully weaned, patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. data logs were not provided. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
35383 | 1220648-2023-02346 | 17191609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191609&pc=OZD | 04/02/2023 | 08/11/2023 |  | 06/23/2023 |  |  | 04/02/2023 | 06/23/2023 |  |  | 11/30/2024 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2023 |  |  |  |  | The user facility reported a 73-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed a large hematoma around the repositioning sheath during impella support. upon explanting the pump, it was discovered that a small gastric branch had been sheared during the initial impella insertion. surgical repair was completed to treat the damage. | [5 empty] | The investigation into the vascular damage has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible and based on the clinical information, the root cause of the vascular damage issue was unable to be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35384 | 1220648-2023-02347 | 17191640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191640&pc=OZD | 03/29/2023 | 08/30/2023 |  | 06/23/2023 |  |  | 03/29/2023 | 06/23/2023 |  |  | 11/30/2024 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295017 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2023 |  |  |  |  | The user facility reported a 67-year-old female with heart failure was implanted with an impella cp device for mechanical circulatory support. that the patient developed ischemia in their right arm following impella implant via axillary access. due to concerns for the right limb, the decision was made to explant the pump. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was most likely patient condition due to patient having small vessels. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35385 | 1220648-2023-02348 | 17191645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191645&pc=OZD | 03/30/2023 | 01/19/2024 |  | 06/23/2023 |  |  | 03/30/2023 | 06/23/2023 |  |  | 12/31/2024 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298022 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2023 |  |  |  |  | The user facility reported a 77-year-old male was implanted with an impella cp for circulatory support. following an acute myocardial infarction, the patient lost pulsatility and required escalation from an impella cp to an impella 5.5 pump for support. the cp pump was left across the valve as the impella 5.5 pump was put into place. however, during this escalation, bleeding was identified at the impella cp graft site. four units of blood were given to counteract the blood loss and the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's access site bleeding has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided.
35386 | 1220648-2023-02349 | 17191660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191660&pc=OZD | 04/05/2023 | 08/11/2023 |  | 06/23/2023 |  |  | 04/05/2023 | 06/23/2023 |  |  | 08/31/2024 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268837 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2023 |  |  |  |  | The user facility reported a 61 year-old male in ventricular fibrillation was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding into their pericardium and around their access site during impella support. once taken to the or for analysis, it was further identified that the patient had poor distal flow. the pump was subsequently removed as a result of the noted conditions. | [5 empty] | The investigation into the access site bleeding and the limb ischemia has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the access site bleeding and the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35387 | 1220648-2023-02350 | 17191668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191668&pc=OZD | 04/13/2023 | 10/20/2024 |  | 06/23/2023 |  |  | 04/13/2023 | 06/23/2023 |  |  | 12/31/2024 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318104 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/13/2023 |  |  |  |  | The user facility reported a 72-year-old male had an impella cp implanted for circulatory support. the pigtail of an impella cp pump was under the papillary muscle or chordae during patient support. repositioning was attempted, however, the issue remained resulting in suction and a drop in mean arterial pressure. the following day, ischemia was noted in the impella leg and support was withdrawn. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Additional information from the complainant reported that an effort to reposition the impella cp was performed without success as the patient¿s condition did not tolerate the process. it was stated that suction alarms were received and the impella pigtail was observed to be under the papillary muscle or chordae. it was reported that the pigtail was unable to be redirected. |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02350 was provided. | The investigation for the reported limb ischemia, low pump flow and positioning issues has been completed. the pump was not returned for investigation. data logs show the impella position unknown and suction alarms with low pump flow at the end of case. the pump pigtail was found under the papillary muscle. the doctor was unsuccessful in repositioning the pump. the patient experienced ischemia in the impella leg. the cause of the positioning issue was most likely use related, which could result in low pump flow. as all the necessary information and returned device were not provided, the root cause of the limb ischemia was not determined. | The investigation for the reported positioning and low pump flow issues have been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the cause of the positioning issue was most likely use related, which could result in low pump flow. this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr (b)(4) was provided in sections b1, b5, d6a, d6b, d7a, h6, and h10.
35388 | 1220648-2023-02351 | 17191684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191684&pc=OZD | 05/27/2023 | 09/18/2024 |  | 06/23/2023 |  |  | 05/27/2023 | 06/23/2023 |  |  | 03/31/2025 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324369 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/27/2023 |  |  |  |  | The user facility reported a 75-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a stroke coinciding with impella support. patient support continued with the implanted pump following the event. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, the root cause of the stroke issue was not determined since product was not returned and insufficient clinical information provided. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
35389 | 1220648-2023-02352 | 17191699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17191699&pc=OZD | 05/28/2023 | 07/26/2023 |  | 06/23/2023 |  |  | 05/28/2023 | 06/23/2023 |  |  | 02/29/2024 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202856 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2023 |  |  |  |  | The user facility reported a 73-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis and two units of packed red blood cells were given. the hemolysis was resolved. per bedside rn, the impella was in good position. after the patient was on support for three days, the family made a decision to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the hemolysis issue was most likely due to improper positioning of the device based on the data log review. the failure mode will be monitored and trended.
35390 | 1220648-2023-02353 | 17192287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17192287&pc=OZD | 06/01/2023 | 08/09/2023 |  | 06/23/2023 |  |  | 06/01/2023 | 06/23/2023 |  |  | 02/28/2023 |  | 06/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024305674 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Coagulation Disorder | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2023 |  |  |  |  | The user facility reported a 52-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the team suspected the device was causing disseminated intravascular coagulation. the impella cp was left for one more day as the patient still required the support. one day later, the patient has successfully been weaned from the cp and survived. | [5 empty] | The investigation into the reported disseminated intravascular coagulation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause of the patient injury was unable to be determined as the failure mode was not related to impella. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35391 | 1220648-2023-02336 | 17181442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17181442&pc=OZD | 03/20/2023 | 08/25/2023 |  | 06/22/2023 |  |  | 03/20/2023 | 06/22/2023 |  |  | 02/29/2024 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP APAC |  |  | ABIOMED, INC. |  |  | US |  | 2023232428 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2023 |  |  |  |  | The user facility reported a 58-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during use of the impella device. as a result, the pump was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis remained undetermined since product was not returned. the failure mode will be monitored and trended.
35392 | 1220648-2023-02337 | 17181476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17181476&pc=OZD | 03/22/2023 | 09/06/2023 |  | 06/22/2023 |  |  | 03/22/2023 | 06/22/2023 |  |  | 10/31/2024 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2023 |  |  |  |  | The user facility reported a 67-year-old female in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient lost pulses distal to the impella insertion site. the decision was made to remove the device, at which time a dissection flap was noted within the patient's artery. a patch repair was completed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia and the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was patient condition related due to pad/pvd vessel conditions; however, the root cause of the vascular damage issue was not determined since product was not returned.
35393 | 1220648-2023-02338 | 17181939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17181939&pc=OZD | 03/26/2023 | 09/28/2023 |  | 06/22/2023 |  |  | 03/26/2023 | 06/22/2023 |  |  | 11/30/2024 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024310426 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 19 YR |  | Male |  |  |  | Injury | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2023 |  |  |  |  | The user facility reported a 19-year-old male on extracorporeal membrane oxygenation and an intra-aortic balloon pump was implanted with an impella cp device for mechanical circulatory support. under imagery, air or fibrous material was noted in the patient's aorta; the physician was unable to confirm. as such, the impella was explanted. | [5 empty] | The investigation into the embolism issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs were returned, and multiple suction events were noticed. per the clinical details and data logs, the root cause of the embolism issue was most likely patient condition related but unrelated to impella. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35394 | 1220648-2023-02339 | 17182465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17182465&pc=OZD | 04/03/2023 | 07/26/2023 |  | 06/22/2023 |  |  | 04/03/2023 | 06/22/2023 |  |  | 10/31/2024 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286794 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2023 |  |  |  |  | The user facility reported a 59-year-old male patient was on impella cp and was to be escalated to an impella 5.5 for continued hemodynamic support. the cp-accessed limb had vascular occlusion that prompted surgical closure and fem-fem bypass. the 5.5 was placed and the patient expired day later while on the 5.5 pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the issue was patient pvd condition related. the failure mode will be monitored and trended.
35395 | 1220648-2023-02340 | 17187382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17187382&pc=OZD | 03/30/2023 | 01/31/2024 |  | 06/22/2023 |  |  | 03/30/2023 | 06/22/2023 |  |  | 07/31/2024 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248938 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2023 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction (ami) and cardiogenic shock (cgs) was implanted with an impella 5.5 device for mechanical circulatory support. the md had difficulty inserting the impella past the subclavian tortuosity. the graft anastomosis was torn leading to excessive bleeding and prompting the md to remove impella and repair the artery to stop the bleeding. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation has been completed. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related since the graft anastomosis was torn.
35396 | 1220648-2023-02341 | 17187617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17187617&pc=OZD | 05/24/2023 | 09/22/2025 |  | 06/22/2023 |  |  | 05/25/2023 | 06/22/2023 |  |  | 12/31/2024 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326303 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Air Embolism | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 05/25/2023 |  |  |  |  | The user facility reported a 30-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. possible air bubbles identified in a patient's aorta under transesophageal echocardiogram (tee) imaging during impella support. concern for air being introduced into the system prompted a purge cassette change out, however, the bubbles persisted. the interventional cardiologist believes turbulent flow due to patient's small aorta may be causing what's seen as bubbles under tee imaging. | Additional information reported the patient was positive for heparin induced thrombocytopenia. during care the patient required 1 unit of packed red blood cells [prbc]. the patient continued on extracorporeal membrane oxygenation support and was later successfully weaned from the impella and the device explanted. | Additional information indicated that, during patient support, the pump experienced low placement signal alarms. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | As noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr 1220648-2023-02341 was provided in sections b1, b5, d6a, d6b, h1, h6 and h10. | The investigation into embolism issue has been completed. the device was not returned for investigation. as per clinical information provided, the cause of the embolism was not determined since insufficient clinical details were provided. unknown relation. "bubbles" in aorta may have been turbulent flow retrograde from aorta to the left ventricle (lv). the failure mode will be monitored and trended. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue could not be determined as insufficient clinical details were provided. the root cause of the thrombocytopenia could not be determined as insufficient clinical details were provided. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem b5 updated with additional information f6 and f8 should have been left blank g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2023-02341 h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-02341.
35397 | 1220648-2023-02342 | 17187622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17187622&pc=OZD | 05/25/2023 | 01/18/2024 |  | 06/22/2023 |  |  | 05/25/2023 | 06/22/2023 |  |  | 05/31/2024 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227200 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2023 |  |  |  |  | The user facility reported a 76-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient developed limb ischemia distal to the impella insertion site. as a result, an emergent trip to the or was taken for a surgical cutdown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the ischemia issue was patient condition related to calcified vessels.
35398 | 1220648-2023-02343 | 17187640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17187640&pc=OZD | 05/23/2023 | 08/09/2023 |  | 06/22/2023 |  |  | 05/23/2023 | 06/22/2023 |  |  | 02/28/2025 |  | 06/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332412 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2023 |  |  |  |  | The user facility reported a 57-year-old female in for ep ablation was implanted with an impella cp device for mechanical circulatory support. the patient exhibited limb ischemia after removal of the impella device. contralateral access was obtained and a balloon was deployed. the ischemia partly improved, however, a femoral dissection was found and required a vascular surgery consult. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia and vascular dissection has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. as per clinical review, thrombus was suspected when patient had ischemia upon explant and perclose was pinching the artery. the cause of the limb ischemia issue was most likely due to thrombus based on clinical review. the cause of the vascular damage issue was most likely use related due to perclose device pinching the artery and unrelated to impella. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35399 | 1220648-2023-02331 | 17171070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17171070&pc=OZD | 03/31/2023 | 07/26/2023 |  | 06/21/2023 |  |  | 03/31/2023 | 06/21/2023 |  |  | 07/31/2024 |  | 06/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259228 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2023 |  |  |  |  | The user facility reported a 58-year-old male with an impella cp was to be escalated to an impella 5.5 pump. the 5.5 pump failed to deliver so the team replaced with an impella cp. the patient then developed ischemia in femoral. the team made the decision to explant the impella cp and instead inset an intra-aortic balloon pump and medical therapy. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia issue was due to patient condition related with patient having pad/pvd vessel condition and being vasculopath. the failure mode will be monitored and trended.
35400 | 1220648-2023-02332 | 17173074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17173074&pc=OZD | 04/03/2023 | 01/16/2024 |  | 06/21/2023 |  |  | 04/03/2023 | 06/21/2023 |  |  | 12/31/2024 |  | 06/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024318446 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2023 |  |  |  |  | The user facility reported a 67-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a gross amount of bleeding from the jackson-pratt drain. a gross hematoma was also noted at the site of the swan-ganz catheter. four units of blood were given to the patient and pressure dressing was applied. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. as per clinical, the patient was bleeding from the impella site around sheath post brillinta administration along with heparin use. activated clotting time (act) levels were noted to be 197 and protamine was given with manual pressure being held to resolve the bleeding. the cause of the access site bleeding issue was anticoagulation management related due to high patient act values post brillinta use along with heparin per clinical review.
35401 | 1220648-2023-02333 | 17174292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17174292&pc=OZD | 05/23/2023 | 08/04/2023 |  | 06/21/2023 |  |  | 05/23/2023 | 06/21/2023 |  |  | 01/31/2025 |  | 06/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306156 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2023 |  |  |  |  | The user facility reported a 49-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed developed ischemia in the impella leg during support. as a result of the noted condition, the decision was made to explant the pump and escalate to an impella 5.5 device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the limb ischemia was most likely due to use issue with introducer peel away sheath left in place with patient during support. the failure mode will be monitored and trended.
35402 | 1220648-2023-02334 | 17174343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17174343&pc=OZD | 03/27/2023 | 07/27/2023 |  | 06/21/2023 |  |  | 03/27/2023 | 06/21/2023 |  |  | 10/31/2024 |  | 06/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292242 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2023 |  |  |  |  | The user facility reported a 42-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia and hemolysis coinciding with the impella support. the pump was explanted and the patient was escalated to impella 5.5 for support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis and the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the hemolysis issue was not determined since product was not returned and insufficient clinical details were provided, and the cause of the limb ischemia issue was most likely patient condition related. the failure modes will be monitored and trended.
35403 | 1220648-2023-02335 | 17174383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17174383&pc=OZD | 05/26/2023 | 09/15/2023 |  | 06/21/2023 |  |  | 05/26/2023 | 06/21/2023 |  |  | 02/28/2025 |  | 06/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335179 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2023 |  |  |  |  | The user facility reported a 74 year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient lost pulse in the impella limb and the foot became cold and mottled. as a result of the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the ischemia issue was patient condition related to the clamped down vessels.
35404 | 1220648-2023-02326 | 17161806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17161806&pc=OZD | 04/16/2023 | 10/27/2025 |  | 06/20/2023 |  |  | 04/21/2023 | 06/20/2023 |  |  | 01/31/2025 |  | 06/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024316439 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/21/2023 |  |  |  |  | The complaint reported an 87-year-old male noted as high risk percutaneous cardiac insertion (hrpci) was implanted with an impella cp device. the study case stated the patient endured a bloodstream infection of mrsa with possible entrance via long procedure implanting the impella and pci for several lesions. the patient was hospitalized due to dyspnea and deterioration of general condition. medication was required. the patient was reported to be stable. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. per the provided clinical information, the root cause of the infection issue was not determined since product was not returned and insufficient clinical information provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿.
35405 | 1220648-2023-02327 | 17161822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17161822&pc=OZD | 05/26/2023 | 12/06/2023 |  | 06/20/2023 |  |  | 05/26/2023 | 06/20/2023 |  |  | 02/28/2025 |  | 06/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320366 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2023 |  |  |  |  | The user facility reported a 74-year-old male (race unknown) in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the team placed an rp 23 fr sheath to the vessel in error. they did not remove the sheath but allowed the support to remain for over 46 hours with groin site access oozing. the ooze was remedied/treated by the infusion of 5 units of red blood cells and 1 unit of platelets. additionally, to close the access a contralateral leg balloon inflation was done. the balloon inflation was for over 20 minutes to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's access site bleeding has been completed. complaint device not returned for investigation. data logs were not provided. per the clinical review, the cause of the access site bleeding issue was use related because a larger size of introducer was used than what was recommended.
35406 | 1220648-2023-02328 | 17163029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17163029&pc=OZD | 04/27/2023 | 08/11/2023 |  | 06/20/2023 |  |  | 04/27/2023 | 06/20/2023 |  |  | 01/31/2025 |  | 06/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024308978 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2023 |  |  |  |  | The user facility reported a 74-year-old female (race unknown) was implanted with an impella cp device for mechanical circulatory support. ischemia was seen after the pump implant. a side-to-side femoral artery bypass was performed and the ischemia was resolved. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was most likely to be patient condition related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35407 | 1220648-2023-02330 | 17166530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17166530&pc=OZD | 05/23/2023 | 07/27/2023 |  | 06/20/2023 |  |  | 05/23/2023 | 06/20/2023 |  |  | 10/31/2024 |  | 06/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266109 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2023 |  |  |  |  | The user facility reported a 49-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the patient was noted as high risk percutaneous cardiac insertion. approximately four hours after implantation, the pump stopped. attempts to restart were unsuccessful, therefore, the impella was removed showing a clot entrained within the outflow cage. the pump was removed and patient remained stable thereafter. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the pump stop was biomaterial. the failure mode will be monitored and trended.
35408 | 1220648-2023-02329 | 17166548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17166548&pc=OZD | 05/22/2023 | 10/15/2024 |  | 06/20/2023 |  |  | 05/22/2023 | 06/20/2023 |  |  | 08/31/2024 |  | 06/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268389 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2023 |  |  |  |  | The user facility reported a 52-year-old male (race unknown) in non-st segment elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site hematoma occurring during insertion of the pump. as a result, the patient's hemoglobin dropped and 2 units of blood products were delivered. the patient also had weak distal pulses, therefore, a distal perfusion catheter placement was required. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. additionally, positioning issues and hemolysis occurred during insertion and support of the pump. no additional information related to the resolution of hemolysis was provided. the pump caught in a papillary muscle. device was attempted to be pulled back, but the pump did not move. repositioning attempts were unsuccessful. patient survived the procedure. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being filed as part of a retrospective review of historical records. the investigation for the reported hemolysis, positioning issues, limb ischemia and access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. product was not returned for investigation. data logs were not returned for investigation as well. per the clinical details, the root cause of the positioning issue was most likely use related to improper technique. without product and data log review, it is not possible to investigate the hemolysis issue. clinical mentions pump was pulled back twice while controlling access site bleeding. pump was not repositioned after confirmed improper position since the pump was caught in pap muscle and it is unknown if hemolysis was resolved. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device using improper technique. clinical mentions that the bleeding happened while moving patient to the icu room. therefore, the root cause of the access site bleeding issue was patient movement. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. as noted in the impella ifu: impella cp with smart assist system section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
35409 | 1220648-2023-02323 | 17159025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17159025&pc=OZD | 05/24/2023 | 02/22/2024 |  | 06/19/2023 |  |  | 05/24/2023 | 06/19/2023 |  |  | 02/28/2025 |  | 06/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311212 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/24/2023 |  |  |  |  | The user facility reported a 56-year-old male noted as high risk percutaneous cardiac inserted was implanted with an impella cp device for mechanical circulatory support. the implanted pump experienced saturated flows and the patient developed an access site bleed during impella support. multiple blood products were administered to counteract the bleed. the pump was subsequently explanted and the site was surgically repaired to resolve the issue. | [5 empty] | Additional updates have been provided for sections b1 and h6 related to the report of saturated flow during impella support. upon investigation closure, a supplemental mdr will be filed. this supplemental report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major and the saturated flow issues has been completed. the device was not returned for investigation. data logs shows the motor current spike followed by the saturated flow from case start. the cause of the saturated flow issue was most likely biomaterial based on log review. however, the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35410 | 1220648-2023-02324 | 17159032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17159032&pc=OZD | 05/26/2023 | 08/11/2023 |  | 06/19/2023 |  |  | 05/26/2023 | 06/19/2023 |  |  | 01/31/2025 |  | 06/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334553 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2023 |  |  |  |  | The user facility reported an 80-year-old female with occlusions in left main coronary artery and mid right coronary artery was implanted with an impella cp device for mechanical circulatory support. the patient developed lower limb ischemia coinciding with impella support. the pump was explanted as a result of the condition, however, a peripheral angio following removal revealed a clot from the right femoral artery (rfa) to the superficial femoral artery. a spyder filter was placed in the rfa and a balloon angioplasty was performed. the patient was considered stable following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia and thrombus has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. based on clinical description, the root cause of limb ischemia & thrombus were most likely due to patient condition. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35411 | 1220648-2023-02325 | 17159114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17159114&pc=OZD | 04/29/2023 | 02/02/2024 |  | 06/19/2023 |  |  | 04/29/2023 | 06/19/2023 |  |  | 12/31/2024 |  | 06/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307363 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2023 |  |  |  |  | The user facility reported a 58-year-old male implanted with an impella cp device for mechanical circulatory support was to be escalated to an impella 5.5 even though the patient had an existing left ventricle thrombus. during support, the impella 5.5 stopped functioning. the pump was subsequently removed and replaced with another impella 5.5 as a result of the reported failure. there was no patient harm. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. pump was not returned for investigation. data logs shows several motor current spike with impella stop alarms, which indicate the biomaterial interaction with the pump impellor. doctor also reported presence of the left ventricle thrombus during the implant. the cause of the pump stop issue was biomaterial based on data log review and provided clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35412 | 1220648-2023-02317 | 17142366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17142366&pc=OZD | 05/25/2023 | 08/30/2023 |  | 06/16/2023 |  |  | 05/25/2023 | 06/16/2023 |  |  | 01/31/2025 |  | 06/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330647 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2023 |  |  |  |  | The user facility reported a 62-year-old male in non-st elevated myocardial infarction was to be implanted with an impella cp device for mechanical circulatory support. the impella cp delivery failed. the team had to abort placement when the sheath (14fr) was occlusive to the vessel and causing no flow and pain to the patient. the team had to enlist the support of the anesthesia team for pain management and sedation. the team removed/explanted the 14fr sheaths and closed the vessel. the contralateral leg additionally was unable to have impella access. | [5 empty] | The investigation into the limb ischemia and the delivery issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia and the delivery issues was patient condition related to the patient¿s small vasculature. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
35413 | 1220648-2023-02318 | 17148086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17148086&pc=OZD | 03/25/2023 | 08/02/2023 |  | 06/16/2023 |  |  | 03/25/2023 | 06/16/2023 |  |  | 11/30/2024 |  | 06/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298913 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2023 |  |  |  |  | The user facility reported a 58-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was on day five of support when a diagnosis of retroperitoneal (rp) bleed was made. the rp bleed was treated with 8 units of blood and impella cp explant. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the access site bleeding issue was most likely due to patient movement. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35414 | 1220648-2023-02319 | 17148127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17148127&pc=OZD | 04/07/2023 | 11/07/2023 |  | 06/16/2023 |  |  | 04/07/2023 | 06/16/2023 |  |  | 01/31/2025 |  | 06/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314467 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2023 |  |  |  |  | The user facility reported a (b)(6) male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient developed limb ischemia in the impella limb. the repo-sheath allowed for no distal perfusion. the team placed an antegrade sheath which allowed dorsalis pedis and posterior tibialis pulses to return. the patient remained on the impella for support until care was withdrawn by the family. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical data reported, the root cause of the limb ischemia issue was patient condition related and due to vasopressors. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35415 | 1220648-2023-02320 | 17148361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17148361&pc=OZD | 09/28/2019 | 06/16/2023 |  | 06/16/2023 |  |  | 03/28/2023 | 06/16/2023 |  |  | 06/30/2021 |  | 06/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1426675 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2023 |  |  |  |  | The complainant reported (via written publication) a 50 year old patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient endured aortic valve damage during impella support with no intervention performed. approximately one year later, the patient was re-admitted to the hospital due to exacerbation of chronic heart failure triggered by bacterial pneumonia. imagery found aortic regurgitation, mitral regurgitation, and transesophageal atrial pacing. the physician attributed the likely cause to be use of the impella device from the previous year, and as such, surgery was required to replace the valve. | [5 empty] | The investigation into reported heart valve damage has been completed. the cause of the heart valve damage was not determined since there are insufficient clinical details. the relation to impella is also unknown due to insufficient clinical details. follow-up with the physician determined the patient received an aortic valve replacement and was doing well. the physician reported since it was a long-term detention, the impella cp touched the valve cusp inevitably. the physician did not know if the impella cp somehow pressed the right coronary cusp or if it deviated due to long-term touch. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
35416 | 1220648-2023-02321 | 17148365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17148365&pc=OZD | 05/26/2023 | 10/06/2023 |  | 06/16/2023 |  |  | 05/26/2023 | 06/16/2023 |  |  | 02/28/2025 |  | 06/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311466 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2023 |  |  |  |  | The user facility reported a 53-year-old male awaiting heart transplant or durable left ventricular assist device was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the access site and was sent to the or for a washout. the impella was explanted and the patient was transferred to the icu. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. as per clinical, impella 5.5 prepped and inserted via right axillary per best practices, but hematoma was observed at access site and washout was done. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35417 | 1220648-2023-02322 | 17148388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17148388&pc=OZD | 04/10/2023 | 07/13/2023 |  | 06/16/2023 |  |  | 04/10/2023 | 06/16/2023 |  |  | 10/31/2024 |  | 06/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported an 89-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient's urine became red. the impella was noted to be 4.9 cm from the atrioventricular (av). the team was made aware to pull the device back as it could be the reason for the red urine. overnight, the patient became confused, disoriented, and lost pulses in both legs. it is unknown whether the lack of pulses is impella related. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia due to the patient history of peripheral vascular disease. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35418 | 1220648-2023-02310 | 17136156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17136156&pc=OZD | 04/11/2023 | 01/19/2024 |  | 06/15/2023 |  |  | 04/11/2023 | 06/15/2023 |  |  | 11/30/2024 |  | 06/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314694 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2023 |  |  |  |  | The user facility reported a 50-year-old male post cardiac arrest was provided impella cp support. the patient lost doppler pulse in the right leg during the support. as a result of the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. no product was returned. as per clinical, there was no flow observed down the impella leg (right femoral) and the pump was explanted which resolved the ischemia. the cause of the limb ischemia issue was most likely patient condition related based on clinical review.
35419 | 1220648-2023-02311 | 17136254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17136254&pc=OZD | 03/17/2023 | 08/02/2023 |  | 06/15/2023 |  |  | 03/17/2023 | 06/15/2023 |  |  | 09/30/2024 |  | 06/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278534 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2023 |  |  |  |  | The user facility reported a 71-year-old female in ventricular fibrillation arrest was implanted with an impella cp device for support. while the patient was moving in bed, an access site bleed occurred. the sheath was adjusted and additional sutures were placed in a forward motion to the access site. the patient was given blood products after the incident and medication was adjusted to help the patient's blood pressure. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the access site bleeding issue was most likely due to patient movement leading to bleeding at groin. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35420 | 1220648-2023-02312 | 17136308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17136308&pc=OZD | 03/15/2023 | 08/30/2023 |  | 06/15/2023 |  |  | 03/15/2023 | 06/15/2023 |  |  | 12/31/2024 |  | 06/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024301674 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2023 |  |  |  |  | The user facility reported a 74-year-old male with congestive heart failure was implanted with an impella 5.5 for support. the patient developed an access site bleed coinciding with impella support. although the bleeding slowed overnight, the patient was provided 1 unit of packed red blood cells (prbc) and 2 units of fresh frozen plasma (ffp) to counteract the condition. patient support continued with the implanted pump following the transfusions. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding issue was not determined since the product was not returned and there was insufficient clinical information. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35421 | 1220648-2023-02313 | 17139928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17139928&pc=OZD | 05/19/2023 | 01/22/2024 |  | 06/15/2023 |  |  | 05/19/2023 | 06/15/2023 |  |  | 12/31/2024 |  | 06/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321429 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported a 59-year-old female in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed signs of hemolysis which prompted early removal and escalation to impella 5.5. post explant, the patient developed a hematoma and loss of pulses distal to the impella site. a fasciotomy was performed, a wound vac was placed, and blood products were delivered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ hemolysis: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported adverse events has been completed. the device was not returned for evaluation. during clinical review, it was noted that the improper pump position in the aorta triggered suction alarms. the root cause of hemolysis was most likely due to the pump was not in the proper position. the root cause of access site bleeding was most likely due to use issue. the root cause of limb ischemia was most likely due to patient condition.
35422 | 1220648-2023-02314 | 17139929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17139929&pc=OZD | 05/18/2023 | 02/14/2025 |  | 06/15/2023 |  |  | 05/18/2023 | 06/15/2023 |  |  | 10/31/2024 |  | 06/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023274589 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2023 |  |  |  |  | The user facility reported a 17-year-old male in cardiogenic shock presented for impella 5.5 device for mechanical circulatory support. there was a difficult delivery through the graft anastomosis. bleeding was noted, therefore, insertion was aborted and an impella cp was delivered via femoral access. multiple units of blood products were required as a result. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed that the pump passed all post sterile inspection checks. as per clinical, device insertion was aborted as the pump pigtail and wire were difficult to cross past over sewn anastomosis after multiple attempts. due to the angle as well as over sewn anastomosis, pump insertion was unsuccessful causing bleeding at axillary site. the cause of the delivery issue was patient condition related due to pump pigtail and wire not being able to pass over-sewn anastomosis after multiple attempts of insertion per clinical. the cause of the access site bleeding issue was most likely use related due to improper angle and over-sewn anastomosis per clinical review.
35423 | 1220648-2023-02316 | 17139944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17139944&pc=OZD | 04/20/2023 | 08/24/2023 |  | 06/15/2023 |  |  | 05/18/2023 | 06/15/2023 |  |  | 12/31/2024 |  | 06/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318838 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2023 |  |  |  |  | The user facility reported a 81-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient endured absent pulses in the impella-inserted limb distal to the impella site. an external bypass was configured in response. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was use issue related as the peel away sheath kept in patient. the retention of the introducer caused the absent pulses so a fem fem bypass was placed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35424 | 1220648-2023-02315 | 17140648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17140648&pc=OZD | 05/19/2023 | 01/19/2024 |  | 06/15/2023 |  |  | 05/19/2023 | 06/15/2023 |  |  | 12/31/2024 |  | 06/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318053 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported 62-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia distal to the impella device insertion point. as a result, the pump was removed. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, patient experienced distal limb ischemia and pump was explanted to manage complications. the clinical team also discussed with the physician on the patient management strategies and educated on best practices for insertion and management of impella device. the cause of the limb ischemia issue was not determined due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
35425 | 1220648-2023-02305 | 17126057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17126057&pc=OZD | 05/21/2023 | 08/31/2023 |  | 06/14/2023 |  |  | 05/21/2023 | 06/14/2023 |  |  | 02/28/2025 |  | 06/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324176 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2023 |  |  |  |  | The user facility reported that a 60-year-old male in st elevated myocardial infarction presented for impella cp support. the patient developed hemolysis coinciding with impella cp support. repositioning of the pump resolved the issue temporarily, however, the hemolysis returned and the decision was made to escalate the patient to an impella 5.5 pump. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis was unable to be determined as no product was returned and insufficient clinical details were provided. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35426 | 1220648-2023-02306 | 17126075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17126075&pc=OZD | 05/22/2023 | 03/10/2025 |  | 06/14/2023 |  |  | 05/22/2023 | 06/14/2023 |  |  | 02/28/2025 |  | 06/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310355 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2023 |  |  |  |  | The user facility reported that a 47-year-old male presented for impella 5.5 support. upon explant of the pump, a blood clot was immediately identified via angiogram in the right axillary. an embolectomy was subsequently performed to treat the clot. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
35427 | 1220648-2023-02307 | 17126086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17126086&pc=OZD | 05/25/2023 | 01/23/2024 |  | 06/14/2023 |  |  | 05/25/2023 | 06/14/2023 |  |  | 04/30/2025 |  | 06/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330413 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2023 |  |  |  |  | A 67 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient experienced bleeding following axillary impella insertion prompting the administration of two units of blood. support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the major access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the root cause of access site bleeding was most likely due to patient condition. correction : the section d1 (brand name) was incorrect in the initial report which has been corrected to this report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35428 | 1220648-2023-02308 | 17126099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17126099&pc=OZD | 04/14/2023 | 01/19/2024 |  | 06/14/2023 |  |  | 04/14/2023 | 06/14/2023 |  |  | 12/31/2024 |  | 06/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298733 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2023 |  |  |  |  | The user facility reported a 44-year-old female whose left anterior descending artery shut down during diagnostic heart cath. was implanted with an impella cp for support. the patient developed ischemia in the impella leg during support. a fem-fem bypass was subsequently performed to return blood flow to the leg. patient support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the limb ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the given clinical information, the root cause of the limb ischemia issue was patient condition related to small/ diseased vessels. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35429 | 1220648-2023-02309 | 17126171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17126171&pc=OZD | 05/19/2023 | 08/07/2023 |  | 06/14/2023 |  |  | 05/19/2023 | 06/14/2023 |  |  | 01/31/2025 |  | 06/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329192 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported an 83-year-old male presented for impella cp support. the patient developed ischemia in the right leg during impella support. it was noted that the patient's right lower extremity was cold with no notable pulse. as a result of the observed condition, the decision was made to explant the pump. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the limb ischemia was not determined. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35430 | 1220648-2023-02303 | 17119017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17119017&pc=OZD | 05/19/2023 | 06/13/2023 |  | 06/13/2023 |  |  | 05/19/2023 | 06/13/2023 |  |  | 02/22/2023 |  | 06/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024303151 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Female |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported a female patient of unknown age/race presented for impella support. after successful implantation, the pump had low flow and suction. the device was swapped as a result. there was no harm to the patient. | [5 empty] | The investigation into the reported malfunction of low flow has been completed. the cause of the issue was fractured impeller blade(s). the fracture was characteristic of an interaction with another device but that was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. instructions for use for the related event: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿.
35431 | 1220648-2023-02304 | 17119793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17119793&pc=OZD | 04/11/2023 | 08/28/2023 |  | 06/13/2023 |  |  | 04/11/2023 | 06/13/2023 |  |  | 11/30/2024 |  | 06/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312710 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2023 |  |  |  |  | The user facility reported a 76-year-old female in non-st segment elevation myocardial infarction presented for impella cp support. the device was placed via the femoral artery. the team observed the femoral to be occluded and so the team placed distal perfusion and performed balloon angioplasty to the artery. the limb's lesion remained and so the vascular surgery team performed a cutdown to achieve hemostasis. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia issue was most likely due to patient condition with femoral artery being occluded per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35432 | 1220648-2023-02297 | 17108057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17108057&pc=OZD | 05/14/2023 | 09/12/2023 |  | 06/12/2023 |  |  | 05/19/2023 | 06/12/2023 |  |  | 01/31/2025 |  | 06/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328980 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported a 60-year-old male with complaints of chest pain presented for impella cp support. the cp was inserted femorally and the limb became ischemic. the vascular surgeon placed fem-fem bypass for perfusion to the limb. the team reported the patient had significant peripheral vascular disease. native condition contributed to the ischemia. the patient was escalated to an impella 5.5. | [5 empty] | The investigation into the limb ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the reversible limb ischemia was pvd vessel patient condition related since fem-fem bypass resolved the issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
35433 | 1220648-2023-02298 | 17108066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17108066&pc=OZD | 05/17/2023 | 01/24/2024 |  | 06/12/2023 |  |  | 05/17/2023 | 06/12/2023 |  |  | 04/30/2025 |  | 06/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330688 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2023 |  |  |  |  | The user facility reported a 55-year-old male in cardiogenic shock presented for impella 5.5 support. the pump was implanted on (b)(6) 2023, and on (b)(6) 2023, the patient was taken for a washout hematoma. it was determined the patient was not actively bleeding. the impella device was explanted on (b)(6) 2023 at which time the patient was bridged to a left ventricle assist device. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
35434 | 1220648-2023-02299 | 17108080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17108080&pc=OZD | 05/20/2023 | 11/08/2023 |  | 06/12/2023 |  |  | 05/20/2023 | 06/12/2023 |  |  | 12/31/2024 |  | 06/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326123 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2023 |  |  |  |  | The user facility reported a 62-year-old female presented with st elevated myocardial infarction with severe hypoxemia requiring bipap. an impella cp was inserted for support. the patient developed limb ischemia during support which prompted early explant after the mitral valve coronary artery disease was treated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the provided clinical information, the root cause of the limb ischemia issue was patient condition related to small/ diseased vessels.
35435 | 1220648-2023-02300 | 17108096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17108096&pc=OZD | 05/19/2023 | 09/12/2023 |  | 06/12/2023 |  |  | 05/19/2023 | 06/12/2023 |  |  | 12/31/2024 |  | 06/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303449 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported a 67-year-old male in st elevated myocardial infarction presented for impella cp support. the patient was then transferred to a tertiary center for continued monitoring. the tertiary center explanted the impella device due to limb ischemia. vascular was consulted and fem-fem discussed and placement of an arterial line. | [5 empty] | The investigation into the limb ischemia issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. the root cause of the issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35436 | 1220648-2023-02301 | 17114808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17114808&pc=OZD | 05/19/2023 | 07/12/2023 |  | 06/12/2023 |  |  | 05/19/2023 | 06/12/2023 |  |  | 10/31/2024 |  | 06/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported an 81-year-old male in acute myocardial infarction/cardiogenic shock presented for impella cp support. the patient developed a hematoma coinciding with an access site bleed during support. pressure was initially applied to treat the hematoma however the condition returned. vascular surgery was subsequently performed to evacuate the hematoma and repair the femoral artery. the patient was provided two units of packed red blood cells as a result of the bleed and the pump was explanted. | [5 empty] | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the cause of the vascular damage was a steep insertion angle use issue. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35437 | 1220648-2023-02302 | 17114810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17114810&pc=OZD | 05/21/2023 | 01/23/2024 |  | 06/12/2023 |  |  | 05/21/2023 | 06/12/2023 |  |  | 09/30/2024 |  | 06/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275026 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2023 |  |  |  |  | The user facility reported a 63-year-old female with left main coronary artery and left anterior descending artery occlusions presented for impella cp support. after the peel-away was removed and reposition sheath was in place, the patient developed a hematoma. two units of blood products were given. hemostasis was achieved by applying manual pressure. the team escalated the patient to an impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. all best practices were followed and no other abnormalities were noted. manual pressure has been applied with protamine to solve the bleeding issue. the cause of the access site bleeding most likely was lack of vessel recoil as the bleeding occurred after switching from the introducer to the repo unit.
35438 | 1220648-2023-02289 | 17102344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17102344&pc=OZD | 05/18/2023 | 08/11/2023 |  | 06/11/2023 |  |  | 05/19/2023 | 06/11/2023 |  |  | 02/28/2025 |  | 06/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024313225 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported a 77-year-old female in acute myocardial infarction and cardiogenic shock presented for impella 5.5 support. the patient experienced an embolic stroke coinciding with the impella support. the patient support continued with the implanted pump following the event. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible to determine the root cause, and the root cause of the stroke/cva was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
35439 | 1220648-2023-02283 | 17102345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17102345&pc=OZD | 05/14/2023 | 01/16/2024 |  | 06/11/2023 |  |  | 05/14/2023 | 06/11/2023 |  |  | 08/31/2024 |  | 06/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265301 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2023 |  |  |  |  | The user facility reported a 53-year-old male with heart disease presented for impella cp support. the patient developed a hematoma coinciding with the support. as a result of the noted condition, the patient underwent a blood transfusion of unknown quantity. the patient was then escalated from an impella cp to an impella 5.5 for ongoing support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the reported access site bleeding has been completed since the original report was filed. per the given clinical details, the root cause of the access site bleeding issue was use related to improper technique.
35440 | 1220648-2023-02285 | 17098890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17098890&pc=OZD | 04/04/2023 | 01/18/2024 |  | 06/09/2023 |  |  | 04/04/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300142 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/04/2023 |  |  |  |  | The complainant reported a 65-year-old male presented in st elevation myocardial infarction and an impella cp was inserted for support. the patient developed hemolysis during the support. the pump was removed as a result of the condition and the patient was escalated to an impella 5.5 for ongoing support. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the issue was not determined since product was not returned or logs were returned, and insufficient clinical details were provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The impella device was not received by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. additional information was received via abiomed's loqi database which revealed a right groin bleed with unknown relationship to the impella, however, the impella was placed in the right femoral artery. an emergent arterial line was placed as intervention.
35441 | 1220648-2023-02286 | 17098922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17098922&pc=OZD | 08/04/2022 | 08/11/2023 |  | 06/09/2023 |  |  | 08/11/2022 | 06/09/2023 |  |  | 04/30/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The complainant reported that a 38-year-old male patient with non-ischemic cardiomyopathy presented for impella support. the patient developed hemolysis and a groin hematoma coinciding with impella cp support. the pump was subsequently removed as a result of the hemolysis and an intra-aortic balloon pump was inserted with plans to implant an impella 5.5 pump at a later time. | [5 empty] | The investigation into the access site bleeding and the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the access site bleeding issue was most likely due to patient movement during transfer to hospital, however the cause of the hemolysis was not determined. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
35442 | 1220648-2023-02288 | 17098985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17098985&pc=OZD | 04/22/2023 | 10/25/2023 |  | 06/09/2023 |  |  | 04/22/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024307533 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/22/2023 |  |  |  |  | The user facility reported a 66-year-old male in post cardiotomy cardiogenic shock and low cardiac output presented for impella 5.5 support. the patient developed hemolysis coinciding with impella support. it was noted that the patient had a large hematoma to the left mediastinum from a previous condition which was applying pressure to the left ventricle. the pump flows were subsequently low and adjustments to the support level were unable to achieve the desired flow rate. the hemolysis did not resolve and the ventricle eventually collapsed from the outside pressure applied by the hematoma. the doctor was unable to manipulate the pump, however, the decision was made to leave the device in place. the patient later began to experience complex tachycardia, coinciding with a loss in pulsatility and flow in the impella. multiple blood transfusions were performed due to the patient¿s condition. the hematoma could not be immediately relieved due to significant clotting in the mediastinum. the existing clots were eventually removed from the patient¿s chest, allowing the doctor to reposition the pump. | [5 empty] | The investigation into the hemolysis and the low pump flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the hemolysis issue and the low pump flow issue was most likely patient condition related to hematoma on the left mediastinum. the failure modes will be monitored and trended. | Section b1 and h6 were updated to include coding for the reported low pump flow. the investigation into the low pump flow issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the cause of the low pump flow and hemolysis issue was most likely patient condition related to hematoma on the left mediastinum. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35443 | 1220648-2023-02290 | 17099009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17099009&pc=OZD | 05/19/2023 | 08/30/2023 |  | 06/09/2023 |  |  | 05/19/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024323365 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported a 59-year-old male admitted with a st elevation myocardial infarction/cardiogenic shock had an impella cp inserted for support. during support, the patient developed limb ischemia on the impella-accessed limb. the team placed a reperfusion catheter for flow to the limb. the team then made a choice to explant the impella device. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the ischemia issue was not determined since the product was not returned and insufficient clinical details provided. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35444 | 1220648-2023-02284 | 17101232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101232&pc=OZD | 03/31/2023 | 09/29/2025 |  | 06/09/2023 |  |  | 03/31/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024300668 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Unspecified Infection | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/31/2023 |  |  |  |  | The user facility reported a 67-year-old male in cardiogenic shock presented for impella support. an infectious disease compromised the impella access site leading to iv antibiotic treatment for the patient. patient support continued; however frequent low placement signal alarms were later noted. the alarms were disabled and care continued as the patient was eventually transitioned to a comfort measures only status. | Additional information received care was eventually withdrawn, and the patient expired. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the infection/sepsis and the optical signal issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. as per clinical, infection assessing impella access site was found and antibiotics were given. the cause of the infection/sepsis was most likely patient condition related per clinical review. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for evaluation. data logs were returned and impella position unknown, placement signal low alarms were observed. there was no variation or impact to optical signal 5s parameters and placement signal low alarms were found with respect to p-level reduction. the cause of the optical signal issue was most likely patient condition related with placement signal low alarms observed during p-level reduction having no impact to 5s parameters based on the data log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b2 updated to death. b5 updated to include patient death narrative. h1 updated to death. h6 updated to include death and placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-02284. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email updated.
35445 | 1220648-2023-02287 | 17101271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101271&pc=OZD | 04/25/2023 | 08/04/2023 |  | 06/09/2023 |  |  | 04/25/2023 | 06/09/2023 |  |  | 01/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187646 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2023 |  |  |  |  | The user facility reported a 47-year-old female presented in cardiogenic shock. an impella cp was inserted for support. the patient developed distal limb ischemia in the right leg following impella implant. bilateral distal limb perfusion sheaths were placed in an attempt to counteract the condition, however, the decision was eventually made to wean and explant the pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the limb ischemia issue was most likely patient condition related. the failure mode will be monitored and trended.
35446 | 1220648-2023-02291 | 17101528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101528&pc=OZD | 05/10/2023 | 03/01/2024 |  | 06/09/2023 |  |  | 05/10/2023 | 06/09/2023 |  |  | 04/30/2025 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2024330300 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2023 |  |  |  |  | The complainant reported a 32-year-old female patient with cardiomyopathy required impella support. during patient support, it was noted that the patient suffered a right hemispheric non-hemorrhagic stroke during impella support. patient support continued with the implanted pump, however, the patient continued to exhibit symptoms of left sided deficits throughout the remainder of therapy. | [5 empty] | The impella device has not been returned by the customer and therefore, evaluation of the device is not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, a stroke is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva event has been completed since the original report was filed. the pump was returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined.
35447 | 1220648-2023-02292 | 17101530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101530&pc=OZD | 05/13/2023 | 07/12/2023 |  | 06/09/2023 |  |  | 05/13/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2023 |  |  |  |  | The user facility reported a 39-year-old male in for implantable cardioverter defibrillator replacement experienced pulseless electrical activity. an impella cp was implanted for support. the patient developed limb ischemia in both the impella and ecmo-inserted extremities. as such, a distal perfusion catheter was configured and the issue was resolved. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was patient condition related, as the patient had limb ischemia in both legs. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35448 | 1220648-2023-02293 | 17101532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101532&pc=OZD | 05/14/2023 | 08/02/2023 |  | 06/09/2023 |  |  | 05/14/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2023 |  |  |  |  | The user facility reported a 69-year-old male with sever peripheral artery disease in acute myocardial and cardiogenic shock presented for impella support. during impella insertion, the patient's iliac artery was dissected. the device was removed and a covered stent was applied to the artery. thereafter, a new impella cp was inserted without further issue. | [5 empty] | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the vascular damage was patient condition related to diseased vessels. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35449 | 1220648-2023-02295 | 17101559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101559&pc=OZD | 05/10/2023 | 01/31/2024 |  | 06/09/2023 |  |  | 05/10/2023 | 06/09/2023 |  |  | 06/29/2022 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Embolism/Embolus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2023 |  |  |  |  | The user facility reported a 59-year-old male of unknown race in the icu with an intra-aortic balloon pump was implanted with an impella cp device for support. the user facility reported a potential relationship between the impella device and patient's pulmonary emboli and thrombocytopenia. the allegation noted potential shearing from the impella may have played a factor in the patient's ailments. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombocytopenia has been completed. the device was not returned. per the provided clinical information, the root cause of the thrombocytopenia issue was most likely patient condition related. the relation to impella is unknown.
35450 | 1220648-2023-02296 | 17101561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101561&pc=OZD | 05/15/2023 | 06/09/2023 |  | 06/09/2023 |  |  | 05/15/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024323357 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/15/2023 |  |  |  |  | The user facility reported a 71-year-old female noted as high risk percutaneous cardiac insertion and with peripheral artery disease presented for impella support. after axillary impella implant, the patient was taken to the icu for the night. as the nurse was settling the patient it was noticed that blood was oozing from the impella cp red handle. it was reported that the graft site was bleeding and the surgeon added more sutures to the anastomosis after the impella was inserted. it was also stated that the surgical needle had possibly stuck the catheter when sewing the axillary graft. the pump was explanted and blood products were given to the patient. | [5 empty] | The investigation into the reported access site bleeding and product damage has been completed. the medical center returned the impella products for benchtop analysis. the evaluation revealed that user error was the most likely cause of the product damage. a hole in the catheter was observed at the 27cm marking on the catheter. the catheter was also found to be stretched between the 31cm and 38cm markings. the cause of the product damage was blood ingress due to a hole in the catheter resulting from improper use. the root cause for access site bleeding remains undetermined since insufficient clinical information was provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35451 | 1220648-2023-02294 | 17101564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17101564&pc=OZD | 05/11/2023 | 08/04/2023 |  | 06/09/2023 |  |  | 05/11/2023 | 06/09/2023 |  |  | 12/31/2024 |  | 06/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322939 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2023 |  |  |  |  | The user facility reported a 93-year-old female in for high risk percutaneous cardiac insertion presented for impella support. the patient was experiencing profound hypotension during support. further investigation found pericardial effusion resulting from an impella left ventricle perforation. the device was removed and the injury repaired, however, the patient required ecmo support due to continued instability. | [5 empty] | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the ventricle perforation was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35452 | 1220648-2023-02273 | 17077637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17077637&pc=OZD | 05/11/2023 | 07/26/2023 |  | 06/07/2023 |  |  | 05/11/2023 | 06/07/2023 |  |  | 01/31/2024 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2023 |  |  |  |  | The user facility reported an 83-year-old male with acute myocardial infarction/cardiogenic shock presented for impella support. the patient developed an access site bleed during support as a result of the patient being combative and sitting up. the patient's hemoglobin levels were low following the event and the patient was transfused two units of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the access site bleeding issue was most likely due to the patient movement. the failure mode will be monitored and trended.
35453 | 1220648-2023-02274 | 17077652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17077652&pc=OZD | 05/08/2023 | 09/21/2024 |  | 06/07/2023 |  |  | 05/08/2023 | 06/07/2023 |  |  | 02/28/2025 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311631 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/08/2023 |  |  |  |  | The study group reported a 72-year-male with congestive heart failure presented for impella support. while on support, the pump stopped due to the presence of a thrombus. the pump was explanted the same day following the event. | Family withdrew care and the procedure outcome was that patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and the thrombus issues has been completed since the original report was filed. the device was not returned. data logs did not report any event related to pump stop or biomaterial interaction. the root cause of the pump stop is undetermined. the cause of the thrombus issue was most likely patient condition related. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, d6, h1, and h6.
35454 | 1220648-2023-02275 | 17077708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17077708&pc=OZD | 05/16/2023 | 08/07/2023 |  | 06/07/2023 |  |  | 05/17/2023 | 06/07/2023 |  |  | 05/31/2024 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231362 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2023 |  |  |  |  | The user facility reported a 69-year-old male having st elevation myocardial infarction presented for impella support. the patient developed an access site bleed coinciding with impella support. it was noted that the patient also had oozing/bleeding from the central line and iv sites, however an exact estimate of blood loss from each site was not available. the patient was given 3 units of packed red blood cells to counteract the blood loss and and the dressing on the site was changed. the patient was considered stable throughout the event and support with the impella continued following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the access site bleeding was most likely patient condition related as the patient had oozing from multiple access sites.
35455 | 1220648-2023-02276 | 17077737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17077737&pc=OZD | 05/18/2023 | 07/19/2023 |  | 06/07/2023 |  |  | 05/18/2023 | 06/07/2023 |  |  | 11/30/2024 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293890 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2023 |  |  |  |  | The user facility reported a 58-year-old female in acute myocardial infarction and cardiogenic shock presented for impella cp support. the patient develop ischemia in the impella leg coinciding with patient support. a stent was placed in the superficial femoral artery (sfa) and a thrombectomy of the sfa and profunda was subsequently performed to treat the ischemia. flow to the limb began to improve, however, ongoing bleeding was noted at the access site. an inspection verified that the origin of the bleed was from the bifurcation of the sfa and profunda and a vascular repair was initiated. a fasciotomy was also performed at that time to the lower left extremity due to compartment syndrome. the pump was explanted as a result of the condition and another impella cp pump was implanted via axillary access for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: thrombus and vessel damage: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ limb ischemia: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was biomaterial. the failure mode will be monitored and trended.
35456 | 1220648-2023-02277 | 17077744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17077744&pc=OZD | 05/17/2023 | 09/26/2024 |  | 06/07/2023 |  |  | 05/18/2023 | 06/07/2023 |  |  | 02/28/2025 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC |  |  | US |  | 2024323000 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/18/2023 |  |  |  |  | The complainant reported a male patient required impella placement. during support, it was documented that the patient suffered an aortic dissection. although a definitive cause for the dissection could not be ascertained, it was noted that difficulty was met during impella implant due to calcification in the abdominal and terminal aorta. the impella remained in place following discovery of the dissection, however, the patient subsequently passed as a result of the condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu, vascular injury and death are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the vascular damage issue was most likely use related with medkit 16fr introducer being used for pump insertion. per ifu 0042-9010-jp only abiomed recommended introducers must be used. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.