5556 | 1220648-2025-46518 | 23063508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063508&pc=OZD | 06/13/2023 | 11/06/2025 |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  | 01/31/2025 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024309067 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional |  |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with ischemic ventricular septal defect (vsd) had an impella 5.5 implanted due to ongoing pulmonary congestion and to allow for rehabilitation prior to vsd repair. per the publication, the impella 5.5 had lower than expected flows despite optimal positioning due to small left ventricle cavity (context: elevated right-sided filling pressures) as well as episodes of increasing ectopy that degenerated into an episode of unstable ventricular tachycardia localized to the same region as the vsd. the impella 5.5 was escalated to veno-arterial extracorporeal membrane oxygenation and the patient survived. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. literature citation: frenkel, m., fajardo, r., dollerschell, j., baber, a., hermsen, j., & xia, y. (2025). ecpella bridge to heart transplantation for a large ischemic ventricular septal defect and refractory ventricular arrhythmias. jhlt open, 9, 100288. https://doi.org/10.1016/j.jhlto.2025.100288. | The investigation of the reported failure modes has been completed. the product was not returned for investigation. arrythmia/tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. low pump flow: data log review showed no motor current spikes. the flow levels are on the lower end of specification or slightly below. the cause of the low pump flow was most likely patient related, which is confirmed by the low pulsatility seen in the data logs. device history review: the device passed all the post sterile inspection checks.
5557 | 1220648-2025-46516 | 23063570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063570&pc=OZD | 08/17/2025 |  |  | 09/16/2025 |  |  | 08/17/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697815 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female | 54 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Multiple Organ Dysfunction Syndrome | Epistaxis | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced bleeding at the access site requiring intervention. subsequently, the patient went into multi-organ failure and expired. after the procedure, the patient was transferred to the unit on impella mcs and bleeding from the groin site was noted. the following day more bleeding was noted, and another suture was placed, which resolved the complication. the next day, the patient underwent a mitral clip procedure. heparin was suspended due to the patient's nosebleed. now three days after start of the impella mcs, the patient was made do not resuscitate. the patient was not a candidate for escalation in care, and it was noted, after the mitral clip procedure, the patient went into multi-organ failure. liver and kidney organs became worse, and partial thromboplastin time would not decrease. the patient expired the next day and care was withdrawn. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed, clinical noted the patient had some more oozing out of impella sore. the cause of the bleeding is determined to be use issue because bleeding was stopped by placing another suture at the place. regarding the organ failure/ epistaxis, the cause of the clinical symptom was not determined due to insufficient clinical details. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation.
5558 | 1220648-2025-46519 | 23063713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063713&pc=OZD | 08/15/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697129 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp was unable to pass through the vascular system. the pump was unable to be inserted. the cause was tortuosity of the patient's blood vessels. an intra-aortic balloon pump was inserted for continued support and there was no patient harm. | [5 empty] | Investigation summary: the investigation has been completed. unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had tortuosity of the blood vessels preventing successful delivery of impella. device history review: the pump passed all the post sterile inspection checks. | A.2 - a.4 are unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
5559 | 1220648-2025-46521 | 23063855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063855&pc=OZD | 08/18/2025 | 09/15/2025 |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026678545 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 64 KG | Non Hispanic |  | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella cp placed to support the patient in acute myocardial infarction with cardiogenic shock (ami/cgs). while on support, it was noted that there was slight oozing at the insertion site, pressure was held, lidocaine with epinephrine injected around site, angle matched and site redressed, activated clotting time at time was 398. additionally, it was noted that the patient was started on continuous renal replacement therapy. furthermore, the team elected to replace the impella due to positioning. the patient survived the explant and remained on impella support. | [5 empty] | Corrections: b1, e4, h6. the investigation of the reported failure mode has been completed. access site adverse event: based on the clinical details, the root cause is due to anticoagulation management. the impella cp ifu states that the "activated clotting time (act) should be maintained at 160 to 180 seconds. device in wrong position: the clinical details state that the pump was replaced due to positioning. there was some positioning alarms seen that were accurately triggered in the logs. due to lack of clinical information and no product returned, the root cause of the positioning issues cannot be determined. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5560 | 1220648-2025-46522 | 23063857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063857&pc=OZD | 08/18/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595444 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient coding on arrival to cath lab from emergengy department. multiple vasoactive and intropic medications started. decision was made to place impella, physician placed impella to right femoral artery (rfa), patient proceeded to loosing a pulse and requiring a second round of cardiopulmonary resuscitation, return of spontaneous circulation obtained, using single access technique physician able to angioplasty and stent the right coronary artery, intravascular ultrasound used. anti-arrhythmia medications given. cardioversion was attempted for increased heart rate. patient had multiple runs of monomorphic ventricular tachycardia, patient was successfully defibrillated. the patient started on continuous renal replacement therapy (crrt) last night and has been tolerating with fluid pulling off. this morning chest x-ray (cxr) showed improvement. right groin with impella did get oozy, nurses will change the dressing when they get the chance. the patient went into atrial fibrillation (afib) after extubation. he was given a bolus of amio for now. the impella was weaned and explanted. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. arrhythmia/ tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
5561 | 1220648-2025-46523 | 23063979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063979&pc=OZD | 08/18/2025 | 10/30/2025 |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707475 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 129 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 the patient¿s heparin was put on hold for now as there was oozing at the axillary site. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b. explant date was provided. investigation was completed and no product was returned for investigation. major bleed: bleeding from axillary site and hemostasis was achieved by pressure dressing. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device history review was completed and passed all post sterile inspection checks.
5562 | 1220648-2025-46524 | 23064022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064022&pc=OZD | 09/03/2025 |  |  | 09/16/2025 |  |  | 08/27/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658693 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 97 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient had brain hemorrhage during impella 5.5 support. the patient's ct (computed topography) head findings noted a large parenchymal hemorrhage within the left temporal lobe. as a result, heparin was discontinued and reversed with protamine. the patient was also transfused two units of blood and taken to surgery for an emergency craniotomy. eventually, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | B.2 added selection that was inadvertently omitted from the initial report that was submitted. b.3 revised date as it was entered incorrectly on the initial report that was submitted. b.7 added information that was inadvertently omitted from the initial report that was submitted. d.4 added model number as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. g.3 date was reported incorrectly on the initial report that was submitted. the date should of reflected 03-sep-2025. h.10 added related report number. this report represents the pump and another report was submitted to represent the automated impella controller. investigation summary: the pump was returned. hemorrhage: there was a reported cerebrovascular accident - a temporal hemorrhage requiring emergent craniotomy. the device was returned and there were no damages noted. there was a thrombus noted which was apparent in the pump run. due to lack of sufficient clinical details, the root cause cannot be determined. device history review: the pump passed all post sterile inspection checks.
5563 | 1220648-2025-46525 | 23064040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064040&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709971 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 72 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the decision for impella 5.5 was made on a patient to support during a cardiac surgery. it was noted that they planned for impella 5.5 via left axillary artery due to permacath in right axillary. cutdown was performed, graft sewn on, and 0.018 wire left in place in left ventricle while coronary artery bypass grafting (cabg) was performed. following cabg, impella 5.5 insertion was attempted, but ultimately failed due to patient anatomy, unable to advance into the aorta from the left axillary. impella cp placed instead, which was able to be advanced and support the patient .no patient harm was reported due to the issue. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. delivery issue: the cause of delivery issue is determined to be patient condition because the pump was unable to pass through left axillary artery. device history review: the device passed all the post sterile inspection checks.
5564 | 1220648-2025-46520 | 23064057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064057&pc=OZD | 01/24/2024 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  | 09/30/2025 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2024400786 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a publication where it was reported that a patient was in acute st-segment elevation myocardial infarction and experienced cardiogenic shock and was implanted with an impella cp for mechanical circulatory support. after implantation, the patient was at high risk for acute limb ischemia, so the decision was made to create a temporary extracorporeal bypass system using femoral artery sheaths and connecting tubing to maintain distal perfusion. the patient remained stable without ischemic complications. additionally, it was noted that the patient received blood for anemia. the impella cp was successfully weaned and explanted after hemodynamic recovery and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ literature citation: bahuva, r., giordano, j., monahan, s., dadi, f., singh, a., fichadiya, h., hayagreev, v., younus, u., & bahekar, a. (2025). temporary extracorporeal bypass system for prevention of acute limb ischemia from percutaneous ventricular assist device. jacc: case reports, 30(18), 103961. https://doi.org/10.1016/j.jaccas.2025.103961.
5565 | 1220648-2025-31474 | 23064238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064238&pc=OZD | 01/02/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 01/02/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1693068 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella support when the optical sensor became defective on the automated impella controller (aic). staff reported there was no issues with the pump or motor. aic was sent in for investigation. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. device analysis: this console was tested after arriving. unable to reproduce the reported failure mode by running the console with the pump connected for more than 24 hours. from the 5s parameter obtained from the optical cavity test ((b)(4)), it was found that the gain value is high (2.6), which does not meet the criteria per section 23 of pm procedure (b)(4). found that the purge door does not open due to the dextrose deposit on the purge door hinge (unrelated to the reported failure). the root cause of the placement signal not reliable alarm was most likely due to the defective lamp. before returning this console back to the customer, instructed to perform the following, replace the optical bench as a precaution (b)(4). remove the purge door hinge and clean the dextrose deposit (b)(4). replace the new spring and purge hinge collar (b)(4) if it is hard to clean. lubricate the purge door hinge and install it to the aic per the assembly procedure (b)(4). perform sections 10, and 23 of the pm procedure (b)(4). perform a full functional check on the console and optical system (b)(4). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5566 | 1220648-2025-46526 | 23064366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064366&pc=OZD | 08/18/2025 | 10/31/2025 |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688103 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 61 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was placed to support a patient admitted for known cardiac disease and a planned high risk percutaneous intervention. the pump was placed and supported for four days. at the time of explant, the team observed signs of a stroke. the team saw a small clot at removal and placed a stent. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The root cause of the stroke was unable to be determined due to insufficient clinical details. the cause of the embolism cannot be determined as insufficient clinical details were provided. the pump passed all the post sterile inspection checks.
5567 | 1220648-2025-31475 | 23064453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064453&pc=PYX | 11/03/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 11/03/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025533880 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an impella rp pump and an impella cp were both placed to support a patient admitted with acute myocardial infarction/cardiogenic shock and biventricular failure. during support, the rp had lost placement signal; however, the pump remained on to support the patient. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. log review showed placement signal drifting up with a simultaneous decrease in pump flow starting on the second day of support. motor current remained stable and reacted to the p-level changes. a dp placement signal not reliable alarm triggered as well. the failure mode was changed from optical signal issue to placement signal issue because the investigation revealed issues with the dp sensor. the root cause of the placement signal issue was mostly likely sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5568 | 1220648-2025-31476 | 23064495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064495&pc=OZD | 08/24/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 08/24/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507293 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump generated a ¿placement signal not reliable¿ alarm. the alarm was disabled. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product was not returned for investigation. data logs were reviewed. data review showed that at the time that the placement signal not reliable alarm did trigger as reported. product was not returned for investigation. the pump did not find any problems. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5569 | 1220648-2025-46527 | 23064496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064496&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695342 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 96 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that on day two of impella 5.5 support, the staff in the intensive care unit was getting the patient back to bed using a hoyer lift. the impella cord got stuck underneath the leg of the hoyer lift pulling out all the wires from the red port causing the pump to stop abruptly. the impella 5.5 was removed and an intra-aortic balloon pump was placed. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for evaluation. pump stop: the data logs were reviewed were consistent with an electrical open pump stop with an alarm #119 and sudden drop in pump flow and speed. the cause of the pump stop was due to a user related issue that caused an electrical open pump stop per clinical details and data log analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
5570 | 1220648-2025-31478 | 23064510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064510&pc=OZD | 09/15/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 09/16/2024 | 09/16/2025 |  |  | 02/28/2026 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025480809 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for post-cardiotomy cardiogenic shock/low cardiac output syndrome. during support, the placement signal generated unreliable alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. data logs were investigated, and as reported, the placement signal not reliable alarm triggered. these signatures fit the pattern of an air gap. this is further supported by clinical details that the complication was resolved by pressing the patient cable plug into the console. the root cause of the optical signal issues was determined to most likely be an air gap between the patient cable plug and the console based on clinical details and data logs. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5571 | 1220648-2025-31477 | 23064513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064513&pc=PYX | 12/20/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 12/20/2024 | 09/16/2025 |  |  | 02/28/2026 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025517712 |  | 10000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. while on support there were placement signal issues. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product not returned. failure mode changed from optical signal issue to placement signal issue because the root cause was found to be related to the dp sensor. (placement signal nor reliable) - psnr alarm #1051 and psnr alarm #1052 correspond with the dp sensor. data logs reviewed and 1 hour into support, ps drifts down to -100 and then spikes to 3000, pump flow increases gradually and then drops to 0, cvp spikes to 300. light, lamp, gain, contrast are stable. the cause of the placement signal issue was sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5572 | 1220648-2025-31479 | 23064537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064537&pc=OZD | 09/22/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 09/22/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512677 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump lost optical signal. support was continued following the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. there were no major changes noted in any of the 5s optical signal parameters. the cause of the optical signal issue was most likely sensor silicone wear, based data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5573 | 1220648-2025-31480 | 23064644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064644&pc=OZD | 10/01/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 10/01/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547737 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 device was implanted for circulatory support. during use, the pump experienced a placement signal stopped. device position was confirmed to be appropriate via transesophageal echocardiography, and support was continued despite the alarm condition. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. log review showed no abnormalities with the 5s parameters. the root cause of the optical signal issue was not determined since product was not returned and insufficient clinical detail was provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5574 | 1220648-2025-31481 | 23064677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064677&pc=OZD | 01/10/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/10/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542857 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction complicated by cardiogenic shock. during therapy, the pump began to drift and exhibited a malfunction of optical signal issue. position was verified via echocardiography, and flow settings were adjusted. support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. log review showed a sudden decrease in placement signal. the root cause of the optical signal issue was most likely sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5575 | 1220648-2025-31482 | 23064705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064705&pc=OZD | 01/16/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/16/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535393 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. following shockwave therapy, the placement signal of the impella cp was lost. an ¿unreliable placement signal¿ alarm appeared, and both the aortic and left ventricular signals were no longer detected. despite this, motor current remained stable and impella flows were unaffected. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed and the placement signal not reliable alarm triggered 55 minutes into support. returned product was investigated and the pump passed the laser continuity test, but when connected to the console, all 5s parameters were out of spec and the placement signal not reliable alarm reproduced. this further supports a damaged sensor. based on all the referenced details, the root cause of the optical signal issue was determined to be a damaged sensor. however, clinical details did not report the proximity with which shockwave was operated with respect to impella. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5576 | 1220648-2025-31483 | 23064761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23064761&pc=OZD | 01/22/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/22/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592240 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 device was implanted for circulatory support. during use, a ¿placement signal unreliable¿ alarm appeared. the staff was advised that the optical sensor might no longer be functioning; however, overall pump performance was not affected. the alarm was disabled, and support was continued with the implanted pump. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed, and a placement signal unreliable alarm occurred. no product was returned. the cause of the optical signal issue was a damaged optical fiber line due to the data log fingerprint of this case. the cause of the fiber line to become damaged is not known due to no product being returned and lack of clinical details. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5577 | 1220648-2025-46528 | 23065005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23065005&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  | 08/31/2025 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398379 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed to support a patient. during support, pulses were lost at the limb and distal flow by external bypass. the impella cp was replaced with an impella 5.5 and the patient survived. additional information was received. the distal perfusion catheter was placed when the patient was initially placed on impella cp support as a preventative measure. when the escalation to an impella 5.5 occurred, it was removed along with explanting of the impella cp. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was returned for investigation. ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5578 | 1220648-2025-31484 | 23065006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23065006&pc=PYX | 03/07/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 03/07/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025535519 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp flex support and during support, there were placement signal issues. the doctors noticed the placement signal over 100 and intermittently jumping off the screen. no change observed during fluctuations. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the returned product was inspected and biomaterial was found in the inlet where optical sensor sits. the 5s parameters were within normal range. the failure mode was changed to 'placement signal issue' as investigation revealed that the issue affected the dp sensor. no product was returned for analysis. the data logs show an ingestion pattern that occurred outside a p-level change. there was a motor current spike with a speed deviation. at the same instant, there was a placement signal shift and a drop in purge flow with a rise in purge pressure. pap was affected at the same instant suggesting biomaterial affecting the dp sensor. the placement signal rose briefly but resolved to normal values. the 5s parameters were stable during this time were stable. the root cause of the optical signal issue was determined to be biomaterial as evidenced by the ingestion pattern in the logs and biomaterial on returned product. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5579 | 1220648-2025-31485 | 23065014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23065014&pc=OZD | 01/25/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/25/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580933 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 device was implanted for circulatory support. while in use, a ¿placement signal unreliable¿ alarm became active. the alarm was disabled, and support was continued with the implanted pump. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was not returned for investigation. data logs were investigated. there was a drift in the placement signal. the placement signal was trending down to negative values until it went out of range and activated the alarm. there were no motor current spikes observed. the 5s parameters were in specification other than contrast and snr which decreased a little. therefore, the root cause of the placement signal issue was most likely sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5580 | 1220648-2025-31486 | 23065040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23065040&pc=OZD | 01/25/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/25/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594026 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 device was implanted for circulatory support. while in use, the placement signal (ps) was lost and a ¿placement signal not reliable¿ alarm occurred. the alarm was disabled, and support was continued with the implanted pump. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. the data logs from the case show that there was placement signal not reliable upon implant with the 5s parameters characteristic of an air gap failure. the issue was troubleshooted and support continued. the root cause of the optical signal issue was most likely due to an air gap between the aic and the patient cable plug as evidenced by the 5s parameter measurements in the data logs and the resolution after troubleshooting. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5581 | 1220648-2025-31487 | 23065141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23065141&pc=OZD | 07/03/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 07/09/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1182294 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced placement signal not reliable no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint, optical placement signal suddenly became unreliable. the issue spontaneously self-resolved, but promptly re-occurred, only to self-resolve again. during testing in danvers, the issue was not reproduced, inspection of the console revealed contamination and minor damage to the optical pump connector, although unrelated to this complaint, if not addressed it might lead to further placement signal issues. the root cause of the console issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5582 | 1220648-2025-31488 | 23065147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23065147&pc=OZD | 10/04/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 11/13/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1576134 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, a "placement signal unreliable" alarm was generated. the audio alarm was disabled, and support was continued. no patient harm occurred. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs data indicates that the unreliable placement signal persisted throughout most of the case, varying in intensity, and this correlated with the low signal not reliable. the console was sent back to field service for further inspection. while the reported issue could not be reproduced, intermittent signal distortions were observed. the root cause of the issue was identified as intermittent distortions in the optical bench signal, likely caused by a malfunction in the optical bench or its components. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5583 | 1220648-2025-46529 | 23065912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23065912&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702510 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was on impella cp support and lost pulses distal to the impella. the patient was taken to the catheterization lab for thrombus removal and fasciotomy. pulses were present after. the patient is stable and still on device support. | The patient was successfully weaned from the impella cp pump. |  |  |  |  | A.4 and a.5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Additional information received: b5 and d6b updated the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia/ thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
5584 | 1220648-2025-46530 | 23066175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23066175&pc=PYX | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590583 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues with an impella rp flex. right femoral vein access was obtained and the peel away sheath was placed. however, the catheter popped out into the right ventricle. at this time, the patient's family made the decision to cease resuscitation efforts. implantation of the impella rp flex was aborted and the patient expired. | [5 empty] | No product was returned for investigation. the cause of the delivery issue was unable to be determined as insufficient clinical details were provided. the pump passed all post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿.
5585 | 1220648-2025-31489 | 23066333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23066333&pc=PYX | 04/06/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 04/06/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025567872 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the sensor on an implanted impella rp flex pump stopped working following implant. the motor current remained stable despite the pressure sensor leading to inaccurate flow readings. support was able to continue with the implanted pump with no interruption to planned support. there was no harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. based on the investigation, the failure mode 'optical signal issue' was changed to 'placement signal issue'. the returned pump was inspected and there were no physical abnormalities. there were no damages observed on the sensor and the 5s parameters were within specification. the pump passed electrical testing. the pump was run in the flow loop and sensor drift reproduced with placement signal not reliable alarms. placement signal and pump flow drifted for about 2 hours before placement signal was lost and flow calculation was disabled. the data logs confirm #1051 placement signal not reliable alarm and show placement signal drifting up and pump flows drifting down while motor current remained stable. the flow calculation was lost about 18 hours after implant. placement signal drift continued while flow was lost before going out of specification. cvp also went out of spec incessantly towards the end of the case. the pattern observed is characterized as a possible sensor drift. the root cause of the placement signal issue was determined to be due to sensor drift as evidenced by data log pattern identified from field logs and drift reproduction on the returned product. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5586 | 1220648-2025-31490 | 23066429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23066429&pc=OZD | 12/31/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 12/31/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539604 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the placement signal and lv waveform disappeared. the impella 5.5 continued to provide support with good flows and normal motor current. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were reviewed and the controller error was confirmed. there were a total of 4 intermittent triggers during the case. returned product was inspected and there were no abnormalities. it passed the laser continuity test and all 5s parameters were within specification when plugged into a console and run in a bucket for 2.5 hours. upon review of data logs and the pump, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5587 | 1220648-2025-31491 | 23066631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23066631&pc=OZD | 12/31/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 12/31/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025551712 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient admitted for acute myocardial infarction/cardiogenic shock. after after a month and half of support, the placement signal was lost. the loss of optical signal did not prompt pump removal or replacement. the patient awaits transplant. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. after 45 days from case start, data logs show a 15-hour gradual decrease trend in the placement signal, followed by a 'placement signal not reliable' alarm and a spike to 3000, which persisted until the end of the case run. there were no major changes noted in any of the 5s optical signal parameters. the cause of the optical signal issue was most likely sensor silicone wear, based data log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5588 | 1220648-2025-31492 | 23066897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23066897&pc=OZD | 01/04/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/04/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580812 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. after five days of support, the pump lost placement signal; however, this did not prompt any intervention or pump explant. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were reviewed and the many (44 total) placement signal low alarms were confirmed early in the case. during that period, pump flows were within specification and all optical signal metrics were stable. that being said, there were 22 suction alarms that were triggering during the same period. the placement signal monitoring is seen to be disabled, but at the time that it is turned back on - around 11:00 am on 1/4 - there is a step up of the placement signal and spike in the pump flows. from this point forward there are no further placement signal or suction alarms. this event aligns with the patient update reporting that a fluid bolus was administered around that same time that day and the following updates reporting improved patient condition. product was not returned for investigation. based on data log review and clinical details reported, the root cause of the optical signal issue was determined to most likely be patient condition related. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5589 | 1220648-2025-31493 | 23066969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23066969&pc=PYX | 01/11/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/11/2025 | 09/16/2025 |  |  | 02/28/2026 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025478749 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. while on support there were placement signal issues that did not affect the patient. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data log review showed the placement signal drifting upward but not going out of range. the pump flow dropped to zero after displaying a downward trend, correlating with the upward placement signal trend. the product was returned and underwent a dp investigation, as the log review indicated an issue with the dp sensor, not the optical sensor. no physical damage was observed on the returned product. there were no abnormalities found during electrical testing. the pump was run in a pressure flow loop test, during which the sensor drift was reproduced within the first hour of operation. the lab reproduction mirrored the field logs, with the placement signal drifting downward before going out of range. the reported failure mode of the optical signal issue was changed to a placement signal issue, as the dp sensor was causing the issue to the placement signal and pump flow. the root cause of the placement signal issue was sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5590 | 1220648-2025-31494 | 23067034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067034&pc=OZD | 01/21/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/21/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557640 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that a patient was implanted with an impella cp for mechanical support. it was reported that just after starting the pump, there was a controller failure alarm and the placement signal wasn¿t visible. the pump was running without any issues. after changing the controller, the alarm remained. it was decided to not remove the pump and support continued. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella cp was returned for evaluation. upon visual inspection, no kinks in the catheter were observed. optical 5s parameters on returned pump was consistent with a fiber break. pump did not pass laser light continuity test. light was coming out of patient cable side of red handle. placement signal not reliable alarm occurred when pump was connected to console. optical fiber break was found on the patient cable side of the red handle. all cables coming out of the patient cable was observed to be twisted. patient cable bond to the red handle was not loose and was not able to translate or twist. data logs were reviewed and lt log of beginning of case shows that upon initiating pump, optical parameters are consistent with a fiber break. alarms during case start for optical system errors for catheter and console were triggered. all optical parameters were 0. clinical details noted that upon pump start up, a "controller error" alarm was triggered but pump was able to run and support patient. the root cause of the optical signal issue is a damaged fiber line. a trend chart was pulled for all impella cp optical complaints with a failure mode of "optical signal issues" and a root cause of "damaged fiber line" from january 2024 to january 2025, and no escalation is needed. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5591 | 1220648-2025-31496 | 23067160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067160&pc=OZD | 12/23/2024 | 09/16/2025 |  | 09/16/2025 |  |  | 12/23/2024 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1215713 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the impella cp pump triggered a placement signal not reliable alarm. hemodynamics were confirmed and the functionality of the pump remained unaffected. the alarm and placement monitoring were subsequently disabled for the course of planned support. there was no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs at this time showed raw optical sensor dropping the console was returned and was found to have significantly low signal not reliable. the root cause of the console issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5592 | 1220648-2025-31497 | 23067201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067201&pc=OZD | 05/03/2023 | 09/16/2025 |  | 09/16/2025 |  |  | 06/12/2023 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1108091 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported placement signal fluctuations on the automated impella controller during impella 5.5 support. despite appropriate positioning, a placement signal not reliable alarm was triggered. the pump was pushed towards the patient's chest wall at the access site and the placement signal normalized. to prevent from having to advance the pump further into the ventricle, the decision was made to place a pressure dressing over the site to aid in maintaining a normalized waveform. patient support continued with the implanted pump. the decision to withdraw care was made, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the issue was not determined since issue could not be reproduced. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5593 | 1220648-2025-31498 | 23067238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067238&pc=OZD | 10/02/2023 | 09/16/2025 |  | 09/16/2025 |  |  | 10/02/2023 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1671526 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the automated impella controller (aic) had an optical sensor error during the impella setup. impella support was initiated without the sensor. the aic was replaced without issue. the patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause for the optical system sensor error was most likely a contamination of the dirty front optical connector. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5594 | 1220648-2025-31499 | 23067242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067242&pc=OZD | 02/02/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 02/02/2025 | 09/16/2025 |  |  | 01/31/2025 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306286 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. while on support there was a placement signal unreliable alarm. the staff disabled the alarm and support with the implanted pump continued uninterrupted. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the returned pump was inspected and there were no abnormalities. the 5s parameters were within range and the pump passed the light continuity test. no problem was found as the root cause of the optical signal issue. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5595 | 1220648-2025-31500 | 23067260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067260&pc=OZD | 01/26/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 02/26/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089163 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal not reliable alarm occurred during implant. the decision was eventually made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the optical signal issues were due to a consistently low signal-to-noise ratio optical bench that intermittently affected placement signal. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5596 | 1220648-2025-31502 | 23067410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067410&pc=OZD | 05/05/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 05/05/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1682677 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by the customer that after repositioning, the impella's the placement signal numbers shot up to 200's/100's. the case was successful with the same device. there was no reported harm to the patient reported. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. console logs - particularly rtlog data - from the reported day of event are consistent with complaint and show that at some point, placement signal values periodically ¿jump¿ to values higher by approx. 110mmhg to those reported moment before (occasionally resulting in alarms # 1048, ¿signal out of range¿), then drop down to correct values. periods of misreported (shifted) values correspond to low values of optical signal . isolated instances of ¿loss of opm data¿ due to optical bench communication protocol fw anomaly were observed (as evidenced by alarms #1045 ¿opm communication invalid crc¿), but were unrelated to this complaint. it¿s been also noted that console time was at first incorrectly reported, but was later manually corrected during the case (it had no effect on console or pump operation). console inspection and testing, performed in danvers, revealed some deficiencies of its optical system: irregularities of the optical bench ccd output (affecting lower pressure values, more that higher pressure values). and debris within optical pump connector (that might have contributed to air gap, thus further lowering signal not reliable). a depleted coin-cell battery on 4-in-1 was discovered, which explains incorrect automated impella controller (aic) time setting after power-up. repair: replace optical bench and optical pump connector as a precaution, replace coin-cell battery on the 4-in-1 carrier, and perform scheduled preventive maintenance of the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5597 | 1220648-2025-31501 | 23067412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067412&pc=OZD | 01/30/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 01/30/2025 | 09/16/2025 |  |  | 09/30/2025 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the optical signal of an implanted impella cp pump disappeared after a third round of shockwave was applied to the patient during a scheduled procedure. support was completed with the implanted pump without interruption. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. data logs were returned and after 2 hours of the pump start, placement signal not reliable (snr) alarm was observed. the placement signal went to 3000, contrast decreased, light was stable, lamp increased to 100 and gain slightly increased. the snr reduced to zero. these characteristics are similar to that of a sensor damage. clinical suggests that the patient was given shockwaves and on the 3rd shock, the sensor went out. as per the ifu 0048-9007 rr the sensor can be damaged when exposed to shockwaves. therefore as per the data log review and clinical information provided, the root cause of the optical signal issue was most likely a damaged sensor due to shockwave interaction. the distance between shockwave and the sensor is unknown. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5598 | 1220648-2025-31503 | 23067473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067473&pc=OZD | 02/12/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 02/12/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606191 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 64 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support, and, during support, there were placement signal issues. placement monitoring was disabled. impella in aorta alarming despite satisfactory position. support continued. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product was not returned for investigation. data log review confirmed several position in aorta alarms, though the available logs started on 2/13 rather than 2/12. through the entire case, all 5s parameters were within specification. the positioning alarms were investigated and the algorithm triggered correctly. the motor current (mc) modulation dropped below the threshold while the placement signal (ps) pulsatility remained above the threshold. however, since placement was confirmed with imaging based on the clinical details, it is not clear what would have caused this to occur. clinical details report that the patient was also on extra corporeal membrane oxygenation support, on (b)(6) - when data logs show a 3rd period of position in aorta alarms - the patient was experiencing aystole/arrhytmias, and care was withdrawn due to patient condition. however, none of these can't be conclusively connected as the cause of the alarms that were reported on 2/12. throughout the case, mc, ps, and pump flows are all variable, however, there are no abnormal spikes that would suggest biomaterial ingestion. since product was not returned for investigation and neither the data logs nor the patient's condition based on clinical details can definitively explain the false alarms, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5599 | 1220648-2025-31504 | 23067478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067478&pc=OZD | 02/20/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 02/20/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612005 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that when a patient on impella 5.5 support coughed or sat up a certain way, the aic alarms 'impella in aorta.' this would resolve on its own. it was reported that the issue would also occur when there was no movement from the patient. echo at bedside confirmed that impella in optimal position measuring 5.8cm and in mid-lv. position monitoring was disabled. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the product was returned and the catheter was cut. the red pump plug was not returned. the data logs were viewed and the 5s parameters were within specification for the entire case. during the impella position in aorta alarms from the 19th to the 20th, the placement signal would rapidly spike while the motor current would lose pulsatility. the alarm was correctly met the triggering criteria. echo confirmed the position in the correct place and the impella position in aorta alarms were intermittently triggering leading to infer that the alarm was falsely triggered. however, the clinical details also state that this alarm occurred while the patient was moving and coughing and also why the patient was still in bed. since it can not be confirmed if the patient was moving or still while these alarms triggered, the cause of this false triggering is not known. due to lack of clinical details, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5600 | 1220648-2025-31505 | 23067648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067648&pc=OZD | 03/03/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 03/03/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574651 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hematoma | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a patient. the team placed the cp and performed percutaneous coronary intervention (pci) with shockwave. the patient's pump lost placement signal (ps) during the pci. the patient also experienced an access site hematoma, which requires a manual hold to resolve. the patient eventually expired while on impella support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. optical signal issue the pump was returned and evaluated. a 5s parameter check showed signal-to-noise ratio (snr) near zero and contrast below spec, but cavity length was within spec. these parameter values indicate sensor damage. there was also biomaterial/blood around the impeller and near the optical sensor. data log review showed 5s parameters within spec for the first ~15 minutes of the data log before changes were seen. snr dropped gradually over the next ~15 minutes before dropping to near zero. at this time, the ps went out of range. clinical details note that the placement signal went out after using shockwave. the distance between the sensor and shockwave in this case is unknown. per cp ifu, "use with shockwave intravascular lithotripsy (ivl) catheter at a distance of less than 20 mm between the optical sensor and ivl device may interfere with or damage the impella optical sensor." the root cause of the optical signal issue was most likely a damaged sensor from shockwave interaction based on clinical details and because the returned product and data logs follow a pattern characteristic of sensor damage. access site bleeding - minor no abnormalities were found with the repositioning sheath on the returned pump. a hematoma developed in left groin on the day of implant. manual pressure was held. no other clinical details were given. the root cause of the access site bleeding was not determined since insufficient clinical details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5601 | 1220648-2025-31506 | 23067674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067674&pc=OZD | 03/09/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 03/09/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025609764 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 121 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 support ongoing for 26 days when the optical signal was lost. the pump remained on for support despite the loss of signal. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. optical signal issue: data logs were reviewed and lt log of case. placement signal repeatedly out of range to 3000 with repeated placement signal not reliable (psnr) alarms. 5s optical parameters were reviewed and signal-to-noise ratio (snr) in lt log is continuously low, light is stable, lamp is slightly variable around 100%, contrast barcoding slightly, and gain is stable. rt log at beginning of case at first instance of psnr alarm shows raw optical sensor railing negative at time of alarms. 30 second window of first rt log shows raw optical sensor going out of range repeatedly. lt log of beginning of case while pump is on the automated impella controller (aic) shows low snr occurring with pump. no psnr alarms were triggered. pump contrast also barcoding slightly. clinical details noted that when psnr alarm occurred, pump position was confirmed and placement signal (ps) waveform was still present on aic at time of alarm. the root cause of the optical signal issue cannot be definitively determined as no product was returned for evaluation and limited clinical details were provided. an additional failure mode of "high purge pressure" was added based on voice of customer. high purge pressure: data logs were reviewed and lt log at time of purge issue shows gradual increase in purge pressure over a period of 6 hours with a drop in purge flow, purge pressures remaining high for approximately 14 hours before a sudden resolve in purge pressure over approximately 5 minutes. no motor current spike was observed in logs prior to rise. clinical details noted that high purge pressure alarms occurred over a month into support and tissue plasminogen activator was administered, however, was not noted if it resolved the issue. the root cause of the high purge pressure cannot be definitively determined as no product was returned to confirm cause of issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5602 | 1220648-2025-31508 | 23067792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067792&pc=OZD | 03/13/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 03/13/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586740 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella 5.5 pump triggered a persistent impella in aorta alarm during patient support. the placement signal alarm was disabled and support continued uninterrupted. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed, and there were 39 positions in aorta alarms, all of which occurred in the first 24 hours of support. the optical signal metrics were all relatively variable throughout the case, but the signal-to-noise ratio (snr) never dropped low enough for a sustained period to trigger a placement signal not reliable (psnr) alarm. an rt log during an early trigger of position in aorta alarms showed the raw optical signal to have an oscillating nature and abnormal waveform, which translated to the calculated placement signal (ps) as well. for this reason, the ps and motor current (mc) were out of phase, causing the position in aorta alarm to trigger. an rt log from (b)(6) 2025 shows the raw and calculated ps to have normalized, matching back up with the mc. it is unclear what caused the optical signal to normalize between these two times. the only clinical details provided were that the patient was put on continuous veno-venous hemofiltration (cvvh) on (b)(6) 2025, they had a small left ventricle and were in critical condition at the time of the issue. product was not returned for investigation. due to limited clinical details, product not being returned for investigation, and data logs not presenting any clear causes for the issue, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5603 | 1220648-2025-31509 | 23067933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23067933&pc=OZD | 04/03/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 04/03/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629952 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 79 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support, and, during support, there were placement signal issues. support continued. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. after two days on support, placement signal not reliable (psnr) alarms occurred. placement signal spontaneously returned while adjusting p-level but disappeared a few hours later. the next morning, the placement signal (ps) reappeared. device used seven more days for support until withdrawal of care occurred. the pump was not returned. data logs were not recovered for the day placement signal issues occurred. the cause of the optical signal issue was not determined since product was not returned, data logs for the day the complaint occurred were not recovered, and insufficient clinical information. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5604 | 1220648-2025-31510 | 23068928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23068928&pc=OZD | 04/03/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 04/03/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162561 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention (hrpci). during support there were placement signal issues that did not prompt intervention. the pump was weaned and explanted at conclusion of the hrpci. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log shows that signal not reliable (snr) and contrast dropped close to 0 and 10 as soon as the pump was inserted into the patient's body . additionally, snr is periodic, but its period does not match the periodicity of the pulsatile placement signal during the case run. the cause of the optical signal issue was not determined since the product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5605 | 1220648-2025-31511 | 23068971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23068971&pc=PYX | 03/04/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 03/04/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025590463 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. while on support there were placement signal issues. audio was disabled. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the 5s parameter trends did not match any known failures related to the pump's optical components. the cause of the optical signal issue was not determined since the product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5606 | 1220648-2025-31512 | 23069096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069096&pc=OZD | 03/06/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 03/06/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590711 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. as flows were ramping up, a sensor unreliable alarm appeared. flows and motor currents were assessed and were appropriate. the alarm was muted. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. log review showed signal not reliable below 1000 upon pump startup, which led to intermittent placement snr alarms on first day of support. the snr varied throughout the case but remained below 1000 for the majority of the 95 hour case. contrast varied but remained just above spec for a large portion of the case. gain, light, and lamp were in spec and relatively stable throughout the case. therefore, there is no indication of a hard failure of the pump. the clinical states the issue presented upon pump startup and pump position and flows were verified, but no other relevant details are known. the root cause of the optical signal issue related to the pump was not determined since product was not returned and insufficient clinical details were provided. upon data log review, it is apparent that the console is a potential cause of the issue. please refer to 25-42234-2 for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5607 | 1220648-2025-31513 | 23069243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069243&pc=OZD | 02/20/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 02/20/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1635221 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a postpartum patient admitted in with cardiomyopathy was evaluated for a transplant. the patient was implanted with an impella cp and then escalated to an impella 5.5 pump. while on support, the 5.5 pump had been on support for 11 days when there were false 'in aorta' alarms and optical signal loss. the patient was cannulated to extracorporeal membrane oxygenation (ecmo). | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. device analysis: console (b)(6) was sent back to field service for further investigation. the reported complaint could not be reproduced. visible light was detected from the optical pump connector in the console, and analog output form the ccd array of the optical bench and distortion was observed without connecting the pump. the fringe pattern was seen damped without connecting the pump. while measuring the "5s" parameter using optical cavity found that the values were very low then the specified values as per the pm procedure. upon further investigating of the data streaming issues and going through the rlm logs, the root cause was likely a corrupted microsd card in the impella connect module/rlm (usdcorruption.png). root cause: the root cause for the optical signal issue which led to positioning alarms was due to the malfunction of the optical bench with low snr. the root cause of the rlm issues was likely a corrupted microsd card. corrective action: before returning the console to the customer, the following actions are instructed to perform: 1. replace the optical bench (0042-1012) due low snr. 2. replace the microsd card (0042-2000) due to corruption. 3. perform section 23 and 24 of the preventative maintenance procedure (0042-7309). 4. perform a full functional check on the console and optical system (0042-7316). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5608 | 1220648-2025-31514 | 23069299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069299&pc=OZD | 03/15/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 03/15/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630103 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. it was reported that the optical sensor was noted to have gone out after shockwave to the left anterior descending artery (lad). the patient was successfully weaned from the pump and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. data log review shows the placement signal going out of range near the end of the case. signal not reliable drops to zero, contrast decreases, and gain and lamp increase. clinical details state that the placement signal went out after shockwave to lad. the pump was not removed due to this issue. per cp ifu (0048-9007 rs), cp use with shockwave intravascular lithotripsy (ivl) catheter at a distance of less than 20 mm between the optical sensor and ivl device may interfere with or damage the impella optical sensor. the distance between shockwave and the optical sensor in this case is unknown. the root cause of the optical signal issue was most likely a damaged sensor since the loss of placement signal seen in the data logs went out after using shockwave according to clinical details. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5609 | 1220648-2025-46532 | 23069875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069875&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512506 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced shortness of breath and chest pain during a diagnostic coronary angioplasty which revealed a severe lesion of the ostial left main coronary artery. injection of the coronary artery caused the patient to go into cardiac arrest and prompted the team to implant an impella cp and treat the vessel in the catheterization laboratory. during impella cp support, the patient lost distal pulses in the impella leg. vascular surgery was consulted for intervention, but it is unknown what treatments, if any, were performed. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | Additional information was received indicating that the patient's pedal pulses returned after the patient was weaned from vasopressor therapy. no additional intervention was required. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | New, additional information provided by the medical team was added to b5.
5610 | 1220648-2025-46533 | 23069878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069878&pc=OZD | 12/05/2024 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506469 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 30 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The medical team in italy published in 2025 on a cp support from 12/2024. the patient had complicated history with marfan syndrome and prior tyron-david and bypass. the pump was placed as care escalation from an iabp. the cp was placed along with ecmo and support with the cp had persistent low flows/suction during the support, but the pump remained on until heart transplant at day 16. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | Investigation summary: the investigation has been completed. low pump flow/suction: clinical details reported the pump was getting low flows and that the patient was on veno-arterial extracorporeal membrane oxygenation support. the product was not returned for investigation. data log review confirmed low flows and suction events were occurring. at the beginning of the case, suction alarms were triggered frequently, and flows were almost at 0 liter per minute at p-1 for the first two days of care. multiple "impella stopped - reverse flow" alarms were triggered, which occur when the pump is not running for greater than 10 seconds. log review showed that at each of these occurrences, the pump was to be turned off. the placement signal also had very low pulsatility throughout the case. there were no motor current spikes outside of the pump being turned back on. on the fourth day, purge pressure momentarily spiked to 1000 mmhg then increased by 400 mmhg, and purge flow dropped. the purge pressure spike event was temporary, and based on logs, it does not factor into the low flow/suction throughout the entire case. clinical details stated that managing this case was difficult because the patient had severe biventricular dysfunction, and the significant ventricular wall thickness causes frequent suction and periods of low flow from the impella. the cause of the low pump flow and suction were mostly likely related to the patients condition, as clinical stated that the patients biventricular dysfunction and ventricular wall thickness led to the suction and low flow events, which was confirmed by the low placement signal pulsatility seen in data log review. device history review: the pump passed all the post sterile inspection checks.
5611 | 1220648-2025-46534 | 23069881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069881&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704303 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Female | 55 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that an impella cp pump was implanted to support a patient admitted with cardiac history and plan for a high risk percutaneous intervention. the impella cp failed to deliver and implantation was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into the reported issue has been completed. no device was received for evaluation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. the device passed all post sterile inspection checks.
5612 | 1220648-2025-46535 | 23069882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069882&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612813 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male | 84 KG |  |  | Injury | Yes | Yes | Difficult or Delayed Positioning | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient on day one of impella support developed hematoma and bleeding from the impella cp access site. heparin was held and manual pressure held for over 2 hours. the bleeding and hematoma were resolved. additionally, transthoracic echocardiogram showed impella under papillary muscle attempted to reposition but unsuccessful. the patient was taken to the cath lab and repositioning of impella was successful. it was further reported that the surgeon feels that the impella cp is causing lysis which is causing increase in lactic acid and preventing the patient from recovering. the impella cp was removed. the patient survived the procedure. | [5 empty] | Patient's race and ethnicity is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Product complaint # (b)(4). investigation summary ==> no product returned. major bleed: bleed from groin and heparin was held. the cause of major bleed was most likely use issue related since heparin was held and hemostasis was achieved. positioning issue: tte showed impella under papillary muscle and pericardial effusion unattributed to impella per surgeon. the pump was repositioned. the cause of the positioning issues was unable to be determined due to lack of clinical details and no product returned. mechanical interaction with blood/hemolysis: surgeon felt pump was causing lhemolysis which caused increase in lactic acid, ldh, and pfhgb. patient remained on crrt after pump explant and was still critically ill. after reviewing the logs, suction alarms were observed starting on 8/18. there was a brief mc spike and speed deviation on 8/19 which all parameter immediately resolve. hours later there is an upward shift in mc and flows. without the product and more clinical details such as troubleshooting and any hemolysis resolution, the issue cannot be investigated further. the cause of the hemolysis cannot be determined due to no product returned and lack of clinical details. device history review ==> device (sn: (b)(6)) passed all post sterile inspection checks.
5613 | 1220648-2025-46536 | 23069888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069888&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707317 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 59 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 it was found to be in mitral sub-annular structures at p6. impella was repositioned at p2 and rotated towards the left ventricle (lv) apex with good positioning. a left ventricle caused by impella repositioning was noted. this was managed by surgical patching. impella repositioned and hemostasis was achieved. it was also reported that the patient was to go to cardiovascular operating room for resternotomy to evacuate large clot near the right ventricle. severe right ventricular (rv) failure overnight with inability to flow above p6/3.8l and more negative lv diastolic pressures, combined with severely depressed papi of 0.5 on average. care team has begun seeking to transfer this patient to a facility which can offer advanced therapies. bedside echocardiogram revealed severely underfilled rv/lv with appropriate rv function despite low papi and a large clot was revealed to be in the pericardial space. a sternotomy and exploration to evacuate clot near rv. large volume clot removed and additional drain placed in the space. upon removal of the clot impella immediately able to be ramped back up to p8 with flows of 5l/min with no suction alarms. there is ongoing thrombocytopenia after cardiac surgery. platelet (plt) count out of proportion low to clinical picture in the setting of several plt transfusions. a hit panel was sent which was reported negative. the patient has undergone a second procedure for chest exploration washout and now has an open chest. low dose heparin was started yesterday for impella related anticoagulation. the patient was weaned successfully and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. thrombocytopenia : the cause of the injury was not determined due to insufficient clinical details. thrombosis : the root cause of the injury was unable to be determined due to insufficient clinical details. vessel perforation : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks.
5614 | 1220648-2025-46537 | 23069890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069890&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695425 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a impella 5.5 was placed via the right axillary at , however, it was positioned towards mitral, which caused suction and episodes of ventricular tachycardia. an attempt was made to reposition pump in the operating room with tee guidance however, this was unsuccessful. decision was made to place the impella 5.5 via left axillary and remove it from the right axillary. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. positioning issues: the product was not returned for investigation. the root cause of the positioning issue was unable to be determined due to insufficient clinical details. pump suction: data log review confirms suction alarms started triggering during the last three days of the case. the motor current showed no spikes during the case. placement signal started to trend downward at the same time suction alarms began. at specific times of suction alarm, the minimum motor current briefly trends up while placement signal has a slight trend down. the cause of the suction was most likely improper positioning as a result of user error, as clinical details state that the pump was improperly positioned and reattempts were unsuccessful. device history review: the device passed all the post sterile inspection checks. the investigation of the reported failure mode has been completed.
5615 | 1220648-2025-46538 | 23069892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069892&pc=OZD | 08/19/2025 |  |  | 09/16/2025 |  |  | 08/19/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702724 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 106 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support with a history of peripheral artery disease, peripheral vascular disease and clotting disorder. the patient lost pulse in the impella leg. the impella was removed. the patient was hemodynamic stable and survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5616 | 1220648-2025-46539 | 23069906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23069906&pc=OZD | 06/05/2025 |  |  | 09/16/2025 |  |  | 06/05/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673983 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support after decompensating after intra-aortic balloon pump removal. the medical team were unable to get pedal pulses with doppler but were able to find strong doppler popliteal pulse. position of the pump was verified, but the left leg continued to have no flow. the patient was weaned from support four days later. the patient survived. | [5 empty] | The investigation has been completed. the device was not returned. the root cause of the injury was unable to be determined due to insufficient clinical details. a device history reviewed was completed and passed all post sterile inspection checks.
5617 | 1220648-2025-46471 | 23050208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23050208&pc=OZD | 09/06/2025 |  |  | 09/15/2025 |  |  | 09/06/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697027 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 75 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support while awaiting a heart transplant. it was reported that on the ninth day of support, there was a purge disc failure on the automated impella controller (aic) during a purge cassette exchange. as a result of the issue, the aic was replaced with a back up unit and support continued. there was no harm reported due to the aic malfunction. additionally, on the 34th day of impella support, it was reported that the patient experienced an ischemic stroke which then converted to a hemorrhagic stroke. the patient was noted to have mild left sided weakness. computed tomography was performed, and the stroke was noted to be stable. the physician stated that they did not believe the stroke to be impella related, however, it could not be definitively ruled out as a contributing factor. it was noted that the patient was not under anticoagulated at the time of the ischemic stroke event. the following day, the impella 5.5 was replaced with a new impella 5.5. there have been no issues reported with the new pump, and the patient remains on impella 5.5 support awaiting heart transplant. this complaint involves an automated impella controller and an impella 5.5. this report specifically addresses the impella 5.5. a separate report will be submitted for the automated impella controller. refer to section h.10 for the related report number. | [5 empty] | B2 revised selection as it was entered incorrectly on the initial report that was submitted. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the pump passed all the post sterile inspection checks. regarding the stroke, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 type of investigation, investigation findings, conclusion codes and component code were updated accordingly based on the completed investigation. | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5618 | 1220648-2025-46472 | 23050361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23050361&pc=OZD | 09/03/2025 | 10/23/2025 |  | 09/15/2025 |  |  | 09/03/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100964 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Operating System Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the automated impella controller (aic) failed to shut down and made a loud buzzing noise for five minutes while stuck on a loading screen. no patient harm was reported. | [5 empty] | H.10 related report number was added. the impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | A.4 - a.6 are unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5619 | 1220648-2025-46473 | 23050522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23050522&pc=OZD | 08/16/2025 |  |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697647 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 92 KG | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a pump stop during impella cp support. the pump was successfully restarted at p2 with flow saturation and intermittent spikes observed in motor current. the patient was on support for 30 days and ultimately expired. | [5 empty] | B.7 added information as it was inadvertently omitted from the initial report that was submitted. h.6 added a medical device problem code. the investigation has been completed. clinical details report a pump stop. the patient was put on support. the pump was able to be restarted, but saturated flows with intermittent motor current (mc) spikes were subsequently noted. care was withdrawn later that day. data logs were reviewed and the pump stop was confirmed. the first temporary pump stop occurred the morning of (b)(6) 2025. at this time, there was a motor current spike above 1400 ma with a speed deviation, triggering a high mc pump stop. the pump was able to immediately restart, but purge pressure stepped up slightly and mc remained elevated compared to prior to the spike. there were a couple more mc spikes between that date and (b)(6) 2025, when the next temporary pump stops occurred. the same characteristics were noted in the mc. purge pressure stepped up and stepped down in between the spikes in mc. after each mc spike, the mc generally trended higher. temporary pump stop: the impella cp design indication is 8 days of use, and the first pump stop triggered beyond 8 days of use. however, the product was not returned for investigation. since the patterns noted in data logs could be indicative of multiple things, the cause of the temporary pump stop could not be determined. saturated flows: as mentioned above, the product was not returned for investigation. additionally, the patterns noted in data logs could be indicative of multiple causes. therefore, the cause of the saturated flows could not be determined. the device history review was completed and the pump passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5620 | 1220648-2025-46474 | 23050691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23050691&pc=PYX | 08/16/2025 |  |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025567901 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Life Threatening | Required Intervention |  |  |  |  | IMPELLA 5.5. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was noted that the patient had progressive heart failure symptoms and underwent a transthoracic echocardiogram that showed left ventricular ejection fraction of 30-50% and moderate mitral regurgitation. the patient underwent an abnormal stress test and a left heart catheterization confirmed the patient had multi-vessel coronary artery disease. the decision was then made for coronary artery bypass graft (cabg) with an elective impella 5.5 placement. during the cabg, the patient was cannulated for bypass and then successfully underwent three cabg's. the impella 5.5 was implanted, support was initiated, and the patient was weaned from bypass. after closing the chest, and during closing of the skin, the patient experienced an episode of ventricular fibrillation. the chest was re-opened but no obvious issues were observed, however it was noted that the right ventricle was stunned. the decision was then made to place an impella rp flex for increased mechanical support. the patient was then closed and sent to the cardiac intensive care unit to recover in critically stable condition. on the third day of bipella support with an impella rp flex and impella 5.5, the patient experienced ventricular tachycardia (vt) arrest which required cardioversion to achieve a return of spontaneous circulation. an echocardiogram was performed to verify the impella devices were well placed. on the fifth day of support, it was noted that the patient experienced vt episodes when the patient was stimulated. there was no change in support and the patient was monitored. the following day, the impella rp flex was weaned and explanted. it was noted the day following impella rp flex removal that the patient has stabilized since the removal with no more runs of vt. it was noted that after impella rp flex removal, the patient experienced pulmonary hypertension and is being treated with inhaled nitric oxide. the patient remained on impella 5.5 support for an additional seven days before being successfully weaned and explanted. the explants of impella rp flex and impella 5.5 were not performed earlier than planned as a result of the complication. the patient survived. this complaint involves two impella devices: an impella 5.5 and an impella rp flex. this report specifically addresses the impella rp flex. a separate report was submitted for the other device associated with this complaint. refer to section h.10 for the related report number. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5621 | 1220648-2025-31427 | 23050828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23050828&pc=OZD | 07/13/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 07/13/2023 | 09/15/2025 |  |  | 03/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024326581 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. wile on support, the patient was having an ongoing placement signal drift, about 20 points higher than arterial line, indicating a placement signaling issue. this procedure was a successful bridge to a left ventricular assist device. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. per the clinical information provided, the root cause of the optical signal issue was not determined since product was not returned and data logs were not returned as well. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5622 | 1220648-2025-46475 | 23050914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23050914&pc=OZD | 08/16/2025 | 10/28/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707345 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that they encountered positioning issues. the impella in aorta" alarm occurred in the morning and there were no preceding events to the alarm. the patient was laying still in bed and the transthoracic echocardiogram showed the impella completely in the aorta. previous measurement was 4.99 centimeters in the left ventricle. the patient was taken back to operating room where the device was removed and a new device was inserted. no patient harm was reported due to the issue. | [5 empty] | A.4 is unknown. the impella device has not yet been received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Impella 5.5 sn (b)(6) returned for investigation. positioning issue: data review: data logs confirmed pump was in aortic area (about 12.25 hours into the case). pump then was stopped manually & disconnected from the console. pump was returned for analysis. visual inspection found no issues on the pump. the cause of positioning issue was unable to be determined as limited clinical details were provided and no problem was found on the pump. pump sn: (b)(6) passed all post sterile inspection checks.
5623 | 1220648-2025-46470 | 23051270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051270&pc=OZD | 07/28/2025 |  |  | 09/15/2025 |  |  | 09/11/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692213 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Network Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller was not connecting to ic.com. after troubleshooting, it is likely that the controller has a bad sd card. the blue cloud light and amber airplane button were blinking simultaneously for more than a few minutes. no patient harm was reported. the device was received for evaluation and it was identified that the device experienced corruption of the microsd disk. | [5 empty] | Section a information was not provided. additional information was received during the device evaluation indicating that the device experienced corruption of the microsd disk during the reported event. in review of this information, the event is considered reportable. the automated impella controller (aic) was received from the customer and an investigation regarding the returned device has been completed. console logs from the reported day of event were reviewed and the complaint condition was confirmed. the logs showed frequent interruptions of data connection from the aic to the remote link module (rlm). the rlm journal could not be retrieved from the device due to an rlm malfunction. during testing of the returned aic, the device was not able to fully boot-up and was continuously rebooting, which prevented it from receiving data from aic as well as from connecting to wi-fi. rlm information monitored by serial number indicated an inability to mount one of the drive partitions, presumably due to file system corruption. preliminary analysis of the microsd card using a third party hardware and software confirmed corruption of the microsd disk. after a known-good microsd card was installed and a connection was established with the aic.
5624 | 1220648-2025-46476 | 23051298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051298&pc=OZD | 08/16/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705226 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that they encountered a placement signal issue. after placement of the impella, suddenly flows dropped to lows 2¿s and team reported a psnr (placement signal not reliable) alarm. once local team was on site, they checked positioning and had physician reposition which did not fix the alarm. impella remained in place for procedure and was removed at completion of procedure. | [5 empty] | The investigation was completed and the device was not returned. optical sensor issue: data logs show a psnr alarm with a spike to 3000 after 30 minutes of pump use. the snr and contrast dropped to 0, and the light and gain remained unchanged during the psnr event. the data log indicates the known pattern of a damaged sensor. as per the clinical information, the iabp turned off after the pump started and remained in place, followed by a reported psnr alarm. shockwave was also utilized during the case run. the cause of the optical signal issue was most likely sensor damage, based on data log review. however, the cause sensor damage could not be confirmed without the returned product and sufficient clinical information. the device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5625 | 1220648-2025-46477 | 23051360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051360&pc=OZD | 08/16/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693855 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 81 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Tachycardia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that a patient implanted with an impella cp experienced sepsis. as a result, the pump was explanted and the patient expired. | The complainant reported a patient with postcardiotomy cardiogenic shock was implanted with impella cp and impella rp flex devices for bipella mechanical circulatory support (mcs). the patient underwent minimally invasive mitral valve replacement one day prior and had a turbulent postoperative course w/ marked bleeding and hypotension. after bipella implants, the patient was taken to the unit on mcs. upon transport to the unit, the patient became increasingly hypotensive (mean arterial pressure was approximately 50). drips were increased to the following doses: epinephrine 0.6 mcg/kg/min, norepinephrine 1 mcg/kg/min, dobutamine 7.5 mcg/kg/min, and vasopressin 0.03 units/min. milrinone was discontinued. after drip changes and 250 albumin, mean arterial pressure improved to greater than 65. the patient continued bipella support. continuous renal replacement therapy (crrt) was planned to improve acidosis. approximately three days later, the impella cp alarmed position in the aorta and was support level was decreased to p-3. continuing support and now four days thereafter, the heparin drip (which was only at <100units/hr.) had been off since one day prior. however, activated clotting time continued to trend up, with current result ¿out of range high.¿ treatment rendered. echo was done yesterday showing improvement in right ventricle function and no right ventricle distension. impella position adequate and left ventricle function still decreased. crrt continued with fluid removal as tolerated, and a tube feeding was started. the next day, approximately eight days after start of the bipella support, the impella rp flex placement signal was noted to be elevated (100s/20s). however, there was no change in purge flow, purge pressure, motor current or pump flow. the nurse noted the patient had received two units of packed red blood cells, fresh frozen plasma, platelets, and cryoprecipitate after draining 170ml blood from chest tube (it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient). a repeat chest x-ray and echocardiogram were completed and showed no change in position from previous placement. activated clotting time (act) was down to 182 following blood products administration. discussion was made with the physician, and the plan was just to monitor cp/rp numbers and hemodynamics and continue current treatment. within about an hour, waveforms returned to baseline. then early morning following the patient being cleaned up and repositioned, the impella rp flex waveforms again elevated (140s/30s). there was no change in pump flow, purge pressure of flow, or motor current, but placement signal pressure remained elevated. repeat chest x-ray showed no change in position of rp flex. act 180 (no systemic anticoagulation at this time). the physician was notified, and plan remained to continue current support and monitor acts every 2 hours. act was noted to have remained greater than 160 seconds. the following day, it was noted the rp flex waveforms returned to baseline. the patient remains on mechanical ventilation. the plan noted was to continue current treatment and recheck echocardiogram the following day and possibly wean the impella as tolerated. with sedation on hold, the patient had not awakened; however, the patient experienced tachycardia and blood pressure and reap rate increased. therefore, fentanyl was restarted. in the afternoon continuous renal replacement therapy was stopped with no plans to restart. the next day, lactate trended up to 6.7. sedation was off and the patient¿s pupils are reactive. however, the patient has remained unresponsive, with no gag or cough. dobutamine still at 5 with no other drips. the following day lactate was noted still increased, and it was believed to be possibly sepsis. the impella rp flex was weaned and explanted this morning. the impella cp was weaned down to support level p-4 and with possible plans to explant later. however, the patient would expire within the next hour as care was withdrawn. |  |  |  |  | No product was returned for investigation. sepsis/ tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: patient was on bipella support and blood was noted in the chest tube. the cause of the bleeding is determined to be patient condition because the patient is having coagulopathy as per clinical details. the device passed all post sterile inspection checks. | Medical safety review was completed and noted the following: the event describes bipella support provided to a patient in a decompensating state. there were complications noted with the impellas and the patient expired while on the impella cp device. the rp flex was explanted approximately one hour or so prior to the demise. impellas events is a serious injury and should be reported. however, with the impella cp, there is not enough to dissociate the device from the demise outcome. in this case, however, the patient's underlying condition likely contributed most to the demise outcome. the patient started in a decompensating state and never fully recovered. the family eventually decided to withdraw care from the patient which appears to have led to the demise outcome. medical safety officer reviewed the event and noted the same in the second impella cp device cannot be dissociated for the outcome and impella rp flex is a serious injury type of reportable event. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | During review of the case and report writing of the related manufacturing device report a correction to the reported event for complete details and events was identified. consequently, the following sections were updated accordingly: b5 event description, b6: tests/lab data including dates, h6 coding (health effect clinical and impact) and h10: related report number. related medical device reports: this impella cp device report is one of two devices associated with the event. there is not enough information to dissociate the impella cp from the demise outcome. however, it should be noted, the patient experienced a "turbulent" postoperative course and remained in critical condition not fully recovering. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex.
5626 | 1220648-2025-46478 | 23051384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051384&pc=OZD | 08/16/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703415 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient placed impella cp support for non-st-segment elevation myocardial infarction (nstemi). there were reported suction events, correlated with periods of ventricular dysrhythmia. the patient was started on amiodarone for dysrhythmias and systemic heparin for impella. later, the pump was successfully weaned and explanted. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: no product was returned. after reviewing the logs, suction alarms were observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5627 | 1220648-2025-31428 | 23051391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051391&pc=OZD | 08/21/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 08/21/2023 | 09/15/2025 |  |  | 05/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355555 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. while on support, intermittent sxn and placement signal low alarms were noted. the team performed serial echo¿s and felt that the position was optimal. the placement signal not reliable alarms persisted. flows, purge flow (pf) and purge pressure (pp) are stable as is the patient. no intervention was needed. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. according to the clinical, on the 39th day of impella 5.5 support intermittent suction and placement signal low alarms were recorded. after optimal pump position was confirmed a placement signal not reliable alarm was also recorded. no attempts were made to resolve the issue. no product was returned for a visual inspection. data log analysis confirmed a drop in signal not reliable, light, contrast, and gain. while the lamp increased during the same period of time. the most likely cause of the optical signal issue was an air gap. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5628 | 1220648-2025-46479 | 23051403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051403&pc=OZD | 08/16/2025 |  |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026716029 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced access site bleed and there were issues with pump suction during impella cp support. anticoagulation was paused, the patient was administered one unit of blood and a femstop was applied. albumin was administered for the suction issues. after 6 days of support, the patient expired. | [5 empty] | Product complaint # (b)(4). investigation summary : product was not returned for review. pump suction: data logs not available. patient went into afib and had suction events which resolved once nsr returned. the cause of the pump suction issue was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: bleeding at all access sites reported so anticoagulation was stopped which resolved the issue. 3 days later it is mentioned the patient is still oozing so blood was given. the cause of the major bleed was not determined as limited clinical details were provided and no product was returned for review. device history lot : device lot #: 1958339. device history batch : subcomponent lot#: n/a. device history review : pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5629 | 1220648-2025-46480 | 23051470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051470&pc=OZD | 08/16/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674817 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella cp lost placement signal but remained implanted and supporting despite the malfunction. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This investigation has been completed. the pump was not returned. optical signal issue - at some point during the case, aorta (ao)-placement signal (ps) was muted because of not reliable ps. ao-ps showed no values and also left ventricle-ps showed no values. a kink in the connector cable and catheter shaft was excluded, also the puncture area. the patient was highly impella dependent. a replacement from automated impella controller or check connector connection was not possible. data logs were not recovered. additional information was requested from the representative but no information was received back. the cause of the optical signal could not be determined due to no product returned, logs not being recovered, and insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
5630 | 1220648-2025-46481 | 23051493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051493&pc=OZD | 09/24/2024 | 10/24/2025 |  | 09/15/2025 |  |  | 09/24/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525862 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was on impella 5.5 due to escalation from a previous impella cp. the patient required blood products to resolve suction alarms, and tpa for low purge flows. due to concerns for pump saturation, dysfunction, and potential failure, the impella was exchanged for another impella 5.5. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. high purge pressure - data log review showed purge pressure rise intermittently while purge flows continually trended down. the cause of the high purge pressure was biomaterial build-up due to the intermittent purge pressure rise with decreasing purge flows and increasing motor current seen in the data logs. high or saturated pump flow - the cause of the saturated flows was biomaterial build-up due to increasing motor current (mc) at the same p-level, decreasing purge flows, increasing purge pressure all without a mc spike. pump suction - the cause of the pump suction could not be determined as the pump was not returned, inconclusive data log review, and insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The investigation for the reported high purge pressure, saturated pump flow and pump suction has been completed. pump was not returned. data log review showed purge pressure rise intermittently while purge flows continually trended down. no mc spike noted but mc did trend up with purge metrics. the root cause of the high purge pressure was biomaterial build-up due to the intermittent purge pressure rise with decreasing purge flows and increasing motor current seen in the data logs. the root cause of the saturated flows was biomaterial build-up due to increasing mc at the same p-level, decreasing purge flows, increasing purge pressure all without a mc spike. data log review showed pump suction continued throughout case even after fluid given and pump repositioned. the root cause of the pump suction could not be determined as the pump was not returned, inconclusive data log review, and insufficient clinical details
5631 | 1220648-2025-46482 | 23051508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051508&pc=OZD | 10/03/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 10/03/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507984 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | Pulmonary Embolism | Stroke/CVA | Thromboembolism | Embolism/Embolus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 due to escalation from a previous impella cp and another impella 5.5 with low pump flows. while on support, the patient was found to have right middle cerebral artery stroke and suction/pump flow issues. a thrombectomy was performed, incidentally finding pulmonary embolism. additional left ventricle thrombus was seen on cat scan, finding a new pulmonary embolism and splenic infarct. the pump had low purge flows and pump saturation. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5632 | 1220648-2025-31429 | 23051801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051801&pc=OZD | 09/24/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 09/24/2023 | 09/15/2025 |  |  | 08/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380527 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. the complainant reported that there was a 'position in aorta' alarm, however, an echo had confirmed the impella was positioned across the valve. the ao placement signal alarm seemed to be erratic. the placement monitoring was disabled. the patient was successfully weaned from the impella. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. according to the clinical, impella position in aorta alarms were recorded, but echo imaging confirmed the position was across the valve. additionally, the ao placement signal seemed to be erratic. the decision was then made to disable placement monitoring. no product was returned for an investigation. data log analysis confirmed multiple false impella position in aorta alarms were recorded up until a placement signal not reliable alarm. analysis of the 5s parameters confirmed a decrease in signal not reliable and contrast, increase in lamp, and stability in light and gain. additionally, the console software version was confirmed to be v11.1. no other abnormalities were observed. the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5633 | 1220648-2025-46483 | 23051826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051826&pc=OZD | 08/16/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142849 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the nurse was changing the purge cassette and the aic would not go to the next screen. they swapped to another aic and there was no further issues. no patient injury was reported. | [5 empty] | Imc logs from the complaint date showed multiple occurrences of the "purge disc not detected" , which aligns with the reported complaint. on (b)(6) 2025, the purge cassette change was initiated and just after the purge cassette and disc were inserted, the "purge disc not detected" alarm was triggered and persisted even after the purge cassette was removed and reinserted multiple times. no alarms or issues related to the lcd display were observed in the imc data logs. device analysis: the console ic2675 was tested after arriving in fs. the issue of purge disc not detected was reproduced by connecting the purge cassette. it was confirmed that the purge flag was missing/broken. additionally, glucose residue was found around the purge assembly. the lcd display was also inspected and found to be functioning correctly and in accordance with the specifications. lastly, found a small crack on the bottom left side on the aic front assembly near lcd display. root cause: the root cause for the reported complaint "aic would not advance screens during cassette change" specifically the purge disc not detected alarm (event set #402) was due to the broken purge flag. no issues related to the display were found within the aic. no problem were found. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
5634 | 1220648-2025-46484 | 23051862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051862&pc=OZD | 08/16/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719304 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support the patient in non-st-segment elevation myocardial infarction. while on support, the impella cp was in the descending aorta due to being pulled back during cardiopulmonary resuscitation/cardiac massage. the physician attempted to reposition the device again, but was also considering extracorporeal membrane oxygenation. the device was explanted. the patient outcome indicated that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1: added serious injury. h6: omit e2403 and added e0602. added codes f1903, f19, b15. updated codes for investigation findings, and investigation conclusions. investigation summary complaint device not returned for analyzing. device in wrong position: the cause of the positioning issue was determined to be an unintentional use error as the pump was in descending aorta due to being pulled back during cardiopulmonary resuscitation (cpr)/cardiac massage. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review the pump passed all post sterile inspection checks.
5635 | 1220648-2025-46485 | 23051877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051877&pc=OZD | 08/17/2025 |  |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712975 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered hematuria and hemolysis during impella cp support, and ischemia post support. the patient was found to have bloody urine and hemolysis confirmed via labs. the pump's p-level was dropped to p6 and the pump was repositioned usingtransthoracic echocardiogram. later, the pump was explanted due to the hematuria. post explant, the nurse reported loss of pedal pulses to the patient's right limb. as a result, heparin drip was resumed and an intra-aortic balloon was placed in the left femoral artery. | [5 empty] | No product was returned for evaluation. the cause of the hemolysis could not be definitively determined as limited clinical details were provided and no product was returned for evaluation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5636 | 1220648-2025-46486 | 23051901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051901&pc=OZD | 08/17/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026700862 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support when intermittent placement signal not reliable alarms occurred. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product or data logs were returned for evaluation. the cause of the optical signal issue could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. the device passed all post-sterile inspection checks.
5637 | 1220648-2025-31430 | 23051917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23051917&pc=OZD | 01/02/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 01/02/2024 | 09/15/2025 |  |  | 09/30/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024397902 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed a 5.5 pump for support of a patient with history of cardiomyopathy and cgs. the pump supported the patient despite loss of placement signal and optical signal issues. the patient was weaned from the impella and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was returned for investigation. data logs shows that after 1 day of implant, the contrast and signal not reliable was spread and "impella position unknown" alarms were observed. after 13 days several "impella position in aorta" alarms were noted. console software version was 11.1.1 during time of use. there was no variation or impact to optical signal 5s parameters. patients condition was also deteriorating as per the clinical information provided. the placement monitoring was disabled and the alarms were not seen after that. therefore, as per the clinical information provided, the root cause of the optical signal issue was patient condition related to false alarm. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5638 | 1220648-2025-31431 | 23052034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23052034&pc=OZD | 12/13/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 12/13/2023 | 09/15/2025 |  |  | 09/30/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397138 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. while on support there were optical signal alarms. the patient was on support for over 64 days when the pump was explanted and patient transitioned to a left ventricular assist device. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. according to the clinical, on the 26th day of impella cp support the first placement signal not reliable alarm was recorded. no attempts were made to resolve the issue. no product was returned for a visual inspection. data log analysis confirmed that the 5s parameters aligned with the sensor wear/drift root cause. the most likely cause of the optical signal issue was sensor wear. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5639 | 1220648-2025-46487 | 23052419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23052419&pc=OZD | 08/17/2025 |  |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643616 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. hematoma developed around impella access site on arrival to icu. attending notified and manual pressure applied for 30 minutes with reduction of hematoma. medication was administered md to inject lido/epi around site for potential collateral vessel bleeding. withdrew care due to lack of neuro exam. impella turned off and patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma/ major bleed: the cause of the bleeding is determined to be use issue because the hematoma was cause due to patient movement. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)".
5640 | 1220648-2025-46488 | 23052454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23052454&pc=OZD | 08/17/2025 |  |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026706996 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 83 KG |  |  | Injury | No | Yes | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a pump stop during impella cp support. during support, there was a sudden pump stop due to high motor current. the pump was replaced with a new impella cp to resolve the issue. the old pump was discarded. no patient harm was reported. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Complaint coding was updated to better describe the reported event. consequently, h6 health effect: clinical/impact and medical device problem coding were repopulated/updated, corrected accordingly. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. temporary pump stop: at approximately 11 days of support, a sudden pump stop was reported due to high motor current. the pump was restarted, but with saturated flow and mc out of range. data log review shows a motor current high pump stop about 1 minute after a motor current spike and shift up in motor current. purge pressure was trending downward near 300-400 mmhg with purge flow reaching 30 ml/hr and then pp began rising shortly before the motor current spike. the cause of the pump stop was not determined since data logs are inconclusive without returned product. high or saturated pump flow: at approximately 11 days of support, a sudden pump stop was reported due to high motor current. the pump was restarted, but with saturated flow and motor current out of range. data log review shows a motor current spike, shift up in motor current, and saturated flow followed by a pump stop about 1 minute later. purge pressure was trending downward near 300-400 mmhg with purge flow reaching 30 ml/hr and then pp began rising shortly before the mc spike. the cause of the saturated flow was not determined since data logs are inconclusive without returned product. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5641 | 1220648-2025-46489 | 23052764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23052764&pc=OZD | 08/17/2025 |  |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686618 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 100 KG |  | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during support of the impella 5.5 that the patient tripped over the white cable when getting out of bed. impella would not restart. the pump was explanted and the patient was placed on intra-aortic balloon pump. | [5 empty] | D6b explant date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6. codes have been updated per investigation findings. the investigation was completed. there was no return of product. pump stop: the pump was not returned for investigation. data log shows the pump stopped on 17-august-2025 (day 21) with pump stop alarms (115 and 119) and a motor current spike reaching >30000 ma. the pump could not be restarted. the pump stop issue was most likely due to an electrical open line caused by an unintended use error, based on the data log review and the provided clinical details. the incident occurred on 17-august-2025, when the patient tripped over the white cable while getting out of bed, which aligns with the timing of the pump stop and motor current spike. a device history review was completed and passed all post sterile inspection checks.
5642 | 1220648-2025-46490 | 23052787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23052787&pc=OZD | 08/20/2025 |  |  | 09/15/2025 |  |  | 08/20/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688995 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient underwent ectopy and suctioning during impella 5.5 support. after a few days of support, the patient had some suction alarms and ectopy over night when laying flat. a bedside echo was done and the pump position was confirmed. the patient also experienced ventricular tachycardia and was treated with milrinone. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. after reviewing logs suction alarms was observed. the cause of the pump suction was not determined due to insufficient clinical details and per log review. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5643 | 1220648-2025-46491 | 23052981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23052981&pc=OZD | 08/17/2025 | 10/14/2025 |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026677223 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 92 KG |  | White | Malfunction | No | Yes | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp and va-ecmo support. on day five of impella support the impella stopped, and motor current was high. the impella was able to restart. the next day the physician decided to remove the cp and continued with va-ecmo. the patient is stable. nopatient harm reported sue to the issue. | [5 empty] | Patient's ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5644 | 1220648-2025-46492 | 23053214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053214&pc=OZD | 08/26/2025 |  |  | 09/15/2025 |  |  | 08/26/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1620355 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 83 KG | Non Hispanic |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support with the impella cp, it was reported that the silver latch on the automated impella controller (aic) where the impella pump connects to the aic was accidentally broken during transport of the patient. the same day, the patient was escalated to an impella 5.5 for increased mechanical circulatory support, and a new aic was used. the device was subsequently weaned and explanted successfully, and the patient outcome was survived. | [5 empty] | H.10 related report number was not added to the initial report that was submitted and was added. | A.6 is unknown. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5645 | 1220648-2025-31432 | 23053241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053241&pc=OZD | 10/24/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 10/24/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558974 |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. a ¿placement signal not reliable¿ alarm occurred. the sensor was working intermittently and normal waveforms and metrics were shown followed by unreliable artifact. the pump remained on until the pump was successfully weaned and explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product not returned. data logs show pump was used on 3 consoles. the 5s parameters are different on each console with no clear optical failures patterns noted. signal not reliable is variable on all 3 consoles with the most widespread on the last console with the noisiest contrast as well. on all 3 consoles there are suction alarms and a few placement signal spikes. the root cause of the optical signal issue was not determined since no product was returned. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5646 | 1220648-2025-46493 | 23053318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053318&pc=OZD | 08/26/2025 |  |  | 09/15/2025 |  |  | 08/26/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1206660 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 107 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient awaiting surgery for a durable left ventricular assist device (lvad) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during the set up for the impella 5.5, the medical team encountered automated impella controller (aic) purge disc issues. the purge disc was inserted during pump set-up, but the purge disc was not detected. a new aic was used, and the same cassette was successfully used in the new aic. there was no reported harm to the patient as a result of the aic issue. the patient remained on impella 5.5 support with the new aic with no further complications until the patient was successfully weaned and explanted during surgery for a lvad. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5647 | 1220648-2025-31433 | 23053349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053349&pc=OZD | 07/30/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 07/30/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420348 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the placement signal not reliable alarm occurred on the automated impella controller. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned and the failure modes were not reproduced during testing. the data log at the time of the placement signal not reliable alarms confirmed that during the times of the alarms, the cause of the intermittent unreliable placement signal and oscillating contrast was most likely due to a malfunctional optical bench lamp. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5648 | 1220648-2025-31434 | 23053473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053473&pc=OZD | 06/10/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 06/10/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1481995 |  | 0042-0040 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. on the console the placement signal was lost. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned and the failure modes were not reproduced during testing. the data log at the time of the placement signal not reliable alarms confirmed that during the times of the alarms. although the unreliable placement signal was not reproduced, the root cause was likely due to a malfunction of the optical lamp, causing intermittent interruptions in light output. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5649 | 1220648-2025-46494 | 23053482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053482&pc=OZD | 08/26/2025 |  |  | 09/15/2025 |  |  | 08/26/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1189530 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered an automated impella controller (aic) issue concerning the pump unable to start. there was no patient involvement as the aic was being used during an intensive care unit in-service training and attempted to start a demo impella cp pump. the error message "impella controller failure" was displayed on the aic and the demo pump would not start. the team then attempted to start the same demo impella cp with a different aic and the pump started successfully without issue. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5650 | 1220648-2025-31435 | 23053489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053489&pc=OZD | 05/19/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 05/19/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1097400 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the console waveforms quickly separated and suction was noted. the pump was successfully weaned and explanted. no patient harm was reported, and the console was not replaced. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console displayed a suction alarm, which persisted throughout the rest of the support period. this confirms the reported failure mode. this console was tested after arriving at fs. the suction alarm could not be reproduced when running the pump simulator. the root cause of the suction alarm could not be determined due to the inability to reproduce it. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5651 | 1220648-2025-31436 | 23053670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053670&pc=OZD | 06/26/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 06/26/2024 | 09/15/2025 |  |  | 12/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450049 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support and they had a hematoma at the access site. staff monitored and support continued. additionally, there were placement signal issues reported. echocardiogram was performed to check placement of pump. motor current remained pulsatile and flows were appropriate for level of support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: on the 14th day of 5.5 support, a "placement signal not reliable" alarm occurred, and they lost the placement signal. the team periodically checked position with echo from that point, and the pump remained in good position. the pump was not removed due to the optical signal issue. upon review of the data logs, it is apparent that the console is the potential cause of the issue. please refer to 24-33757-2 for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5652 | 1220648-2025-31437 | 23053752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053752&pc=OZD | 07/24/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 07/24/2025 | 09/15/2025 |  |  | 02/28/2026 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025473893 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. the complainant reported there was a placement signal not reliable alarm. the audio was disabled. the patient was successfully weaned from the pump and the device explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no pump returned for investigation. data logs show that 254 hours into support on 7/11 there is a spike in ps to 300 mmhg with placement signal not reliable alarms. signal not reliable drops to 0. optical parameters remain like this for 70 hours and then resolves for the rest of the case. this is indicative of an automated impella controller (aic) failure. upon review of the data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5653 | 1220648-2025-46495 | 23053768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053768&pc=OZD | 08/17/2025 | 11/10/2025 |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716873 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 76 KG | Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. in icu multiple suctions alarms overnight occurred and low flows on impella. transthoracic echocardiogram was done the following day and impella was repositioned. the flows were back up to p-8. the patient also developed gastrointestinal bleed, and one unit of blood was administered. it was further reported that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. clinical symptoms (hemorrhage/bleeding): the cause of the gi bleed was not determined due to insufficient clinical details. low or blocked pump flow and pump suction: the data log shows that the pump was unable to achieve good flow on a higher p-level without triggering suction alarms after 5:30 am on (b)(6) 2025. there was one instance of a "pump position in ventricle" alarm during the low pump flow event. the issue was resolved and pump set to p-8 on 5:30 am. overall downward trend in the placement signal was observed during the time of the low pump flow event. however, no motor current spikes were reported during the case run. according to the clinical details, the physician repositioned the device, successfully resolving the low pump flow and suction issues; however, there is no conclusive evidence in the data log that the pump was in the wrong position during the low pump flow event. the cause of the low pump flow and pump suction issues was not determined, as the product was not returned and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5654 | 1220648-2025-46497 | 23053770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053770&pc=OZD | 08/17/2025 |  |  | 09/15/2025 |  |  | 08/17/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 90 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp in aorta alarm occurred while supporting a patient. pigtail was under the papillary muscles. the impella was repositioned. the patient remained in stable condition. the patient continued with support and was weaned and explanted as planned. | [5 empty] | The investigation of the reported failure mode has been completed. device in wrong position: the data log shows several "pump position in aorta" alarms with spiked and dampened motor current waveforms, which remained consistent until the end of the case. the pump was explanted 6 hours later. the cause of the pump positioning issue was not determined as the product was not returned for investigation, and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5655 | 1220648-2025-46498 | 23053773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053773&pc=OZD | 08/18/2025 |  |  | 09/15/2025 |  |  | 08/18/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707348 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5. after four days of support the patient suffered a hemorrhagic cerebrovascular accident. the patient was successfully weaned from the pump and survived. | [5 empty] | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. stroke : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5656 | 1220648-2025-46499 | 23053781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23053781&pc=OZD | 08/21/2025 |  |  | 09/15/2025 |  |  | 08/21/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1114572 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 124 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for over two months awaiting transplant. the support was ongoing when the automated impella controller (aic) alarmed. the aic was replaced to resolve the issue. no patient harm was reported. | The complainant reported the patient was implanted with the impella 5.5 device for mechanical circulatory support (mcs) and experienced complications. the patient would ultimately expire. the patient was initially admitted with advanced heart failure and was subsequently transferred to the unit status post intra-aortic balloon pump for cardiogenic shock. the patient was brought to operating room for impella 5.5 placement and would continue with orthotopic heart transplantation (oht) work-up. the impella 5.5 successfully placed, and the patient returned to the unit to recover. after approximately 2.5 months of support, a sudden controller error alarm occurred, and an aic-to-aic transfer was completed. the patient continued on support. approximately four days later, intermittent placement signal not reliable (psnr) alarm occurred (with impella 5.5 and automated impella controller-aic #2). the intermittent psnr issue self-resolved. additionally noted this day was the purge flow rate issue occurred and was closely monitored. the patient continued to wait for oht but would still have intermittent psnr. instructions were given to disable alarm. two occurrences again (10 days and 15 days later), purge flow issues were encountered. tissue plasminogen activator (tpa) was discussed. however, it is unknown if tpa was added to the purge system. support continued however four days after the latest purge flow rate issue. however, the patient expired on support while waiting for heart & kidney transplant, despite being elevated to status 1. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Correction: b5. event description and b7. medical history/preexisting condition were corrected to provide all relevant details regarding the complaint. medical safety review. medical safety reviewed the event. the event describes an automated impella controller (aic) failure resulting in aic-to- back up aic transfer. thereafter on the second aic and with the impella 5.5 device, intermittent placement signal not reliable (psnr) was encountered. additionally involving the impella 5.5 device, there was high purge pressure and tissue plasminogen activator (tpa) was possibly added to the purge system. the patient would eventually expire on support while waiting for heart and kidney transplant, despite being elevated to status 1. the first and second aics had a device malfunction and did not contribute to the demise outcome. although the first aic failure resulted in temporary loss of support due to aic-to-aic exchange, this was a planned switch that occurred approximately 21 days prior to the patient's demise. regarding the intermittent psnr, positioning was confirmed and support continued. based on the information at hand, the patient expired for reason unrelated to the aics despite the product malfunction. for the impella 5.5 device, high pressure at first self-resolved but noted to have continued on three occasions. tissue plasminogen activator (tpa) was possibly added to the purge. however, there was no report or indication of inadequate support. mechanical circulatory support (mcs) support continued. in this case, the patient's underlying condition likely contributed most to an outcome of death. the patient required both heart and kidney transplant which did not materialized although at status 1. conservatively, the death will be reported against the impella 5.5 device as this device was in use at time of expiration. the impella 5.5 cannot be ruled out but highly unlikely due patient condition. despite the purge concerns the pump remained supporting the patient. related medical device reports: this impella aic device is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the second aic device and the impella 5.5, which is associated with the demise outcome.
5657 | 1220648-2025-46500 | 23057067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057067&pc=OZD | 08/27/2025 |  |  | 09/15/2025 |  |  | 08/27/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1097261 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient presented to the hospital with an anterior in st elevation myocardial infarction after feeling uncomfortable for several days following an unrelated medical procedure. the patient was found to have a flush occlusion of their left anterior descending artery (lad) and ejection fraction of 20%. the physician then elected to place the impella cp to treat the lad. it was reported that while the patient was in the operating room there was an alarm for air in purge and low purge pressure. it was noted that the medical team attempted to de-air the purge without success. the purge cassette was changed and the low purge pressure and air in purge alarms were resolved. it was also reported that on (b)(6) 2025, the patient experienced an episode of atrial fibrillation with no information on interventions that were performed to resolve the issues. additionally, the same day the patient had bleeding at the impella access site and as a result of the bleeding heparin was withheld and manual compression was held. the following day, the bleeding complication persisted, and the patient was transfused with two units of red cells and a purse string closure was placed at the impella access site. the patient was then escalated later that day to an impella 5.5 for increased hemodynamic support. the patient experienced a hematoma at the axillary graph anastomosis site that was tunneled with no further complications and bleeding at the impella access site that was sandbagged to resolve. additionally, there were false alarms for impella in the aorta when the patient coughed. the patient¿s care was withdrawn, and the patient expired on impella 5.5 support. this complaint involves one impella cp, one purge cassette used during impella cp support, one impella 5.5, and one automated impella controller which alarmed during impella 5.5 support. purge cassette lot number is unknown. impella cp with serial number (b)(6). impella 5.5 with serial number (b)(6). automated impella controller with serial number (b)(6). this report specifically addresses automated impella controller with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. refer to section h.10 for the related report numbers. | [5 empty] | A.4 - a.5 is unknown. based on the information available, the impella cp cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety review was completed and noted the following: the event may have cascaded from the impella cp. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the outcome of the bleed for the patient requiring two units of prbcs and is a serious injury (as well as the arrhythmic event) that should be reported. the patient was escalated to an impella 5.5 and experienced a hematoma at the access site that was sandbagged. the patient would eventually expire on the impella 5.5. regarding the automated impella controller (aic) sensing issue there was no report or indication of aic-to-aic exchange contributing to a disruption in or lack of support. the patient expired for reasons unrelated to the aic (this type of event is a product malfunction). impella 5.5 pump positioning was confirmed, and support had continued. there is no reason to believe that this hematoma is associated with the demise outcome given the information at hand. however, the impella 5.5 device was present at the time of the demise outcome and dissociation cannot be made. however, the patient's underlying condition (cardiogenic shock, inability to wean from ecmo and not a transplant candidate) most likely contributed to the outcome and the care was withdrawn which appears to have led to the demise outcome. medical safety officer reviewed the event and noted the same; the impella 5.5 device cannot be dissociated from the outcome and impella cp is a serious injury and aic is a product malfunction type of reportable events.
5658 | 1220648-2025-46501 | 23057097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057097&pc=OZD | 08/19/2025 |  |  | 09/15/2025 |  |  | 08/20/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678078 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient presented to the hospital with anterior in st elevation myocardial infarction after feeling uncomfortable for several days following an unrelated medical procedure. the patient was found to have a flush occlusion of their left anterior descending artery (lad) and ejection fraction of 20%. the physician then elected to place the impella cp to treat the lad. it was reported that while the patient was in the operating room, there was an alarm for air in purge and low purge pressure. it was noted that the medical team attempted to de-air the purge without success. the purge cassette was changed and the low purge pressure and air in purge alarms were resolved. it was also reported that on (b)(6) 2025, the patient experienced an episode of atrial fibrillation with no information on interventions that were performed to resolve the issues. additionally, the same day the patient had bleeding at the impella access site and as a result of the bleeding heparin was withheld and manual compression was held. the following day, the bleeding complication persisted, and the patient was transfused with two units of red cells and a purse string closure was placed at the impella access site. the patient was then escalated later that day to an impella 5.5 for increased hemodynamic support. the patient experienced a hematoma at the axillary graph anastomosis site that was tunneled and treated with vistaseal with no further complications and bleeding at the impella access site that was sandbagged to resolve. additionally, there were false alarms for impella in the aorta when the patient coughed. the patient¿s care was withdrawn, and the patient expired on impella 5.5 support. this complaint involves one impella cp, one purge cassette used during impella cp support, one impella 5.5, and one automated impella controller which alarmed during impella 5.5 support. purge cassette lot number is unknown. impella cp with serial number (b)(6), impella 5.5 with serial number (b)(6), automated impella controller with serial number (b)(6). this mdr specifically addresses impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. refer to section h.10 for the related report numbers. | [5 empty] | A.4 - a.5 is unknown. based on the information available, the impella cp cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Medical safety review was completed and noted the following: the event may have cascaded from the impella cp. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the outcome of the bleed for the patient requiring two units of prbcs and is a serious injury (as well as the arrhythmic event) that should be reported. the patient was escalated to an impella 5.5 and experienced a hematoma at the access site that was sandbagged. the patient would eventually expire on the impella 5.5. regarding the automated impella controller (aic) sensing issue there was no report or indication of aic-to-aic exchange contributing to a disruption in or lack of support. the patient expired for reasons unrelated to the aic (this type of event is a product malfunction). impella 5.5 pump positioning was confirmed, and support had continued. there is no reason to believe that this hematoma is associated with the demise outcome given the information at hand. however, the impella 5.5 device was present at the time of the demise outcome and dissociation cannot be made. however, the patient's underlying condition (cardiogenic shock, inability to wean from ecmo and not a transplant candidate) most likely contributed to the outcome and the care was withdrawn which appears to have led to the demise outcome. medical safety officer reviewed the event and noted the same; the impella 5.5 device cannot be dissociated from the outcome and impella cp is a serious injury and aic is a product malfunction type of reportable events. change in reportability: after review of the case by the medical safety officer, it was determined the impella 5.5 is a death type of reportable event. sections b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect - clinical and impact codes) have been updated, corrected accordingly. the event story also involves two other devices. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller and impella cp.
5659 | 1220648-2025-31439 | 23057291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057291&pc=OZD | 03/30/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 05/08/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1616687 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal stopped functioning. review of the device log showed low signal not reliable alarms. this finding indicates a potential issue with the automated impella controller (aic) as the source of the alarm. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs confirmed the issue. the console (b)(6) was returned to field service for further investigation. the device issues could not be reproduced. the aic was run with a known good pump and purge cassette without any issue observed. the root cause for the placement signal not reliable alarm was due to the intermittent failure of the optical bench. additional issue was found during investigation aic - loss of opm data the root cause of the controller error alarm was due to a firmware issue (optical bench fw anomaly). a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5660 | 1220648-2025-31438 | 23057344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057344&pc=OZD | 04/03/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 05/05/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1481995 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced a loss of placement signal. review of the device log later showed an oscillating contrast pattern that triggered the ¿placement signal not reliable¿ alarms. this finding indicates a potential issue with the automated impella controller (aic) as the source of the alarm. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs show a corresponding reduction in signal not reliable measurements and confirms the reported failure mode. the console was returned. the pump connector optical fiber was inspected and did not have any damage or contamination that would cause an unreliable placement signal. the 5s parameters of the optical fiber were checked and all values were within the specification. although the unreliable placement signal was not reproduced, the root cause was likely due to a malfunction of the optical lamp, causing intermittent interruptions in light output. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5661 | 1220648-2025-31440 | 23057427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057427&pc=OZD | 04/05/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 05/19/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, a placement signal not reliable alarm. the investigating engineer attributed the issue to the automated impella controller (aic). no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs were retrieved and showed oscillating contrast that resolved during the case. the root cause of the oscillating contrast could not be determined as console was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5662 | 1220648-2025-31441 | 23057768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057768&pc=OZD | 04/07/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 05/08/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1197229 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, an intermittent ¿placement signal unreliable¿ alarm occurred. device positioning was verified by echocardiography, and support continued without interruption. a later review of the logs revealed an oscillating contrast pattern, suggestive of a potential automated impella controller (aic) bulb failure. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs confirms that all signals received from the optical bench were confirmed. this console was tested and the bulb power and the optical bench parameters are within specifications, the analog output from the optical bench was distorted. the root cause of the alarm was the malfunction of the optical bench. additional failure mode was found with the console for the controller error and loss of placement signal is due to an optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5663 | 1220648-2025-31442 | 23057909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057909&pc=PYX | 03/04/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 03/04/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella rp flex support and, during support, there were placement signal issues. audio was disabled. no changes in motor current or flows occurred. the investigating engineer found that this is likely an issue with the automated impella controller. the pump was eventually weaned and explanted. no known patient harm or injury has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned and evaluated. data analysis: ¿ the case associated with the reported complaint was initiated on mar 01, 2025, with the pump impella rpflex (rpij) sn: (b)(6). ¿ on the complaint date, mar 03, 2025, the console displayed ¿placement signal not reliable (rp dual sensor, optical bench invalid signal condition) ¿ alarm,¿ (psnr) event #(b)(4) (imc_ic6395.pdf). this confirms the reported failure mode. ¿ the rt log shows that the contrast was oscillating between -3276 and 3274 (rt_log_1055_ic6395.png). ¿ the lt log confirms the psnr was never resolved (lt_log_ic6395_contrast.png & lt_log_ic6395_snr.png). ¿ after disconnecting the pump, the log recorded abnormal optical parameters, especially light: 0.45v and snr: 2.9. device analysis: this console was tested after arriving at field service. upon running the optical test pump for 12 hours, unable to reproduce the reported failure mode psnr alarm and the oscillating contrast. no issue was found with the fringe pattern. the optical parameters (contrast, gain, snr, and bulb power) were within specifications per the pm procedure (0042-7309). found the front panel optical connector contaminated, unable to remove the debris upon cleaning (most likely unrelated to the reported failure mode and clinical procedure). during testing, it was found that the rotary knob had a mechanical malfunction. upon a normal push of the rotary knob, it was unable to select the menu options intermittently, it must be pushed hard to select the options if it doesn¿t respond to a normal push (unrelated to the reported failure mode and the clinical procedure). root cause: the root cause of the placement signal not reliable alarm was the oscillating contrast. the cause of the oscillating contrast was not determined due to the inability to reproduce, it is most likely related to either the lamp or the optical bench malfunctioning. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5664 | 1220648-2025-31443 | 23057938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057938&pc=OZD | 03/06/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 03/27/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 120405 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 55 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that impella cp was inserted in normal fashion while connected to an automated impella controller. flows were ramping up, and a sensor unreliable alarm appeared. flows were assessed and were appropriate, and motor current was also appropriate. the alarm was muted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data analysis: on 3/6, after pump is started to run on the automated impella controller (aic), the user acquired a placement signal not reliable alarm as alluded to in the complaint. additionally there were signal-to-noise ratio (snr) out of range logs immediately beforehand. on a previous case on ic1661 running pump 565604, there was similar behavior by the console. the lt logs from pump show very low snr for the duration of the case and sporadic placement spiking coinciding with the low snr decreasing more. similarly, pump had the same low snr with spiking placement signal across the entire case which is what triggered the placement signal not reliable alarm. the rt logs confirmed the optical signals of distinctly low snr throughout with intermittent spiking placement signal. device analysis: the console was returned and a 5s test was performed and the snr was found to be 1149 well below the minimum threshold of passing. root cause: the cause of the optical signal abnormalities was a defective optical bench with low snr. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5665 | 1220648-2025-31444 | 23057963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057963&pc=OZD | 03/13/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 03/13/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089161 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had several instances when the placement signal was lost. the issue occurred with the same aic on different devices. no further intervention was mentioned, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data analysis: the case started on (b)(6) 2025. about 2.5 hours after starting the pump, a suction alarm went off, and there was an alarm for an unknown impella position. the case continued until on (b)(6) 2025. lt and rt logs did not indicate any irregularities with the placement signal or other optical signals. further investigating the previous case with the pump sn: (b)(6) also did not reveal any issues with the placement signal. additionally, while reviewing the logs, found an unexpected controller reboot on (b)(6) 2024 after a day of pump operation, following a partial reset of the watchdog and a missed acknowledgment by process sau_motor_1. after the system rebooted, the watchdog started again and the pump resumed at the previous p-level. the caused of the unexpected controller reboot was most likely due to a software anomaly. this issue is unrelated to the complaint. device analysis: the console was returned to field service for further investigation. a visual inspection of the internal circuitry of the console did not reveal any issues. obtained the analog output from the optical bench without any noise or distortion with pump and without pump connected as well as 5s parameter values were within the specification as per the pm procedure. see attached 5sparameter.jpg. the console was tested with a known good pump and purge system, and no anomalies were observed during testing. root cause: the root cause for the reported lost of placement signal not related to any problems with the aic. no issues were found within the aic. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5666 | 1220648-2025-31445 | 23057996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23057996&pc=OZD | 05/04/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 05/04/2023 | 09/15/2025 |  |  | 11/30/2024 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023281072 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during therapy, the placement signal disappeared unexpectedly, triggering a placement signal unreliable alarm. troubleshooting confirmed that positioning was correct, and all cables were properly seated without visible kinks. as the issue persisted, the user was advised to silence the alarm and monitor motor current during support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. there is no data available on constellation for the second half of the case during which the complaint occurred. the optical signal issue occurred on day 63 of support. the 5s parameters were of the returned pump were tested and were within specification and the placement signal not reliable alarm reproduced. light came out at the sensor in the optical continuity test. the pump was pressure tested in the uson and the pump did not hold the correct pressure. the sensor was removed and there was etching of the silicone observed. the cause of the optical signal issue was sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5667 | 1220648-2025-31446 | 23058106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058106&pc=OZD | 04/27/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 04/27/2023 | 09/15/2025 |  |  | 12/31/2024 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024301345 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. data logs were returned and investigated and there were no motor current spikes observed in the data logs. no major abnormalities were found. per the given clinical information, the root cause of the optical signal issue was most likely patient condition related. all pertinent information available to abiomed has been submitted at this time.
5668 | 1220648-2025-31447 | 23058118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058118&pc=OZD | 07/24/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 07/24/2025 | 09/15/2025 |  |  | 01/31/2026 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025452605 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. the complainant in japan had a 5.5 pump support ongoing for a patient admitted in with a left ventricle rupture. the patient was escalated to the 5.5 from intra-aortic balloon pump and extracorporeal membrane oxygenation (ecmo). after 49 days of successful support, the optical signal was lost. the patient was successfully weaned from the pump and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. on the 51st day of support, the "placement signal not reliable (psnr) alarm triggered." data logs were investigated, and on the reported date, the psnr alarm was confirmed. all the optical parameters went out of specification for the remainder of the case. returned product was investigated and it passed all optical sensor testing: 5s parameters were all within specification, the pump passed the laser continuity test, and when the pump was ran in a bucket in the lab, all the optical parameters remained normal. there was no problem found with the pump regarding the optical signal issues. the data logs displayed common signatures for a potential automated impella controller (aic) issue. please refer to 24-34348-2 for further investigation into the console. regarding the concern about using meylon in the purge fluid: visual inspection found that there was some biomaterial or foreign material in the cannula near the impeller on returned product. when that object was removed, the pump was able to run smoothly at p9. this, however, is unrelated to the failure mode as even before the object was removed, the pump's optical parameters were all measured to be within specification in the lab. additionally, material in the cannula should not effect or damage the optical sensor. finally, the pump did not show any evidence of biomaterial ingestion in the logs, so this object was likely picked up during explant. as for using meylon in the purge, it is not one of the abiomed recommended fluids for the purge, so it is best practice to use approved fluids per the ifu. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5669 | 1220648-2025-31448 | 23058136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058136&pc=OZD | 05/07/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 05/07/2023 | 09/15/2025 |  |  | 12/31/2024 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307956 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during therapy, the patient experienced frequent ¿impella in aorta¿ alarms. impella positioning was verified under echocardiography and confirmed to be appropriate; therefore, placement signal monitoring was disabled. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were returned and several impella position in aorta were observed. there was no variation or impact to optical signal 5s parameters. the root cause of the optical signal issue was most likely patient condition related based on the data log review. all pertinent information available to abiomed has been submitted at this time.
5670 | 1220648-2025-31449 | 23058167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058167&pc=OZD | 07/12/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 07/12/2023 | 09/15/2025 |  |  | 02/28/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310324 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs shows a 'placement signal not reliable' alarm and hard failure of all 5s optical sensor parameters right after case start. based on the pattern of the 5s parameters, the root cause of the optical signal issue is most likely due to a torqued fiber line. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5671 | 1220648-2025-31450 | 23058335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058335&pc=OZD | 07/24/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 07/24/2025 | 09/15/2025 |  |  | 02/28/2026 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467989 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient transitioning to a heart transplant. after just under a month of support the optical signal was lost. despite the placement signal issue, the pump remained on, to support the patient. the patient was successfully bridged from impella 5.5 to the heart transplant. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for analysis. the data logs confirm placement signal not reliable alarm. the data logs show a spike in placement signal and a minor increase in variability of contrast. there was a decrease in lamp and signal not reliable during this. this occurrence lasts for about 3 days until all 5s parameters recover. there is no other "placement signal not reliable" alarm for the rest of the case. the root cause of the optical signal issue was not determined as no product was returned for analysis and limited clinical information was provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5672 | 1220648-2025-31451 | 23058336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058336&pc=OZD | 08/19/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 08/19/2023 | 09/15/2025 |  |  | 04/30/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024339436 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, sensor damage was noted, and the pump exhibited placement signal issues but continued to provide adequate support. impella position was confirmed with echocardiography, and the alarm was disabled. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no pump was returned for investigation. data logs were returned and investigated. all optical parameters were normal/stable. per the data logs and clinical information provided, the root cause of the optical signal issue was patient condition related. all pertinent information available to abiomed has been submitted at this time.
5673 | 1220648-2025-31452 | 23058374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058374&pc=OZD | 09/13/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 09/13/2023 | 09/15/2025 |  |  | 04/30/2024 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336966 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issues cannot be determined due to insufficient information and lack of product or data logs returned. all pertinent information available to abiomed has been submitted at this time.
5674 | 1220648-2025-31453 | 23058411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058411&pc=OZD | 09/18/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 09/18/2023 | 09/15/2025 |  |  | 05/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345503 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal not reliable alarms. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product and data logs were not returned for investigation and there are minimal clinical details. the root cause of the optical signal issue was not determined since product and data logs were not returned. all pertinent information available to abiomed has been submitted at this time.
5675 | 1220648-2025-31454 | 23058532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058532&pc=OZD | 02/19/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 02/19/2024 | 09/15/2025 |  |  | 09/30/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403635 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited fiber optic issues, and the placement signal did not display any metrics. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. the logs were checked and there were no issues found. the root cause of the optical signal issue is most likely a damaged fiber line during transition from the first console to the second. the location of the damage is unknown as product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5676 | 1220648-2025-31455 | 23058544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058544&pc=OZD | 07/24/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 07/24/2025 | 09/15/2025 |  |  | 11/30/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431470 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported a 16 year old patient received escalated care from an impella cp to an impella 5.5 pump while awaiting a heart transplant. the 5.5 lost placement signal after 4 days of support. the patient was successfully weaned from the 5.5 and received the heart transplant. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for investigation. after 85 hours from the start of the case, the data logs show an 11-hour gradual decrease trend in the placement signal, followed by a 'placement signal not reliable' alarm and a spike to 3000, which persisted until the end of the case run. there were no major changes noted in any of the 5s optical signal parameters. the returned pump did not have any physical abnormalities and passed the laser continuity test for the optical signal. biomaterial was observed in the outflow case, and it was released during the test run. the pump was running as expected during the test in a bucket of water, but the 'placement signal not reliable' alarm was reproduced. according to the clinical data, the placement signal was not matching the arterial line reading and was lost during the run. the cause of the optical signal issue was most likely sensor silicone wear, based on returned product and data log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5677 | 1220648-2025-31456 | 23058592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058592&pc=OZD | 04/08/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 04/08/2024 | 09/15/2025 |  |  | 12/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025445427 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the aortic placement signal was lost and left ventricle waveforms were lost but it resolved automatically without any troubleshooting. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was not returned for evaluation. data logs were reviewed, and "impella position unknown" alarms are continuous throughout case. optical parameters were not affected with positioning alarms. data logs also indicate "controller error" alarms occur twice during case. no problem was found with the pump, a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5678 | 1220648-2025-31457 | 23058626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058626&pc=OZD | 06/11/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 06/11/2024 | 09/15/2025 |  |  | 12/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450021 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the optical signal was lost. good pump position was verified during the alarm event indicating impella position in the aorta. the physician disabled placement signal monitoring. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs were not provided for investigation. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5679 | 1220648-2025-31458 | 23058660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058660&pc=OZD | 07/03/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 07/03/2024 | 09/15/2025 |  |  | 02/28/2026 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467032 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal generated unreliable alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for an investigation. data log analysis confirmed that after 26 days of operation, the optical placement signal suddenly became unreliable. the issue spontaneously self-resolved, but promptly re-occurred, only to self-resolve again. all 5s parameters other than lamp had a decreasing trend. upon review of the data logs, no problem was found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5680 | 1220648-2025-31459 | 23058748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058748&pc=OZD | 07/13/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 07/13/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492236 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal generated unreliable alarms. support was continued following the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no data logs available. pump returned but cut at the catheter and missing pump body. the cause of the optical signal issue was not determined since sufficient product and data logs were not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5681 | 1220648-2025-31460 | 23058812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058812&pc=OZD | 07/22/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 07/22/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507367 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump lost optical signal. support was continued following the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product was not returned for investigation since the pump is still running and supporting the patient. data logs were reviewed at the time of failure and it was observed that the placement signal was seen slightly going in a downward trend. there were numerous "impella position unknown" alarms observed through the time the pump was active. "impella position in aorta" alarms were also observed. with consistent "impella in aorta" alarm and since echo confirmed pump was in correct position, the alarms were false and most likely due to the patients condition. since the pump has not been returned for investigation, as per the data log review, the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5682 | 1220648-2025-31461 | 23058849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23058849&pc=OZD | 11/07/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 11/07/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503423 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported a patient was implanted with an impella 5.5 pump while awaiting a heart transplant. the 5.5 lost placement signal after 17 days of support; however, the pump remained on to support the patient. the patient was successfully weaned from the 5.5 and received the heart transplant. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. as per the clinical details, the doctor reported a false position in the aorta alarm and disabled the placement signal monitoring on 11/7. also placement signal was corelating with a-line. the pump continued to be used. the data log shows several random temporary drops in signal not reliable to 0 and contrast below 15 during the reported placement signal discrepancy, with no correlation to motor current. this trend in optical signal parameters caused the spike in placement signal to 3000 and the psnr (placement signal not reliable) alarm during the case run. additionally, there were several pump in aorta alarms followed by the disabled positioning motor. the false position in the aorta alarm or placement signal discrepancy related to the pump could not be confirmed with the trend observed in the data log. the cause of the optical signal issue was not determined since product was not returned sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5683 | 1220648-2025-31462 | 23059042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059042&pc=OZD | 12/19/2024 | 10/24/2025 |  | 09/15/2025 |  |  | 12/19/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580200 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. while on support, the patient had inconsistent changes in waveforms not relating to monitor ao pressure. diastolic flows significantly decreased intermittently. the impella position was unknown, and it was recommended to get an echo when issue arrives to see where pump position is. the impella 5.5 was a successful bridge to a left ventricular assist device for the patient. | It was reported by the customer that the device exhibited sensor placement issues. |  |  |  |  | G3: date corrected. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. as per the data logs, the signal not reliable was a little spread corresponding to the impella position unknown alarms but the other parameters were stable. there were numerous of "impella position unknown" alarms but no placement signal not reliable alarms were observed. position was mentioned to be good in the clinical. no hard failures of the pump were identified on the data logs. the root cause of the optical issue was not determined since product was not returned and data logs were inconclusive. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D6b: corrected explant date.
5684 | 1220648-2025-31463 | 23059065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059065&pc=PYX | 01/22/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 01/22/2024 | 09/15/2025 |  |  | 10/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024428124 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the complainant reported that the placement signal was lost during support. as the motor current and flow rates remained stable, support was able to continue with the implanted pump. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs show cvp metrics starting to drift downwards and then going out of spec with a placement signal not reliable (psnr) alarm about 90.5 hours into the case. upon investigation of the pump, it passed electrical and the laser continuity test. the pump was unable to hold pressure properly at 200 mmhg during uson testing and reproduced the psnr alarm. the root cause of the optical signal issue was determined to be sensor silicone wear. this is the second occurrence of optical sensor wear on rp flex out of 300 usages (0.67%) as of the date of this complaint 1/22/2024. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5685 | 1220648-2025-31464 | 23059072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059072&pc=OZD | 10/14/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 10/14/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552338 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during shockwave therapy, the optical sensor was damaged; however, support continued without interruption. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data log shows the placement signal not reliable alarm along and all 5s parameter was down which indicate hard failure of the sensor. placement signal was lost immediately after the shockwave device used. the cause of the optical signal issue was most likely damaged sensor due to the shockwave device interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5686 | 1220648-2025-31465 | 23059146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059146&pc=PYX | 05/30/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 05/30/2024 | 09/15/2025 |  |  | 11/30/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024434880 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed an impella rp flex along with an impella 5.5 to support a patient in need of bipella support for cardiogenic shock and heart transplant rejection. the rp supported for over 15 days despite the loss of optical signal. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs show that throughout support the motor current, placement signal (dp sensor), and cvp (optical sensor) constantly fluctuate. all 3 metrics behave similarly. the 5s optical parameters are in specification. product not returned. the cause of the placement signal issue was most likely patient condition related as all data log metrics are following a similar pattern indicating they are reacting to their environment and the patient was a post transplant requiring ecmo, flex, and impella 5.5 support. the fm was changed from optical signal issue to placement signal issue as calculated flows and pa were reported to be inaccurate; the dp sensor is used to calculate the flows. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5687 | 1220648-2025-31466 | 23059172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059172&pc=PYX | 10/13/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 10/13/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025551100 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the complainant reported that an intermittent 'placement signal not reliable' alarm occurred. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were investigated and the placement signal not reliable (psnr) alarms were confirmed. there are indications that the optical signal was out of range. clinical details report that the patient presented to the hospital with a pulmonary edema, there blood pressures were low and required increased doses of epinephrin, the pump outlet was not past the pulmonary valve on 10/14, and the patient had a gi bleed. all of these factors would cause noise that would impair the sensor signal quality throughout the case. product was not returned for investigation. based on data logs and clinical details, it was determined that the root cause of the optical signal issues was most likely patient condition. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5688 | 1220648-2025-31467 | 23059228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059228&pc=PYX | 10/29/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 10/29/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025561678 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the complainant reported that the impella rp had a placement signal unreliable alarms. the alarm was disabled. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the fm was changed from optical signal issue to placement signal issue as the customer stated that flow was also unreliable indicating this is the dp sensor. the data logs were observed and drift was seen to start about 35 hours into the case on 10/29. the placement signal drifts down while the flow drifts up for about 25 hours before going out of range. the product was returned and there were no damages found to the device or dp sensor. the electrical resistances were in range. the pump was tested in a pressurized flow loop for 48 hours and dp sensor drift was seen about 55 minutes into starting the pump. the placement signal drifted down while the flow drifted up for about 25 minutes before completely going out of range. the root cause of the placement signal issues is due to sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5689 | 1220648-2025-31468 | 23059252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059252&pc=PYX | 12/19/2024 | 09/15/2025 |  | 09/15/2025 |  |  | 12/19/2024 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025561524 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient admitted in for coronary artery bypass grafting. the patient had an impella rp and 5.5 support ongoing. the rp lost optical signal but remained on for support. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. data logs show placement signal not reliable alarm dp out of range is then triggered and flow calculation is disabled. it seems to slightly resolve by the end of the case. the root cause of the placement signal issues was determined to be dp sensor drift. the fm was changed to placement signal issues based on the investigation indicating an issue with the dp sensor not the optical sensor. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5690 | 1220648-2025-31469 | 23059304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059304&pc=PYX | 02/01/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 02/01/2023 | 09/15/2025 |  |  | 07/31/2024 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023229871 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. during support, there was loss of placement signal. the team decided to not replace the pump and the patient remained on support. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for analysis and the 5s parameters showed a hard optical sensor failure. a light continuity test was performed and the light leaked out the sensor face indicating that the optical fiber was intact. the optical sensor was imaged while still on the pump and an irregularity was noted in the sensor face. the cannula was removed to get better visualization of the sensor face. a puncture could be seen in the silicone, the source of the small punctures likely external to the sensor and punctured inward toward the cavity. the data logs confirm a hard optical failure on the 6th day of support. the root cause of the optical issue was determined to be a damaged sensor face due to external material puncturing the sensor face. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5691 | 1220648-2025-31470 | 23059326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059326&pc=PYX | 12/05/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 12/05/2023 | 09/15/2025 |  |  | 01/31/2025 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395779 |  | 0046-0035 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient admitted for mitral valve and right ventricle failure. the rp had lost optical signal within hours of placement. the patient was additionally implanted with an impella 5.5 pump. the rp remained on for patient support until explantation. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. the data log shows several #1053 and #1055 placement signal not reliable (snr) alarms at the start of the case run, with low snr and contrast observed in the optical signal parameters. as per the clinical details, the pump had intermittent psnr alarms at the start of the case. some blood was given to the patient. the cause of the optical signal issue was most likely patient condition related based on data log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5692 | 1220648-2025-31471 | 23059448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059448&pc=OZD | 09/07/2023 | 09/15/2025 |  | 09/15/2025 |  |  | 09/07/2023 | 09/15/2025 |  |  | 12/31/2024 |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024318861 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. while on support, there was an aorta alarm falsely triggered with high motor current for p3. the impella positioning was confirmed via an echocardiogram. the patient was weaned from the impella cp and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. according to the clinical, shortly after beginning impella cp support false impella in aorta alarms with high motor current were recorded. no attempts were made to resolve the issue. the next day the patient was successfully weaned off of support. no product was returned for a visual inspection. data log analysis confirmed all 5s parameters were within normal bounds, the presence of no motor current spikes, and the presence of false impella in aorta alarms. the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5693 | 1220648-2025-46502 | 23059854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059854&pc=OZD | 08/16/2025 | 09/15/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662742 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 104 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient experienced access site bleeding and thrombocytopenia during impella cp support. access site bleeding was noted from oozing at the axillary site and was treated by packing and placing a drain. additionally, thrombocytopenia was found and as a result the extracorporeal membrane oxygenation was decannulated. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. minor bleed: the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
5694 | 1220648-2025-46503 | 23059856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059856&pc=OZD | 08/16/2025 | 11/03/2025 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688019 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 78 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support experienced runs of ventricular tachycardia. lidocaine was given, amiodarone bolus, and 1 gram calcium chloride given. rhythm stabilized with no ectopy. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the tachycardia could not be determined.
5695 | 1220648-2025-46504 | 23059857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059857&pc=OZD | 08/16/2025 | 01/17/2026 |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673677 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella cp placed to support the patient in st-segment elevation myocardial infarction. while on support, a thrombectomy to the left anterior descending artery was performed. throughout the time in the cath lab, the patient went into ventricular fibrillation and required at least 15 shocks. additionally, upon arrival to the intensive care unit (icu) from cath lab, the right groin dressing was saturated with blood and oozing out from the sides. there was no hematoma. per the intensivist (who was present in cath lab), this was an issue in lab that the doctor was aware of. there was oozing from hub site, which did not improve with addition of two sutures. surgical was placed at the site before it was dressed and the doctor gave icu the directive to not change the dressing. afterwards, it was noted that there was no more hemolysis or groin bleeding. the patient is no longer on impella support. | [5 empty] | Product complaint # ==>(b)(4). ventricular fibrillation/ thrombosis : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5696 | 1220648-2025-46505 | 23059875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059875&pc=OZD | 08/16/2025 |  |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683275 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced malpositioning of the device and arrhythmic events. the patient was planned for transferred in for higher level care. however, prior to the transfer the patient went into ventricular fibrillation arrest with return of spontaneous circulation (rosc), was intubated and then pulseless ventricular tachycardia arrested with rosc. the transfer via airflight was completed, and the patient was taken directly to the catheterization lab for impella cp placement and percutaneous coronary intervention (pci). two stents were placed in the ostial mid-left anterior descending artery (lad). during the procedure, the physician was unable to appropriately position the inlet in the left ventricle. the inlet remained at the level of the aortic valve with intermittent ¿impella in aorta¿ alarms. during an attempt to advance the impella, it jumped deep into the left ventricle and caused sustained ectopy and then jumped back to the level of the aortic valve. after the pci was completed, the patient remained off pressors. the impella cp device was slowing weaned and explanted with a survived patient outcome. hemostasis was achieved after perclose x2. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 added selection as it was inadvertently omitted from the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. h.6 investigation type, findings, conclusion and component code were updated accordingly based on the completed investigation. additional information was received and confirmed the pump is unavailable for return. in discussion with the physician, it was agreed upon that the pump was likely caught in the papillary. he was pleased with the support that the pump provided during percutaneous coronary intervention and elected not to place a new pump as the patient's hemodynamics were stable off pressors. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history review confirmed the device passed all the post sterile inspection checks. regarding, tachycardia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pertaining to the positioning issue, they were unable to position the inlet in appropriately in the left ventricle. the inlet remained at the level of the aortic valve, and it would jump deep into left ventricle when advanced. the cause of the positioning issue cannot be determined since insufficient clinical details were provided.
5697 | 1220648-2025-46507 | 23059899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23059899&pc=PYX | 08/16/2025 |  |  | 09/15/2025 |  |  | 08/16/2025 | 09/15/2025 |  |  |  |  | 09/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026652185 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Tachycardia | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock was implanted with impella cp and impella rp flex devices for bipella mechanical circulatory support (mcs). the patient underwent minimally invasive mitral valve replacement one day prior and had a turbulent postoperative course w/ marked bleeding and hypotension. after bipella implants, the patient was taken to the unit on mcs. upon transport to the unit, the patient became increasingly hypotensive (mean arterial pressure was approximately 50). drips were increased to the following doses: epinephrine 0.6 mcg/kg/min, norepinephrine 1 mcg/kg/min, dobutamine 7.5 mcg/kg/min, and vasopressin 0.03 units/min. milrinone was discontinued. after drip changes and 250 albumin, mean arterial pressure improved to greater than 65. the patient continued bipella support. continuous renal replacement therapy (crrt) was planned to improve acidosis. approximately three days later, the impella cp alarmed position in the aorta and was support level was decreased to p-3. continuing support and now four days thereafter, the heparin drip (which was only at <100units/hr.) had been off since one day prior. however, activated clotting time continued to trend up, with current result ¿out of range high.¿ treatment rendered. echo was done yesterday showing improvement in right ventricle function and no right ventricle distension. impella position adequate and left ventricle function still decreased. crrt continued with fluid removal as tolerated, and a tube feeding was started. the next day, approximately eight days after start of the bipella support, the impella rp flex placement signal was noted to be elevated (100s/20s). however, there was no change in purge flow, purge pressure, motor current or pump flow. the nurse noted the patient had received two units of packed red blood cells, fresh frozen plasma, platelets, and cryoprecipitate after draining 170ml blood from chest tube (it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient). a repeat chest x-ray and echocardiogram were completed and showed no change in position from previous placement. activated clotting time (act) was down to 182 following blood products administration. discussion was made with the physician, and the plan was just to monitor cp/rp numbers and hemodynamics and continue current treatment. within about an hour, waveforms returned to baseline. then early morning following the patient being cleaned up and repositioned, the impella rp flex waveforms again elevated (140s/30s). there was no change in pump flow, purge pressure of flow, or motor current, but placement signal pressure remained elevated. repeat chest x-ray showed no change in position of rp flex. act 180 (no systemic anticoagulation at this time). the physician was notified, and plan remained to continue current support and monitor acts every 2 hours. act was noted to have remained greater than 160 seconds. the following day, it was noted the rp flex waveforms returned to baseline. the patient remains on mechanical ventilation. the plan noted was to continue current treatment and recheck echocardiogram the following day and possibly wean the impella as tolerated. with sedation on hold, the patient had not awakened; however, the patient experienced tachycardia and blood pressure and reap rate increased. therefore, fentanyl was restarted. in the afternoon continuous renal replacement therapy was stopped with no plans to restart. the next day, lactate trended up to 6.7. sedation was off and the patient¿s pupils are reactive. however, the patient has remained unresponsive, with no gag or cough. dobutamine still at 5 with no other drips. the following day lactate was noted still increased, and it was believed to be possibly sepsis. the impella rp flex was weaned and explanted this morning. the impella cp was weaned down to support level p-4 and with possible plans to explant later. however, the patient would expire within the next hour as care was withdrawn. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella rp flex report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome. | Product complaint # (b)(4). investigation summary: no product returned optical signal issue: the data logs show no psnr alarms. clinical reported intermittent elevated placement signal values which resolved once after patient received pbrc and blood drained from chest tube. the logs confirm elevated placement signal readings and show noisy motor current with stable 5s parameters. a closer look at the logs revealed the first instance of elevated placement signal, during the chest washout, had pap and cvp elevate to 100s to 200s. the second instance occurred at about 4am when patient was being cleaned and repositioned. similar pattern was observed however cvp and pap spiked higher and was elevated for a longer duration. it is likely that the chest washout contributed to the incident occurrence and also likely that other devices such as crrt and a pa catheter could have interfered with the pump sensor causing elevated readings. no targeted troubleshooting efforts were reported. the cause of the issue could not be definitively determined as no product was returned and limited information was provided major bleed: the cause of the injury was not determined due to insufficient clinical details. tachycardia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. sepsis: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot number: 1901669. device history batch: subcomponent lot: n/a. device history review: the pump sn: (b)(6) passed all post sterile inspection checks. | Investigation team reviewed the complaint and from this review noted coding updates to align with voice of the customer as tachycardia as captured in the event description (per the initial report). consequently, h6: health effect-clinical and impact coding was updated, repopulated. medical safety review was completed and noted the following: the event describes bipella support provided to a patient in a decompensating state. there were complications noted with the impellas and the patient expired while on the impella cp device. the rp flex was explanted approximately one hour or so prior to the demise. impellas events is a serious injury and should be reported. however, with the impella cp, there is not enough to dissociate the device from the demise outcome. in this case, however, the patient's underlying condition likely contributed most to the demise outcome. the patient started in a decompensating state and never fully recovered. the family eventually decided to withdraw care from the patient which appears to have led to the demise outcome. medical safety officer reviewed the event and noted the same in the second impella cp device cannot be dissociated for the outcome and impella rp flex is a serious injury type of reportable event. related medical device reports: this impella rp flex device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome.
5698 | 1220648-2025-46440 | 23046242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046242&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700763 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient went into ventricular fibrillation arrest. cpr was performed and the patient was shocked once with return of spontaneous circulation. furthermore there were suction alarms due to low blood pressure and volume. the patient was reported to be tachycardic with a heart rate of 160-170 beats per minute. blood pressure dropped so vasopressin and epinephrine were started. they were turned off when the heart rate decreased to 115 with amiodarone, 1 unit of packed red blood cells administered, albumin and cryoprecipitate given. the gastrointestinal department was consulted for blood in the stool. the device was explanted and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. product was not returned for review. suction. there were no motor current (mc) spikes outside of p-level changes. no suction alarms on the reported day of the issue (8/17). the clinical details do not align with the dates of the data log review. the cause of the pump suction issue was not determined due to lack of clinical information and product was not returned. vt / tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension / gi bleed: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5699 | 1220648-2025-46441 | 23046243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046243&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587321 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male | 70 KG | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing for the patient admitted in with a st-segment elevation myocardial infarction. the pump was placed but during support an arrhythmia was noted, and an echocardiogram was performed to check whether the impella pigtail was touching the myocardium, but no contact was found. as a precaution, the pigtail was pulled out toward the aorta and repositioned, but it ended up being pulled further toward the aorta than expected. because the pigtail was on the left ventricle, we considered inserting the impella pump catheter into the left ventricle, but delivery along the wire was not possible, and there were areas of thin myocardium due to the presence of a left ventricular aneurysm, so there was a risk of delivery in an unexpected direction. therefore, it was decided to switch from impella to intra-aortic balloon pump. the patient was noted to be in stable condition and survived. | [5 empty] | Correction: b1. the investigation of the reported failure mode has been completed. no product was returned. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause for position issues was patient related due to patient having lv aneurysm with thin myocardium. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5700 | 1220648-2025-46442 | 23046244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046244&pc=OZD | 08/15/2025 | 09/12/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604917 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemoptysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received stating the cessation of systemic anticoagulation successfully led to the end of systemic bleeding. | The user facility reported that during support of the impella 5.5 on (b)(6) 2025, the patient in icu, had sudden large, black tarry stool concern for gastrointestinal bleed gib and significant hemoptysis. per icu doc, concern for systemic bleeding in setting of systemic anticoagulation before and in or as well as anti-plt therapy ongoing surgical procedure lad stent. optimize i's/o's. heparin is being withheld for a few hours longer. patient survived. |  |  |  |  | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of major bleed could not be determined as insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | B5 updated with additional information received from the clinical site. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5701 | 1220648-2025-46444 | 23046245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046245&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707191 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 78 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 the patient having multiple large melanotic bowel movements overnight, 1 packed red blood cells given and bival on standby with concern of gastrointestinal (gi) bleed, computed tomography angiography this afternoon to rule out. the patient was found to have a gi bleed (anti coags on hold). patient will be taken to operating room today for an embolization for duodenal ulcer after failed esophagogastroduodenoscopy cauterization on friday. care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5702 | 1220648-2025-46443 | 23046246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046246&pc=OZD | 08/15/2025 | 10/29/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681847 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Prime | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the placement purge fluid not detected alarm noted multiple times despite taking the tubing in and out of the console / disconnecting and connecting the yellow connection. a new purge cassette was opened and attempted with the same aic. same alarms noted. device would not purge. second aic powered on, connected and same priming process attempted with cp and second purge tubing with no luck. md said to open a new catheter and abort placing current cp: this fixed the purge problem. cp was implanted successfully. | The complainant reported a patient in cardiogenic shock (stage e) from an acute myocardial infarction was planned for impella cp device implant for mechanical circulatory support and experienced purge fluid not detected alarms. a second impella cp device was eventually placed. however, the patient expired. the patient came in as an outpatient. low ejection fraction of 19% noted prior to angiogram. distal left main and proximal left anterior descending disease was noted. the physician decided to pause and bring patient back for intervention to plan the case. the patient developed 10 of 10 chest pain after ivus run. thrombus was noted in distal left main during the picture and the patient rapidly decompensated, was emergently intubated and started on levophed. mean arterial pressures were still 65 despite pressors. the patient started to code, and the physician decided to place an impella device and complete percutaneous coronary intervention to coronaries. an impella cp was primed, and the purge fluid not detected alarm noted multiple times despite taking the tubing in and out of the console / disconnecting and connecting the yellow connection. a new purge cassette was opened and attempted with the same automated impella controller (aic). same alarms were noted. device would not purge. second aic powered on, connected and same priming process attempted with cp and second purge tubing without success. the physician requested a new impella cp catheter and aborted placing current impella cp. the impella cp pump 2 successfully primed. during cardiopulmonary resuscitation, successful insertion of impella cp pump 2 was achieved. return of spontaneous circulation was achieved, and the percutaneous coronary intervention was able to be completed. one stent was placed in the left main and old stents were ballooned in the proximal left anterior descending artery. the patient was noted to be doing well and staff started weaning levophed. prior to transferring to the unit, anesthesia was called to better sedate the patient. after propofol was administered, patient coded and could not be revived. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: analysis summary: purge fluid not detected: the pump and purge cassette were returned for investigation. no physical damage was observed on the returned product. the pump was then primed using the returned purge cassette without issue. the purge fluid was seen coming out from the purge gap and maintained steady pressure and flow rate for about 10 minutes during a test run. data logs shows several "purge fluid not detected" alarms during case start, with purge pressure readings around 30 mmhg and pump flow at 0. no issues were found on the returned product. the cause of the purge fluid not detected issue was most likely due to unintentional user error during case start. device history review: the pump sn (b)(6) passed all post sterile inspection checks. | Sections b5 event description, b6. tests/lab data including dates and b7: medical history/preexisting condition were updated accordingly to provide complete details regarding the reported event. patient-specific information a4 weight is unknown. medical safety review: medical safety reviewed the event. this event describes a priming failure resulting in an alarm of "purge fluid not detected." a different impella cp was successfully primed and implanted. the first impella cp malfunctioned and delayed hemodynamic support for approximately 10-15 minutes. this is a death type of reportable event as dissociation cannot be made given the delay. however, the patient was in a pre-demise state at the time of the support and highly likely contributed most to an outcome of death. regarding the second impella cp, it did not malfunction, and there was no report or indication of an associated adverse event. the second impella cp was in use during the patient's expiration but would not have materially contributed to the patient's demise. the patient doing well post procedure per the physician. to better sedate and after administration of propofol is when the patient coded and was unable to be revived. change in reportability. based on medical safety review, the type of reportable event was changed to death. consequently, the following fields: b1 adverse event and/or product problem, b2 outcomes attributed, h1 type of reportable event, and h6 health effect clinica/impact codes and medical device problem code were updated or repopulated to demonstrate alignment with medical safety review accordingly. note: h6 investigation: type, findings and conclusion codes and h6 component code remain unchanged.
5703 | 1220648-2025-46445 | 23046247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046247&pc=OZD | 08/15/2025 | 09/12/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669153 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp placed for the support of a patient found down from a witnessed arrest. the patient was transferred from community to a tertiary medical center. the cp was placed and coronary microvascular disease assessed as well as lack of neurological status. the cp was explanted with signs of hemolysis after 31 hours. patient survived event. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h6: added code to medical device problem code. h11: updated based on the results of the investigation. investigation summary: no product returned for investigation. no data logs available. hemolysis/mechanical interaction with blood: patient¿s creatinine is continuing to climb with blood in his urine. hemolysis began during cp support and patient not responding to stimulants, so pump was explanted. the cause of the hemolysis was not determined due to lack of product and data logs available for investigation. device history review: the pump passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5704 | 1220648-2025-46446 | 23046248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046248&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026692517 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Backflow | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu had a cp pump placed to support the patient admitted in with ami/cgs. the pump had supported for just over 8 days when there were retrograde flow alarms that prompted the team to troubleshoot, and this lead to a pump stop. the pump was explanted and not replaced as patient was noted to be stable. | [5 empty] | Patient's age, race and ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for evaluation. insufficient data logs were returned. logs at time of reported issue were not returned for evaluation. retrograde flow/pump stop: clinical details noted that a "retrograde flow" alarm was noted. pump was attempted to be restarted multiple times with motor current increase and immediate pump stops. the cause of the pump stop could not be definitively determined as insufficient logs were returned and no product was returned for evaluation. the cause of the retrograde flow could not be definitively determined as limited clinical details were provided and insufficient data logs were returned. device history review: the device passed all the post sterile inspection checks.
5705 | 1220648-2025-46447 | 23046249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046249&pc=PYX | 08/15/2025 | 10/30/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026686616 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Difficult to Advance | [5 empty] | Cardiac Arrest | Cardiac Perforation | Pericardial Effusion | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Impella rp inserted over wire till the tip was appropriately positioned in pulmonary artery. wire pulled out and impella device turned on achieving satisfactory flows. peelaway sheath was removed and repositioning sheath was advanced to skin level. the patient shortly after had hemodynamic instability and echocardiogram confirmed pericardial effusion, (suspected due to iatrogenic injury during advancing the repositioning sheath). surgeon came and completed emergent sternotomy, after which, emergent arterial and common atrial cannulation was done and cpb. after systematic exploration of cardiac chambers, injury rv apex observed and repaired using 4-0 prolene. chest was closed and pump weaned. | The complainant indicated that a patient suffering from right heart failure had an impella rp flex device implanted for mechanical circulatory support. during the support period, the pump advanced into the right pulmonary artery (rpa) before retracting to the main pulmonary artery (pa). a 23 fr peel-away sheath was extracted, and transesophageal echocardiography (tee) revealed that the outlet was shallow within the main pa. the physician advanced the rp flex and repositioning sheath without direct visualization. it was observed that the motor current/waveform spiked on the automated impella controller (aic). upon fluoroscopic examination, it was noted that the pump/inlet was situated in the right ventricle (rv) while the tip/outlet was located in the left pulmonary artery (lpa). the pump was subsequently retracted, and a significant pericardial effusion was detected via tee. the patient experienced cardiac arrest, prompting the initiation of cardiopulmonary resuscitation (cpr) and the opening of the chest. the patient was placed on cardiopulmonary bypass (cpb). surgical exploration uncovered a perforation at the rv apex. the perforation was then surgically repaired, and the rp flex was further retracted and repositioned under fluoroscopic guidance. subsequently, the patient ultimately expired while on support. |  |  |  |  | The investigation for the reported positioning issue and cardiac arrest has been completed. the cause of the positioning issue was determined to be an unintentional use error since the pump pulled out during 23 fr peel-away sheath removal and repositioning unit advancement. the distal tip/outlet came back slightly in the main pa and was re-advanced without visualization. the root cause of the cardiac arrest was not determined due to insufficient clinical details. | Correction: g2 additional information received regarding procedure, b5 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5706 | 1220648-2025-46448 | 23046251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046251&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705096 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 69 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella cp the patient was without distal pulses the physician placed retrograde sheath down left groin but did not begin fem-fem bypass due to poor prognosis. the patient was transported to intensive care unit. however, the next day the patient remains without distal pulses, medical doctor was aware. the patient was made do not resuscitate, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary : ischemia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot : device lot: 1945899. device history batch : subcomponent lot: n/a. device history review : device (sn: (b)(6)) passed all post sterile inspection checks.
5707 | 1220648-2025-46449 | 23046252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046252&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720164 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 70 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with undergoing high-risk percutaneous coronary intervention (pci) was preoperatively implanted with an impella cp device for mechanical circulatory support. the patient was noted as too high risk for coronary artery bypass graft surgery. after the pci the patient was transferred to the unit with impella mcs in place. the next afternoon, the patient began to desaturate with abdominal swelling and tenderness. patient was then paralyzed. labs were drawn and with results, there was concern for internal bleeding. the patient had a computed tomography scan of the abdomen/pelvis area to rule out access site bleed and assess if surgical intervention and repair was necessary. the ct scan showed retroperitoneal hematoma with no active extravasation, arterial perforation, or evidence of access-site complication. impella access site was soft and intact without evidence of hematoma or bleeding. the timing of bleeding symptoms with coagulopathy suddenly in the afternoon was felt to be more consistent with spontaneous/small-vessel bleed in the abdomen. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. the patient was stabilized and pressors weaned down after blood transfusion. impella mcs remained in place for three additional days before the patient was successfully weaned and device explanted in the catheterization lab with a survived patient outcome. two perclose devices were used with successful hemostasis to the right groin. the left groin was used to visualize an extravasation. no extravasation was found throughout the iliac or femoral arteries bilaterally. the access site was clean, dry, and intact with no signs of hematoma. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the major bleed, the cause of the injury was not determined since product was not returned and insufficient clinical details provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5708 | 1220648-2025-46451 | 23046253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046253&pc=OZD | 08/15/2025 | 10/22/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026683506 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experiencing acute myocardial infarction and cardiogenic shock was implanted with an impella cp device to provide mechanical circulatory support. approximately one hour following the insertion, the placement signal (ps) was lost, and shortly thereafter, a purge system blocked alarm was triggered. the team made efforts to de-air and replace the cassette; however, the new cassette failed to prime. an exchange to a new console was not performed due to the potential risk of losing the ps. consequently, the pump was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported purge pressure and optical signal issue has been completed. the data logs for the purge pressure show a gradual rise over 7 minutes in purge pressure before alarms were triggered. purge pressure was elevated through till explant. the root cause of the purge pressure issue was likely biomaterial build up. the logs for the optical signal issue show a sudden loss of placement signal. at that instant there was a drop in snr, light and gain and an increase in amplitude of contrast. lamp remained at 100. the pattern identified is characteristic of a fiber break but insufficient information was provided to determine cause of fiber break.
5709 | 1220648-2025-46452 | 23046255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046255&pc=OZD | 08/15/2025 | 09/12/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701035 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 87 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support experienced thrombocytopenia and blood products was administered. the patient also had intermittent rate controlled atrial fibrillation and amio was given. it was further reported that the patient had a hematoma at the right axillary artery requiring a washout. the patient continued to be stable and was explanted as planned. | [5 empty] | Correction: e4. the investigation of the reported failure modes has been completed. no product was returned for investigation. paroxysmal atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma: the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5710 | 1220648-2025-46454 | 23046257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046257&pc=OZD | 08/16/2025 | 09/12/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700200 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 63 YR |  | Male | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella 5.5 as an escalation of therapy until receipt of transplant. the impella site sleeve was noted to be compromised. the medical staff were advised to clean with betadine, cover with sterile gauze, and a tegaderm. also recommended recording the impella device to the patient with less tension. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product or images were returned for evaluation. the cause of the damage to the anti-contamination sleeve could not be determined as limited clinical details were provided and no product were returned for evaluation. device history review: the device passed all the post sterile inspection checks.
5711 | 1220648-2025-46453 | 23046258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046258&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026685315 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they had a single access issue. the patient's anatomy was very tortuous and calcified. it was difficult to pass the impella cp until bifurcation, and then insertion was possible. additionally, the patient experienced bleeding which was stopped with compression. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5712 | 1220648-2025-46456 | 23046260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046260&pc=PYX | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025529618 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella rp underwent a laparoscopic cholecystectomy. during the procedure, the pump might have pulled back as the team was having a hard time visualizing the outlet of the pump with echo. a decision was made to place another impella rp in the contralateral groin. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5713 | 1220648-2025-46455 | 23046261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046261&pc=OZD | 08/15/2025 | 10/24/2025 |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657789 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 144 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that on day one of impella cp support the patient¿s labs were drawn twice and both were hemolyzed. the impella p-level was lowered to 6. the patient continued on impella support and was explanted as planned. patient survived | [5 empty] | The investigation for the reported hemolysis has been completed. no product was returned. after reviewing the logs, suction alarms along with a trend in purge and placement signals were observed. the cause of hemolysis was most likely patient condition related per the review of logs. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5714 | 1220648-2025-46458 | 23046262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046262&pc=OZD | 08/22/2025 |  |  | 09/13/2025 |  |  | 08/22/2025 | 09/13/2025 | [6 empty] | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300700 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient diagnosed with cardiomyopathy received an impella 5.5 device for mechanical circulatory support. the complainant also mentioned an issue with the automated impella controller (aic) regarding battery power. the aic (b)(6) had been disconnected from ac power for about three minutes. consequently, the decision was made to postpone the patient's ambulation, and the aic was reconnected to the same outlet. following this, the aic promptly issued a low battery power notification, with the indicator displaying 50%. the team then switched to backup aic (b)(6). the (b)(6) was promptly taken to the biomed department, so the alarm history was not reviewed. no serious or critical alarms were found in the connect log. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This follow-up report is being submitted to correct d2a and for investigation conclusions. d2a. common device name has been corrected. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. the cause of the battery level low alarm was a communication anomaly between the battery and power battery management board.
5715 | 1220648-2025-46457 | 23046264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046264&pc=OZD | 08/27/2025 |  |  | 09/13/2025 |  |  | 08/27/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1076153 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 97 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered automated impella controller (aic) issues. it was reported that the aic encountered a "placement signal not reliable" alarm and so the placement monitoring was disabled and the audio was turned off. monitoring via motor current and echo for placement was to be checked, and he physician was also made aware of situation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5716 | 1220648-2025-31426 | 23046489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23046489&pc=OZD | 08/01/2025 | 09/13/2025 |  | 09/13/2025 |  |  | 08/01/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456617 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, placement signal not reliable alarms were triggered. despite these alarms, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data log shows confirm at the two times there were placement signal not reliable alarms. the console was returned for investigation. the console's 5s had low signal not reliable and failing contrast though this would not have caused the failure mode seen in the field. a console was run with purge and pump but the failure mode was not reproduced. the root cause of the intermittent unreliable placement signal and oscillating contrast was a defective optical bench because oscillating contrast is indicative of inadequate light production which would be a failure of the lamp. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5717 | 1220648-2025-46460 | 23047399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047399&pc=OZD | 08/15/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647523 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed to support the percutaneous coronary intervention procedure. at the end of the procedure, the patient had ventricular tachycardia run requiring to be shocked twice. the patient continued on cp support until successful weaning of the pump. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5718 | 1220648-2025-46462 | 23047404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047404&pc=OZD | 08/16/2025 |  |  | 09/13/2025 |  |  | 08/16/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1760215 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 164 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the battery percentage of the automated impella controller (aic) dropped to 50% despite being plugged into ac power the entire time. attempting alternate outlets did not resolve the issue. did not. a new aic was used for patient support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: data analysis: console logs from the reported day of event are consistent with complaint and show that 2.5 hours after case start, console presented with battery failure alarm (#360), and after ac power was disconnected, the console kept operating but showed a battery level low alarm (#23). after reconnecting ac cord, battery level low alarm resolved, but battery failure alarm was unresolved. pump was transferred to a backup console, and ic10987 was powered down. battery and power data embedded in ltlog files showed that at the time of event, battery 2 had diminished capacity, and individual cell voltages for battery 2 showed lover voltage for cell 4 (cell imbalance) due to failure of once of cells. device analysis: battery failure was confirmed during testing. results of batttery diagnostic tool indicated battery cell imbalance of 212mv and battery 2 was internally shut down, as evidenced by its battery status reading ¿terminate_charge + terminate_discharge¿. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d.9 and h.6 were updated accordingly due to the new information. d.3 manufacturer name should not have had a number behind it on the initial manufacturer report that was submitted. e.1 initial reporter phone was added as it was inadvertently omitted from the initial manufacturer report.
5719 | 1220648-2025-46461 | 23047405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047405&pc=OZD | 08/17/2025 |  |  | 09/13/2025 |  |  | 08/15/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631991 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | Perforation of Vessels | Pericardial Effusion | [8 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was implanted with an impella cp device to provide mechanical circulatory support. during the pump support, the patient experienced bleeding complications in the lungs. one unit of packed red blood cells (prbcs) was administered to address intermittent suction alarms that resolved following repositioning. the impella device was adjusted at the bedside under the guidance of echocardiography (echo). a slight pericardial effusion was observed on the echo. heparin administration was paused, and the physician ultimately opted to proceed with d5w and sodium bicarbonate. additionally, the patient lost distal pulses in the left lower extremity (lee) at the site of the impella device. a vascular consultation was sought, and a fasciotomy procedure was discussed. all extremities were noted to be dusky and cold. an electroencephalogram and computed tomography scan were conducted, revealing a stroke; however, the timing of the stroke could not be established. the patient's family made the decision to withdraw care, and the patient subsequently passed away. | [5 empty] | Correction g4, h6 health effect - clinical code the investigation of the reported failure mode has been completed. no product was returned. major bleed: the root cause of major bleed could not be determined as insufficient clinical details were provided. stroke, ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. vascular injury: the root cause of vascular injury cannot be determined since the product was not returned and insufficient clinical details were provided. positioning issue: the cause of the positioning issue was unable to be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
5720 | 1220648-2025-46463 | 23047406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047406&pc=OZD | 08/16/2025 | 09/12/2025 |  | 09/13/2025 |  |  | 08/16/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702696 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 115 KG | Non Hispanic | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had hemolysis. the patient's lab came back hemolyzed and the urine output was dark. no further information was noted. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary hemolysis: no product was returned. after reviewing logs, no issues were observed. the root cause of hemolysis could not be determined as insufficient clinical details were provided. device history review device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5721 | 1220648-2025-46464 | 23047408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047408&pc=OZD | 08/16/2025 | 09/12/2025 |  | 09/13/2025 |  |  | 08/16/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678138 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had an impella cp placed to support the patient in st-segment elevation myocardial infarction. while on support, the doctor was concerned with bleeding around the repositioning sheath and, despite troubleshooting, decided to wean the impella, remove it, and hold pressure. the doctor sent the patient to the unit. patient survived | [5 empty] | E.1 added initial reporter phone and zip code extension as they were inadvertently omitted from the initial report that was submitted. e.4 added selection as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturing site name should have been left blank on the initial report that was submitted. the investigation was completed and no product was returned. major bleed: the root cause of major bleed could not be determined as insufficient clinical details were provided. the device history review was completed and the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5722 | 1220648-2025-46465 | 23047409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047409&pc=OZD | 08/16/2025 |  |  | 09/13/2025 |  |  | 08/16/2025 | 09/13/2025 |  |  | 01/31/2026 |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025477202 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced oozing because the mattress suture was not tied prior to the impella sheath. the sutures were clipped and re-tied, manual pressure was held for 30 minutes, and a pressure dressing was placed to successfully resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.6 code f1901 was removed and a new impact code was added. the investigation was completed and no product was returned for investigation. major bleed: bleeding during the time of implantation at access site around the sheath before placing the mattress sutures. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. sutures were clipped and retied. hemostasis was achieved by placing the art line and manual pressure held for 30 minutes and pressure dressing. the device history review was completed and the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5723 | 1220648-2025-46467 | 23047410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047410&pc=OZD | 08/16/2025 | 11/03/2025 |  | 09/13/2025 |  |  | 08/16/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674441 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 64 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Bradycardia | Tachycardia | Asystole | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella cp flex on (b)(6) 2025, impella was inserted and support initiated. the patient had ventricular tachycardia (vt) arrest x1 during prep process. peel-away sheath removed, exchanged for repositioning sheath. after the exchange process, the patient lost all pulsatility on the impella, had multiple suction alarms, low flows on cp, and had a rhythm change to bradycardia, followed by asystole. cardiopulmonary resuscitation (cpr) initiated, tvp placed, unable to capture. cpr continued, unable to regain rhythm. time of death called at 2050. | [5 empty] | The investigation for the reported tachycardia, low pump flow, and asystole. . the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the tachycardia, low pump flow, and asystole could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5724 | 1220648-2025-46466 | 23047411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047411&pc=OZD | 08/16/2025 |  |  | 09/13/2025 |  |  | 08/16/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710428 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 164 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp was suspected of having an infected kidney stone and exhibited symptoms of distributive shock. the patient was treated with calcium chloride, epinephrine, and sodium bicarbonate. additionally, there were frequent suction events. ultimately, care was withdrawn and the patient expired. | [5 empty] | No product was returned. after reviewing the logs, suction alarms were observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5725 | 1220648-2025-46468 | 23047412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047412&pc=OZD | 08/16/2025 | 09/12/2025 |  | 09/13/2025 |  |  | 08/16/2025 | 09/13/2025 |  |  |  |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696145 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 78 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 on (b)(6) 2025, the impella inserted smoothly with no issues, initial flow per dr. (b)(6) around 5 lpm. drops being slowly weaned. patient in stable condition and sent back to cardiovascular intensive care unit to await percutaneous coronary intervention. the patient had two ventricular tachycardia (vt) codes this afternoon with continued decompensation. however, the patient required more pressor support with additional runs of vt. after discussion with family, decision made to make patient comfort care. impella support was discontinued. prayers of comfort for family. the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. tachycardia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5726 | 1220648-2025-46469 | 23047413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23047413&pc=OZD | 08/17/2025 |  |  | 09/13/2025 |  |  | 08/17/2025 | 09/13/2025 |  |  | 12/31/2025 |  | 09/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448891 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 75 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ischemia; there was poor distal flow with the repositioning sheath in place. this complication was managed with an external 6 french femoral-femoral bypass. later, care was withdrawn and the patient expired. | [5 empty] | B2 added selection that was inadvertently omitted from the initial report that was submitted. e1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." investigation summary: ischemia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5727 | 1220648-2025-31385 | 23034800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23034800&pc=OZD | 07/25/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 07/25/2025 | 09/12/2025 |  |  | 03/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024329246 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported a patient was implanted with an impella 5.5 for circulatory support. during implantation there was sensor failure that did not require the pump to be explanted. the patient was successfully weaned and the pump was later explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the product was received with the pump already cut open. no investigation could be performed. no data logs being provided. the cause of the issue was not determined since product was not returned. all pertinent information available to abiomed has been submitted at this time.
5728 | 1220648-2025-31386 | 23034830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23034830&pc=OZD | 08/24/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 08/24/2023 | 09/12/2025 |  |  | 04/30/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335647 |  | 0048-0024 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during impella support, the pump experienced abnormal position waveform but resolved changing the p-level. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for a visual inspection. no data logs were returned for analysis. the cause of the optical signal issue was not determined because no product or logs were returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5729 | 1220648-2025-31387 | 23034841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23034841&pc=OZD | 04/24/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 04/24/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390491 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump intermittently triggered the alarm message 'impella in aorta.' device positioning was subsequently verified via echocardiography, and support was maintained. there was no patient harm associated with the issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs show that signal not reliable got very noisy in the middle of the case, but otherwise had good optical signals. the console was returned for analysis but the failure mode was not reproduced during testing. the root cause of the optical signal issue could not be determined since it was not reproduced. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5730 | 1220648-2025-46392 | 23036536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23036536&pc=OZD | 08/15/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 08/15/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695276 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp placed to support a patient in acute myocardial infarction with cardiogenic shock. during implantation, the 14 french peel-away was unable to advance so the doctor elected to use a 16 french gore and was able to advance the impella cp. the pump was unable to pass the aortic arch. the doctor believed the patient had a porcelain aorta. the decision was made to crash the patient onto extracorporeal membrane oxygenation and the impella cp was aborted. | [5 empty] | A.4 - a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of the reported failure mode has been completed. unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had porcelain aorta preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
5731 | 1220648-2025-46393 | 23036653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23036653&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126837 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported there was a battery failure identified when the automated impella controller (aic) was turned on for a potential case. the aic had been plugged in overnight and a battery failure alarm presented the next day in the morning. there was no patient involvement with this event. | [5 empty] | Additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. | Section a - no patient was involved. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5732 | 1220648-2025-46391 | 23036807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23036807&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704276 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp had limb ischemia. distal flow to the right femoral artery was occlusive from the impella cp sheath. antegrade and ipsilateral access was obtained, and a balloon angioplasty and mechanical thrombectomy were performed to restore flow to the distal foot. the patient was then transferred to the intensive care unit. additionally, during support, the patient had absent right leg pulses and a reperfusion sheath was put in place. the patient was successfully weaned from the pump and survived. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation has been completed. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks.
5733 | 1220648-2025-46394 | 23036952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23036952&pc=OZD | 08/15/2025 | 10/14/2025 |  | 09/12/2025 |  |  | 08/15/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713217 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 62 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp placed to support a patient with non-st-segment elevation myocardial infarction. when the impella cp was removed, there was a clot that encapsulated the pigtail. the patient was therapeutic during length of procedure and therapy. the patient was stable post removal. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | H6 the impact code in the initial manufacturer report was incorrect and a new code was added. the investigation for the reported thrombosis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombosis could not be determined.
5734 | 1220648-2025-46395 | 23036991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23036991&pc=OZD | 08/13/2025 | 01/30/2026 |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1306449 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that during impella cp support, the automated impella controller (aic) had an alarm. details on the alarm were not provided. it was noted that the alarm did not cause patient harm and there was no impact on vital signs due to the alarm. to resolve the issue, the aic was replaced. the patient remains on impella cp support with the new aic. | [5 empty] | A.4 is unknown. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | Medical safety reviewed the event. this event describes an automated impella controller (aic) failure and per the physician patient's condition remained "normal" and there was no impact on vital signs. there is no report or indication of this aic failure requiring an aic-to-aic transfer and no hemodynamic decompensation in the patient. in this case, the patient's underlying condition likely contributed most to the demise outcome. however, this device malfunction if occurring in another patient was repeated (as in this case) could result in significant hemodynamic decompensation, and therefore, the event should be reported. with this event there was no impact to the patient, and the type of reportable event is a malfunction.
5735 | 1220648-2025-31388 | 23037415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23037415&pc=OZD | 02/15/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 02/15/2024 | 09/12/2025 |  |  | 08/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374324 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, the device experienced an optical sensor malfunction. echocardiography confirmed that the impella was in good position. despite this, the pump remained operational throughout the support period until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. upon review of the returned pump, the pump passed the laser continuity test. a clot was observed around the impeller. the 5s parameters were within normal specification. there were no catheter kinks. data logs show there are no optical related alarms. the root cause of the optical signal issue was most likely patient condition related due to the patient's deteriorating condition from the previous aortic dissection requiring surgical intervention while on this device. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5736 | 1220648-2025-46389 | 23037428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23037428&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181834 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complaint reported that while prepping for patient support, the automated impella controller (aic) had a purge disc failure. the cassette was replaced but the issue did not resolve. no patient harm was reported. | [5 empty] | Section a - patient information was not provided. e.1 initial reporter title, first and last name are unknown. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5737 | 1220648-2025-46396 | 23037554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23037554&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701566 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 108 KG | Non Hispanic |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted for high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. during implanting, the medical doctor went to reposition the impella cp and pulled back across the aortic valve and could not re-advance. the medical doctor made the decision to remove the impella cp and continue with the procedure. no patient harm was reported due to the issue. | [5 empty] | A.6 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary complaint device not returned for analyze. device in wrong position: the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled back across aov during repositioning and was not able to re-advance by the md. device history lot device lot#: 1942861 device history batch sub-component lot #: n/a device history review pump sn (b)(6) pass all post sterile inspection criteria.
5738 | 1220648-2025-46398 | 23038022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23038022&pc=OZD | 08/15/2025 |  |  | 09/12/2025 |  |  | 08/15/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161467 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 95 KG | Hispanic | White | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. post the impella cp removal from the patient and disconnecting the impella cp from the automated impella controller (aic), three red alarms displayed: impella stopped retrograde flow, impella stopped motor current high and complete procedure. all three alarms would not clear although the impella cp was unattached to the aic. there was no harm to the patient as a result of this event as this event occurred after successful wean and explant of the pump from the patient. | [5 empty] | The automated impella controller device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The product with event logs were received and the investigation was completed. service history review noted there are no quality notifications associated with the complaint device's most recent service report. subcomponent lot review revealed there are no other product malfunction complaints in the subcomponent lot related to the complaint failure mode. in summary, there were no issues related to the complaint discovered when reviewing the product history of console ic3191. the case associated with the reported complaint was initiated on aug 10, 2025, with the pump (sn: 605964). during support at p-level p1, on the complaint date august 15, 2025, the log recorded ¿15/08/25 17:03:15 processsampleforopminvalidsignalerror: calculator placement signal 1036.¿ at this timestamp, rt log confirms that the pump was disconnected, as the signal not reliable (snr) dropped. although the pump was disconnected, the console did not recognize the disconnection and displayed the following three alarms, which confirm the reported failure mode. ¿impella stopped (mechanical/electrical issue) ¿ alarm,¿ ¿impella stopped (motor current high) [detected by epc] ¿ alarm,¿ issued, and ¿impella stopped. retrograde flow. - alarm issued.¿ when the user tried to shut down the console, it displayed ¿shutdown requested by user, but pump is connected,¿ most likely due to a malfunction in the impellatronic disconnection mechanism. since the pump disconnection was not recognized, the console continued to calculate the placement signal without the pump, resulting in the ¿placement signal not reliable (out of range - optical)¿ alarm. the console was then forcefully shut down and restarted four times. after each restart, the console displayed the ¿unexpected controller shutdown¿ alarm. device analysis: this console was tested. the pump disconnection issue could not be reproduced. following the same sequence of actions as on the complaint date - while running the pump at p-level p1 on battery power¿the console recognized the pump disconnection when the pump was disconnected. the root cause of the console not recognizing the pump disconnection was most likely a malfunction in the impellatronic disconnection mechanism. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Additional information noted the automated impella controller was returned and not discarded, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, d.9 device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. d.4 expiration date should of been left blank on the initial manufacturer report as the automated impella controller does not have an expiration date. d.6a and d.6b these fields should of been left blank on the initial manufacturer report that was submitted as the automated impella controller is not implanted.
5739 | 1220648-2025-46400 | 23038618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23038618&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  | 08/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387746 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 114 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had multiple runs of ventricular tachycardia and was shocked six times. an echocardiogram was performed and the impella was free and clear in the mid ventricular space. the patient also had heparin-induced thrombocytopenia. all heparin was removed and bivalirudin was used. it was further reported that care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia, thrombocytopenia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5740 | 1220648-2025-31389 | 23038902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23038902&pc=OZD | 08/28/2023 | 10/27/2025 |  | 09/12/2025 |  |  | 08/28/2023 | 09/12/2025 |  |  | 04/30/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347823 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, ¿position in aorta¿ alarms were observed. despite this, the pump remained operational throughout the support period until weaning and explantation. no patient harm was reported. | [5 empty] | D4 unique identifier (udi) # corrected. d4 expiration date corrected. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for review. the root cause of the optical signal issue was most likely patient condition as the patient has a small left ventricle on echo with frequent positioning issues throughout the duration of support. all pertinent information available to abiomed has been submitted at this time.
5741 | 1220648-2025-46401 | 23038925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23038925&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026676217 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | Yes | Backflow | Malposition of Device | Difficult to Advance | Adverse Event Without Identified Device or Use Problem |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, slight oozing was noted from the access site after closing. anastomosis was not bleeding when closing. pressure dressing was applied and a 3 point fixation was applied. additionally, the p-level was turned down to p3 but the pump gave ¿preventing retrograde flow¿ alert so speed was increased. at this time, the patient is still on impella support. | [5 empty] | Explant date is added. no product returned, only event logs were returned and investigation was completed. based on the clinical information, the root cause of the access site adverse event is due to user error. the data logs show that there were no large discrepancies in motor current, but there were spikes in ps before the alarm occurred. since the pump was at p-3, the flow was calculating off of the ps. the pump flow calculation was less than 0.1 l/min at points triggering the minimum motor speed to be increased. due to a lack of clinical details and no product surrounding what was occurring during this time, the root cause of the retrograde flow cannot be determined. | Correction: h6. the investigation of the reported failure mode has been completed. no product was received for evaluation access site adverse event: the root cause of the access site adverse event is due to use error as there was improper anticoagulation management. retrograde flow: the data logs show that there were no large discrepancies in motor current, but there were spikes in ps before the alarm occurred. since the pump was at p-3, the flow was calculating off of the ps. the pump flow calculation was less than 0.1 l/min at points triggering the minimum motor speed to be increased. due to a lack of clinical details and no product surrounding what was occurring during this time, the root cause of the retrograde flow cannot be determined. difficult to deliver or advance: the root cause of the difficult to deliver or advance cannot be determined. device in wrong position: due to a lack of sufficient information, the root cause cannot be determined. device history review: the device passed all the post sterile inspection checks. | Updated information: d4 catalog number updated. h6 impact code f01, f2301, and f19 removed and replaced with f26. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5742 | 1220648-2025-46402 | 23038955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23038955&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026699476 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. post explant groin bleed requiring protamine and fem stop. checked on the patient after interventions above done and the groin looked much better. patient is stable. | [5 empty] | Product complaint # (b)(4). investigation summary: major bleed: the root cause of major bleed could not be determined as insufficient clinical details were provided. device history lot: device lot: n/a. device history batch: subcomponent lot: n/a. device history review: phr review is not possible since the serial number of device is unknown. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5743 | 1220648-2025-31390 | 23039030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039030&pc=OZD | 10/17/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 10/17/2024 | 09/12/2025 |  |  | 02/28/2026 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490865 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the impella cp lost aortic and left ventricular waveforms. purge flow and purge pressure remained within range. despite this, the pump remained operational throughout the support period until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cut pump was returned for investigation. biomaterial was found wrapped on the impeller. no further investigation was performed on the returned product. as per the data log, a loss in placement signal was reported with all 5s parameters temporarily. upon review of data logs, no problem was found with the pump and it is apparent the console was the potential cause of the issue. please refer to complaint (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5744 | 1220648-2025-46403 | 23039071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039071&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703213 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant of the impella, the patient went into supraventricular tachycardia (svt) and as a result cardioversion was performed successfully. additionally, the patient had a site bleed with hematoma. | [5 empty] | The investigation of the reported failure mode has been completed. arrhythmia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
5745 | 1220648-2025-46382 | 23039072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039072&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026700150 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 49 KG |  |  | Malfunction | No | Yes | Crack | Restricted Flow rate | Device Sensing Problem | Material Split, Cut or Torn |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was supporting a patient for close to a month while the patient was waiting for a heart transplant. on day two of support, leaking/condensation was noted inside the purge sidearm. purge flows and pressures remained within normal ranges. the purge sidearm seemed to leak when the purge was upright. when the sidearm was laid flat some leaking was noted. the decision was made to keep the impella in place and to closely monitor for any changes. on day seven of support, the anticontamination sleeve tore and was repaired with a tegaderm. two days later, the patient's central venous pressure (cvp) dropped resulting in a suction alarm. the patient was given albumin which resolved the suction issue. thirteen days later, placement signal disappeared but reappeared after the provider repositioned the catheter. the plastic sheath was torn while repositioning the pump. the patient remains on impella support. | [5 empty] | A.5 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represent the pump and the other report represents the automated impella controller. refer to section h1.0 for the related report number.
5746 | 1220648-2025-31392 | 23039109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039109&pc=OZD | 07/21/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 07/21/2024 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1211522 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed placement signal unreliable. the audio was disabled. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. console logs from the reported day of event are consistent with complaint. console was inspected in danvers, and it¿s optical system tested. debris and contamination were found on the optical pump connector, which also failed ¿5s¿ specification. the root cause of the aic issue was a defective blue receptacle. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5747 | 1220648-2025-46404 | 23039110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039110&pc=OZD | 08/13/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026702469 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | Ventricular Fibrillation | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on impella cp support for high-risk percutaneous coronary intervention (hrpci). during intervention, the patient became hypotensive, went into ventricular tachycardia, and required multiple shocks, rounds of cardiopulmonary resuscitation (cpr), and advanced cardiac life support medications and return of spontaneous circulation (rosc) was achieved. the patient remained hypotensive during the duration of the case and flows decreased. the pump was repositioned multiple times without success. the patient's hemoglobin was low and required 2 units of packed red blood cells. within 20 minutes of arrival to the intensive care unit, the patient went into ventricular fibrillation, requiring multiple rounds of cpr until rosc was achieved. the patient coded again later and the family withdrew care. the patient expired within 3 hours of admission to the intensive care unit. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. vt/vf: the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. low pump flow: no conclusive biomaterial ingestion trend or other abnormalities found per log review. the cause of the low flow cannot be determined since the complaint device not returned for analyzing, inconclusive log evidence and due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5748 | 1220648-2025-31393 | 23039250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039250&pc=OZD | 04/10/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 04/10/2024 | 09/12/2025 |  |  | 01/31/2026 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460800 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during support, the placement signal was unreliable. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned and upon review of data logs, it is apparent that the console is the potential cause of the issue. no problem was found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5749 | 1220648-2025-46381 | 23039266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039266&pc=OZD | 09/02/2025 |  |  | 09/12/2025 |  |  | 09/02/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1713948 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 49 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the placement signal of an impella 5.5 disappeared but reappeared after the physician repositioned the catheter. no patient harm was reported. the clinical consultant believes the placement signal was the pump rather an automated impella controller issue. this report was created to represent the aic as it is not for certain that the signal issue was a pump issue. | [5 empty] | A.5 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represent the automated impella controller and the other report represents the pump. refer to section h1.0 for the related report number. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5750 | 1220648-2025-46405 | 23039289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039289&pc=OZD | 08/14/2025 | 10/30/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716839 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella cp on (b)(6) 2025, that the patient was still oozing from all access sites. manual pressure and new dressings applied. patients activated clotting time still high from cath and aggrastat continued. the patient oozing from sheaths and art line attempted access. manual pressure held on sites and sites redressed. aggrastat to be held. received 2 units of blood. the pump was repositioned at bedside. volume is being given and levo is being weaned. new labs are pending. the device was explanted. | [5 empty] | The investigation was completed and no product was returned. the cause of major bleed was most likely use issue related since the act was high and out of recommended normal range. the device history review was completed and passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5751 | 1220648-2025-31391 | 23039354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039354&pc=OZD | 02/09/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 02/09/2025 | 09/12/2025 |  |  | 11/30/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435819 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during shockwave lithotripsy to the left main was performed on impella support and the placement signal was lost and never returned for the remainder of the case. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. the pump was returned for investigation. data log shows the placement signal not reliable alarm along and all 5s parameter was down which indicate hard failure of the sensor. the root cause of the optical signal issue was most likely damaged sensor after the shockwave device interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5752 | 1220648-2025-31394 | 23039358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039358&pc=OZD | 01/22/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 01/22/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559896 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and during support, the placement signal was not correlating to patient's arterial line. position was confirmed and support continued. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the failure mode "placement signal issue" was changed to "optical signal issue" as issue is in relation to the optical sensor. no product was retuned for analysis. the data logs were reviewed and there were no placement signal alarms. there were positioning alarms for "impella in aorta" and "impella position unknown". clinical details mention position was confirmed. the optical parameters throughout the case were within specification. there were no patient condition related issues noted. the root cause of the optical signal issue cannot be definitively determined as no product was returned for analysis and limited clinical details were provided on patient's condition at time of issue. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5753 | 1220648-2025-46385 | 23039439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039439&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711588 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 86 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted because the patient's vessels were too small for insertion and the pump was unable to pass through the anastomosis/graft. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | Corrections: d9, h6. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. delivery issue: no issues were found with the returned product. the cause of the delivery issue was determined to be patient condition as the patient had vessel anatomy too small for insertion preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
5754 | 1220648-2025-31395 | 23039462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039462&pc=OZD | 06/19/2024 | 10/10/2025 |  | 09/12/2025 |  |  | 06/19/2024 | 09/12/2025 |  |  | 01/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024307838 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during shockwave therapy, the pump lost placement signal; however, support continued without interruption. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data log shows that the placement signal was unreliable, and the trend on the 5s parameter indicates a hard failure of the sensor. according to the clinical details, the placement signal was lost when the shockwave device was used. the root cause of the optical signal issue was most likely damaged sensor due to the shockwave device interaction, based on given clinical details and data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D4: added information for udi. h4: corrected manufacturer date.
5755 | 1220648-2025-46399 | 23039498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039498&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026684997 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male | 64 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted for high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, urine output was less than 30 milliliters per hour and lactate dehydrogenase was slightly elevated. the patient was on p-9 and will continue to monitor for hemolysis. the patient was given one unit of packed red blood cells. a company representative reported that blood was administered because of bleeding from traumatic intubation, not as a result of impella usage. the p-level was turned down to 7 due to the urine becoming more red. the device was explanted and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5756 | 1220648-2025-31396 | 23039532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039532&pc=OZD | 07/20/2023 | 10/10/2025 |  | 09/12/2025 |  |  | 07/20/2023 | 09/12/2025 |  |  | 03/21/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024329654 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | D4: serial number is unknown. this report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. the root cause of the optical signal issue is most likely due to patient movement. after the patient sat up the alarm occurred around 15 hours after the start of the case. all pertinent information available to abiomed has been submitted at this time. | D4: additional device information added. h4: added manufacturer date.
5757 | 1220648-2025-31397 | 23039576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039576&pc=OZD | 11/06/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 11/06/2023 | 09/12/2025 |  |  | 07/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024379647 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant in japan reported a patient was implanted with an impella 5.5 for circulatory support. on day 51 of support, the position waveform stopped appearing and a "placement signal not reliable" alarm occurred. issue with the optical sensor was suspected. the patient was weaned from the impella and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. according to the clinical, the positioning waveform was no longer displayed and a placement signal not reliable alarm was recorded. echo imaging was then used to confirm the pump's position. however, no attempts were made to resolve the issue as the patient was registered on the heart transplant list and was scheduled to be reviewed soon. no product was returned for an investigation. no data logs were returned for analysis. the cause of the optical signal issue was not determined because no product or logs were returned and not enough clinical information was given. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5758 | 1220648-2025-46388 | 23039592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039592&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699764 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced hematuria. proper pump positioning was confirmed via echocardiogram and the medical team declined to further reposition the pump. the plan was to trend haptoglobin and see if urine clears. the next day, the urine continued to be red and the patient with hemolysis. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The product was not returned for review. the cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the hematuria cannot be determined as insufficient clinical details were provided. the pump passed all post sterile inspection checks.
5759 | 1220648-2025-31398 | 23039699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039699&pc=OZD | 03/20/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 03/20/2024 | 09/12/2025 |  |  | 01/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458206 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the complainant reported on going placement signal not reliable alarm. the audio was disable. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. upon review of the data logs, oscillating contrast issue observed with placement signal not reliable alarm relevant to definitive failure pattern suspecting the console. no problem was found with the pump and it is apparent the console was the potential cause of the issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5760 | 1220648-2025-31399 | 23039868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039868&pc=OZD | 01/21/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 01/21/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597947 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 device was implanted for circulatory support. during use, the placement signal disappeared and a ¿controller error¿ alarm was generated. despite this, the motor signal, flow, and purge system continued to function normally. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for investigation. no physical abnormalities were reported for the returned product. the pump passed the laser continuity test for the optical signal. additionally, all the 5s parameters were within specification for the returned product. the data log suggests that some of the 5s optical signal parameters were reported to have a value of 16384, accompanied by a controller failure alarm. there was no issue found during the case. the device functioned/operated to specification. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5761 | 1220648-2025-31400 | 23039999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23039999&pc=OZD | 09/29/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 09/29/2023 | 09/12/2025 |  |  | 07/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374316 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, a change was noted in the placement signal waveform; however, the alarm resolved and waveforms returned to normal. the intermittent ¿impella in aorta¿ alarms occurred but self-resolved. device position was verified by echocardiography and confirmed to be appropriate. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was not returned for investigation. data logs shows the several "impella position in aorta" alarm. there was no variation or impact to optical signal 5s parameters except during the catheter clamp. the root cause of the optical signal issue was most likely patient condition related based on the data log review. all pertinent information available to abiomed has been submitted at this time.
5762 | 1220648-2025-46406 | 23040053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23040053&pc=OZD | 08/15/2025 |  |  | 09/12/2025 |  |  | 08/15/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1272370 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | False Alarm | Failure to Read Input Signal | Application Program Problem | Device Sensing Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered an automated impella controller (aic) issue concerning false alarms. after the team plugged in the impella cable into the aic they got the following alarm: optical sensor defect. they checked the light in the aic like recommended. however, it was on. they plugged the impella cable into the aic again and tried to start the impella. there was no placement signal and there was an alarm: "controller fault". they changed the aic and used another pump. they could not switch off the aic, because the aic said the impella is still connected to the aic, which it was not. the issue was solved via emergency shut down. there was no patient harm. | [5 empty] | The device was returned. the logs confirmed the reported failure mode. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the controller error and loss of placement signal was due to an aic software issue. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A.2 - a.6 are unknown. the automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this is one of two related reports. this report represents the automated impella controller and the other report represents the pump. refer to section h.10 for the related report number.
5763 | 1220648-2025-31401 | 23040055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23040055&pc=OZD | 04/29/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 04/29/2023 | 09/12/2025 |  |  | 12/31/2024 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024305022 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp pump for circulatory support. it was reported that due to ongoing false position alarms, the placement signal was disabled. patient support continued with the implanted impella despite the disabled alarm. the patient later received a heart transplant. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the lt log was reviewed and the position in aorta alarms was confirmed. 5s parameters were steady. no spikes in motor current. some trending in placement signal. the cause of the false position in aorta alarm was patient condition related since the software version was v11.1. all pertinent information available to abiomed has been submitted at this time.
5764 | 1220648-2025-46407 | 23040093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23040093&pc=OZD | 08/15/2025 | 10/06/2025 |  | 09/12/2025 |  |  | 08/15/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026663772 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Backflow | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they suspected a potential for pump stop and retrograde flow during impella cp support. after the impella cp was placed and started, the automated impella controller (aic) alarmed for pump stop and retrograde flow. as a result, the impella cp was replaced and used on another aic. the patient was reported to be in stable condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. pump stop/retrograde flow: data log review showed that upon the first connection of the pump, the signal not reliable (snr) was below the minimum, due to an airgap based on the optical signal metrics. the pump was disconnected and reconnected as reported, however, the user disconnected again quickly at this point while the pump was still being initialized. because the pump was disconnected while the pump was being initialized and the epm/opm were attempting to communicate, it stopped all further communication, including when the pump was reconnected. since optical communication was stopped, snr was never calculated. after the pump was started and run for a period of time on the third connection, the user dropped the pump to p-0 manually and disconnected, but the console never recognized the disconnection because snr was not being calculated, and therefore, it couldn't drop below the minimum value to indicate that a pump was disconnected. 10 seconds later, when the user was attempting to shut down the console to troubleshoot, the "impella stopped. retrograde flow." alarm triggered and the console wouldn't allow the user to shutdown or start a new case since it believed a pump was still connected. the returned pump was investigated and passed electrical testing of the motor cables and all optical signal 5s parameters were within specification. when connecting the pump to a console, there were no issues with the optical signal nor starting, stopping, or disconnecting the pump. based on the clinical details provided and review of data logs, the cause of the "impella stopped. retrograde flow." alarm was determined to be caused by the console. device history review: the device passed all the post sterile inspection checks. | A.2 - a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represents the pump and the other report represents the automated impella controller. refer to section h.10 for the related report number.
5765 | 1220648-2025-31402 | 23040129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23040129&pc=OZD | 07/24/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 07/24/2025 | 09/12/2025 |  |  | 10/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411514 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the complainant reported that a placement signal not reliable alarm occurred. the audio for the alarm was disabled. the implant was reported to be a successful bridge to a open heart transplant for the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product and data logs were returned for investigation. the cause of the optical signal issue was not determined since product and logs were not returned. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5766 | 1220648-2025-31403 | 23040290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23040290&pc=OZD | 02/13/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 02/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1132985 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, placement signal issues occurred. support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs showed console displayed the ¿impella position unknown¿, and the ¿placement signal low - alarm¿ throughout the case. these alarms confirm the reported failure mode. this console was tested. upon running the optical test pump for 22 hours, unable to reproduce the reported failure mode. the root cause of the optical signal issue could not be determined due to the inability to reproduce the issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5767 | 1220648-2025-31404 | 23040319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23040319&pc=OZD | 03/10/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 03/10/2024 | 09/12/2025 |  |  | 10/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418238 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user reported that during the priming process the automated impella controller alarmed multiple times with ¿purge fluid not detected¿. the purge cassette was replaced to resolve the issue. additionally, during impella cp support, the placement signal disappeared. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the priming issue and the optical signal issue has been completed. the priming issue could not be reproduced. upon review of the data on the device it was determined that the automated impella controller is the product with the potential optical signal issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5768 | 1220648-2025-46408 | 23040332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23040332&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662739 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 100 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs). the following day, a hematoma was noted at the left groin (impella leg). information regarding treatment is unknown. the patient continued on impella cp mcs for an additional three days before being bridged to an impella device for escalation in care with a survived outcome after end of total support. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hematoma, the cause of the injury was not determined since product was not returned and insufficient clinical details provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5769 | 1220648-2025-31405 | 23041012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041012&pc=PYX | 08/12/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 08/12/2024 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during use, the placement signal was unexpectedly lost, and a "placement signal unreliable" alarm was generated. no patient harm reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs confirmed placement signal not reliable alarm triggers again but does not resolve for the rest of the case. the console was confirmed to have an intermittently failing lamp in the cases following and testing, though the specific oscillating contrast symptom was not reproduced. the root cause of the unreliable placement signal was a defective optical bench lamp. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5770 | 1220648-2025-46410 | 23041248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041248&pc=OZD | 08/14/2025 | 09/11/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697705 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 97 KG |  | Black Or African American | Death | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp pump was unable to pass through the patient's vasculature. after abortion of an impella 5.5 implant, a new cp was opened and attempted to advanced it but got stuck at same point as the impella 5.5 did so team transitioned cp back to groin right femoral artery (rfa). | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). additional information was received and noted the device was explanted with a survival at explant reflecting "expired." accordingly, the following sections were updated/repopulated as a result of the additional information. b1 adverse event product problem, b2 outcomes attributed, d6b: explant date, h1 type of reportable event, and h6 health clinical/impact coding and h6 medical device problem code. medical safety review. medical safety reviewed the event. the first event describes an acute kidney injury requiring intervention and was not believed to be due to the impella cp pump 1. the acute kidney injury is most likely a result of the patient condition, and the impella device was unlikely a contributing factor. the patient had poor ejection fraction (10%) low cardiac output state requiring impella support. the patient was critically ill and possible delay in considering escalation to the impella 5.5. with this information alone, it is very unlikely the impella device caused the acute kidney injury. however, it cannot be excluded based on the information at hand and is a serious injury. the impella 5.5 inability to be inserted due to calcification in the axillary artery should be reported as a malfunction type of reportable event. the patient was therefore placed on another impella cp via the right femoral artery after failure to delivery through the axillary artery and ecmo. the devices not being able to be inserted through the axillary artery are also most likely the result of the noted patients preexisting calcification which is a contraindication to insertion. after approximately 9 days on ecpella the patient's care was withdrawn. there is not enough information to dissociate the demise out from the impella cp device pump 3. cause of death is unknown. however, although this dissociation cannot be made, due to the duration of support and the fact that care was electively withdrawn, it is unlikely that cp (pump 3) was associated to the patients cause of death. the patient's underlying condition likely contributed most to an outcome of death. related medical device reports: this impella cp device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical devices reports for the other two impella devices. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the deliver issue was determined to be patient condition as the patient had severe calcified axillary artery preventing successful delivery of impella. regarding, hemodynamic instability, the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. correction(s): event description was updated to provide complete details of the reported event.
5771 | 1220648-2025-31406 | 23041311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041311&pc=OZD | 10/18/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 11/26/2024 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200205 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump the team observed optical signal loss. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs confirmed significantly low signal not reliable. the console was returned for investigation and the failure mode was reproduced. running a 5s test on the console, the snr was confirmed to be very low. the root cause of the optical system sensor error and placement signal not reliable alarms was a defective optical bench with low signal not reliable. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5772 | 1220648-2025-31408 | 23041801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041801&pc=OZD | 06/02/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 06/02/2023 | 09/12/2025 |  |  | 03/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329507 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an ongoing placement signal low alarm during impella cp support. the alarm was silenced, and proper positioning was verified via echocardiography. the patient deteriorated and was made do-not-resuscitate. the patient eventually expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed. the cause of the optical signal issue was most likely patient condition related due to several amount of placement signal low alarms observed during the entire case with p-level change. all pertinent information available to abiomed has been submitted at this time.
5773 | 1220648-2025-46412 | 23041955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041955&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720603 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 69 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced a groin hematoma with bleeding. the patient was transfused two units of packed red blood cells and treated with a femstop and aggrastat. ultimately, care was withdrawn and the patient expired. | [5 empty] | A.4 added as it was inadvertently omitted from the initial report that was submitted. b.2 revised date of death as it was entered incorrectly on the initial report that was submitted. d.3 added manufacturer fax number as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter address line 1, initial reporter city, initial reporter state and initial reporter phone number as they were inadvertently omitted from the initial report that was submitted. h.4 revised device manufacture date as it was entered incorrectly on the initial report that was submitted. the investigation has been completed. hematoma/ major bleed : the cause of the injury was not determined since the product was not returned and insufficient clinical details were provided. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5774 | 1220648-2025-46411 | 23041967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041967&pc=OZD | 08/13/2025 | 10/29/2025 |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654478 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 172 KG |  |  | Malfunction | No | Yes | Pumping Stopped | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support (mcs) as a bridge to a left ventricular assist device (lvad). prior, the patient was on veno-arterial extracorporeal membrane oxygenation (va ecmo) and intra-aortic balloon pump (iabp) mcs. once the impella 5.5 was implanted the iabp was removed. the patient was sent to the unit on ecmo and impella mcs. on day nine of impella mcs, the patient jumped out of bed and ripped out intravenous lines and 3-point fixation of the impella. the impella pump stopped alarm started, and the care team was unable to restart the pump. the team tried switching consoles. however, the pump still was unable to start. consequently, the patient was taken to the catheterization lab and another impella device (impella cp) was successfully implanted. the patient continued on impella cp mcs for an additional six days before being successfully bridged to the lvad. | [5 empty] | A.5 and a.6 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The product with data logs was received, and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. as per the clinical details, the patient jumped out of bed and ripped out the ivs and fixation of the impella, which was followed by a pump stop alarm. bleeding was noted on the catheter, white cord, and red plug due to the ivs being ripped out. regarding, clinical symptoms hemorrhage/bleeding, the cause of the injury was not determined due to insufficient clinical details. pertaining to the pump stop and device in wrong position, the following was noted: pump stop: a minor kink was noted on the returned product. no other visual damage was observed on the pump. the pump could not be restarted, and all three motor cable lines were found to be open during the electrical resistance test. the red handle was inspected internally, and all motor cable lines were found to be stretched and fractured inside the red handle on the catheter side. data log shows several pump stop 119 alarms, with motor current spikes reaching approximately 30,000 at the end of the case run. the cause of the pump stop issue was electrical open lines inside the red handle due to unintentional patient movement, based on returned product investigation and provided clinical details. device in wrong position: no positioning issue¿related physical abnormalities were observed on the returned product. data logs showed no alarm related to improper position during case run. the cause of the device in the wrong position issue was due to unintentional patient movement, based on returned product investigation and provided clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5775 | 1220648-2025-46413 | 23041970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041970&pc=OZD | 08/14/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712066 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 70 KG |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented for weaning from cardiopulmonary bypass (cpb) and sternotomy aortic valve replacement with impella 5.5 placement with heparin in saline . the patient was successfully weaned from cpb and impella 5.5 was placed. during support, it was noted that the patient was taken to the operating room for a wash out. found a small vessel bleed. the patient was doing well however, impella was having flow issues due to placement. the pump was repositioned with no change in metrics. the physician was aware the pump was not performing at normal standards however was ok leaving at p 3-4 for duration of support. the impella was successfully weaned and explanted. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. product evaluation: returned pump visually inspected. biomaterial observed around the impeller, biomaterial removed and pump run for 10 mins in the lab at p9 and flow was 6.2 l/min which pre removing the biomaterial pump was unable to run and still some biomaterial left on the top of impeller blades. dry foreign material observed during evaluation looks attached during wrapping the device after use and removed easy. log analysis: many suctions occurred during impella support, multi- motor current (mc) spikes occurred on the date low flow issue reported. no evidence for other issue observed. low or blocked pump flow: the root cause of the impella low and blocked flow was biomaterial ingestion and mc spikes confirmed the root cause. major bleeding: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5776 | 1220648-2025-46414 | 23041983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23041983&pc=OZD | 08/14/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered an automated impella controller (aic) placement signal issue. placement signal was off screen upon start up of cp. eventually placement signal unreliable alarm became present. decision was made to switch to a new aic. new aic seemed to resolve issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This is a duplicate report. the issue has been captured in mfr# 1220648-2025-46427.
5777 | 1220648-2025-46409 | 23042071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042071&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686371 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs). approximately 10 hours after start of support, the patient was experiencing hemolysis and signs of low flow. the team intervened with continuous renal replacement therapy. echocardiogram was completed, and the impella positioning was noted as appropriate. the hemolysis was noted likely from high systemic vascular resistance due to pressors. the situation was monitored as impella mcs continued. three days later, the patient was noted to have "drastically" improved with no vasoactive drips and was down to support level p-5. the patient was successfully weaned and the device was explanted with a survived patient outcome. | [5 empty] | Product complaint # (b)(4). investigation summary- no product was returned. hemolysis, low pump flow: after reviewing the logs, suction alarms along with a trend in purge and placement signal was observed. the cause of hemolysis and low pump flow was most likely patient condition related per the review of logs. device history lot- device lot: 1923195. device history batch- subcomponent lot: n/a. device history review- device (sn: (b)(6) passed all post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
5778 | 1220648-2025-31409 | 23042117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042117&pc=OZD | 08/17/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 08/17/2023 | 09/12/2025 |  |  | 02/28/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337422 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a loss of placement after one day of impella cp support. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed. no problem was found with the pump, and it is apparent the automated impella controller was the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5779 | 1220648-2025-31407 | 23042144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042144&pc=OZD | 06/23/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 06/23/2023 | 09/12/2025 |  |  | 04/30/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337081 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported had a placement signal not reliable alarm while the patient was walking. no known patient harm or injury occurred. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. as per clinical, placement signal went off while patient was walking. the data logs were returned and found that the signal-to-noise ratio dropped down to zero, and contrast was decreasing while there was huge rise in gain to 2.6 with lamp increased to maximum. the cause of the optical signal failure was most likely an air gap from between the automated impella controller and the patient cable plug due to patient movement based on clinical and data log review. all pertinent information available to abiomed has been submitted at this time.
5780 | 1220648-2025-31410 | 23042240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042240&pc=OZD | 11/05/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 07/25/2025 | 09/12/2025 |  |  | 02/28/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319214 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal did not appear on the automated impella controller following impella cp implant. support was able to continue without the noted signal; however, that patient eventually succumbed to their pre-existing complication and passed away. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue was not determined since product and data logs were not returned. all pertinent information available to abiomed has been submitted at this time.
5781 | 1220648-2025-31411 | 23042271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042271&pc=OZD | 02/17/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 02/17/2024 | 09/12/2025 |  |  | 01/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306352 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and there were intermittent placement signal unreliable alarms. position was confirmed and support continued. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5782 | 1220648-2025-31412 | 23042446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042446&pc=OZD | 06/11/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548998 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported upon the automated impella controller (aic) had an alarm for suction/yellow alarm. the aic was due for preventative maintenance. there was no patient involvement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5783 | 1220648-2025-31413 | 23042634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042634&pc=OZD | 11/28/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 11/28/2024 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564666 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the technician mishandled an impella cp pump during preparation and damaged the optical sensor. upon implant, no placement signal was transmitted. good flows and motor current were noted, and the decision was ultimately made to leave the pump in position for patient support. the patient later succumbed to their condition and passed away with no indication of device association. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue was most likely due to a damaged fiber line. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5784 | 1220648-2025-46415 | 23042690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042690&pc=OZD | 08/13/2025 | 10/24/2025 |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026674405 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was found to have a thrombus in the left ventricle. as a result, the device was explanted and replaced with an intra-aortic balloon pump (iabp). patient is stable post explant. | [5 empty] | The investigation for the reported thrombosis and minor bleed has been completed. no device was returned. the root cause of minor bleeding was most likely use issue related since the bleed started occurring after the transfer of patient to intensive care unit. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5785 | 1220648-2025-46416 | 23042695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042695&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026648700 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient placed on impella cp while chest compressions were performed. the patient had no pulsatility and was cannulated for extracorporeal membrane oxygenation. there was suction present at p2 and the left circumflex was ballooned. the second day of support, a large right ventricle thrombus was found. care was withdrawn and the patient ultimately expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: after the review of logs, several suction alarms #14 & #73 were found during log review. the cause of suction was patient related as patient was on ecmo support and suction alarms were found for pump running above p-1 (p-level). device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5786 | 1220648-2025-46384 | 23042705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042705&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1098588 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 91 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella automated impella controller (aic) was not charging despite attempting multiple outlets. the display was showing a battery level low alarm. patient care continued with a replacement aic. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D4 expiration date should of been left blank on the initial report that was submitted. d4 lot number, serial number and primary udi number were updated per investigation information. d6a and d6b should of been left blank on the first follow-up report that was submitted. investigation summary: data analysis: the data logs were consistent with the reported complaint. the case started on 13th aug 2025 and shortly after the automated impella controller (aic) was disconnected from the alternating current (ac) power supply. this disconnection triggered the ac power alarm, followed by a low battery level alarm approximately two hours later. notably, neither alarm was cleared subsequently. additionally, the aic was forcefully rebooted by turning off and then on the battery transporter switch, yet the alarms persisted. device analysis: the console (b)(6) was returned for further investigation, during which the failure mode was successfully reproduced. visual inspection revealed that the internal white (neutral) wire was loose. upon opening the ac plug, it was evident that the wire had broken. however, the clamping marks on the outer sheath indicated that the wire had been cut, consistent with the vendor¿s wiring instructions. it appears that the torque applied when tightening the terminal screws was excessive, which likely contributed to the copper wires breaking more easily. note: the supplier quality team was made aware of the complaint and the investigation findings. the aic was listed as (b)(6) but according to an email the serial number was updated from (b)(6). root cause: the reported issue of aic not charging despite attempting multiple outlets and displaying battery level low alarm was due to the defective ac power cord set (broken white wire). | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5787 | 1220648-2025-31414 | 23042938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23042938&pc=OZD | 01/29/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 01/29/2023 | 09/12/2025 |  |  | 06/30/2024 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234809 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal unreliable and position in the aorta alarms occurring on and off over the weekend for a patient on impella 5.5 support. echocardiography showed good position, and staff elected to disable audio. the device was not replaced or removed, and no console was switched. the patient eventually expired on support unrelated, to the alarms. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal failure was most likely an air gap from between the automated impella controller (aic) and the patient cable plug. the issue resolved when the aic was replaced and the patient cable was reconnected. all pertinent information available to abiomed has been submitted at this time.
5788 | 1220648-2025-31415 | 23043054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23043054&pc=OZD | 01/23/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 01/23/2024 | 09/12/2025 |  |  | 07/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377022 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. about nine days into support, a placement signal not reliable alarm was triggered. both the placement signal and lv waveform were no longer displayed. the alarm was disabled and the staff was aware of how to continue monitoring the pump via other indicators. support was maintained with the implanted pump. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. data logs were returned and investigated. the 5s parameters were found to be incorrect. in the start of the event, snr dropped to 0, the light decreased and at the same time the lamp went to 100. a decrease in contrast was observed with an increase in gain. change of console did not resolve the placement signal issues. these 5s parameter failure trend indicate likely fiber break. therefore, per the data log and clinical information, the root cause of the optical signal issue was most likely a damaged optical fiber line. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5789 | 1220648-2025-31416 | 23043155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23043155&pc=OZD | 04/12/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 04/13/2023 | 09/12/2025 |  |  | 12/31/2024 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024295419 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The placement signal of the pump always was incorrect. hemolysis was detected but did not require any intervention. the pump could only run on p-4 at maximum, due to suction alarms. position was always correct, confirmed with echocardiography and trans-esophageal echocardiography. care was eventually withdrawn, and the patient expired due to multi-organ failure. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the issue was most likely related to the patient¿s aortic valve insufficiency condition. all pertinent information available to abiomed has been submitted at this time.
5790 | 1220648-2025-31417 | 23043157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23043157&pc=OZD | 07/24/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 07/24/2025 | 09/12/2025 |  |  | 08/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024391410 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. the complainant reported that the placement signal became not reliable during patient support. the issue began after a cross clamp was applied to the motor/outflow area during a surgical procedure. upon removing the clamp, the alarm subsided, but the readings remained inaccurate. patient was successfully weaned and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. data logs show that on the day when the pump was put into surgical mode, there is an immediate placement signal (ps) spike to 3000 mmhg with placement signal not reliable alarms. after the procedure, ps returns but at higher values. the root cause of the optical signal issue was most likely a damaged sensor due to clamping near the outflow cage / optical sensor. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5791 | 1220648-2025-31418 | 23043184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23043184&pc=OZD | 07/24/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 07/24/2025 | 09/12/2025 |  |  | 12/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438018 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal unreliable alarm appeared following an impella 5.5 implant. it was noted that the delivery had been difficult due to the patient's anatomy. the alarm was silenced and support with the implanted pump was maintained. the patient was weaned from the pump and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. upon review of data logs, it is apparent that the console ic3614 is the potential cause of the optical signal issue. no problem was found with the pump, and it is apparent the console was the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5792 | 1220648-2025-31419 | 23043188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23043188&pc=OZD | 06/28/2023 | 09/12/2025 |  | 09/12/2025 |  |  | 06/28/2023 | 09/12/2025 |  |  | 09/30/2024 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023263387 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 support encountered a placement signaling alarm. the alarm persisted despite changing the automated impella controller. there was no harm to the patient as a result of this incident, however, the patient eventually expired during support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was sensor silicone wear, which happened beyond the design life of 60 days. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5793 | 1220648-2025-31420 | 23043511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23043511&pc=OZD | 03/19/2024 | 10/27/2025 |  | 09/12/2025 |  |  | 03/19/2024 | 09/12/2025 |  |  | 10/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415926 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. the complainant reported that a placement signal not reliable alarm occurred. the patient trended well and was managed by icu without the placement signal. the patient was successfully weaned from the pump and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. clinical team reported placement signal not working since priming and that ic4510 has had similar problem in the past. upon reviewing the data logs, after about 26 hours of support, there were multiple placement signal not reliable alarms and the 5s parameter analysis showed a rise in lamp to 100, a rise in gain, and a drop in signal not reliable. the data logs of the 2 most recent cases where this console: ic4510 was sued, on (b)(6) 2023, were also evaluated and there were no optical signaling issues for the former case and the latter had placement signal issues that resolved during the case. these cases did not have similar issues as the complaint pump. insufficient product was returned for analysis of this complaint. the product returned had a severed catheter with the cannula/motor not returned, limiting investigation. the returned portion of the device passed the light continuity test, indicating that the issue was likely isolated to the optical fiber/sensor face that was not returned. the 5s parameters from the lab analysis are therefore not accurate representations of optical signaling for this pump. the root cause of the optical signal issue was not determined as sufficient product was not returned. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D6b: corrected explant date. | G3 ¿date received by manufacturer¿ corrected.
5794 | 1220648-2025-46417 | 23044373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044373&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690145 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient implanted with an impella cp encountered access site bleeding, noted by the oozing by the repositioning sheath. as a result, vascular surgeons sutured and closed the bleeding. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The pump and purge cassette were returned for investigation. no issues were found with the returned product. the cause of major bleed was most likely use issue related since hemostasis was achieved by closing the bleed with additional suture. the pump passed all post sterile inspection checks. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5795 | 1220648-2025-46418 | 23044376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044376&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700815 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Unspecified Infection | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was reported that the medical stuff decide to end with therapy, due to poor general condition, no problems with impella reported. | The complainant had a impella 5.5 pump running for the patient admitted in with cardiogenic shock and cardiomyopathy. it was reported that there was huge drop in purge-flow and increasing pressure and started lysis immediately. purge flow increased after lysis and no issues with the pump. later the doctor reported infection/sepsis unknown treatment. it was also reported that there was a known thrombus in the left atrium may caused suction of material which disturbed the purge flow. staff is aware of the problem and will keep an eye on the pump and purge. the patient remains on support. |  |  |  |  | B2. additional information was received. b5. additional information was received. d6b. additional information was received. h6. impact code added due to new information received . the investigation has been completed. purge pressure high: the product was not returned for investigation. data log was not provided or available on the remote server. the cause of the high purge pressure issue could not be determined, as the product was not returned and sufficient clinical information, including data logs, was not provided. clinical symptoms(infection/thrombosis): the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review has been completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5796 | 1220648-2025-46419 | 23044377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044377&pc=PYX | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026687352 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 48 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella rp flex pump placed to support the patient¿s right ventricle and help with recovery from inferior st-segment elevation myocardial infarction. while on support, lower than normal flows were noted when the p-level was decreased. the decision was made to explant the device. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation for the reported event, low or blocked pump flow, was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the returned rp flex pump was visually inspected. biomaterial around impeller observed. no cannula kink or broken blade observed. log analysis showed: mc spikes observed at p6 with shifting placement signal greater than 80mmhg, simultaneously the flow dropped from 3l/m to less than 1l/m. post the mc spike mc shifted upward high to 300ma. the root cause of the low flow issue most likely was biomaterial ingestion.
5797 | 1220648-2025-46390 | 23044382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044382&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026696262 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 failed as the pump was unable to advance past the distal valve struts. several attempts were made and multiple alternative access options were discussed, before the case was ultimately aborted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the root cause of delivery issue is determined to be patient condition because due to calcification. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5798 | 1220648-2025-46421 | 23044383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044383&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692793 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump placed to support the patient during extracorporeal membrane oxygenation weaning. while on support, there were concerns of gastrointestinal bleeding. four units of red blood cells were given and anticoagulation was stopped. additionally, there were purge high/blocked alarms present. the device was explanted due to complication and the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5799 | 1220648-2025-31421 | 23044391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044391&pc=OZD | 04/04/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 04/04/2024 | 09/12/2025 |  |  | 01/31/2025 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025457815 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that a patient was on an impella 5.5 for circulatory support. there were placement signal issues; however, motor current and flows were stable. support continued. the patient was successfully weaned and the device explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for review. upon review of data logs, it is apparent that the console is the potential cause of the issue on 4/04. on 4/02 the pump was stopped, unplugged and re-plugged in. during this time the optical parameters change and placement signal not reliable #1048 occurs. the signal not reliable decreases, lamp is stable, contrast decreases, gain increases and the light is variable. based on these parameters, the root cause of the optical signal is most likely due to an air gap/use issue. all pertinent information available to abiomed has been submitted at this time.
5800 | 1220648-2025-46422 | 23044395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044395&pc=OZD | 08/14/2025 | 10/17/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707343 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the decision for impella 5.5 placement was made as a bridge to left ventricle assisted device. however, during insertion, it was noted that once into brachiocephalic, significant difficulty in getting wire down the ascending aorta due to tortuous vasculature. left ventricle ultimately accessed. decision made to utilize a buddy wire due to difficulty with anatomy. impella primed and backloaded over a .025 plat plus wire. also coated with rotaglide. impella advanced into subclavian, and initially unable to pass suspected calcium. multiple attempts made to advance impella, but unable to pass. impella removed and heavily irrigated. multiple dilators inserted through anatomy to dilated the vessel and suspected area of stenosis. impella coated in rotaglide and attempted insertion. impella made it to the level of the distal edge of the medtronic tavr valve. despite seeing the wire through the center of the valve, impella stuck on distal edge of valve struts. impella removed and catheter noted to be kinked due to attempted insertion, decision made to attempt final time with a second pump. | [5 empty] | The investigation for the reported delivery issue and catheter kink has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of delivery issue and catheter kink most likely was patient condition related due to tortuous vasculature and patient's transcatheter aortic valve replacement valve. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. complaint device not returned for analyzing. clinical states initially unable to deliver pump due to suspected calcification and tortuosity. multiple dilators used, pump coated, and able to make it to distal edge of patient's tavr valve. pump got stuck on distal edge of valve struts. catheter kink observed once pump was removed and second pump attempted. delivery issue: the cause of delivery issue most likely was patient condition related due to tortuous vasculature and patient's tavr valve. catheter kink: the cause of the catheter kink was most likely patient condition related due to tortuous vasculature and patient's tavr valve. device history review was completed and passed all post sterile inspection checks.
5801 | 1220648-2025-46423 | 23044399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044399&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026723093 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female | 63 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during procedure md unable to deploy transcatheter aortic valve replacement( tavr) valve, patient acutely decompensated requiring pacing, intubation, and cpr. impella cp was inserted emergently during cpr via 18 fr., gore tavr sheath. md observed bubbles on transesophageal echocardiogram. md reported no difficulties with priming or inserting impella. no leaking or air observed in impella purge tubing. discussed removing current impella and placing new impella cp through left groin due to ¿bubble¿ and securement concerns since impella was placed through a non-peel away sheath. advised mds proposed method of securement is not within ifu recommendations, mds opted to leave current impella in place. care was withdrawn and the patient expired. the patient was dnr prior to tavr procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Product complaint # (b)(4). investigation summary : impella cp sn (b)(6) + purge tubing returned for investigation. air embolism: no issues were found with the returned product. the cause of the injury was not determined due to insufficient clinical details. device history lot :device lot: 1967164. device history batch : subcomponent lot: n/a. device history review : device (sn: (b)(6)) passed all post sterile inspection checks.
5802 | 1220648-2025-46424 | 23044402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044402&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678477 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support while awaiting a heart transplant. it was reported that on the 9th day of support there was a purge disc failure on the automated impella controller (aic) during a purge cassette exchange. as a result of the issue, the aic was replaced with a back up unit and support continued. there was no harm reported due to the aic malfunction. additionally, on the 34th day of impella support, it was reported that the patient experienced an ischemic stroke which then converted to a hemorrhagic stroke. the patient was noted to have mild left sided weakness. computed tomography was performed, and the stroke was noted to be stable. the physician stated that they did not believe the stroke to be impella related, however it could not be definitively ruled out as a contributing factor. it was noted that the patient was not under anticoagulated at the time of the ischemic stroke event. the following day, the impella 5.5 was replaced with a new impella 5.5. there have been no issues reported with the new pump, and the patient remains on impella 5.5 support awaiting heart transplant. this complaint involves one automated impella device and one impella 5.5 device: automated impella controller with serial number (b)(6) impella 5.5, with serial number (b)(6). this mdr specifically addresses automated impella controller with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | Information was received that the patient was successfully weaned and explanted from impella 5.5 during a successful heart transplant. the patient outcome of procedure was survived. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received in regards to the impella 5.5 used with this automated impella controller on this patient. a separate report was submitted for the impella 5.5. refer to section h.10 for the related report number. | Additional information added to b.5 with patient outcome. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
5803 | 1220648-2025-46425 | 23044404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044404&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656893 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Cardiogenic Shock | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient who was supported with an impella 5.5 developed an infection at the graft site after explantation. the patient was treated with antibiotics and survived. | [5 empty] | No product was returned for investigation. infection/ cardiogenic shock: in order to make a determination of the root cause of the infection/cardiogenic shock was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | Product complaint # (b)(4).
5804 | 1220648-2025-46426 | 23044417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044417&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686482 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 the patient stated some impella site pain and some swelling. the care team monitored and later the patient was weaned successfully. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5805 | 1220648-2025-46427 | 23044420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044420&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when the impella was started it worked perfectly until our flows got up to 3.5. at that point the pressure wave forms went away completely although the flow was running at 3.9 l and the motor current was perfect. the staff dropped the flows to p6 and the pressure wave forms came back. it was decided to trade out the consoles. as soon as the catheter was switched to a new console it worked perfectly giving 3.9 with normal wave forms. the case was completed and the patient was transferred to the unit with normal wave forms issues. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is 1 of the 2 reports where this report is for aic and related report is for pump,
5806 | 1220648-2025-46428 | 23044421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044421&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675680 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. on the second day of support, the patient was noted to have an acute kidney injury (aki). continuous renal replacement therapy was initiated due to the aki. the physician stated that they did not believe the impella was the cause of the aki. the same day, the patient was taken to the operating room to be escalated to an impella 5.5 for increased hemodynamic support. during the impella 5.5 implant, the physician was unable to advance the impella and deliver to the left ventricle due to calcification in the axillary artery. the impella 5.5 was aborted, and the impella cp was attempted in the same axillary artery access. the impella cp was also unsuccessfully delivered to the left ventricle, and the axillary artery access was aborted. the impella cp was then placed successfully into the right femoral artery, support was initiated, and the patient was taken back to their hospital room to recover. additionally, the same day, the patient was cannulated for veno-arterial extracorporeal membrane oxygenation support. the patient continued impella cp support, however it was noted that on (b)(6) 2025 support was withdrawn. status of impella support and patient is unknown at this time. this complaint involves three impella devices: pump 1: impella cp with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). pump 3: impella cp with serial number (b)(6). this mdr specifically addresses impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary. pump sn (b)(6) and purge cassette returned. unable to deliver or advance: no issues were found with the returned product. the cause of the delivery issue was determined to be patient condition as the patient had severe calcified axillary artery preventing successful delivery of impella. device history lot. device lot: 1924526. device history batch. subcomponent lot: n/a. device history review. device (sn: (b)(6)) passed all post sterile inspection checks. | Medical safety reviewed the event. the first event describes an acute kidney injury requiring intervention and was not believed to be due to the impella cp pump 1. the acute kidney injury is most likely a result of the patient condition, and the impella device was unlikely a contributing factor. the patient had poor ejection fraction (10%) low cardiac output state requiring impella support. the patient was critically ill and possible delay in considering escalation to the impella 5.5. with this information alone, it is very unlikely the impella device caused the acute kidney injury. however, it cannot be excluded based on the information at hand and is a serious injury. the impella 5.5 inability to be inserted due to calcification in the axillary artery should be reported as a malfunction type of reportable event. the patient was therefore placed on another impella cp via the right femoral artery after failure to delivery through the axillary artery and ecmo. the devices not being able to be inserted through the axillary artery are also most likely the result of the noted patients preexisting calcification which is a contraindication to insertion. after approximately 9 days on ecpella the patient's care was withdrawn. there is not enough information to dissociate the demise out from the impella cp device pump 3. cause of death is unknown. however, although this dissociation cannot be made, due to the duration of support and the fact that care was electively withdrawn, it is unlikely that cp (pump 3) was associated to the patients cause of death. the patient's underlying condition likely contributed most to an outcome of death. related medical device reports: this impella 5.5 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical devices reports for the other two impella devices in which one is associated with the demise outcome.
5807 | 1220648-2025-46429 | 23044443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044443&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702719 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient presented to the catheterization lab during st-segment elevation myocardial infarction. high left ventricular end-diastolic pressure was observed along with multivessel disease. cardiothoracic surgery was consulted. the patient was intubated, and the impella cp was placed in the right groin. a temporary pacer wire was also utilized for support. the patient was thereafter transferred to the unit in critical condition. upon arrival at the unit, ¿heavy¿ oozing at the insertion site was noted and a stitch was added. suction occurred soon thereafter arrival as well, and 2l of albumin was given resolving the suction. latest hemoglobin levels were low, and four units of packed red blood cells and two units of fresh frozen plasma were administered. the following day, the patient remained in critical condition. however, it was noted hemoglobin had improved since the transfusion, and there was no more bleeding from the insertion site. mcs would continue for one additional day before the patient expired. the patient was without neurological response, and the family decided to withdraw care. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. log analysis: suctions occurred, no motor current (mc) spike observed. at p9 impella low flow alarms seen on same day suction reported followed by impella in aorta alarms, pump suction: the cause of the suction issue most likely was patient related. access site adverse event: the root cause of oozing issue most likely was use related since the physician at bedside added a stitch to site and oozing has since improved. device history review: the device passed all post sterile inspection checks. | Complaint coding was updated to better reflect the reported event. consequently, section h6 health effect clinica/impact codes and medical device problem codes were updated/corrected and repopulated accordingly. note: h6 investigation type/findings/conclusion remain unchanged as previously reported | Patient-specific information a4: weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5808 | 1220648-2025-46430 | 23044468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044468&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669823 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support for bridge to transplant. approximately 62 days after start of the support, placement signal issues were encountered. the placement signal dropped out and triggered multiple suction and impella placement signal low alarms. customer inquired about the alarms tones, which may have been disabled because of the event. the patient would continue on support and 17 days later, the automated impella controller read of pump flow was assumed to be inaccurate. additionally, this day, the impella pump was previously noted to be malpositioned posteriorly on transthoracic electrocardiogram, and the medical team was unable to reposition the pump due to placement monitoring being previously disabled. the patient was noted to be in systemic organ failure. a shock call was placed and decision made to cannulate for venoarterial extracorporeal membrane oxygenation (va ecmo) until transplant. the impella 5.5 device was left in place as vent on support level p-3, with discussions to replace the pump. however, the team noted there was no reasonable access to replace the impella due to the implantable cardioverter defibrillation left in the left axillary space. support continued and the following day the impella support level was increased to p-4. however, there was still no placement monitoring. latest update noted the patient was listed as status 1 for transplant and continues of va ecmo and impella mcs. | [5 empty] | Additional information was received and noted the impella 5.5 device was explanted with a survived patient outcome. section d6b was updated accordingly. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. however. the event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. subcomponent lot review was completed and noted there were other product malfunction complaints in this subcomponent lot related to this failure mode. regarding the cardiovascular insufficiency, due to a lack of clinical details provided, the cause cannot be determined. pertaining to device in wrong position, the clinical details state that the pump was noted to be malpositioned but they were unable to reposition as placement signal monitoring was disabled. veno-arterial extracorporeal membrane oxygenation (va-ecmo) was placed and the impella 5.5 was left in place to vent on p-3. there were unreliable impella metrics with worsening patient status. due to lack of clinical details provided, it is unknown how the pump became malpositioned. pertaining to the optical sensor issue, the clinical details state that on the 14th the placement signal dropped out and triggered positioning and suction alarms. the automated impella controller (aic) was also reading flows of 3.2-4l on p-3, so it was assumed that the placement signal was inaccurate. the data logs showed that about 1323 hours into support the placement signal begins to gradually drift down with a change in snr but stable 5s parameters. it drifts about 25mmhg but then an offset of +25mmhg is applied to try to counteract the drift. an additional offset is applied 23 hours later as it begins to drift again of approximately 15mmhg so another +25mmhg is applied. there was a significant drift down to negative placement signal values and then the ps became unreliable, in which offset was then applied to the sensor again, but it did not correct it. the placement signal remains out for most of the case but comes back displaying negative numbers to the user at multiple points. the root cause of the optical signal issue is due to sensor wear which occurred at 55 days which is within the design life. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the aic. | Patient-specific information a4: weight is unknown. device status: the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related reports: refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller.
5809 | 1220648-2025-46431 | 23044472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044472&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1635221 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | [24 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support for bridge to transplant. approximately 62 days after start of the support, placement signal issues were encountered. the placement signal dropped out and triggered multiple suction and impella placement signal low alarms. customer inquired about the alarms tones, which may have been disabled because of the event. the patient would continue on support and 17 days later, the automated impella controller read of pump flow was assumed to be inaccurate. additionally, this day, the impella pump was previously noted to be malpositioned posteriorly on transthoracic electrocardiogram, and the medical team was unable to reposition the pump due to placement monitoring being previously disabled. the patient was noted to be in systemic organ failure. a shock call was placed and decision made to cannulate for venoarterial extracorporeal membrane oxygenation (va ecmo) until transplant. the impella 5.5 device was left in place as vent on support level p-3, with discussions to replace the pump. however, the team noted there was no reasonable access to replace the impella due to the implantable cardioverter defibrillation left in the left axillary space. support continued and the following day the impella support level was increased to p-4. however, there was still no placement monitoring. latest update noted the patient was listed as status 1 for transplant and continues of va ecmo and impella mcs. | [5 empty] | Additional information was received and noted the impella 5.5 pump connected to the automated impella controller (aic) was explanted with a survived patient outcome. device status: other additional information noted the aic device was received, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related medical device reports: this aic device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5. | Patient-specific information a4: weight is unknown. device status: the automated impella controller device was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed. related reports: refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5 device.
5810 | 1220648-2025-46432 | 23044474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044474&pc=OZD | 08/13/2025 |  |  | 09/12/2025 |  |  | 08/13/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701575 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 143 KG | Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention was preoperatively implanted with an impella cp device for mechanical circulatory support which was noted successfully completed. post procedure the impella was weaned and explanted with a survived outcome. however, there was a ¿small¿ hematoma successfully managed with manual pressure and femostop. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hematoma, the cause was unable to be determined as limited clinical details were provided, and no product was returned for review. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5811 | 1220648-2025-46433 | 23044477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044477&pc=OZD | 08/14/2025 | 10/10/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026723017 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Respiratory Tract Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump supported and during the support of 6 days there was new infection noted, blood cultures were positive and antibiotics were begun. when the pump was successfully weaned and pump was explanted, but at the time of explant. after hemostasis was achieved they saw that there was a lack of flow in the left femoral artery. they had to use a thrombectomy device and stent the right popateal. patient is stable post explant. | [5 empty] | The investigation for the reported ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is 1 of the 2 reports where this report is for pump and related report is for aic.
5812 | 1220648-2025-46434 | 23044479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044479&pc=OZD | 08/14/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026662756 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 213 KG | Non Hispanic | White | Death | Yes | Yes | False Alarm | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  |  | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella cp support, urine was starting to hemolyze. the impella was repositioned earlier as aorta alarms were occurring, but the patient was still having dark urine. the next day, labs were coming back hemolyzed and the impella was deep in the left ventricle. team was aware and were discussing repositioning. additionally, the a-line was not correlating with placement signal. it was further reported that the patient¿s family decided to withdraw care as there was no ability to escalate care or intervene. the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5813 | 1220648-2025-46435 | 23044480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044480&pc=OZD | 08/11/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  | 02/28/2026 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494016 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Long-term outcome and quality indicator (loqi) impella registry reported that a patient experienced hemolysis with a ¿probable¿ relationship with the impella cp device during impella support. the patient was in acute myocardial infarction/cardiogenic shock and was implanted with an impella cp device for mechanical circulatory support during a percutaneous cardiac intervention (pci) following stent placement in the left circumflex coronary artery (lcx). it was noted that the patient was found down, with agonal breathing and brought to the hospital by emergency medical services. after the patient arrived, the patient was found to be in complete heart block with the lcx occluded. after a stent was placed during pci, the patient¿s mixed venous oxygen saturation was 51% and cardiac index was 1.4, and the decision was made to place an impella cp for additional support. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5814 | 1220648-2025-46436 | 23044483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044483&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706517 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had an 8 french distal perfusion cannula implanted to assist with distal flows. later, the patient was successfully weaned from the pump and survived. | [5 empty] | B.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the investigation has been completed. ischemia : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5815 | 1220648-2025-46437 | 23044484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044484&pc=OZD | 08/14/2025 | 10/30/2025 |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026666672 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | Asystole | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient arrived in operating room (or) and after he was transferred onto or table, he became asystole requiring brief cardiopulmonary resuscitation. return of spontaneous circulation achieved immediately. additionally, a large hematoma noted to right neck - old right internal jugular vein s/g was removed this morning in unit secondary to patient being febrile. vascular surgery consulted on right neck. right groin noted with ecchymoses which extends to right flank. all heparin was removed from purge lines and systemic protamine given. post procedure patient had bleeding from rfa. plan to admit to coronary care unit, labs echo, groin mgt. impella was removed and patient was escalated to 5.5. patient survived the procedure. | [5 empty] | The investigation for the reported major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5816 | 1220648-2025-46438 | 23044497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044497&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677832 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was inserted via direct aorta to support coronary artery bypass graft. after the impella 5.5 was put in while closing the skin the patient had multiple episodes of ventricular fibrillation which prompted the team to add the impella rp flex for additional support. later it was reported that the patient had ventricular tachycardia (vt) and was shocked once therefore return of spontaneous circulation achieved. later it was reported that the patient again had episodes of vt. the physican wanted to leave at current support for now and continue to wean pressors as tolerated. the patient continued on impella support until successful weaned and explanted. the patient survived. | [5 empty] | D.10 concomitant product was added. h.10 related report number was added. the investigation for ventricular fibrillation, arrhythmia, tachycardia, cardiac arrest has been completed. the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5817 | 1220648-2025-46439 | 23044499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044499&pc=OZD | 08/14/2025 |  |  | 09/12/2025 |  |  | 08/14/2025 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698825 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Impairment | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. on the second day of support, the patient was noted to have an acute kidney injury (aki). continuous renal replacement therapy was initiated due to the aki. the physician stated that they did not believe the impella was the cause of the aki. the same day, the patient was taken to the operating room to be escalated to an impella 5.5 for increased hemodynamic support. during the impella 5.5 implant, the physician was unable to advance the impella and deliver to the left ventricle due to calcification in the axillary artery. the impella 5.5 was aborted, and the impella cp was attempted in the same axillary artery access. the impella cp was also unsuccessfully delivered to the left ventricle, and the axillary artery access was aborted. the impella cp was then placed successfully into the right femoral artery, support was initiated, and the patient was taken back to their hospital room to recover. additionally, the same day, the patient was cannulated for veno-arterial extracorporeal membrane oxygenation support. the patient continued impella cp support, however it was noted that on (b)(6) 2025 support was withdrawn. status of impella support and patient is unknown at this time. this complaint involves three impella devices: pump 1: impella cp with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). pump 3: impella cp with serial number (b)(6). this mdr specifically addresses impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event. the first event describes an acute kidney injury requiring intervention and was not believed to be due to the impella cp pump 1. the acute kidney injury is most likely a result of the patient condition, and the impella device was unlikely a contributing factor. the patient had poor ejection fraction (10%) low cardiac output state requiring impella support. the patient was critically ill and possible delay in considering escalation to the impella 5.5. with this information alone, it is very unlikely the impella device caused the acute kidney injury. however, it cannot be excluded based on the information at hand and is a serious injury. the impella 5.5 inability to be inserted due to calcification in the axillary artery should be reported as a malfunction type of reportable event. the patient was therefore placed on another impella cp via the right femoral artery after failure to delivery through the axillary artery and ecmo. the devices not being able to be inserted through the axillary artery are also most likely the result of the noted patients preexisting calcification which is a contraindication to insertion. after approximately 9 days on ecpella the patient's care was withdrawn. there is not enough information to dissociate the demise out from the impella cp device pump 3. cause of death is unknown. however, although this dissociation cannot be made, due to the duration of support and the fact that care was electively withdrawn, it is unlikely that cp (pump 3) was associated to the patients cause of death. the patient's underlying condition likely contributed most to an outcome of death. related medical device reports: this impella cp device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical devices reports for the other two impella devices in which one is associated with the demise outcome. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5818 | 1220648-2025-31422 | 23044501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044501&pc=OZD | 07/24/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 07/24/2025 | 09/12/2025 |  |  | 02/28/2026 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467911 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was implanted in a patient for circulatory support. the patient was on support with the impella pump for 3 days, when the optical signal was lost. the pump remained on for support for another 9 days without any reported harm or injury to the patient. the patient was successfully weaned from the pump and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the impella 5.5 pump and data logs were not returned for investigation. the cause of the optical signal issue was not determined since product and data logs were not returned for review. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5819 | 1220648-2025-31423 | 23044504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23044504&pc=OZD | 07/24/2025 | 09/12/2025 |  | 09/12/2025 |  |  | 07/24/2025 | 09/12/2025 |  |  | 02/28/2026 |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474826 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an implanted impella 5.5 pump began to drift coinciding with suction alarms during patient support. the pump was repositioned successfully, but the placement signal continued to drift. the signal was adjusted manually multiple times until a placement signal not reliable alarm was triggered. support with the implanted pump remained ongoing and intervention was not required for the noted issue. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. on day 30 of support, clinical reported placement signal drift. troubleshooting efforts were carried out but did not resolve the issue and placement signal not reliable alarm was triggered. no product was returned. the data logs confirm the sensor wear pattern on 6/9 with decreasing signal not reliable values aligning with clinical details reported. the placement signal went out of range and did not recover for the rest of the procedure. the logs also confirm placement signal not reliable alarm. the root cause of the optical signal issue was most likely due to a sensor wear of the optical sensor. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5820 | 1220648-2025-31424 | 23045673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23045673&pc=OZD | 07/10/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 07/10/2024 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505883 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was implanted in a patient for hemodynamic support. the pump support allowed for a laser atherectomy and shockwave interventions, however placement signal was lost. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs show placement signal spiking to 3000 mmhg with placement signal not reliable alarm coinciding with when the time laser atherectomy is initiated per clinical details, correlating with the sensor break metrics per the optical failure handbook. product not returned. the cause of the optical sensor damage was most likely a damaged sensor per data log review. exact cause of the sensor damage was unable to be definitively determined but there was a potential interaction with laser atherectomy per clinical details. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5821 | 1220648-2025-31425 | 23045674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23045674&pc=OZD | 10/14/2024 | 09/12/2025 |  | 09/12/2025 |  |  | 10/14/2024 | 09/12/2025 |  |  |  |  | 09/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509770 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was implanted in a patient for circulatory support. while on support, the placement signal "not reliable alarm' occurred. the placement signal alarm was disabled. there was no reported harm or injury to the patient. | [5 empty] | D6: explant date are unknown. this mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the returned product was inspected and there were no abnormalities. there was no kink, the 5s parameters were within range, the pump passed the light continuity test. insufficient logs were returned for analysis. the rt logs between (b)(6)2024 and (b)(6)2024 were not returned. the pump was implanted on (b)(6)2024 and the complaint occurred on (b)(6)2024 beyond the 14 day indication for use for the 5.5 pump. the data logs show a durability sensor wear pattern with 5s parameters within range and a slight decrease of contrast and signal not reliable. as the failure occurred beyond the indication for use, the root cause of the optical signal issue was most likely sensor wear. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5822 | 23025307 | 23025307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23025307&pc=OZD | 08/07/2025 | 09/05/2025 | 09/05/2025 | 09/11/2025 | 09/05/2025 | 09/11/2025 |  | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [6 empty] | 40 YR |  | Female |  |  |  | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | An axillary impella cp stopped working after 13 days. it had to be replaced with a femoral impella. manufacturer response for cath cp impella 14f, cath cp impella 14f (per site reporter). abiomed confirmed that the device and console are under review.
5823 | 1220648-2025-46358 | 23025355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23025355&pc=OZD | 08/12/2025 |  |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147694 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. a nurse called in to discuss a yellow controller error alarm that resulted in loss of aorta and left ventricle placement signals. motor current and flow estimate were unaffected. the alarm spontaneously resolved and placement signals were restored. a backup console was advised. no patient harm was reported due to the issue. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. this is one of two reports. this report represents the automated impella controller and the other report represents the pump. refer to section h.10 for the related report number.
5824 | 1220648-2025-46359 | 23025372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23025372&pc=OZD | 08/12/2025 |  |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688317 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. there were intermittent loss of placement signals which spontaneously resolve without change in motor current or flows during impella cp support. there was no harm to the patient reported and support continued until the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of the two reports. this report represents the pump and the other report represents the automated impella controller. refer to section h.10 for the related report number. | The investigation was completed. no product was returned for evaluation. optical signal issue: clinical details noted a "controller error" alarm with the loss of ao and lv placement signals. no change in motor current or pump flows. no hemodynamic changes noted. alarm spontaneously resolved and issue was intermittent throughout remainder of support. data logs were reviewed and at time of "controller error" alarms, placement signal spiked to 3000 and all optical parameters spiked out of spec then resolved to normal values. upon review of returned data logs, it is apparent that the console is the potential cause of the issue. no problem was found with the pump upon review of data logs and clinical details. device history review was completed and passed all post-sterile inspection checks.
5825 | 1220648-2025-46361 | 23025621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23025621&pc=OZD | 08/12/2025 |  |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701040 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A user facility reported that a patient was prepped for placement of an impella 5.5 for tune up coronary artery bypass graft that was scheduled. the physician attempted to advance the impella 5.5 over a .025 wire. the impella 5.5 was unable to deliver past the mid subclavian artery. the catheter kinked while attempting to advance down the subclavian artery. the catheter was removed, and an angiogram was performed. subclavian stenosis was noted on angiogram. a decision was made to perform an angioplasty on the subclavian artery. the impella 5.5 was reinserted with no success as the impella 5.5 was still unable to deliver past mid subclavian. after several attempts to maneuver the catheter, the impella 5.5 shaft kicked. the case was aborted. no patient harm was reported due to the issue. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the root cause of delivery issue is determined to be patient condition. delivery was not success as the impella was still unable to deliver past mid subclavian. material kink/twist: the root cause of the catheter kink was determined most likely to be patient condition as the patient had resistance in axillary artery resulting in buckling/kinking of catheter. device history review: the device passed all the post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
5826 | 1220648-2025-46362 | 23025643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23025643&pc=OZD | 08/12/2025 |  |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710691 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hematoma | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during support of the impella 5.5, the physician decided to go into the right subclavian artery and was found to have a stent measured to be about 1 mm of blood flow via ultrasound. the decision was made to go through the left subclavian artery. once the artery was opened, and graft was sewn on, there was trouble trying to pass the wire through the aortic arch due to severe calcification. wire was eventually passed with difficulty and exchanged to abiomed 0.018¿ wire. impella was back loaded onto the abiomed wire and attempted to advance into the left subclavian and aortic arch. severe resistance was met, and heavy calcifications were noticed. impella was through the graft but kinking in the arch. a decision was made to abort the case due to the severe calcifications and high risk of stroke. case aborted due to severe calcification in the left subclavian artery and aortic arch. there was no reported patient harm. | Additional information was received from the long-term outcome and quality indicator (loqi) impella registry about an adverse event. left pectoralis hematoma was reported with relationship to impella listed as "possibly". loqi stated that patient's lue (left upper extremity) experienced weakness and numbness for several days (exact date is unknown). neuro was consulted and mri showed left pectoralis hematoma along left subclavian vein. team believed this was due to traumatic/compressive left brachial plexopathy from impella insertion. there was no active bleed. however, medical team monitored site and gave total of 4 units of blood to the patient. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received, the following fields were updated based on the information received from the clinical site. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint #: (b)(4). investigation summary: difficult to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had severe calcification in the left subclavian artery and aortic arch preventing successful delivery of impella. device history lot: device lot: n/a. device history batch: sub-component lot: n/a. device history review: the necessary materials were not returned for analysis so the serial number could not be identified to complete a device history review.
5827 | 1220648-2025-46363 | 23025649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23025649&pc=OZD | 08/11/2025 | 10/13/2025 |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709947 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the patient developed hematoma at the surgical site. manual pressure was held, pressure dressing changed, and anticoagulants paused. one unit of blood product was administered. the next day, the access site was bleeding and anticoagulants continued to be held. it was further reported that the patient had some non-sustained runs of ventricular tachycardia. echo performed and no adjustments were made. the patient also had atrial fibrillation with rapid ventricular response (rvr). no known intervention. | It was discovered during head ct small silent strokes in multiple parts of the brain. the patient was successfully bridged to left ventricular assist device (lvad). |  |  |  |  | The investigation for the reported arrhythmia and major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia and major bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 included additional information regarding pump explant and d6b was updated. b5 also included a correction where the mention of stroke was inadvertently omitted on the manufacturer device report. h6 health effect - clinical code e0133 was inadvertently omitted on the manufacturer device report.
5828 | 1220648-2025-46365 | 23026108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23026108&pc=OZD | 08/12/2025 | 10/07/2025 |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698914 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male | 64 KG |  | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during insertion of an impella cp, the cannula kinked at the motor housing. this was visible via fluoroscopy. the cannula was not able to push the device deeper through the artery to the heart. the device was removed and exchanged. no patient harm was reported. | [5 empty] | The investigation of the reported failure mode has been completed. the pump returned. delivery issues: during insertion of the pump, it became kinked at the motor housing and could not advance throughout the artery to the heart. the product was returned for investigation and a kink was observed on the cannula near the motor housing. the cause of the delivery issue was determined to be patient condition as the patient had tortuous vessel conditions preventing successful delivery of the impella. product damage (cannula kink): during insertion of the pump, it became kinked on the cannula at the motor housing and could not advance. the product was returned for investigation and a kink was observed on the cannula near the motor housing. the cause of the cannula kink was most likely an adverse event related to the patient condition, and it occurred when trying to advance the pump through the patients anatomy, and the patient was reported to have tortuous vessel conditions. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
5829 | 1220648-2025-46364 | 23026109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23026109&pc=PYX | 08/08/2025 |  |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1618084 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) had a system self check error during preparation for use. an attempt was made to restart the device, but the message appeared again, so the device was replaced with another aic. | [5 empty] | A2 and a4 are unknown. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The root cause of the ¿system self check failure¿ was determined to be power manager corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
5830 | 1220648-2025-46366 | 23026439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23026439&pc=OZD | 08/13/2025 |  |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628464 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. prior to impella cp insertion, the patient presented to the emergency room complaining of chest pain and an electrocardiogram revealed a st elevation myocardial infarction. left heart catheterization was performed and revealed mitral valve disease and the patient became hypotensive. the patient was then started on norepinephrine and an impella cp was implanted and the patient was transferred to a large hospital in the network for further management and a possible coronary artery bypass graft. it was reported that on the second day of support, the p level was weaned to p4, and following the drop in p level the patient went into atrial fibrillation and the mitral valve prolapsed. the p level was turned back up to p6 and normal sinus rhythm was established. on 9-aug-2025, it was reported that the patient continued to have episodes of atrial fibrillation, and the medical team discussed changing the amiodarone drop to dobutamine in attempt to resolve the arrythmia and if unsuccessful then the patient may need cardioversion. the following day, the patient was taken to the catheterization laboratory for a high-risk percutaneous cardiac intervention due to decompensation and requirements of high dose pressor and inotrope support. it was noted that the patient was cardioverted twice that morning for ventricular tachycardia and once the previous day for atrial fibrillation. after the patient¿s procedure, it was noted that the impella cp had to be repositioned as it was found to be deep. the impella was measured at 5.27 centimeters (cm) in the ventricle and pulled back to 3.53 cm. on 10-aug-2025, it was noted that the patient remained in atrial fibrillation and was non-pulsatile. it was reported later that night, the purge system on the impella cp required a bypass due to a leak. on (b)(6) 2025, the medical team decided to escalate the patient to an impella 5.5. the patient was brought to the operating room and the right axillary artery was prepped for impella 5.5 access. during insertion of the impella 5.5, resistance was met in the axillary artery and the impella 5.5 could not be advanced. the impella was removed, and the axillary artery was shockwaved and a balloon was inserted and inflated to open the vessel. the impella 5.5 was still unable to pass through the artery. the artery was shockwaved again and the impella 5.5 was still unable to be advanced through the patient¿s artery due to the patient's anatomy. the decision was then made to abort the impella 5.5 insertion and to insert a new impella cp into the axillary artery. the impella cp was able to be delivered with no resistance. two days later, the patient experienced ventricular tachycardia and was successfully defibrillated to resolve. later that day, the decision was made to withdraw the patient¿s care. the patient¿s care was withdrawn, and the patient expired. this complaint involves three pumps: pump 1, impella cp with serial number (b)(6). pump 2, impella 5.5 with serial number (b)(6). pump 3, impella cp with serial number (b)(6). this report specifically addresses the third pump. separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | A.4 is unknown. based on the information available, the impella cp cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety review was completed and noted the following: the patient expired while on impella cp support. the patient implanted with an impella cp pump 1 and experienced arrhythmic events requiring support level adjustment, possible change in medication, defibrillation, and pump repositioning. this device's purge system was bypassed due to system leak. the impella cp pump 1 was explanted and an impella 5.5 was unsuccessfully attempted. however, another impella cp pump 2 was successfully placed with no loss of support. the failure to deliver the impella 5.5 is a product malfunction and the first impella cp device event resulted in serious injury and should be reported. while on the second impella cp, the patient continued to have ventricular tachycardia (vt) runs that were successfully defibrillated; however, the patient subsequently expired as care was withdrawn. given the information at hand, there is not enough information to clearly dissociate the demise out from the second impella cp device. however, although this dissociation cannot be made, it is the care that was withdrawn that appears to have led to the patient's demise; vt runs were successfully defibrillated. medical safety officer reviewed the event and noted the same in the second impella cp device cannot be dissociated for the outcome and the first impella cp pump and impella 5.5 is a serious injury and product malfunction respectively type of reportable event. investigation: the product was not return (as it was discarded), and therefore an evaluation of the device was not possible. a device history review was completed and noted the device passed all sterile inspection checks. the investigation was completed and indicated in order to make a root cause determination for the tachycardia; essential clinical details are required. the root cause of the injury was unable to be determined due to insufficient clinical details. related medical device reports: this impella cp device report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device report on the first impella cp and the impella 5.5.
5831 | 1220648-2025-46367 | 23026774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23026774&pc=OZD | 08/13/2025 | 09/10/2025 |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1566737 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported an automated impella controller (aic) issue with the placement signal. a placement signal system error occurred when connecting the pump. the problem was resolved after replacing the aic. the same issue had occurred with the same aic in the past, so team received a request for inspection. | [5 empty] | Section a - no information was available. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5832 | 1220648-2025-46368 | 23026874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23026874&pc=OZD | 08/12/2025 |  |  | 09/11/2025 |  |  | 08/12/2025 | 09/11/2025 |  |  | 08/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393030 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp for mechanical circulatory support. the patient experienced a large hematoma at the insertion site. vascular surgery was working on closing the vessel in the operating room. the impella cp was removed due to the hematoma at the insertion site and patient stability. the patient is stable post explant. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. hematoma: large hematoma ta access site. the root cause of the injury was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. vascular dissection: vascular surgery working on closing vessel in operating room. the root cause of the vascular dissection was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. device history review: device with passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
5833 | 1220648-2025-46369 | 23027053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23027053&pc=OZD | 08/13/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701042 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Multiple Organ Dysfunction Syndrome | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella 5.5 support had decreased in liver function. elevated lactate dehydrogenase (ldh) and alanine aminotransferase (alt). it is unknown if the liver enzyme increase is related to the impella or the liver failure. the team suspected potential hemolysis from the impella. an echocardiogram showed no movement in the impella 5.5 and appropriate position. the impella 5.5 was removed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. mechanical interaction with blood (hemolysis): the pump was not returned for investigation. data logs show several suction alarms and impella position unknown alarm during the day of hemolysis issue and two days prior. there is no confirmed evidence that the pump was in an improper position. no abnormalities were reported in motor current. the placement signal and pump flow showed fluctuating trends while running at a steady p-level. the cause of the hemolysis was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks.
5834 | 1220648-2025-46370 | 23027960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23027960&pc=OZD | 08/13/2025 |  |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691969 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced oozing and limb ischemia requiring intervention. the patient would subsequently expire. the patient was found down at a restaurant and after extensive resuscitation efforts, return of spontaneous circulation was achieved. the patient arrived at the hospital and was emergently taken to the catheterization lab where left main artery disease was noted. they were unable to stent. the physician completed ballooned angioplasty. an impella cp was placed. post procedure, the patient was taken to the unit on impella mcs. the following day, the repositioning sheath site was bleeding and femostop was applied. this helped the oozing. however, pulses were absent with signs of mottling in the foot. an ultrasound was planned and the femostop was adjusted. heparin was still ongoing. pulses were found in the popliteal artery and the next day it was noted there had been no change in the patient¿s extremity. the patient continued on support. urine output was decreased, and continuous renal replacement therapy was started. following the team declined any revascularization option. approximately two days thereafter, a gastrointestinal bleed was noted and system heparin was suspended. it could not be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the outcome of the bleed for the patient. the patient was administered two units of blood product. additionally, it was noted heparin was stopped due to melena in the flexiseal and hemoglobin drop. a discussion was held with the family and care was eventually withdrawn. the patient expired two days later. | [5 empty] | Additional information was received and noted the following: there was no confirmed loss of pulses on the impella site. there were confirmed popliteal pulses ad with improvement in pulsatility there were later distal doppler pulses attainable. the patient who was transported with a femostop in place due to site oozing, and despite the presence of the compression device pulses were available. the gastrointestinal (gi) bleed was believed to be a lower gi bleed. in regard to the device and its availability, the additional information noted it was likely discarded after withdrawal of care. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ischemia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. regarding the major/minor bleeding, the cause was not determined since product was not returned and insufficient clinical details provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5835 | 1220648-2025-46372 | 23028000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028000&pc=OZD | 08/13/2025 | 11/06/2025 |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026694309 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was taken to the cath lab as a code acute myocardial infarction . upon arrival patient was found to have an occlusion of the left circumflex artery. a percutaneous coronary intervention (pci)was performed. impella cp was placed post pci due patient needing further hemodynamic support. during support it was noted that the urine was red tinged, so impella was repositioned. additionally there were notifications and a call for ongoing suction. provider who stated that they are now maxed on multiple inotropes/pressors and stated that the likely reason for suction is the patient's poor cardiac function. the patient also had a patient has a positive disseminated intravascular coagulation panel. there was bleeding orally, rectally and insertion site. the patient was given 2 units of packed red blood cells and 1 unit of cryoablation. heparin drip was stopped. ultimately care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure modes has been completed. major bleed: the cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. pump suction: the cause of the suction issue was not determined due to not product or data logs available for investigation. hematuria: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5836 | 1220648-2025-46373 | 23028167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028167&pc=OZD | 08/13/2025 |  |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705935 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during repositioning, the impella 5.5 blue button was extremely difficult to depress and stay depressed. the doctor commented "there is something wrong with the spring, it was impossible to depress". additionally, stress led to the sterile repositioning sheath to rip in the process. the patient is stable post explant. | [5 empty] | The investigation of the reported failure mode has been completed. product was not returned for investigation, and data logs are not relevant to the reported issue. difficulty positioning: since product was not returned for investigation, and insufficient clinical details were provided, the cause of the difficulty positioning could not be determined. product damage (sterile sleeve): based on the clinical details provided, the sterile sleeve tore due to excessive force being applied by the user. for this reason, the cause of the product damage was determined to be a use issue. device history review: the device passed all the post sterile inspection checks. | A.4 and a.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5837 | 1220648-2025-46376 | 23028482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028482&pc=OZD | 08/13/2025 | 10/07/2025 |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716833 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the impella cp was placed in standard fashion via right femoral and support was initiated without difficulty. percutaneous coronary intervention was attempted but unsuccessful. the patient was transferred to another facility and was scheduled for an emergent coronary artery bypass graft (cabg) and valve replacements; the cp was pulled out of the aortic valve into the ascending aorta and put in surgical mode. the impella was clamped. the cp was reinserted after emergent surgical intervention. when the pump was turned back on the cp would not flow higher than p3 without suction events. the impella under echocardiogram (echo) looked perfect with appropriate volume. the right ventricle looked okay as well. the pump then started alarming pump in the aorta along with suction events, but echo had not changed, and position was appropriate. team decided to then remove impella and put the patient on intra-aortic balloon pump( iabp) support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrected information: e4. the investigation of the reported failure mode has been completed. the pump, cassette, and data stick were returned. low pump flow/suction: data logs were reviewed, and the low flows and suction alarms were confirmed. initially, pump flows were in specification for p-9 while running on the first console but started to trend down after 2 hours. once transferred to the second console, flows were in specification the first 14 minutes of the case. at this point, the ps began to show values 45-280 mmhg, indicating cpr was being done. after this, the pump was increased to p-9, and flows were below specification for the rest of the case. just before the pump was put into surgical mode, there was also a motor current (mc) spike to 1400 ma. suction alarms post-procedure were systolic suction alarms indicating ample ps pulsatility (>20 mmhg) but low mcmax values. returned product had damage to several impeller blades. the pump was run in a bucket of water and low flows did reproduce, with 2.4 l/min running at p-6 (specification: 2.5-2.9 l/min). there was evidence of a minor kink in the cannula, however, it is unlikely to be related to the low flows/suction based on the finding of damaged impeller blades and low flows reproducing. based on the damage to impeller blades on returned product with low flows reproducing, the cause of the low pump flows and suction was determined to be damaged impeller blades. there were several potential causes of the damages impeller blades reported (cpr/cabg/mvr), however, it could not be definitively linked to any one based on the information provided. false alarm: data logs were reviewed and the position in aorta alarm was confirmed, and it was triggering correctly. optical signal metrics were within specification during the case and were particularly strong during the time of the reported alarm. the mc modulation was low while ps pulsatilitity remained ample, which can be indicative of the pump being out of position. however, the pump was also in intermittent suction at the time, which has the potential to affect the reliability of the alarm. additionally, it was confirmed that pump position was good. returned product was investigated and there were no abnormalities noted related to the optical signal. all 5s parameters were in specification and the pump passed laser continuity testing. since pump position was confirmed to be good, data log review was inconclusive, and there were no abnormalities on returned product, the cause of the false alarms could not be determined. device history review: the device passed all the post sterile inspection checks.
5838 | 1220648-2025-46375 | 23028484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028484&pc=OZD | 08/13/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661066 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 65 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump support ongoing for the patient with admission for acute myocardial infarction with cardiogenic shock. during the support, on the second day of support, there was hematuria that prompted assessment of pump position/volume. the pump was weaned and explanted as planned on day 3. patient survived . | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. hematuria: the cause of the hematuria was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5839 | 1220648-2025-46377 | 23028503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028503&pc=OZD | 08/14/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 08/14/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707228 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Inflammation | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a 5.5 pump placed to support the younger (30yo) patient with cardiomyopathy. the patient was placed on the transplant list. on day 6 of the support the pump was seen to have pain at the access site, there was tenderness that was controlled by pain meds and eventual removal of the jp drain. the pump remains on for support, currently on day 33 while awaiting transplant. | [5 empty] | Product complaint # (b)(4). no device received: the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5840 | 1220648-2025-46378 | 23028510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028510&pc=OZD | 08/14/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 08/14/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025606189 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had cp placed a to support the patient admitted in with an myocardial infarction, ventricular tachycardia and in cardiogenic shock. the pump was placed but explanted after 22 minutes due to an access site bleed/ooze and hematoma. despite all measures the bleeding would not resolve after perclose and manual hold. the bleed was persistent and as after explant the patient was hemodynamically stable the cp was not replaced. patient survived. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. hematoma/ major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5841 | 1220648-2025-46379 | 23028524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028524&pc=OZD | 08/14/2025 | 10/30/2025 |  | 09/11/2025 |  |  | 08/14/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705928 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had an infection and frequent ventricular tachycardia. after impella implant the patient encountered continued issues with arrhythmias requiring shocks eight times, increased lidocaine levels, conversion, multiple defibrillations, and repositioning of the impella. additionally, during support the patient had positive sputum culture and was given antibiotics as a result. patient survived. | [5 empty] | The investigation for the reported arrhythmia and positioning issue has been completed. product was not returned for review. the root cause of the arrhythmia and positioning issue was unable to be determined as insufficient clinical details were provided. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5842 | 1220648-2025-46371 | 23028743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028743&pc=OZD | 08/12/2025 | 10/07/2025 |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571192 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that during an impella cp procedure, the pump did not start. the procedure was successfully completed with a new impella cp. there was no harm to the patient reported. | [5 empty] | The investigation of the reported failure mode has been completed. the pump and purge cassette returned. pump did not start: returned pump did not have any damage upon visual inspection. pump was connected to console and was able to be primed and run at p-9 for approximately 5 minutes without issue. no data logs were returned for evaluation. no issues were found with pump during case as no defects were found on returned product. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A.2 - a.6 are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5843 | 1220648-2025-31355 | 23028896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23028896&pc=OZD | 12/11/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 12/11/2024 | 09/11/2025 |  |  | 02/28/2026 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED, INC. |  |  | US |  | 2025562863 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during the impella cp insertion placement signal not reliable alarm occurred. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were not available and the pump was not returned for investigation. clinical details report that the placement signal not reliable alarm was active intermittently from case start. the only time the alarm would resolve was when the pump was running at p-7. it is not reported what devices were used during the pci procedure. due to limited clinical details and product not being returned, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5844 | 1220648-2025-46374 | 23029309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23029309&pc=OZD | 08/11/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an access site bleed during impella cp support with the repositioning sheath in place at the access site. the team had to stich the femoral artery and place a pressure bandage and femostop. hemostasis was accomplished. | [5 empty] | Additional information was received that this report is a duplication of report number 1220648-2025-46417. all information will be reported under report number 1220648-2025-46417. | Section a is unknown. d.4 lot number, serial number, expiration date and primary udi number are unknown. e.1 initial reporter title, first and last name are unknown. e.3 h.4 device manufacture date is unknown. the impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5845 | 1220648-2025-46380 | 23030177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030177&pc=OZD | 08/14/2025 |  |  | 09/11/2025 |  |  | 08/14/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1229940 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Product Quality Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) that power button/housing loose and was pushed into the case/housing. there was no patient involvement. the aic needs to be replaced. | [5 empty] | Section a - no patient was involved. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5846 | 1220648-2025-31356 | 23030452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030452&pc=OZD | 06/13/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 06/13/2024 | 09/11/2025 |  |  | 12/31/2025 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025500553 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a 68-year-old male noted as high-risk coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the impella had multiple failure alarms. there was a placement signal low, placement signal not reliable and position in aorta alarms. all these despite good pulsatibility in all curves. the physician manipulated and shook impella cable, and the problem was resolved. the impella position improved and confirmed via tte. the device was not removed as a result of the issue. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. a purge cassette was returned for evaluation. no pump was returned for evaluation. no visual abnormalities noted on purge cassette. cassette was connected to console and purged for approximately 8 minutes and no leaks were observed. data logs from field were reviewed and at time of placement signal not reliable alarm, all 5s parameters were consistent with an air gap and were able to all fully recover after about 2 hours. clinical details noted that psnr alarms were resolved when physician shook impella cable. the root cause of the purge leak cannot be determined as no imc logs from field on day of purge leak were returned and the purge cassette serial number on day of issue cannot be confirmed. the root cause of the optical signal issue was most likely due to an airgap as it was resolved with manipulation of impella catheter. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5847 | 1220648-2025-31357 | 23030467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030467&pc=OZD | 07/03/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 07/03/2024 | 09/11/2025 |  |  | 11/30/2025 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434934 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during implant procedure, the impella cp experienced a placement signal that was not reliable and sensor error since the beginning of the procedure. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. upon impella cp case start, a "placement signal not reliable" alarm occurred. a sensor error was noted as a result. correct positioning was confirmed by echo. neither the pump nor console were replaced due to the issue. the pump was explanted after 5 hours of support. product was returned and evaluated. the 5s parameters were all in spec and no abnormalities were found with the pump. the "placement signal not reliable" alarm did not reproduce. no logs were returned. the root cause of the optical signal issue was not determined since the issue did not reproduce with the returned product and no logs were returned. impella cp has optical smart assist sensor and any placement signal metrics issue would be addressed as optical signal issue. so, the as investigated fm was changed to optical signal issue. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5848 | 1220648-2025-46386 | 23030600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030600&pc=OZD | 08/13/2025 |  |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1107790 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the automated impella controller (aic) had yellow controller error alarm. motor current waveform visible, motor spinning and pump flows displayed, but placement signal missing during error alarm. spontaneous resolution of alarm/controller error with placement signal visible again. | [5 empty] | Patient's race and ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5849 | 1220648-2025-46383 | 23030684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030684&pc=OZD | 08/15/2025 | 10/24/2025 |  | 09/11/2025 |  |  | 08/15/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706486 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella 5.5 placed to support a patient post coronary artery bypass graft procedure. while on support, the patient was in a rhythm that strayed from the baseline. the patient was in atrial fibrillation/atrial flutter. it was noted that an atrial electrocardiogram was pending. the device was explanted and the patient survived. | [5 empty] | The investigation for the reported arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5850 | 1220648-2025-31358 | 23030717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030717&pc=OZD | 02/14/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 02/14/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584295 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that placement signal was inactive on the same day of impella cp implantation, however the patient was stable. eventually the patient was successfully weaned and the impella was explanted. no patient harm reported. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. device implanted (b)(6) 2025 without issue. after 1 day on support, placement signal reported to be inactive. device was removed after 2 days on support due to pt successfully weaned off. the pump was returned. data logs could not be recovered. the returned pump showed no physical abnormalities and showed nominal 5s parameters. additionally, it passed all testing related to electrical wiring. the cause of the optical signal issue was not determined since returned product showed no abnormalities but field logs were unable to be recovered. lack of information from the field could not determine the cause of the placement signal inactive. changed fm to optical signal issue from placement signal issue as cpc9 only has an optical sensor. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5851 | 1220648-2025-31359 | 23030764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030764&pc=OZD | 01/16/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 01/16/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551668 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the patient jerked, and the placement signal was lost. staff was instructed to monitor pump motor current and flows, as well as the patient¿s hemodynamics, for the remainder of the case. the pump remained operational until weaning and explant. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned product; the pump passed the laser continuity test, and the alarm was reproduced during pump testing. the pump ran as expected. the data log shows a damage sensor trend on the 5s optical parameter. as per the clinical notes, the placement signal was lost after the shockwave, but the pump was kept in place. the cause of the optical signal issue was a damaged sensor, based on returned product and data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5852 | 1220648-2025-31360 | 23030904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030904&pc=OZD | 05/07/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 05/07/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534739 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. the optical pressure sensor of the impella cp was potentially damaged during insertion, resulting in an unreliable placement signal alarm that was initially displayed but later resolved. subsequent waveforms appeared appropriate. despite these issues, the pump remained operational throughout support until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs show that there was low signal not reliable starting the case and the placement signal was very low. placement signal unreliable alarms were triggered but then resolved. the pump was returned and the 5s parameters were within specification. the pump had a high and stable signal not reliable. there was no problem found from the pump side regarding the optical signal issue as the returned product had no issues. please see the (b)(4) against the console for a potential cause into the optical signal issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5853 | 1220648-2025-46387 | 23030972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030972&pc=OZD | 08/13/2025 |  |  | 09/11/2025 |  |  | 08/13/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684898 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Perforation | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient on impella cp support for high-risk percutaneous coronary intervention. during impella explant, two percloses failed, ripping out a piece of the artery. the patient's activated clotting time (act) drifted down to below 150. the medical staff held manual pressure successfully and there were no additional access site complications. the patient survived. | [5 empty] | H6 health effect clinical code e2114 was erroneously entered on the initial manufacturer device report number and updated to the correct code e0511. the investigation was completed and no product was returned. mechanical interaction with tissue: the cause for dissection would be use since perclose failed because perclose failed x2 ripping out a piece of the artery. a device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5854 | 1220648-2025-31361 | 23030973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23030973&pc=PYX | 08/04/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 08/04/2024 | 09/11/2025 |  |  | 10/31/2024 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023261740 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during impella support, the placement signal was lost, and the placement signal alarm proved unreliable. x-ray evaluation confirmed appropriate impella positioning; the alarm was deactivated, and regular position checks with x-ray were recommended. despite these issues, the pump remained operational throughout support until weaning and explant. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product/logs were returned. the root cause of the placement signal issue was unable to be determined since no product/logs were returned and insufficient clinical details were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5855 | 1220648-2025-31362 | 23031241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031241&pc=OZD | 03/28/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 03/28/2024 | 09/11/2025 |  |  | 05/31/2025 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024357393 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the patient experienced a loss of pulsatility and required delivery of a shock. during shock delivery, the placement signal to the impella was lost. optical sensor issue. despite this, the pump remained operational throughout the support period until weaning and explant. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cp pump was not returned. data logs were returned and investigated. contrast dropped, lamp increased and gain dropped. these are the characteristic of damaged optical sensor. also, shockwave was used in conjunction to pump which is known to damage the sensor. the root cause of the optical sensor issue was most likely a damaged sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5856 | 1220648-2025-31363 | 23031404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031404&pc=OZD | 06/05/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 06/05/2024 | 09/11/2025 |  |  | 11/30/2025 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439356 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the placement signal to the pump was lost during shockwave therapy. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. no data logs were returned. the clinical details confirm that the use of shockwave coincided with the loss of placement signal. the root cause of the optical signal issue was most likely a damaged sensor due to shockwave interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5857 | 1220648-2025-31364 | 23031438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031438&pc=OZD | 09/13/2024 | 10/24/2025 |  | 09/11/2025 |  |  | 09/13/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532583 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, optical signal loss occurred following patient cable handling. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | G3: corrected date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were investigated, and the placement signal not reliable alarm was confirmed. returned product was investigated and it was noted that there was a kink in the catheter just outside the kink protector. then, the pump failed the laser continuity test as light leaked at the previously mentioned kink location in the catheter. based on clinical details, data logs, and returned product, the root cause of the optical signal issue was determined to be a damaged fiber in the catheter just outside the kink protector. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D6b: corrected the explant date.
5858 | 1220648-2025-31365 | 23031569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031569&pc=OZD | 07/14/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 07/25/2025 | 09/11/2025 |  |  | 12/31/2025 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443046 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a 5.5 pump support ongoing for a patient escalated up from the cp and ecmo. the 5.5 pump lost the optical signal during month 3 of the support. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump not returned. data logs show psnr alarm on 7/14-7/15 with ps going out of range for 20 hours and then resolves. snr drops to 0, contrast oscillates +/- 3000, light drops, gain drops, lamp is already maxed at 100. the console was switched on 7/20 and there was no issue. upon review of the data logs it is apparent that the console is the potential cause of the issue. please refer to (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5859 | 1220648-2025-31366 | 23031609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031609&pc=OZD | 07/26/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 07/26/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1339213 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and there were placement signal issues. the placement signal was absent with dash marks. the impella support continued, and the placement signal returned. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. console logs from the reported day of event are consistent with complaint and show ¿placement signal not reliable¿ (psnr) condition onset before and after pump start, and later multiple ¿invalid sample due to snr below minimum¿ errors. console was tested in danvers after it arrived for investigation (b)(4), and initially the issue was not reproduced. upon visualizing analog output of the ccd from optical bench, we observed intermittent distortions of the signal (envelope / fringe) with increasing frequency, the longer the bench was operating. when not distorted, measured ¿5s¿ parameters are within specification. because of the intermittent nature of this failure mode, console ic6002 passed functional test on 2025-07-10 (after prior event on 2024-04-12 investigated in (b)(4)) with ¿5s¿ parameters within specification. the cause of the automated impella controller (aic) issue was a defective optical bench. repair: replace optical bench assembly (0042-3015) and perform functional check of the console. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. please review dhr checklist for indications of: "22-11697-1: battery failure alarm due to disengaged battery switch 22-10874-1: purge disk not detected due to broken flag" this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5860 | 1220648-2025-31367 | 23031802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031802&pc=OZD | 08/08/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 08/08/2024 | 09/11/2025 |  |  | 02/28/2026 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471165 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that a patient was on impella 5.5 support. while on support, aorta (ao) placement signal was way off from non-invasive blood pressure (bp) with ao placement signal pressure maps at 40, patient was asymptomatic and noninvasive pressures were within limits. the issue was resolved by recalibrating the ao placement signal with the non-invasive blood pressure. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. failure mode was changed from placement signal issue to optical signal issue since the pump is 5.5 s2 which has an optical signal. pump was not returned for investigation. data logs were investigated. the pump was ran for about 2010 hours which is over the life cycle of the pump and the optical sensor. as per the impella 5.5 dtm 0550-9900. the pump was used on 2 consoles ic1963 and ic5320. snr and contrast was widespread since the start of the case but no placement signal related alarms observed. after 78 days, placement signal issue was reported. data logs show the placement signal drifting downward similar to a sensor wear trend. the ps was later offset and after 8 days of the reported issue, the pump was explanted. as per the impella dtm 0550-9900 the pump has a life cycle of 60 days after which the sensor wear is observed. this is a misuse case since the pump was ran well over that suggested time of use. therefore, the root cause of the optical signal issue was sensor silicone wear due to overuse of the pump. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. refer to dhr checklist for indication of: "the deviation 210026925 was applied to the pump. the deviation calls out that the sensor can be changed at the time if testing in case of sensor failure. this is unrelated to the failure mode." this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5861 | 1220648-2025-31368 | 23031841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031841&pc=OZD | 08/11/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 08/11/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485375 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that a patient was on impella cp support due to occluded left anterior descending artery (lad) an optical sensor alarm was present. the patient was brought to intensive care unit, automated impella controller (aic) was switched, and placement signal was obtained. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was not returned for investigation. data logs were reviewed, and the lt log showed that when plugged into the console, an "optical system error controller" alarm triggered. in reviewing the imc logs, the pump was recognized when plugged in, the automated impella controller (aic) started to sample the optical signal, and an "opm signal invalid" alarm triggered. they tried unplugging the pump and starting over, and the same alarm triggered again. the team then plugged in a new pump (sn (b)(6) into the same aic and the same alarm persisted twice. they start the pump anyways and the "placement signal not reliable alarm" triggers and persists until they can plug the pump into a new aic (ic7734). with the new console, the issue does not occur based on lt logs. additionally, for the pump used prior to the 24-35133-1 pump, the placement signal not reliable alarm triggered midway through the case and never recovered. signal not reliable (snr) dropped to zero, contrast increased, lamp increased, gain and light remained the same. based on review of data logs, it is apparent that the console is the potential cause of the issue. please refer to 24-35133-2 for an investigation into the aic. the failure mode was changed from placement signal issue to optical signal issue because it is a left-sided device, so the only sensor is an optical (not dp). a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5862 | 1220648-2025-31369 | 23031993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23031993&pc=OZD | 09/11/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 09/11/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532561 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after the impella cp was inserted and started on auto the cp was placed on p6. good aortic (ao) waveform and lv waveform present on the automated impella controller (aic). while holding pressure on the femoral artery after removing the peel away sheath, the ao placement signal and lv waveform disappeared. motor current remained pulsatile, flows within acceptable range, and pump was in a good position on fluoroscope. placement signal monitoring was disabled. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump and data logs were not provided for investigation. according to the clinical details, the pump had a good placement signal at the start, but this signal disappeared while the sheath was being removed. the pump was in a good position, and the doctor decided to keep it in place. the placement signal values were derived from optical signal parameters in cp pump; therefore, the failure mode related to the placement signal issue has been changed to an optical signal issue. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5863 | 1220648-2025-31370 | 23032243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23032243&pc=OZD | 12/17/2023 | 09/11/2025 |  | 09/11/2025 |  |  | 12/17/2023 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1402375 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed a optical signal was not working, which was resolved by replacing the controller. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product has not been returned for investigation. the console logs from the reported day of event are consistent with complaint and showed the ¿placement signal not reliable¿ condition occurred. the root cause of the optical signal failure was an air gap, most likely from between the aic and the patient cable plug. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5864 | 1220648-2025-31371 | 23032344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23032344&pc=OZD | 04/13/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 07/16/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1339213 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed optical signal and low flows were observed. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show ¿placement signal not reliable¿ condition and unreliable placement signal was present throughout the case. console was tested and initially the issue was not reproduced. when not distorted, measured ¿5s¿ parameters are within specification. the root cause of the aic issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5865 | 1220648-2025-31372 | 23032465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23032465&pc=OZD | 02/23/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 02/23/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1469546 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed a placement signal unreliable. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show that 8 days after support started, the placement signal suddenly became unreliable. console and its optical system was tested, but the issue was not reproduced. the behavior of opm signals in lt/rt log files, however, is consistent with a failure of optical bench or its lamp assembly. the root cause of the automated impella controller issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5866 | 1220648-2025-31373 | 23032621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23032621&pc=OZD | 07/14/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 07/26/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129855 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed optical signal issues. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show that after a week of operation the optical signal became unreliable, with abnormal characteristics, followed by placement signal not reliable alarm. console was tested, but the issue was not reproduced. the root cause of the aic issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5867 | 1220648-2025-31374 | 23032819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23032819&pc=OZD | 07/05/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 07/05/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080191 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed placement signal unreliable alarms. the audible alarm was turned off. support was continued, and the patient was successfully weaned off the device. an engineer evaluated the logs and determined that this was an aic issue, not a pump related issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show ¿placement signal not reliable¿ alarm appearing 40 minutes after pump start. this pattern is consistent with loss of light on the optical bench detector. console was returned. the device failed ¿5s¿ test due to contamination of optical connector. the root cause of the aic issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5868 | 1220648-2025-31375 | 23033180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23033180&pc=OZD | 08/23/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 08/23/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200770 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during the third month of support, the pump lost its optical signal. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs confirm that oscillating contrast occurred leading to the placement signal not reliable alarm. the consoles internal circuitry was inspected with no observable anomaly. the root cause for the placement signal not reliable alarm caused by oscillating contrast was most likely a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5869 | 1220648-2025-31376 | 23033462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23033462&pc=OZD | 01/09/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 01/09/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200770 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs and console was returned. console logs from the reported day of event are consistent with complaint and show 2 ¿placement signal not reliable¿ alarms. console inspection also revealed slight misalignment of optical bench communication cable, which most likely caused alarms. the root cause of the issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5870 | 1220648-2025-31377 | 23033570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23033570&pc=OZD | 09/08/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 09/08/2024 | 09/11/2025 |  |  | 01/31/2026 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025469102 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal value was excessive and did not match the patient¿s arterial blood pressure. pump support continued without interruption, and there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. neither data logs nor product were returned for investigation. very little information is provided via clinical details that could explain why the placement signal issue could have occurred or if it ever resolved. due to the lack of product return and data logs, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5871 | 1220648-2025-31378 | 23033612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23033612&pc=OZD | 08/25/2023 | 09/11/2025 |  | 09/11/2025 |  |  | 08/25/2023 | 09/11/2025 |  |  | 06/30/2025 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367590 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump displayed a ¿placement signal not reliable¿ alarm. the position alarm was disabled after the pump position was confirmed multiple times. pump support continued without interruption, and there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned for investigation. the log shows consistent 5s parameters. on the second day of support there were false impella position in aorta alarms with impella position confirmed. these alarms are confirmed in the data logs. there are no "placement signal not reliable" alarms. the cause of the optical signal issue and the false "impella position in aorta" alarms are most likely patient condition related. there are no patterns in the 5s to suggest a failure of the optical system. all pertinent information available to abiomed has been submitted at this time.
5872 | 1220648-2025-31379 | 23033801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23033801&pc=PYX | 06/06/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 06/06/2024 | 09/11/2025 |  |  | 01/26/2026 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025461561 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a rp flex and cp bipella support ongoing for a patient admitted in with a st elevation myocardial infarction and cardiomyopathy. the rp flex lost the placement signal; however, the pump continued to remain on to support the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. data logs show that ps starts around 2000+, drops to 0 over period of 5 minutes, and then spikes to 3000 mmhg. rt logs show that raw optical is stable while raw ps is drifting downwards in the 2000s mmhg. mc is stable. flow is displayed as 0 while at p-6. there are psnr epm sensor failure and psnr dp sensor out of range alarms throughout. after 1 hour, the ps is displayed maxed at 3000 mmhg for the rest of the case (14 hours) with no placement signal not reliable (psnr) alarms. upon investigation of the pump, the pump failed signal and voltage electrical testing (too high) but passed the 3 motor phase testing. returned 5s parameters were within specification. the ps issues were reproduced when run in a bucket and when run in the pressurized flow loop. both times psnr alarm was also reproduced. the root cause of the placement signal issue was an open electrical line causing high resistances in voltage and signal. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5873 | 1220648-2025-31380 | 23033818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23033818&pc=OZD | 06/14/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 06/14/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481982 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the optical signal malfunctioned. despite this issue, the pump continued to provide uninterrupted support, and the patient experienced no harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump has not been returned. upon review of field data logs, it is apparent that the console is the potential cause of the optical signal issue. no problem was found with the pump. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5874 | 1220648-2025-31381 | 23033951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23033951&pc=OZD | 03/18/2025 | 10/27/2025 |  | 09/11/2025 |  |  | 03/18/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622191 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. no placement signal was received from the start of implantation, and subsequent placement signal alarms were not reliable. the plug was confirmed to be correctly inserted, and there were no unusual findings during preparation of the pump. despite the signal issue, the pump remained operational throughout support until weaning and explantation. no patient harm was reported. | [5 empty] | D6a: corrected the implant date. d6b: corrected the explant date. | G3: corrected date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product returned for investigation. no data logs returned for analysis. the root cause of the optical signal issue was not determined as no product was returned for investigation and no data logs returned for analysis. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5875 | 1220648-2025-31382 | 23034009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23034009&pc=OZD | 03/22/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 03/22/2025 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025526079 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. the pump experienced a placement signal issue and generated retrograde flow alarms. the retrograde flow alarms were subsequently disabled. despite these issues, the pump remained operational throughout support until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. data log review showed no alarms during this case. they disabled the preventing retrograde flows alarms and did not replace the pump/console due to this issue. the root cause of the optical signal issue was not determined since product was not returned and there was no conclusive evidence from data logs on the failure pattern. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5876 | 1220648-2025-31383 | 23034099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23034099&pc=OZD | 11/20/2024 | 09/11/2025 |  | 09/11/2025 |  |  | 11/20/2024 | 09/11/2025 |  |  |  |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025571036 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during procedure a patient's map was 20/30 mmhg lower than the art line. repositioning the impella was attempted with no change; however, the motor current and flows were optimal. decision was made to proceed with procedure as the patient was well supported during the case. hemostasis was achieved. the pump was weaned and explanted with no reported harm to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the failure mode was changed from placement to optical signal issues because the pump in question only has an optical sensor. data logs were not returned for investigation. returned product was investigated and there were no issues noted. the pump passed the laser continuity test. additionally, the pump was connected to the console and all 5s parameters were within specification. then, the pump was left to run in a bucket overnight and there were no alarms or issues with the optical signal. since data logs were not returned, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5877 | 1220648-2025-31384 | 23034241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23034241&pc=OZD | 01/30/2025 | 09/11/2025 |  | 09/11/2025 |  |  | 01/30/2025 | 09/11/2025 |  |  | 01/31/2025 |  | 09/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299933 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that a patient was implanted with an impella cp for circulatory support. there was a loss of placement signal during the procedure; however, it was reported that the physician was not concerned with loss of placement signal and felt that the flows were adequate to complete the procedure. the patient was successfully weaned and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the failure mode was changed from "placement signal issue", which only applies to failures of the dp sensor, to "optical signal issue" as the impella cp pump only has an optical sensor and not a dp sensor. no product was returned for evaluation. data logs were reviewed and lt log at time of issue shows a ¿placement signal not reliable¿ alarm with a sudden increase in ps to 3000 and remains high for remainder of case. the root cause of the optical signal issue was most likely due to a damaged sensor face due to shockwave interaction. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5878 | 1220648-2025-46314 | 23014183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014183&pc=OZD | 08/11/2025 | 10/29/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703381 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support (mcs) and experienced access site hematoma requiring intervention. the patient was initially placed on an impella cp for mcs (along with veno-arterial extracorporeal membrane oxygenation) for approximately three days before being successfully escalated to an axillary impella 5.5 on the right. while on the impella 5.5, day 10 of support, a hematoma at the right axillary access site was noted. the patient required and received two units of packed red blood cells and cryoprecipitate. the patient had a computed tomography scan confirming a pseudoaneurysm at the axillary site. the patient was thrashing her right arm due to inability to speak with tracheostomy. approximately two days later, the impellla 5.5 was successfully weaned, and the device explanted. the patient survived. the removal was via a surgical cutdown. when the impella 5.5 was removed, clots were observed (but not in the graft) and were removed. the graft looked fine. stimlan beads were used and a surgical stapler to close the site with the graft trimmed and a small amount of the graft was left for now to limit the risk for an infection. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the major bleed, a small hematoma at right axillary site was noted. the cause of the major bleed was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. regarding the thrombosis, in order to make a root cause determination, all of the clinical details outlined would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5879 | 1220648-2025-46315 | 23014351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014351&pc=OZD | 08/11/2025 | 10/24/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668966 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 101 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. significant difficulty when trying to pass the impella 5.5 through the axillary artery was noted. when the pump reached the innominate, it would not advance. multiple efforts were made but were unsuccessful. at this point, due to insertional strain on the axillary artery, the decision was made to abort impella 5.5 insertion and perform vascular repair to axillary artery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported vascular injury has been completed. the root cause of the injury is most likely use related since the patient had a large body habitus. significant resistance met and surgeon was unable to pass through the innominate. due to insertional strain on the axillary artery vascular repair to axillary artery performed.
5880 | 1220648-2025-46316 | 23014412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014412&pc=OZD | 08/11/2025 |  |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548298 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 90 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and the following day post veno-arterial extracorporeal membrane oxygenation cannulation (va ecmo) and upon arrival back to the unit, it was noticed the patient had developed a hematoma at the impella access site. it was noted the patient had been moved several times from and to different beds and was placed in a semi upright position. the decision was made to explant the impella cp. the patient survived the explant and was continued on va ecmo support. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Product complaint#: (b)(4). investigation summary: hematoma: the cause of the injury is most likely use related since patient was moved several times from/to different beds and was sat up semi upright. device history lot: device lot: 1839218. device history batch: sub-component lot: n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, hematoma, the cause of the injury is most likely use related since the patient was moved several times from and to different beds and was positioned semi upright. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5881 | 1220648-2025-46317 | 23014473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014473&pc=OZD | 08/11/2025 | 10/29/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686819 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experience a positioning issue resulting in device explant. the patient underwent stenting to the left anterior descending artery which was complicated by coronary perforation. the patient acutely decompensated and coded shortly after the perforation. a covered stent was utilized to correct the perforation. a small pericardial effusion was noted and protamine was given. the patient remained hemodynamically unstable, and the decision was made to place the patient on an impella cp for mcs. the patient was then planned for and underwent an emergent coronary artery bypass graft (cabg) surgery. after arrival back to the unit post-cabg, a decision was made to pull the impella cp into the descending aorta. overnight, the patient decompensated requiring an increase in pressor support. the surgeon attempted to advance the impella cp back into the left ventricle unsuccessfully. the plan thereafter was to explant the device which was completed with a survived patient outcome. there was no harm to the patient as a result of this event. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the positioning issue was unable to be determined as insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5882 | 1220648-2025-46318 | 23014582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014582&pc=OZD | 08/11/2025 |  |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697737 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced access site bleeding requiring intervention. the patient went into cardiac arrest in a parking lot and cardiopulmonary resuscitation was performed for approximately 20-30 minutes with shocks three times for ventricular fibrillation/ventricular tachycardia. return of spontaneous circulation was achieved. the patient was transferred for higher level of care and emergently transported to the catheterization lab. the patient was on levophed and amiodarone drips and heparin drip was suspended. activated clotting time was 423 seconds. an impella cp was successfully placed pre percutaneous coronary intervention to the distal and mid left anterior descending artery via two stents with post dilatation performed. subsequently, the patient was transferred to the unit on impella mcs. on arrival to the unit, it was noted that the impella site was noted clean, dry and intact. no bleeding or hematoma at this time and pedal pulses were confirmed on doppler. the patient was intubated and sedated and had urine output less than 30 milliliter per hour and was red. bedside echocardiogram was completed for impella position check. the impella cp was repositioned, and tuohy borst was locked and tightened at 90 centimeters. when getting the patient settled and prepared for an a-line/swan-ganz placement, a hematoma and active bleeding at the impella site were noticed. manual pressure was applied and the impella cp was repositioned with new forward sutures. the decision was made to place a femostop to help with hemostasis, which was achieved. knee immobilizer was placed on the left leg and impella mcs continued. the following day, the patient was bridged to an impella 5.5. under fluoroscopy of left femoral artery, the impella cp revealed a very low stick. the physician performed a surgical cutdown and repair of left femoral artery. the incision was closed with sutures and staples. there was no bleeding or hematoma noted and pedal pulses were present. the patient was transferred back to the unit in stable condition for continued monitoring. | [5 empty] | Product complaint # (b)(4) investigation summary: hematoma/ major bleed: the cause of the injury is most likely use related since doctor repositioned impella aoi and secured with new forward sutures. manual pressure was held and a fem-stop placed to maintain hemostasis. device history lot: device lot: 1946918, device history batch: sub-component lot: n/a, device history review: the pump (b)(6) passed all post sterile inspection checks. qn (b)(4) was for the batch 1946918 for missing yellow esd label for sn # (b)(6) which was reworked. this is unrelated to this failure mode. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review noted batch was reworked for a quality notification (missing yellow esd-electrostatic discharge-label) unrelated to the failure mode. the review, however, confirmed the device passed all the post sterile inspection checks. regarding, hematoma/ major bleed, the cause of the injury is most likely use related since doctor repositioned impella and secured with new forward sutures. manual pressure was held and a fem-stop placed to maintain hemostasis. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5883 | 1220648-2025-46285 | 23014637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014637&pc=OZD | 07/12/2025 | 10/30/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696980 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received an impact study (b)(4) regarding a patient who had a right-sided hemothorax with a definite relationship to the impella procedure. a large right-sided complex hemothorax, extensive volume loss in the right lung, large right effusion, small right pneumothorax, and small left effusion with significant anemia the night before 4.9 hemoglobin that required transfusion (packed red blood cells). the patient was taken for re-exploration/washout. two liters of blood was removed from the right chest, and small bleeding site was noted at fatty tissue right chest. nothing else in the right chest was identified as the culprit. there was a drastic improvement in hemodynamics and respiratory status. the patient survived the procedure. | Additional information received from study patient 225-001 of thrombocytopenia. the relationship to pump is listed by abiomed safety as "possible". the pi has not yet entered the relationship. study form states: patient had washout/re-exploration for right sided hemothorax (ae001) on (b)(6) 2025. patient platelets were at 85 afterwashout/re-exploration. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the bleeding was determined to be patient condition because the bleeding was reported at non-impella site. the cause of the injury / anemia was not determined due to insufficient clinical details. device with sn: (b)(6) passed all post sterile inspection checks.
5884 | 23014848 | 23014848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014848&pc=OZD | 08/08/2025 | 09/04/2025 | 09/04/2025 | 09/10/2025 | 09/04/2025 | 09/10/2025 |  | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMART ASSIST |  |  | OZD |  |  | 0550-0008 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No |  |  | Not Available | [6 empty] | User Facility report | No | Risk Manager | [6 empty] | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | No Display/Image | Incomplete or Inadequate Connection | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The display of the impella device stopped showing lv [left ventricle] and ao [aorta] tracing. fiberoptic cable malfunction within device.
5885 | 1220648-2025-46319 | 23014888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23014888&pc=OZD | 08/11/2025 |  |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624430 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 88 KG |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock underwent implantation of an impella cp for mechanical circulatory support. the complainant noted that the impella cp was advanced over the wire. however, it could not be fully deployed into the left ventricle. the pigtail of the impella cp was positioned across the aortic valve but could not be advanced any further. multiple attempts were made, all of which were unsuccessful. consequently, a decision was reached to abort the impella cp attempt and to replace it with an intra-aortic balloon pump. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | Product complaint # (b)(4). investigation summary unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had stenosis in the aorta preventing successful delivery of impella. device history lot device lot : 1888643 device history batch sub-component lot : n/a device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
5886 | 1220648-2025-46321 | 23015639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23015639&pc=OZD | 08/10/2025 | 11/03/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700153 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Pulmonary Edema | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced heparin-induced thrombocytopenia (hit) requiring intervention. the patient arrived at the emergency room and went into cardiac arrest. cardiopulmonary resuscitation was performed with return of spontaneous circulation. some time thereafter, the patient began to decompensate requiring increased levophed drip. the patient was taken to the catheterization lab and prior to percutaneous coronary intervention, the impella cp was implanted. percutaneous transluminal coronary angioplasty to the ostial and proximal left anterior descending artery was performed. the patient was sent to the unit on mcs. approximately three days after start of the support, the patient was heparin-induced thrombocytopenia positive. a switch was made from systemic heparin to argatroban, and the patient was administered two units of packed red blood cells. this sent the patient into a flash pulmonary edema and lasix was started. additionally, on the following day, suction alarms were encountered a support level p-2. on arrival, the patient had been maintaining hemodynamics with the impella at p-2 and suction alarms. albumin was ordered for volume and the physician decided to explant the device which was successfully completed. hemostasis was achieved with manual pressure for 45 minutes and fem-stop was applied to pressure for precaution as the patient was actively coughing and moving. pedal pulses were confirmed per doppler. the access site was clean dry and intact. no bleeding or hematoma was noted. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. suction: clinical states suction at p-2 on last day of support. volume was given and the decision to explant the pump was made. data logs showed intermittent suction alarms triggered throughout the entire case. there were no motor current spikes outside of p-level changes or positioning issues observed. the cause of the suction issue was unable to be determined as limited clinical details were provided and no product was returned for review. thrombocytopenia/pulmonary edema: the patient was heparin-induced thrombocytopenia positive and went into pulmonary edema during support. the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5887 | 1220648-2025-46322 | 23015761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23015761&pc=OZD | 08/12/2025 |  |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710643 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 66 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella cp there was concern for malposition of the impella cp due to inability to flow above p3-p4 without getting suction alarms. the impella cp was pulled back and attempted to reposition multiple times with the same result. on fluroscopy, the impella cp inflow appeared close to the mitral valve. repositioning was attempted. however, flow issues did not resolve. the decision was made to return the patient to the cardiovascular intensive care unit for formal echocardiogram to better assess and reposition if needed. the patient received blood. oozing at the site was noticed the next day and redressed with angle matching. weaning of the impella cp may occur later. the impella cp was turned to p2. the left ventricle function improved and the patient was taken to the operating room for explant and cutdown repair of artery. right ventricle function was still poor and required support. hemodynamics were documented due to being inaccurate related to the proteck flow. the pump was explanted and the patient survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. device in wrong position: the cause of the positioning issue was determined to be an adverse event related to the anatomy of the patient due to the patient's small lv. low or blocked pump flow: data log review showed many suction alarms observed in first 24 hours of impella support. no mc spike, no placement signal trend observed. positioning alarms observed. majority of suction alarms resolve by 8/13, which is after clinical states blood was given and pump was sitting mid ventricular. the cause of the suction issue was determined to be an adverse event related to the anatomy of the patient due to small lv leading to the pump having to be run shallow and issue improving with volume. peripheral artery dissection: minor bleed: oozing at access site on 2nd day of support, redressed with angle matching. the cause of the minor bleed was not determined due to lack of clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. h6 medical device problem codes were updated | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
5888 | 1220648-2025-46325 | 23016029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016029&pc=OZD | 08/12/2025 | 10/07/2025 |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710522 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 67 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered placement signal issues. during the case, shockwave was used and placement signal went up to 300 systolic and diastolic went negative. the placement signal went to dashmarks for a few seconds and then back to having diastolic placement signal that went negative. there was a lot of artifact shown on the placement signal. this resolved at the end of the case when the impella was dropped to p2 for weaning and explant. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the device and data logs were returned for investigation. the data logs show that about an hour and 29 minutes into support, the placement signal (ps) spikes to 200mmhg and signal not reliable (snr) drops to 0, lamp steps up, contrast decreases, and gain and light are relatively stable. this continues to occur for about 43 minutes and then ps stabilizes and the 5s parameters improve. the product was returned and the 5s parameters were within specification. the sensor face did not have any artifacts on it and the diaphragm was not fractured. laser continuity did not reveal any damages to the fiber line. no psnr alarms were produced. the root cause of the optical signal issue is most likely due to an interaction with another device based on the clinical details and data log/returned product analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5889 | 1220648-2025-31333 | 23016039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016039&pc=OZD | 08/21/2023 | 09/10/2025 |  | 09/10/2025 |  |  | 08/21/2023 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090363 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed a placement signal unreliable alarm with loss of placement signal. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs verified from the day of the event, align with the complaint. the logs show that after three days of support, the optical signal became unreliable, and a "placement signal not reliable" alarm appeared. the console was sent back to field service for further investigation. the reported complaint could not be reproduced. the optical system of the console was tested and confirmed to be functioning properly. the root cause of the unreliable placement signal was due to a intermittent malfunction of the optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5890 | 1220648-2025-46327 | 23016224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016224&pc=OZD | 08/12/2025 |  |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702636 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 89 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Great Vessel Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day three of impella cp support vascular injury occurred requiring a fasciotomy. the patient survived. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned for investigation. vessel perforation (peripheral): the root cause of the vascular damage/vessel perforation was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. ischemia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1936149. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5891 | 1220648-2025-31334 | 23016259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016259&pc=OZD | 09/13/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 09/13/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517001 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the placement signal not reliable alarms. despite this issue, the pump continued to provide uninterrupted support, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned. the alarms were confirmed on the 8th and 9th days of support. there were no mentions of other devices being used at this time that could have affected the placement signal. the cause of the shift in placement signal is unknown. the root cause of the optical signal issue was not determined since the pump was not returned. upon review of the logs, it is apparent that the console is the potential partial cause of the issue. please refer to 24-36806-2 for an investigation into the console. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5892 | 1220648-2025-31335 | 23016520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016520&pc=OZD | 09/30/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 09/30/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539596 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. on day 9 there was a pump in aorta alarm. despite this issue, the pump continued to provide uninterrupted support, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for analysis. the data logs show impella position unknown alarms and impella position in aorta alarms soon after implant before placement signal monitoring was disabled. there were no motor current spikes. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5893 | 1220648-2025-46328 | 23016545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016545&pc=OZD | 08/12/2025 | 10/30/2025 |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673723 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 83 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a physician felt that there were vascular complications with a patient on impella cp support. the patient was brought back to the catheterization lab to insert a new impella cp via the left groin and had to have vascular surgery to remove previous impella cp in the right groin. additional information was received. the interventional cardiologist suspected possible vascular complications and as a precaution consulted the vascular surgery team. the patient was taken to an operating room for further evaluation. upon surgical exploration, no vascular complications were identified. the vascular surgeon proceeded to remove the impella device from the right femoral groin. prior to this removal, a new impella device was successfully placed in the left femoral groin. | [5 empty] | The investigation for the reported vessel perforation (peripheral) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vessel perforation (peripheral) could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5894 | 1220648-2025-31336 | 23016562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016562&pc=OZD | 10/16/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 10/16/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025523311 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. on day 27, the pump lost its optical signal. despite this issue, the pump continued to provide uninterrupted support, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data log were not provided for this case. as per the clinical, placement signal not reliable was observed during case run. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5895 | 1220648-2025-31337 | 23016694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016694&pc=OZD | 11/02/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 11/02/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509842 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the aortic placement signal began to drift, producing readings that did not correlate with the patient¿s actual blood pressure and triggering suction alarms. normal pump position was confirmed via echocardiography. the patient remained awake, alert, and asymptomatic. sensor drift was confirmed 71 days into the pump run, and the aortic placement signal was re-zeroed using an external reference. following this adjustment, the false suction alarms resolved and the displayed metrics again correlated with the patient¿s clinical presentation. no reported harm to patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs show that throughout the case, placement signal and flows are noisy. the rest of the 5s parameters are within specification. clinical states re-zeroing sensor resolved the issue. the cause of the optical signal issues was most likely patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5896 | 1220648-2025-46329 | 23016699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23016699&pc=OZD | 08/12/2025 |  |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712059 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 61 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed for the elderly patient admitted in with cardiac history and ventricular fibrillation arrest. the impella cp was placed but despite all troubleshooting and repositioning, there was suction. additionally, there was a hematoma on the non-impella limb. the patient got red blood cells (rbcs) infused for the dropping hemoglobin and hematocrit and the rbcs could have benefitted the suction as well as the access site. after 74 hours, care was withdrawn and patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. the complaint device not returned for investigation. suction and low flow: data log reviewed, many suction alarms occurred during impella support. no motor current (mc) spike, no placement signal issue observed. a few purge spikes seen. flows were in minimum side of specified range to the last 8hrs where flow was below 1l/min at p2. no positioning issue seen. the root cause of reported issue cannot be determined since the complaint device not returned for analyzing and no indication observed from log review. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5897 | 1220648-2025-31338 | 23017069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017069&pc=OZD | 11/07/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 11/07/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025534803 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a 5.5 pump support ongoing for a younger (27y0) patient admitted in with cardiomyopathy, pulmonary edema, and hypoxia. the pump and ecmo were placed to support but after 2 days the pump lost placement signal. the optical signal loss did not cause any harm or injury nor pump replacement. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the product was not returned. during the time of the placement signal not reliable alarms , there was a slight spike in motor current, spike in placement signal, decrease in flow and slight increase in purge pressure. the 5s parameters were within specification. the root cause of the optical signal issue is most likely due to biomaterial interaction on the sensor face. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5898 | 1220648-2025-31339 | 23017178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017178&pc=OZD | 11/29/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 11/29/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563622 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, placement signal not reliable alarm occurred. audio was disabled. no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the pump was not returned for investigation. the data log shows the placement signal was lost on 11/29, with all optical signal parameters indicating a hard failure trend related the damaged optical signal fiber line. clinical also confirms that a large accordion kink was noted on the catheter at the red plug. the cause of the optical signal issue was most likely damaged optical fiber line, based on given clinical details and data log review. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5899 | 1220648-2025-31340 | 23017339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017339&pc=OZD | 12/14/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 12/14/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580860 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Eu complainant reported the patient was on impella 5.5 support and during support, the pump lost placement signal. the placement signal and left ventricle curve disappeared. motor current was pulsatile and the pump position was confirmed. the pump flows were good so support continued. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. no product was returned for analysis. the console number was not provided therefore data logs for this case could not be obtained. the root cause of the optical signal issue was not determined as no product or data logs were returned for analysis. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5900 | 1220648-2025-46333 | 23017353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017353&pc=OZD | 08/11/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685126 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female | 75 KG |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after the impella cp was inserted and started, there was an alarm saying preventing retrograde flow- preventing pump to start. the pump was unplugged and replugged- which did not work. aic was restarted and pump unplugged again and re-plugged and the pump restarted, waveforms separated, however not suctioning. after adjusting left ventricle signal, the pump began to run normally, with waveforms. the procedure was successfully completed and the impella was weaned and explanted. patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the pump was not returned for investigation. data logs show the pump stopped with a 13-high motor current alarm after being placed into the patient¿s body. no other abnormalities were confirmed in the data log. the pump was successfully restarted based on subsequent log evidence. according to clinical reports, the pump was able to restart after restarting the console. the cause of the pump stop issue was not determined, as the pump was not returned for investigation and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks.
5901 | 1220648-2025-31341 | 23017700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017700&pc=OZD | 12/23/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 12/23/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569294 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 support, a placement signal not reliable message appeared. positioning and functionality of the pump were verified and support with the pump remained uninterrupted. there was no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. data logs were reviewed, and as reported, the placement signal not reliable alarm triggered on 12/23. at this time, (dsnr) dropped to zero, contrast began to barcode, lamp remained maxed out at 100%, light dropped, and gain became variable. none of the optical metrics or placement signal returned for the remainder of the case. as reported, when the pump was connected to a new console. these are all indications of an optical fiber break, however, product was not returned for investigation. based on data log review, the root cause of the optical signal issue was determined to most likely be a damaged optical fiber but could not be determined. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5902 | 1220648-2025-46335 | 23017704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017704&pc=OZD | 08/22/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/22/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1590813 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 73 KG |  |  | Malfunction | No | Yes | Degraded | Material Separation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after device removal and powering off console, it was noticed that insulation around plug is starting to show wires within. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The data logs are not relevant to this complaint. however no issues were observed while going through the imc logs. device analysis: the console ic9046 was returned to fs for further investigation. during visual inspection, it was observed that the wire covering or outer sheath was improperly cut, as per the vendor wiring instructions. also, the light clamping marks were visible on the outer sheath, which likely contributed to the wire becoming loose. lastly, heavy clamp marks were seen on internal wire root cause: the reported issue of the internal wires were seen and outer sheath becoming loose was determined to be caused by defective ac power cord set.
5903 | 1220648-2025-46336 | 23017716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017716&pc=OZD | 08/11/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1402264 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female | 75 KG |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that after the impella cp was inserted and started, there was an alarm saying preventing retrograde flow- preventing pump to start. the pump was unplugged and replugged- which did not work. aic was restarted and pump unplugged again and re-plugged and the pump restarted, waveforms separated, however not suctioning. after adjusting left ventricle signal, the pump began to run normally, with waveforms. the procedure was successfully completed and the impella was weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5904 | 1220648-2025-31342 | 23017726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017726&pc=PYX | 05/01/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 05/01/2024 | 09/10/2025 |  |  | 01/31/2026 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025458236 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed a rp flex for support along with a 5.5 pump. the rv was distended and pressures were elevated. the rp was placed but, during the support the placement signal was not reliable. the loss of optical signal did not prompt pump explant or any further intervention. no report of patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. no product was returned for evaluation. data logs were reviewed and lt logs showed a placement signal drift with mean pump flow dropping to 0 when "psnr (dp signal out of range)" alarm occurred. 4 hours into support, placement signal starts to drift, then 8 hours into support, placement signal not reliable alarm was triggered and placement signal goes out of spec to 300mmhg for approximately 2 hours with disabled flow calculations. placement signal and flow return for the duration of the case. motor current and p-level were constant throughout alarm. the root cause of the placement signal issue was due to sensor drift. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5905 | 1220648-2025-46337 | 23017787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017787&pc=OZD | 08/11/2025 | 12/24/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713250 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | Hemodynamic instability | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that the decision for impella cp placement on a patient was due to hemodynamic instability and lethal arrhythmias. the cp was successfully inserted and started. during support it was reported of a small hematoma with notable bleeding around repositioning sheath. manual pressure was applied. patient was severely hypertensive in lab prior to leaving. patient sedated on continuous propofol infusion and given one dose of versed. sandbag applied above site to maintain manual pressure. additionally, point-of-care ultrasound showed impella shallow, so was advanced. in addition to the patient receiving 250mls each albumin and lactated ringer's for suctioning and sequentially dropped to p4. the impella was successfully weaned and explanted. patient survived | [5 empty] | The investigation for the reported major bleed and pump suction has been completed. overnight suction was present. after reviewing the logs, suction alarms along with a trend in purge and placement signal was observed. the root cause of pump suction was most likely patient condition related per the review of logs. the root cause of bleeding could not be determined as insufficient clinical details were provided. | Correction: h6 health effect - clinical code e2343 and medical device problem code a24 were inadvertnaly omitted. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of bleeding could not be determined as insufficient clinical details were provided. pump suction: overnight suction was present. after reviewing the logs, suction alarms along with a trend in purge and placement signal was observed. the cause of pump suction was most likely patient condition related per the review of logs. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5906 | 1220648-2025-31343 | 23017876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017876&pc=PYX | 06/20/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 06/20/2024 | 09/10/2025 |  |  | 01/31/2026 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025513777 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that impella rp had placement signal not reliable alarm. the audio was disabled. no report of patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the pump was not returned for investigation. data logs show several 1051 placement signal alarms after a downward trend in the placement signal. the cvp 300 was active during the event, and the placement signal returned to the normal range after 7 hours. failure indicates dp sensor drift. no issues were noted with the optical signal parameters. the cause of the placement signal issue was sensor drift, based on data log review. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. please refer to dhr checklist for indications of: "the qn #: (b)(4) associated with this device lot #1783817 was iso/pe issue which is unrelated to this failure mode." and "the qn #: (b)(4) associated with pressure sensor parts lot #2024320712 was related to a scratches issue, which was corrected with a 100% visual inspection." this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5907 | 1220648-2025-46339 | 23017885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017885&pc=OZD | 08/11/2025 | 10/21/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026707974 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. he cp was prepped, inserted and initiated on auto and they proceeded with percutaneous coronary intervention ( pci) successfully. during cp support the nurse reported ¿impella position in aorta¿. echocardiogram showed inlet above aorta valve. decision was made to take the patient to the cath lab. upon arrival, additionally it was noted that the patient had red urine, oozing at groin site and significant hematoma development at insertion site. decision made to remove impella, hemodynamics stable. patient is stable post explant. | [5 empty] | The investigation for the reported positioning issue, hematoma, and hematuria has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue, hematoma, and hematuria could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5908 | 1220648-2025-46341 | 23017888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017888&pc=OZD | 08/12/2025 |  |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302710 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella aic had a controller failure alarm. the aic was changed to a new aic. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown.
5909 | 1220648-2025-46340 | 23017921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23017921&pc=OZD | 08/12/2025 | 09/25/2025 |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676991 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male | 83 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient escalated from an impella cp to an impella 5.5 pump as a bridge to recovery due to cardiomyopathy. while on support, a ¿placement signal unreliable¿ alarm was noted with patient movement. it resolved when the patient settled. furthermore, there was a plan to remove the impella, but the patient coded. therefore, the impella was left in place. at this time, the patient is still on support. | [5 empty] | Correction to h6 impact code from f26 to f12. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data logs were reviewed and the complaint was confirmed. the cause of the cardiac arrest cannot be determined as insufficient clinical details were provided. the console is the potential cause of the optical sensor issue. the device passed all post-sterile inspection checks.
5910 | 1220648-2025-46332 | 23018003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018003&pc=OZD | 08/11/2025 |  |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624406 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Decrease in Pressure | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | Renal Impairment | [10 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. continuous renal replacement therapy (crrt) was started early this morning with evidence of blood in the collecting bags. the patient had hemolyzed labs. the performance was adjusted from p9 to p8. a swan-ganz pulmonary artery catheter (pac) or a simple central venous pressure (cvp) line were in the plans for hemodynamic support. the impella placement signal was linear and consistent with arterial lines and only produced a flow of 3.0-3.1 liters per minute. the patients condition worsened resulting in further acidosis and severe biventricular failure. considerations for escalation were not entertained as there was no revascularization strategy. care was withdrawn and the patient expired. based on the information available, the impella cp cannot be excluded as a possible contributing factor to the patient¿s outcome. | [5 empty] | H6: added appliable codes to health effect-impact code, health effect-clinical code, medical device problem code. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. hemolysis: clinical details report bloody collection in continuous renal replacement therapy (crrt) bag the morning of case start, which was confirmed by hemolyzed labs. the pump was noted to be positioned close to the posterior wall and a papillary, but no obstructions were noted at the inflow cage. no blood products were administered. it was suspected that the performance level was simply too high, per follow up with the field representative. no further details were provided, and the patient passed away the following morning. the only data logs available were from the first console (b)(4), which was replaced due to purge related issues during case start. therefore, no data logs were available for the time that the patient was on support and experienced hemolysis. product was not returned for investigation. since neither product nor data logs were available for investigation and insufficient clinical details were provided, the cause of the hemolysis could not be determined. device history review: pump passed all post-sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this is one of two reports. this report represent the pump and the other report represents the purge cassette. refer to section h.10 for the related report number.
5911 | 1220648-2025-46344 | 23018361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018361&pc=OZD | 08/21/2025 |  |  | 09/10/2025 |  |  | 08/21/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient presented to the hospital with acute decompensated heart failure with reduced ejection fraction in the process of an open-heart transplant work-up and was brought to the operating room for an impella 5.5 placement. it was reported on (b)(6) 2025 during impella 5.5 support there were alarms for suction due to the patient¿s low volume and the patient was treated with intravenous albumin to resolve the issue. additionally, on (b)(6) 2025 the nurse reported there was intermittent brief disappearances of the placement signal with no alarm present on the automated impella controller (aic). as a result of the issue, the nurse swapped the aic for another console and resolved the issue. the impella device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. this complaint involves one automated impella controller and one impella 5.5 device: impella 5.5 with serial number (b)(6). automated impella controller with serial number (b)(6). this mdr specifically addresses automated impella controller with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
5912 | 1220648-2025-46346 | 23018683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018683&pc=OZD | 08/23/2025 |  |  | 09/10/2025 |  |  | 08/23/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1079575 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered automated impella controller (aic) failed system self-check issue. a patient was transferred from another facility to another facility on impella support. after the console was returned the console was turned on and the failed system check alert appeared on the screen. the console was turned off and turned on again with the same alert appearing on the screen. no patient was involved in this event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
5913 | 1220648-2025-31344 | 23018761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018761&pc=OZD | 11/13/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 11/13/2024 | 09/10/2025 |  |  | 12/31/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450978 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the device experienced intermittent spiked in placement signal readings, which would then self-resolve. the placement signal became inconsistent with the patient¿s actual arterial pressure. despite these issues, the pump continued to provide uninterrupted support, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. complete data logs not available for this case. the cause of the optical signal issue was not determined since product and logs were not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5914 | 1220648-2025-46272 | 23018776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018776&pc=OZD | 08/09/2025 |  |  | 09/10/2025 |  |  | 08/09/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676796 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient diagnosed with postcardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos) was implanted with an impella cp to provide mechanical circulatory support. during the period of impella cp support, the patient experienced bleeding at the access site and episodes of torsades. it was noted that the patient had several occurrences of torsades, which were resolved through cardioversion after the second shock. furthermore, amiodarone was administered to reduce ectopy, and an echocardiogram was performed to verify the placement of the pump. the following day, the patient exhibited oozing and swelling at the site of the impella insertion. the access site appeared puffy and swollen, with no hematomas observed. laboratory tests were conducted to check hemoglobin levels, and the pump was removed the next day. unfortunately, care was withdrawn and the patient ultimately passed away. | [5 empty] | The medical safety officer reviewed the event and the impella cp device cannot be dissociated from the demise outcome. the event describes a delivery issue with the impella 5.5 due to the patient anatomy in which the case was aborted; switch was made to the impella cp which was successfully placed. the delivery issue is a product malfunction and should be reported. additionally, regarding the impella cp, the event describes runs of torsades and the patient was cardioverted out, addressed with medication. the next day the patient coded again with return of spontaneous circulation. however, the patient would eventually expire as the family withdrew care. there is not enough information to dissociate the impella cp present at the time of the patient's demise. however, the patient's underlying condition likely contributed most to the demise outcome, and the family decided to withdraw support which appears to have led to the patient's demise. related medical device reports: this impella cp report is one of two devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5. | Product complaint # (b)(4). investigation summary : no product was returned for investigation. major bleed: oozing was noted at the insertion site. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. arrhythmia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot : device lot: 1934621. device history batch : subcomponent lot: n/a. device history review : device with sn: (b)(6) passed all post sterile inspection checks. | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5915 | 1220648-2025-46330 | 23018784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018784&pc=OZD | 08/08/2025 | 09/09/2025 |  | 09/10/2025 |  |  | 08/08/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709962 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 113 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella 5.5 pump electively placed to support the patient during transfer. while on support, persistent suction alarms occurred. the impella¿s position was verified by bedside ultrasound and an albumin bolus was given. additionally, the impella 5.5 was deep in the mitral valve. the pump was unable to be positioned appropriately and ultimately came out through the aortic valve which the patient did not tolerate. the patient crashed into bypass. the pump was explanted due to the complication. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. suction: based on the clinical description and data logs review, the root cause of the suction is most likely patient related as volume resolved the alarms. unable to reposition: due to lack of clinical information provided, it is unknown how the pump became deep and the reason it was unable to be appropriately positioned again, so the root cause of the unable to be repositioned cannot definitely be determined. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5916 | 1220648-2025-31345 | 23018872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018872&pc=OZD | 07/29/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 07/29/2024 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507081 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that a patient was on impella 5.5 support to bridge recovery. while on support, optical sensor not reliable alarm present, no kinks on white cable noted. flows/motor current wave forms were present. no further information provided. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. data logs at time of alarm were reviewed and contrast was found to be oscillating, signal not reliable dropped to 0, light and gain both decreased, and lamp remained at 100%. the reported failure mode for this pump was "placement signal issue", however, as the impella 5.5 does not have a dp sensor, the failure was investigated as "optical signal issue". upon review of the data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. no problem was found with the pump. the device functioned/operated to specification. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5917 | 1220648-2025-46125 | 23018890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23018890&pc=OZD | 08/02/2025 | 10/13/2025 |  | 09/10/2025 |  |  | 08/03/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687342 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 84 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella 5.5 support was oozy from the access site. two days later, the patient was taken to the operation room for a washout at the impella site due to a large hematoma because the wound vac clotted off. the patient is stable. | [5 empty] | The investigation for the reported access site bleeding has been completed. the device nor sufficient clinical information was returned for evaluation. the cause of access site bleeding was most likely patient condition related since patient was oozing from multiple sites. | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5918 | 1220648-2025-46345 | 23019156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23019156&pc=OZD | 08/19/2025 |  |  | 09/10/2025 |  |  | 08/19/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658381 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient presented to the hospital with acute decompensated heart failure with reduced ejection fraction in the process of an open-heart transplant work-up and was brought to the operating room for an impella 5.5 placement. it was reported on 19-aug-2025 during impella 5.5 support there were alarms for suction due to the patient¿s low volume and the patient was treated with intravenous albumin to resolve the issue. additionally, on 21-aug-2025 the nurse reported there was intermittent brief disappearances of the placement signal with no alarm present on the automated impella controller (aic). as a result of the issue, the nurse swapped the aic for another console and resolved the issue. the impella device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. this complaint involves one automated impella controller and one impella 5.5 device: impella 5.5 with serial number (b)(6). automated impella controller with serial number (b)(6). this mdr specifically addresses impella 5.5 with serial number (b)(6) which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: no product was returned. pump suction - after 42 days on support, pt experienced occasional suction alarms that resolved with giving albumin iv. data log review showed suction alarms across multiple days. suction alarms turn off after p-levels are reduced. pump position noted to be ok. the cause of the pump suction is patient condition related to a lack of volume that resolved when given fluids and reducing the p-level. device history lot: device lot: 1917208. device history batch: subcomponent lot: n/a. device history review: (b)(4) was opened for device lot 1917208 for label legibility issues related to s/n (b)(6). this pump was reworked with no impact to pump s/n (b)(6), therefore, pump s/n (b)(6) passed all post sterile inspection checks.
5919 | 1220648-2025-31346 | 23019293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23019293&pc=OZD | 05/25/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 05/16/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132981 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump implanted for circulatory support. during support, the console displayed a placement signal failure, which resolved following a purge cassette change. the pump remained functional throughout therapy. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console logs are consistent with complaint and show that approximately 12 days after pump start, console presented with a placement signal not reliable alarm, that self-resolved after 1.5 hours. console was tested in danvers, but the issue of light loss was not reproduced. the root cause of the console issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5920 | 1220648-2025-31348 | 23019442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23019442&pc=OZD | 03/08/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 03/08/2024 | 09/10/2025 |  |  | 11/30/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430013 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, placement signal issues were observed, with the placement signal reading significantly lower than the arterial line. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cp pump was returned for investigation. visual inspection found no major abnormalities. optical bench 5s parameters were verified and observed to be in normal range. no alarms reproduced. light continuity test was done and passed. pump was connected to console and ran for 35 to 40 mins and no major abnormalities observed with the pump optical 5s parameters. data logs were reviewed, and placement signal readings observed to be lower with declining trend post 1 hour of pump start. placement signal low alarms are observed with no major impact to optical 5s parameters. no other issues were found. the cause of the optical signal issue was most likely patient condition related based on the return product analysis and log analysis. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5921 | 1220648-2025-46347 | 23019599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23019599&pc=OZD | 07/24/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622086 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was unable to be transitioned to a left ventricular assist device and therefore, was continued on the impella 5.5 even after the expected usage period has expired. there were multiple placement signal not reliable alarms. optical signal was lost and not recovered. the patient survived. | [5 empty] | Product complaint # (b)(4). investigation summary: optical sensor issue: the pump was not returned for investigation. data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed on 24-jul. at this time, psnr drops to zero, lamp maxes out at 100%, light decreases, gain decreases, and contrast oscillates between +/-3,000. the pump was remained in use during psnr. optical signal parameters and placement signal was seen fully returned after 35 days of case run. upon review of data logs, it is apparent that the console is the potential cause of the issue. device history review: the pump sn574965 passed all post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5922 | 1220648-2025-31349 | 23019676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23019676&pc=OZD | 02/09/2024 | 10/27/2025 |  | 09/10/2025 |  |  | 02/09/2024 | 09/10/2025 |  |  | 02/28/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335169 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that high risk percutaneous coronary intervention procedure was being performed with impella support when sensor on impella went out and placement signal was lost. motor current and flows were stable, so the physician continued with the intervention. the hrpci procedure was completed successfully with impella support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no pump was returned for evaluation. clinical details noted that shockwave device was in use at the time placement signal stopped working. motor current and pump flow remained consistent. the data logs were reviewed and at the time shockwave device was said to be in use, it was found that the snr dropped to zero and raw optical signal dropping into the -3000s with the lamp increasing to 100%. additionally, the contrast stepped down in 15% with the placement signal going to 3000mmhg. these are all characteristic 5s parameters of a sensor break due to interaction with a shockwave device. the root cause of the optical signal issues is due to a sensor break as a result of interaction with shockwave device. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D6b: corrected the explant date. | G3 ¿date received by manufacturer¿ corrected.
5923 | 1220648-2025-46323 | 23019686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23019686&pc=OZD | 08/11/2025 | 10/07/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED INC. |  |  | US |  | 1398838 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump stopped and was unable to be restarted successfully. the pump was explanted. there was no patient injury reported. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. there are two associated devices for the reported event. the pump is addressed under manufacturer report number: 1220648-2025-46289. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: the real time (rt) logs confirmed that the pump stops kept occurring while the pump was attempting to be run on this console. device analysis: the failure mode was not reproduced during testing with a known good engineering pump on the complaint console, despite running the pump in the same manner as the clinical situation. root cause: the cause of the pump stop was not related to the console. device history review: the device passed all post sterile inspection checks.
5924 | 1220648-2025-31347 | 23019927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23019927&pc=OZD | 11/18/2023 | 09/10/2025 |  | 09/10/2025 |  |  | 11/18/2023 | 09/10/2025 |  |  | 08/31/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392964 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the placement signal of the pump became unreliable. despite the signal issue, the pump remained operational throughout support until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was not returned for investigation. data logs were investigated and no alarms were observed. "placement signal low" alarm was observed along with suction alarms but no motor current spikes. the placement signal was seen trending downwards. 5s parameters of the pump were stable. these conditions are seen when the patient condition is not optimal. therefore, the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5925 | 1220648-2025-31350 | 23020112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23020112&pc=OZD | 03/11/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 03/11/2024 | 09/10/2025 |  |  | 07/31/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440326 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that while doing groin dressing, a team member noticed that the placement signal had disappeared from the automated impella controller (aic). after a few minutes the placement signal returned. patient stable throughout. no changes to the plan of care were made | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. upon review of the data logs and pump, there was no problem observed with the return product and logs indicate definitive automated impella controller (aic) failure pattern suspecting the console. no problem was found with the pump, and it is apparent the console ic8030 was the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5926 | 1220648-2025-31351 | 23020366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23020366&pc=OZD | 03/14/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 07/25/2025 | 09/10/2025 |  |  | 06/30/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431061 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal not reliable alarm occurred. all connections were secure. patient on impella cp support was hemodynamics stable and motor current is unchanged. the placement monitoring was disable and audio was turned off. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data log review shows that the placement signal (ps/psnr) goes out of spec to 3000 mmhg with psnr and ps low alarms at the same time that light drops, snr drops to 0, contrasts increases in altitude, gain increases, and lamp goes to 100%. no product was returned for investigation. the root cause of the optical signal was most likely a damaged optical fiber line, however, the cause could not be determined. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5927 | 1220648-2025-31352 | 23020468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23020468&pc=OZD | 03/17/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 07/25/2025 | 09/10/2025 |  |  | 05/31/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358121 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support for 60-year-old male patient, there was a permanent placement signal low alarm. a transthoracic echocardiogram was done and the pump was confirmed to be in good position. there was no reported patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella cp was returned for evaluation. upon visual inspection, a kink was observed near the 45 cm mark on the pump catheter. a laser light continuity test was performed and no fiber breaks were observed along the catheter. the 5s parameters were obtained and all parameters were found to be within range. pump was plugged into automated impella controller (aic) and "placement signal low" alarm did not reproduce. data logs were reviewed and at time when "placement signal low" alarms first were triggered, placement signal drops down from a max/min of 64 mmhg/32 mmhg to a max/min of 54 mmhg/27 mmhg. at the same time, motor current and motor speed both lose some pulsatility. clinical details noted that pump position was checked with tee and no obstructions were found near inflow or outflow cage. the root cause of the optical issue triggering "placement signal low" alarms was most likely due to patient condition related. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5928 | 1220648-2025-31353 | 23020508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23020508&pc=OZD | 04/26/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 04/26/2024 | 09/10/2025 |  |  | 11/30/2025 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434312 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal kept alarming despite correction position on echocardiogram was reviewed. the placement signal was disabled. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. data logs show low pump flow while running on p8, followed by saturated pump flow and several pump position alarms in the aorta. there were no motor current spikes or suction alarms reported during the case run. the doctor disabled the placement signal monitoring after verifying the good pump position. the cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5929 | 1220648-2025-31354 | 23020560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23020560&pc=OZD | 05/21/2024 | 09/10/2025 |  | 09/10/2025 |  |  | 05/21/2024 | 09/10/2025 |  |  | 02/28/2026 |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481249 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp was implanted successfully but upon start up placement signal not reliable alarm occurred. pci procedure was completed with no issues. impella flows and motor current were present and normal for patient support. patient¿s act values remained therapeutic throughout the pci procedure. impella weaned down and explanted without incident. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. data logs show a 'placement signal not reliable' alarm after the pump started and persisted throughout the case. the optical signal not reliable (psnr) values dropped to zero, and the contrast value decreased to 7. gain, light, and lamp parameters remained constant during the psnr issue. according to the clinical details, the doctor witnessed the opening and insertion of the impella with no abuse or noticeable damage occurring during the opening or handling of the pump prior to insertion. the cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5930 | 1220648-2025-46342 | 23022597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022597&pc=OZD | 08/12/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698914 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male | 64 KG |  | White | Injury | Yes | No | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during support of the impella cp on (b)(6) 2025, during insertion of impella device cannula kinked at the motor housing (visible via fluoroscopy) and it could not push the device deeper through the artery to the heart. the impella was removed and peel away sheath, exchange and use of new device. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5931 | 1220648-2025-46349 | 23022601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022601&pc=OZD | 08/13/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/13/2025 | 09/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697971 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that ic5115 did not recognize the cp sn (b)(6). the pump was unplugged and plugged in once again without the aic recognizing the pump. the staff inspected the prongs inside of the impella plug and no damage was noted. as a result, a loaner console ic3767 was brought into the room and purge cassette and tubing were transferred to the loaner console. impella was successfully plugged in and ic3767 recognized pump sn (b)(6). the doctor proceeded with implanting the impella cp and when the pump was started, suction alarms occurred immediately. impella was repositioned without any success. flow was changed from auto to p-2, however, suction alarms remained. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. this file represents the pump. refer to section h.10 for the related report number 1220648-2025-46348, which represents the controller. | Product complaint # (b)(4). investigation summary: impella cp sn 609305 + purge cassette was returned for investigation. pump not detected: clinical details state that when pump sn 609305 was connected to ic5116, it could not be recognized. no damage was observed. the pump was then connected to ic3767 without issue. the console was later switched back to ic5116 due to suction alarms on ic3767 and the pump was able to be connected successfully to ic5116. data log review showed the first attempt to connect to ic5116 to be unsuccessful after the console informed the user to connect the pump and there was no activity. the pump was then connected to ic3767 without issue. the pump was transferred back to ic5116 due to suction issues and connected successfully. the pump was returned and evaluated. it was able to be connected to the lab console without issue. no defect was found with the pump. upon review of the returned pump and imc logs, it is apparent that the console is the potential cause of the issue. please refer to the investigation of ic5116 under pi-17574359159696458. suction: clinical details stated that when the pump was started on ic3767, suction alarms occurred immediately. impella was repositioned without any success. flow was changed from auto to p-2 with suction alarms. the pump was removed and no clot was noted. pump was flushed, connected to ic5116, re-inserted, and run without suction alarms. the pump was returned and evaluated. there was no evidence of biomaterial. evidence of a cannula kink was found. the pump was run in a bucket and pump flow was in range. data log review showed a pulsatile suction alarm that started shortly after pump startup and remained active throughout the whole case along with low pump flow. when the pump was running on ic3767, the max mc was lower than expected, which correctly triggered the suction alarm. there were no abnormalities with the placement signal and no mc spikes. there were no suction alarms present in the data logs from (b)(4) and the pump appeared to run without issue as noted by the clinical. the cause of the suction was likely a cannula kink as seen on the returned product along with no signs of biomaterial or placement signal abnormalities in the data logs. the cause of the cannula kink is unknown. device history lot: device lot: 1945882. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all post-sterile inspection checks.
5932 | 1220648-2025-46348 | 23022602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022602&pc=OZD | 08/13/2025 | 01/30/2026 |  | 09/10/2025 |  |  | 08/13/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288947 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 72 KG | Non Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that they encountered an automated impella controller (aic) issue regarding a pump not being detected. it was reported that ic5115 did not recognize the cp sn (b)(6). the pump was unplugged and plugged in once again without the aic recognizing the pump. the staff inspected the prongs inside of the impella plug and no damage was noted. as a result, a loaner console ic3767 was brought into the room and purge cassette and tubing were transferred to the loaner console. impella was successfully plugged in and ic3767 recognized pump sn (b)(6). | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. this file represents the introducer. refer to section h.10 for the related report number 1220648-2025-46349, which represents the pump. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5933 | 1220648-2025-46351 | 23022604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022604&pc=OZD | 08/13/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/13/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438016 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) placement signal issue. the hospital received a transfer patient already on impella cp. when staff completed aic to aic swap to (b)(6), the placement signal and lv waveforms were drastically higher than the previous aic¿s. the staff retrieved a different aic and performed another aic-to-aic swap, and these numbers were back to baseline. for this reason, local biomed is requesting (b)(6) be evaluated. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown.
5934 | 1220648-2025-46354 | 23022614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022614&pc=OZD | 08/08/2025 |  |  | 09/10/2025 |  |  | 08/08/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708747 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp access the pre-close failed. it was noted that the patient had a substantial panus to navigate. they successfully managed by placing a long sheet from the right radial artery and a balloon occlusion was attempted while trying to deploy percloses. percutaneous coronary intervention was successfully completed and the impella was weaned. however during explant, hemostasis was achieved by balloon occlusion, manual pressure and a covered stent was required. | [5 empty] | No product was returned for investigation. major bleed: the cause of the bleeding was determined to be patient condition because of patients large panus and difficult to control the bleeding and to get an effective pre-close as deployed. device with sn: (b)(6) passed all post sterile inspection checks. | Patient's race is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5935 | 1220648-2025-46355 | 23022795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022795&pc=OZD | 08/11/2025 |  |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658508 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 79 KG |  |  | Injury | Yes | Yes | False Alarm | Inaccurate Flow Rate | Device Sensing Problem | Patient Device Interaction Problem |  |  |  | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support (mcs) and experienced signal issues, high pump flow and hemolysis suspected from thrombus. the device was eventually explanted and replaced. the patient was admitted and received left ventricular assist device (lvad) workup again to assess candidacy. the impella 5.5 was therefore placed as a bridge to a lvad. the impella 5.5 was slowly ramped up to support level p-8 achieving flows of 5.2 liters per minute. the patient was sent to the unit on impella mcs. approximately six days after start of the impella mcs, an impella position in aorta alarms was experienced and determined to be false as the position was verified under echocardiogram. subsequently, there was a sudden motor current spike and left ventricular waveforms climbed higher than placement signal waveforms. the impella 5.5 flows also exceeded normal values. additionally, the patient developed hemolysis this day and the physician suspects it was due to thrombus. a new impella device was requested. the patient was successfully weaned and the impella 5.5 was explanted and replaced. the patient would continue with impella 5.5 mcs until successful bridge to a lvad approximately 28 days later. | [5 empty] | Product complaint # (b)(4). investigation summary- no product was returned. saturated flow: clinical details state that impella flows exceeded normal values after a motor current spike. data log review showed saturated flows immediately after a motor current spike. there was another smaller mc spike about 11 hours later with a step down in pump flow after. there were also suction alarms occurring after the large mc spike. placement signal also stepped down from ~80 to ~60 mmhg at the time of the large mc spike. the cause of the saturated flow was biomaterial ingestion as data logs showed a motor current spike followed by saturated flow. hemolysis: clinical details state that the patient developed hemolysis after they observed a motor current spike and high flows. dark urine was reported. after replacing the pump, yellow urine was reported. data log review confirmed the mc spike following by saturated flows and suction alarms. there was also a drop in ps at the time of the mc spike. there were impella in aorta alarms after the mc spike but good position was verified with echo. the cause of the hemolysis was biomaterial ingestion since data logs showed a motor current spike followed by saturated flow and there were no more reports of hemolysis after pump was replaced. false alarm/buzzer: clinical details state that there were false impella in aorta alarms overnight. position was verified with echo. data log review showed several impella in aorta alarms beginning immediately after the large mc spike. these alarms were consistent until the smaller mc spike about 11 hours later. placement signal also stepped down from ~80 to ~60 mmhg at the time of the large mc spike and increased at the time of the smaller mc spike. the cause of the false alarm was biomaterial ingestion since data logs showed a motor current spike followed by impella in aorta alarms, which resolved later after another smaller mc spike. device history lot- device lot: 1917232. device history batch- subcomponent lot: n/a. device history review- the pump sn (b)(6) passed all post-sterile inspection checks. | A.5 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
5936 | 1220648-2025-46356 | 23022879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022879&pc=OZD | 08/11/2025 | 10/29/2025 |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693716 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Pulmonary Edema | [12 empty] | Required Intervention | Death |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient who underwent coronary artery bypass, valves replacement, left atrial appendage and maze surgeries was placed on an impella 5.5 electively before coming off bypass. the impella 5.5 was inserted through the axillary sheath and visualized crossing the aortic valve via small aortotomy. central venous pressure reading was 2 mmhg. volume was administered in the form of products and intravenous fluid. the impella 5.5 flows improved as volume was infused. the chest was closed, and the patient had remained in the operating room to briefly assess chest tube outputs. during this time, the patient had a brief episode of flash pulmonary edema resulting in partial pressure of oxygen reading in the 50¿s. however, this was corrected prior to the patient leaving the operating room. the patient was eventually transferred to the unit on impella mechanical circulatory support (mcs). the following day, the nurse reported a ¿purge system blocked¿ alarm and consequently, changed the purge fluid out for tissue plasminogen activator. impella mcs was continued. two days later, a bronchoscopy was performed, and it was noted that the team was able to get a significant amount of clot out. the patient was transferred to another facility and evaluated for extracorporeal membrane oxygenation (ecmo) in which the patient was eventually cannulated. impella mcs was continued now with ecmo and approximately four days later, the patient¿s lactic was 9.0, which fluctuated up and down over the past few days per the report. platelets count was 18, and consequently, the patient had received seven units of platelets within the last 24 hours. due to the patient¿s inability to clear lactic, reevaluation in 24 hours was planned along with family discussion regarding care goals. however later this day, it was noted the patient expired. care was withdrawn. additional information was received and noted bleeding concerns were unrelated to impella placement. the patient experienced "traumatic" bronchoscopies that caused bleeding inside the bronchi. | [5 empty] | Initial reporting of the event, concomitant device was inadvertently omitted from section d.10 concomitant medical products and therapy dates and has been, therefore, captured accordingly. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the high purge pressure, clinical stated that on the 12th there was purge blocked alarms and tpa was started. the patient had been off systemic heparin for 2 hours prior. about 9 hours later the alarms resolved. the data logs show that the purge pressure was stable until about 10 hours into support it gradually began to start trending up and then triggered purge pressure high/blocked alarms after about 2 hours of trending up with no motor current spikes occurring before this. it remained high for about 10 hours and then it slowly decreased and the purge pressure was within expected values. based on the data log analysis and clinical information, the root cause of the purge pressure high is due to biological material buildup. regarding the pulmonary edema, the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5937 | 1220648-2025-46357 | 23022907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022907&pc=OZD | 08/08/2025 |  |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696253 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 70 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support (mcs) and experienced a middle cerebral artery infarct. the patient was admitted for acute heart failure management and the decision was made to place an impella 5.5 as a bridge to a left ventricular assist device (lvad) versus an orthotopic heart transplant. the impella 5.5 was successfully placed via right axillary artery. the patient was extubated in the operating room and then transferred to the unit on impella mcs. however, approximately eight days after start of the support, the patient was underwent a bridge to a lvad. post impella 5.5 explant and lvad implantation surgery, it was noted that the patient had developed a "large" middle cerebral artery infarct. a computed tomography angiography showed non-occlusive thrombus involving the right innominate and subclavian arteries. no further information was noted other than survival outcome at explant noted as survived. the patient's status following the event is unknown. | [5 empty] | Product complaint #: (b)(4). investigation summary: no product was returned for investigation. stroke/ thrombosis: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1937662. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer as per additional information the device was discarded in the operating room during the bridge to the left ventricular assist device. therefore, an evaluation/analysis of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5938 | 1220648-2025-46338 | 23022934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022934&pc=OZD | 08/08/2025 | 09/10/2025 |  | 09/10/2025 |  |  | 08/12/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316950 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during preparation, the console had a pump failure alarm. it was noted that after completing the impella support setup and priming the purge cassette, and connecting the pump catheter, the system returned to the initial startup screen and a "pump failure" alarm occurred. the system displayed messages such as "do not use this catheter" and "replace with a new pump catheter." the serial number was not recognized by the control device. after replacing the control device, the device functioned without any problems and support is currently being provided to the patient. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. data analysis: on 8/7, case start wizard was began on ic5657 but the console could not proceed through step 3 as there was an error reading eeprom (all crc values were continuously 65535). this happened after all four pump connections with 597138 (pump number was retrieved from salesforce as the console was unable to read the pump. the console had successfully read all pumps during previous cases and field service testing. device analysis: the failure mode was unable to be produced as the console was repeatedly able to read eeproms. root cause: the eeprom was unable to be read while running a pump in the field but that failure was not reproduced during testing. device history review: q1) service history review - are there an qns associated with the complaint device¿s most recent service report? no. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? n/a. q3) complaint history review - have there been any other complaints against this console? no. summary: there were no issues related to this complaint discovered when reviewing the product history of console ic5657. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5939 | 1220648-2025-46278 | 23022984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23022984&pc=OZD | 08/11/2025 |  |  | 09/10/2025 |  |  | 08/11/2025 | 09/10/2025 |  |  |  |  | 09/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601494 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. prior to impella cp insertion, the patient presented to the emergency room complaining of chest pain and an electrocardiogram revealed st elevation myocardial infarction. left heart catheterization was performed and revealed mitral valve disease and the patient became hypotensive. the patient was then started on norepinephrine and an impella cp was implanted and the patient was transferred to a large hospital in the network for further management and a possible coronary artery bypass graft. it was reported that on the second day of support, the p level was weaned to p4, and following the drop in p level the patient went into atrial fibrillation and the mitral valve prolapsed. the p level was turned back up to p6 and normal sinus rhythm was established. on (b)(6) 2025 it was reported that the patient continued to have episodes of atrial fibrillation and the medical team discussed changing the amiodarone drop to dobutamine in attempt to resolve the arrythmia and if unsuccessful then the patient may need cardioversion. the following day, the patient was taken to the catheterization laboratory for a high-risk percutaneous cardiac intervention due to decompensation and requirements of high dose pressor and inotrope support. it was noted that the patient was cardioverted twice that morning for ventricular tachycardia and once the previous day for atrial fibrillation. after the patient¿s procedure, it was noted that the impella cp had to be repositioned as it was found to be deep. the impella was measured at 5.27 centimeters (cm) in the ventricle and pulled back to 3.53 cm. on (b)(6) 2025 it was noted that the patient remained in atrial fibrillation and was non-pulsatile. it was reported later that night, the purge system on the impella cp required a bypass due to a leak. on (b)(6) 2025, the medical team decided to escalate the patient to an impella 5.5. the patient was brought to the operating room and the right axillary artery was prepped for impella 5.5 access. during insertion of the impella 5.5, resistance was met in the axillary artery and the impella 5.5 could not be advanced. the impella 5.5 was removed, and the axillary artery was shockwaved and a balloon was inserted and inflated to open the vessel. the impella 5.5 was still unable to pass through the artery. the artery was shockwaved again and the impella 5.5 was still unable to be advanced through the patient¿s artery due to the patient's anatomy. the decision was then made to abort the impella 5.5 insertion and to insert a new impella cp into the axillary artery. the impella cp was able to be delivered with no resistance. two days later, the patient experienced ventricular tachycardia and was successfully defibrillated to resolve. later that day, the decision was made to withdraw the patient¿s care. the patient¿s care was withdrawn, and the patient expired. this complaint involves three pumps: pump 1, impella cp with serial number (b)(6). pump 2, impella 5.5 with serial number (b)(6). pump 3, impella cp with serial number (b)(6). this report specifically addresses pump 2, impella 5.5 with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | H.10 one of the related report numbers was left off the initial manufacturer report number 1220648-2025-46278 and was added. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned for investigation. delivery issue: the cause of delivery issue is determined to be patient condition because impella 5.5 was unable to be inserted into left ventricle due to calcification in the vessels. later, the impella cp was inserted successfully. device history review: the device passed all post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety review was completed and noted the following: the patient expired while on impella cp support. the patient implanted with an impella cp pump 1 and experienced arrhythmic events requiring support level adjustment, possible change in medication, defibrillation, and pump repositioning. this device's purge system was bypassed due to system leak. the impella cp pump 1 was explanted and an impella 5.5 was unsuccessfully attempted. however, another impella cp pump 2 was successfully placed with no loss of support. the failure to deliver the impella 5.5 is a product malfunction and the first impella cp device event resulted in serious injury and should be reported. while on the second impella cp, the patient continued to have ventricular tachycardia (vt) runs that were successfully defibrillated; however, the patient subsequently expired as care was withdrawn. given the information at hand, there is not enough information to clearly dissociate the demise out from the second impella cp device. however, although this dissociation cannot be made, it is the care that was withdrawn that appears to have led to the patient's demise; vt runs were successfully defibrillated. medical safety officer reviewed the event and noted the same in the second impella cp device cannot be dissociated for the outcome and the first impella cp pump and impella 5.5 is a serious injury and product malfunction respectively type of reportable event. related medical device reports: this impella 5.5 device report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device report on the first impella cp and second impella cp, which is associated with the demise outcome.
5940 | 1220648-2025-46267 | 23000744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23000744&pc=OZD | 08/09/2025 |  |  | 09/09/2025 |  |  | 08/09/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710455 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). difficulty positioning the pump was encountered; however, ultimately the physician achieved a satisfactory position for impella mcs. the patient was brought to the catheterization lab with levo running and intubated. two stents were successfully placed into the left anterior descending artery. there was not much movement in the anterior apical wall. an impella cp device was successfully placed post percutaneous coronary intervention. on attempting to get access for a swan-ganz catheter, the impella cp was accidentally pulled back into the aorta. the physician was able to get the pump back into the left ventricle. however, there were multiple suction alarms. venous access was made, and swan was placed. the patient was transferred to the unit on mcs. support level was p-6, and the patient had suction alarms. device repositioning was attempted and completed. however, the following day, an echocardiogram was completed, and it appeared the pump was now under a papillary muscle in the left ventricle. the impella cp device was pulled back a couple times and readvanced under echocardiogram. however, the physician was unable to remove the impella cp device from under the papillary muscle. the physician found a stable placement in the left ventricle, however, with no suction alarms at support level p-6. physician decided to leave the pump in place and depending on how the patient did, the physician would bring the patient back to the catheterization lab to attempt to reposition the impella cp device pump again or place with a new pump in the correct position. ejection fraction at that time was noted 20-25%, the patient was not on pressors and stable. the patient remained on support for two additional days noted to be doing well before being successfully weaned and device explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary ==> device out of position: the cause of the positioning issue was determined to be an unintentional use error as the pump was accidentally pulled back into aorta on attempting to get access for a swan. device history lot: device lot: 1949021. device history batch: subcomponent lot: n/a. device history review: device (b)(6) passed all post sterile inspection checks. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the device out of position: the cause of the positioning issue was determined to be an unintentional use error as the pump was accidentally pulled back into aorta on attempting to get access for a swan. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5941 | 1220648-2025-46235 | 23001536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23001536&pc=OZD | 08/08/2025 | 10/07/2025 |  | 09/09/2025 |  |  | 08/08/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668059 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | Device Sensing Problem | Priming Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support for cardiomyopathy. placement signal not reliable popped up briefly, but then returned to normal/self-resolved. the nurse checked the cable and it appeared fine. the left ventricle signal was significantly higher than the aorta placement signal, and flows for given p-levels were significantly higher than expected. there was a noted drop in purge flow and an increase in mean motor current. after further investigation, it seemed that the pump was heading towards ¿pump saturation¿. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported optical signal issue, purge pressure, and purge flow has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. data log shows the placement signal was not reliable, with an alarm on (b)(6) for about 11 minutes. data logs for purge pressure shows approximately a 3-day gradual increase in purge pressure, with no recorded motor current spike during the case run. the pump flow and motor current were also seen to gradually increase during the purge pressure rise event. data log also show an air in purge alarm accompanied by a drop in purge flow. the issue was resolved after a successful de-airing procedure. the cause of the optical signal issue was most likely related to an air gap caused by unintentional user error, based on data log review. the cause of the high or saturated pump flow and purge pressure was most likely related to biomaterial buildup, based on data log review. the cause of the air in purge alarm was not determined, as the product was not returned for investigation and sufficient clinical information was not provided.
5942 | 1220648-2025-46268 | 23002874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23002874&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677408 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient experienced bleeding at the impella access site. manual pressure was held. the site was dressed multiple times, and the site was injected with lidocaine and epinephrine to resolve the bleeding. the patient remains on impella cp support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Investigation results: no product was returned for investigation. major bleed: bleeding caused during insertion in cath-lab when the act values were super therapeutic and the bleeding was managed by applying pressure and redressing multiple times. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
5943 | 1220648-2025-46269 | 23003040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23003040&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670810 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 62 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella cp, there was groin site oozing overnight after a turn which started to lok better. activated clotting times were supratherapeutic at that time and are therapeutic now. later, it was reported that the patient skin tears on the patient¿s arms and legs which are attributed to the need for multiple blood transfusions. two units of red blood cells were given. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. the device was not received for investigation. major bleed : the cause of the injury is most likely use related since activated clotting times were supratherapeutic at that time of the bleed and are therapeutic later on. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5944 | 1220648-2025-46270 | 23003093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23003093&pc=OZD | 08/10/2025 | 10/24/2025 |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685944 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 104 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed access site bleeding. manual pressure was held and femstop was placed but the impella cp was weaned down and removed due to the bleeding. | [5 empty] | The investigation for the reported major bleed has been completed. the root cause of the major bleed is most likely use related since during the bleeding the patient had very high act values above 400. | A.5 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5945 | 1220648-2025-46271 | 23003337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23003337&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699532 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that during impella cp support, a left ventricle thrombus was revealed with echocardiography. the decision was made to place the patient on veno-arterial extracorporeal membrane oxygenation and to keep the impella cp in place for left ventricle unloading. the patient continues on impella cp support. additional information was received. the patient was transferred to another facility. the patient received a durable left ventricular assist device. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | D6b: impella explant date added.
5946 | 23003714 | 23003714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23003714&pc=OZD | 07/17/2025 | 09/05/2025 | 09/05/2025 | 09/09/2025 | 09/05/2025 | 09/09/2025 |  | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 1000100 |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [6 empty] | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Device malfunction, needed replacement. manufacturer response for impella, kit ventricular assist device pump impella 5.5 smartassist (per site reporter). in review.
5947 | 1220648-2025-46273 | 23003719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23003719&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636389 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient classified as high-risk for percutaneous coronary intervention received an impella cp for mechanical circulatory support. the complainant also noted that the impella cp was unintentionally retracted into the aorta. the patient continued to be hemodynamically stable. subsequently, the impella cp was extracted. | [5 empty] | The investigation of the reported failure mode has been completed. the product was not returned for investigation. the cause of the incorrect positioning was most like use related, as clinical reported the pump being mistakenly pulled back. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5948 | 1220648-2025-46274 | 23003846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23003846&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713083 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the physician decided to place an impella cp on a septic patient. the impella cp was inserted with all best practices. the patient did not have pedal pulses on the right foot (impella site). the patient was in septic shock and coronaries were clear. as per the physician, the device was left in only to give the patient enough flow to do continuous renal replacement therapy (crrt). the patient had been on a high dose of vasopressors and crrt. the patient coded multiple times and passed away on support. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia/cardiac arrest; the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A.4 and a.5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5949 | 1220648-2025-46275 | 23004075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23004075&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710418 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 70 KG |  | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Atrial Fibrillation | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support (mcs) and experienced atrial fibrillation, hematuria and malpositioning of the device requiring intervention. post coronary artery bypass graft and aorta valve replacement surgeries, the patient remained in the operating room for several hours related to bleeding and difficulty achieving hemostasis (related to the surgeries, not the impella device). the patient was eventually taken to the unit on impella mcs. related to the surgeries bleeding event, the patient received 30 plus units of blood products and continued to be anemic. the following morning, the patient was given additional blood products. no source of active bleeding was visualized. no heparin products were received. the patient was just given nova7. the patient was taken for exploratory surgery related to the bleeding and additional blood products were given. the next day in the morning, it was noted that the patient went into atrial fibrillation and was treated with amiodarone and converted back to sinus rhythm. the patient complained of very blurry vision and the patients right eyebrow was higher than the left. the patient was evaluated. no further blood products were administered to the patient on this day. the results of the exploration surgery were unnoted. the patient continued on impella mcs. the following day, the patient was out of bed in the chair still with atrial fibrillation. one day later, the patient was noted again out of bed in the chair, and it was also noted the pump was withdrawn about one centimeter earlier in the morning of this day and the patient shortly began to experience hematuria. an echocardiogram noted the pump was still deep. next steps were pending. there was no definitive word on timing for waiting or explant. cardiothoracic surgery planned for a slow wean because of the patient¿s age and ejection fraction. over the next couple of days, the patient was noted hemodynamically stable on support level p-3. the patient was successfully weaned, and the device was explanted with a survived outcome. | [5 empty] | Complaint record coding of the reported event was updated. therefore, this correction follow-up (1) report is being submitted accordingly for recoding of the reported event via h6 health effect: clinical and impact codes. note: h6 component code and h6 investigation (type, findings and conclusion) remain unchanged. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, hematuria, the cause was not determined due to insufficient clinical details. device in wrong position, the cause of this complaint was not determined as well since it was unknown exactly when the pump went out of position. and regarding, paroxysmal atrial fibrillation, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5950 | 1220648-2025-46276 | 23004233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23004233&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601491 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 54 KG |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos) was initially implanted with an impella 5.5 followed by an impella rp flex for bipella mechanical circulatory support. the patient would experience complications and subsequently expire. patient had removal of mediastinal mass, mitral and tricuspid valves replacement and maze (modified atrial fibrillation surgery with electrophysiological zero ablation) on two days prior to the impella 5.5 device implant and was unable to come off pump. the patient was, therefore, placed on central extracorporeal membrane oxygenation (ecmo). the patient underwent chest washout and closure. impella 5.5 was placed and ecmo decannulated. the patient began to have suction events, which led to increased vasopressor requirement and moderate right ventricular dysfunction. the impella rp flex was implanted via left subclavian vein for bipella support. following, it was noted the patient was severely coagulopathic. numerous blood products were given to correct. then impella rp flex flows decreased, mean pulmonary artery pressure and motor current increased from baseline. the patient was noted to be well supported with mixed venous of 68 with plans to start heparin in 6 hours if bleeding was stable. (note: it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient.) over the night, the patient continued to bleed requiring additional products and ultimately was unable to flow right to left despite the rp flex support. the patient would soon thereafter expire. the physician commented and suspected the lungs were not allowing flow back into the left atrium due to pulmonary embolism. | The complainant indicated that a patient diagnosed with postcardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos) was implanted with an impella 5.5 for mechanical circulatory support. the complainant stated that suction events were experienced during the operation of the impella 5.5. it was reported that these suctioning events led to an increased requirement for vasopressors. furthermore, it was observed that the patient exhibited moderate dysfunction of the the right ventricle. consequently, an impella rp flex was implanted. |  |  |  |  | Section b5 was corrected, updated to include complete event details. medical safety review: medical safety officer reviewed the event and noted there is not enough evidence to exclude the impella rp flex as a contributing factor in the demise outcome. the physician suspected that the impella rp flex threw pulmonary embolism and prevented flow from right to left as they were unable to flow on the impella 5.5. the suction event and actions occurring with the impella 5.5 is a device malfunction and should be reported. change in reportability: after review of the case by the medical safety officer, it was determined the impella 5.5 is a serious injury type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this impella 5.5 report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex, which is associated with the demise outcome. | Corrections b1, h1, h6 the investigation of the reported failure mode has been completed. the product was not returned for investigation. bleeding: the cause of the injury was not determined due to insufficient clinical details. suction: data log review showed suction alarms triggering throughout support. flow levels were slightly below expected p-level ranges. the p-level was adjusted to suspend suction alarms - the patient could not tolerate higher p-levels without triggering suction. an rp pump was added, but good flow still couldn't be established from right to left, aligning with rv dysfunction mentioned in the clinical. logs show a spike in placement signal at the time of a placement signal not reliable alarm. this spike was most likely attributed to an air gap, as the motor current, snr (signal to noise ratio), gain, and contrast, all drop slightly. placement signal had low pulsatility, indicating decline in the patients condition. no motor current spikes were observed. the cause of the low pump flow was most likely patient condition related, as clinical reported moderate rv dysfunction, which is in alignment with the low placement signal pulsatility seen in data logs. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5951 | 1220648-2025-46277 | 23004374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23004374&pc=OZD | 08/11/2025 | 09/08/2025 |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636241 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the decision to place impella cp was made when a patient presented in cardiogenic shock. the cp was successfully placed. during support it was noted that the patient's urine was pink/dark (hemolysis) and albumin was given and p level was decreased. the patient was escalated to a 5.5, however was noted that was off of anticoagulants due to oozing post op. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5952 | 1220648-2025-46281 | 23005220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23005220&pc=OZD | 08/12/2025 | 10/24/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702574 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 45 KG |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered access site bleeding and suspected acute kidney injury. the patient was oozing from the groin site due to the blanket resting on the impella site. the care team as a result inserted a foley catheter. additionally, due to contrast, the doctor suspected that the patient had acute kidney injury. the device was then explanted later that same day. the patient survived. | [5 empty] | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported acute kidney failure and major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the acute kidney failure and major bleed could not be determined.
5953 | 1220648-2025-46280 | 23005262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23005262&pc=OZD | 08/11/2025 | 10/16/2025 |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686445 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Sepsis | Tachycardia | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 the patient experienced bleeding from right axillary cutdown site. hemoglobin downtrend was noted. the patient was administered two units of packed red blood cells and pressure dressing was applied. the bleeding resolved the following day. additionally, the patient experienced fever overnight. antibiotics were administered and the fever resolved. later, it was reported that the patient experienced ventricular tachycardia which reportedly resolved. the patient was also reported to be septic. the patient continues on support. | Additional information was received that the patient expired and the pump was explanted. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2, b.5 and d.6b additional information was received and these fields have been updated. b.7 added information that was inadvertently omitted from the initial report that was submitted. d.7a revised as it was previously entered incorrectly on the initial report that was submitted. e.4 added as selection was inadvertently omitted from the initial report that was submitted. h.6 added impact code due to new information that was received. the investigation has been completed. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. tachycardia/sepsis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. fever: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the pump passed all the post sterile inspection checks.
5954 | 1220648-2025-46279 | 23005264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23005264&pc=OZD | 08/11/2025 |  |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624406 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had a impella cp placed to support a patient who had been admitted in with st-segment elevation myocardial infarction and arresting with biventricular failure. the pump was placed but the team did not remove/peel away the introducer nor tighten the tuohy borst. this allowed the pump to have positional issues and hemolysis occurred. the hemolysis was observed with labs hemolyzing. the patient expired when care was withdrawn after over 27 hours of support. additional information was received. the physician believes that the position was ok and that the hemolysis was likely related to the high performance they were running the impella during a period of high after load. eventually, the hemolysis resolved and it was not at all a factor for pump removal. | [5 empty] | H.6 medical device problem code was added. the investigation was completed. the pump was not returned for investigation. data logs were not provided or could not be retrieved from the remote server. mechanical interaction with blood/hemolysis: as per the clinical details, the patient experienced high afterload during hemolysis event, which was managed by adjusting the impella flow and monitoring plasma-free hemoglobin levels. the cause of the hemolysis issue was most likely related to the patient condition related, based on the provided clinical details. device in wrong position: the cause of the positioning issue was most likely determined to be unintentional user movement during pump use. the device history review was completed and the pump passed all post sterile inspection checks. | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being submitted to correct a duplicate mdr. it has been determined that mdr and mdr 1220648-2025-46332 report the same event. this report is being closed as a duplicate of the initial, valid report (1220648-2025-46332).
5955 | 1220648-2025-46282 | 23005282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23005282&pc=OZD | 07/26/2025 |  |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133094 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was implanted with an impella cp device to provide mechanical circulatory support. the complainant noted issues with the automated impella controller (aic) placement signal. it was reported that the placement signal vanished from the console, and an unreliable alarm for the placement signal was present. troubleshooting procedures were followed, and the audio was turned off. additionally, there was a discussion about the signs indicating the need for pump replacement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b pump explant date provided. d9 updated. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. corrections: e4, g4.
5956 | 1220648-2025-31313 | 23005342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23005342&pc=OZD | 05/12/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 05/12/2024 | 09/09/2025 |  |  | 11/30/2024 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297489 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, a sensor failure occurred. however, the patient remained hemodynamically stable and continued to receive effective pump support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for investigation. pump had no biomaterial. signal not reliable were in specification. no placement signal not reliable alarms observed. the failure was not reproduced. the data logs show placement signal low alarm. the root cause of the optical signal issue was not determined since product and data logs investigation were inconclusive. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5957 | 1220648-2025-46283 | 23005559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23005559&pc=OZD | 08/12/2025 | 10/29/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710493 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues while trying to implant an impella cp. impella access was obtained and the right iliac showed chronic disease with left iliac showing chronic occlusion. the right iliac was treated with shockwave and balloon angioplasty and the short impella sheath was placed. the impella cp was advanced but ultimately would not pass the aortic bifurcation. several attempts were made to treat the vessel and advance the impella cp but ultimately the decision was made to stop and close the site. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. no product has been returned for review. the root cause of delivery issue is determined to be patient condition because right iliac showed chronic disease with left iliac showing chronic occlusion and the pump was unable to pass aortic bifurcation despite multiple attempts. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
5958 | 1220648-2025-46284 | 23005652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23005652&pc=OZD | 08/11/2025 |  |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657741 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 to support an ablation patient. the pump was placed but the site developed a hematoma that was treated by blood product infusion and adjustment of the kinked jp drain. the pump remained on for just under seven days and was then explanted. the patient is stable post explant. | [5 empty] | Product complaint # (b)(4). investigation summary- hematoma: the cause of the injury is most likely use related since the patient had high act values, patient act kept supra-therapeutic (>300) . device history lot- device lot: 1916518. device history batch- sub-component lot: n/a. device history review- the pump s/n: (b)(6) passed all the post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5959 | 1220648-2025-46286 | 23006118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006118&pc=OZD | 08/12/2025 | 10/29/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663211 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient encountered access site bleeding noticed by blood at the groin site. as a result, pressure was held for more than two hours off and on after an additional suture was placed. the patient did state that they needed to urinate so a foley was placed to relieve the bladder pressure and the bleeding from the groin stopped. the patient was taken back to the lab to explant the impella per doctors request. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported major bleed has been completed. no product was returned for review. the root cause of the major bleed was most likely use issue related since the bleed started occurring during the transport of patient.
5960 | 1220648-2025-46287 | 23006204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006204&pc=OZD | 08/22/2025 |  |  | 09/09/2025 |  |  | 08/22/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1637920 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complaints team was notifed by email from field service regarding a customer from a biomedical engineer reached out noting a controller error red alarm on a loaner console. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5961 | 1220648-2025-46291 | 23006593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006593&pc=OZD | 08/11/2025 |  |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026697715 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella placed via radiofrequency ablation, preclosed x2. no culprit lesion found but diffuse multi-vessel disease noted. opted to leave peel-away in place, initially was able to doppler/palpate pulses distally. the leg became cold and lost the ability to palpate/doppler pulse. additionally, the patient had foley placed after impella placement and noted red/pink urine output. decision made by md to remove device. patient is stable post issue. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. hemolysis: data log review showed no spikes in motor current or any positioning issues during the case. the cause of the hemolysis was not determined due to insufficient clinical details and lack of product returned. device history review: the device passed all post sterile inspection checks. | Patient race and ethnicity is unknown. - the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
5962 | 1220648-2025-31314 | 23006643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006643&pc=OZD | 11/16/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 11/16/2024 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496146 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the optical signal was intermittent. the alarm was disabled, and pump support continued without interruption. there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs were reviewed and the placement signal not reliable alarms were confirmed. they were intermittently triggering throughout the case: 4 triggers of placement signal not reliable and placement signal low each. pump flows were within specification throughout the entire case and no suction alarms triggered during the periods that the placement signal monitoring was active. this suggests the patient's condition nor environmental factors were not causing the noise in the signal. the product was not returned for investigation. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5963 | 1220648-2025-46290 | 23006653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006653&pc=OZD | 08/11/2025 | 09/08/2025 |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337720 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 112 KG | Non Hispanic | White | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed on an impella cp to support a patient admitted in with cardiac history and syncopal episode and arrest and were getting cardiopulmonary resuscitation. the impella cp was placed after the team placed extracorporeal membrane oxygenation and reperfusion catheter. the impella cp and automated impella controller (aic) were supporting for six days when the aic was replaced due to a controller alarm. the local team recommend switching out the aic for a new one. on day nine, the team escalated to an impella 5.5. the impella 5.5 supported for a total of 18 more days til the patient expired while decisions being made for a left ventricular assist device or transplant. no allegation was made of the aic malfunction causing or contributing to the death outcome. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5964 | 1220648-2025-46288 | 23006702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006702&pc=OZD | 08/11/2025 | 11/06/2025 |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708764 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was fitted with an impella cp to provide mechanical circulatory support. the complainant further noted that the day after the impella cp was inserted, it had migrated into the left ventricle, close to the mitral valve. attempts to reposition it under echocardiographic guidance were unsuccessful. the patient's condition remained stable at p2 level. additionally, it was reported that there were no signs of improvement following the placement of the impella cp. consequently, a decision was made to escalate the support to an impella 5.5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle. no other issues were found on the logs. product was not returned for review. the cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all the post sterile inspection checks.
5965 | 1220648-2025-31315 | 23006787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006787&pc=OZD | 01/14/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 01/14/2023 | 09/09/2025 |  |  | 09/30/2024 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023275509 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal became unreliable. the alarm was disabled, and pump support continued without interruption. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs showed that sudden placement signal spikes began to occur around 54 hours into support and occurred intermittently throughout the remainder of the case. the onset of the spikes were accompanied by the optical 5s parameters shifting out-of-specification; the 5s parameters were most closely characteristic with an air gap. an console issue was ruled out since the pump used with the same console following this case did not experience any relevant optical signal issue. no damage or abnormalities were found on the returned pump. the pump passed optical continuity test. when the device was connected to a lab console, the optical 5s parameters were within specification; additionally, the "placement signal not reliable" did not reproduce. in conclusion, it was determined that the cause of the optical signal issue was most likely an air gap between the console and the patient cable plug. all pertinent information available to abiomed has been submitted at this time.
5966 | 1220648-2025-31316 | 23006823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006823&pc=OZD | 03/22/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 03/23/2023 | 09/09/2025 |  |  | 08/31/2024 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259607 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the optical signal failed. pump support continued without interruption, and there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the optical sensor failed on the 42nd day of support. no data logs were returned for investigation. the pump was plugged into a console and the "placement signal not reliable" alarm was reproduced. the 5s parameters were within specification. the sensor was also removed from the pump and imaged and silicone sensor wear was observed with silicone etching at the edges of the sensor. the cause of the optical signal issue was most likely sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5967 | 1220648-2025-31317 | 23006876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23006876&pc=OZD | 06/26/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 06/26/2023 | 09/09/2025 |  |  | 06/30/2024 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229203 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was lost. the alarm was disabled, and pump support continued without interruption. there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs showed pump was started & run without issue. the logs also showed the placement signal was disabled after placement signal issue occurring & multiple suction alarms. per clinical description, the pump was potential interaction in graft with clamp & might affect to the optical signal of the pump. however, the root cause of the optical signal issue was unable to be determined as limited clinical details were provided and no product was returned for review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5968 | 1220648-2025-31318 | 23007274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007274&pc=OZD | 03/07/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 03/07/2023 | 09/09/2025 |  |  | 12/31/2024 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024286900 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was not functioning after a difficult insertion through the anastomosis. pump support continued without interruption, and there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. the data logs show the "placement signal not reliable" alarm triggered; the placement signal does not recover for the remainder of support. the optical sensor 5s parameters at the time of the failure were characteristic of an optical fiber break due to torque. in conclusion, when also considering the clinical information, it was determined that the cause of the optical signal issue was most likely a damaged optical fiber line. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5969 | 1220648-2025-46292 | 23007314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007314&pc=OZD | 08/11/2025 |  |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680607 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 52 KG | Hispanic |  | Malfunction | No | Yes | Backflow | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. on the third day of impella cp support, the impella cp was explanted as it was preventing retrograde flow. | [5 empty] | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary: low or blocked pump flow: data log shows several suction alarms and low pump flow while running at steady p-levels (p-9), with a decreasing motor current trend during the start of the case. an improving trend in placement signal and pump flow was noted when the pump was set to lower p-level p-5, two hours into the case run. no suction or low pump flow was reported later during the case. overall, there was fluctuation in the placement signal with some upward trend throughout the case run. before used pump's data log was reviewed and found to show low pump flow for the entire case, with active suction alarms and no motor current spikes. clinical details suggest afterload management with multiple inotropes and pressors. the patient had severely elevated left and right filling pressures, as well as diffuse coronary vasospasm, prior to switching to the complaint pump. the cause of the low pump flow issue at the start of the case was most likely related to the patient condition, based on the data log review and provided clinical details. retrograde pump flow: data log shows reverse flow control notification for about 5 minutes on 11-aug-2025 at 2:10. later, it was active from 02:27 to 07:53. there was some motor speed and motor current ramp-up noted during the time of the reported retrograde flow notification. pump flow did not fluctuate significantly during the retrograde flow event. the cause of the retrograde pump flow issue could not be determined without the returned product and sufficient clinical information. device history lot: device batch: 1931491. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks.
5970 | 1220648-2025-46293 | 23007388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007388&pc=OZD | 08/11/2025 |  |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540892 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 99 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported during an impella cp implant to a patient with acute myocardial infarction/cardiogenic shock for mechanical circulatory support (mcs) the impella cp device was unable to connect to the automated impella controller (aic) despite multiple attempts. per the reporter, the impella cp device was connected according to standard practice. once the pump was connected yellow to yellow and was primed, the white cable was connected to the aic. however, the aic did not detect the pump. several disconnects with reconnects were executed without success. care was taken to assess the connection cable for moisture as well as the pump terminal, and both were dry, appeared to be in standard configuration. the case was cancelled and restarted. the purge tubing was reinserted; however, again the pump was not recognized. consequently, the impella cp device was replaced, connected to the aic and was recognized in standard fashion without an issue. there was no patient harm as a result of this event. | [5 empty] | E.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. h.10 related report number was added. the product with event logs were received and the investigation was completed. device history review was completed and the pump passed all post-sterile inspection checks. clinical details stated that when connected to the automated impella controller (aic), the impella cp was not recognized. no alarm was reported. a new pump was connected successfully. log review showed the console telling the user to connect the pump but there was no further activity. the new pump was connected successfully instead. the pump was returned and evaluated. there was no visible damage to the plug or rest of pump. the pump was successfully recognized by two separate lab consoles. the pump was connected under different scenarios i.e. bending the cable/plug before connection, plugging in during case start, plugging in outside of case start, and repeatedly plugging and unplugging. the pump was detected without issue each time. no defect was found with the pump. upon review of the returned pump and logs, it is apparent that the console is the potential cause of the issue. h6: investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the report on the automated impella controller. | Additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
5971 | 1220648-2025-46297 | 23007402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007402&pc=OZD | 08/12/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689309 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had access site bleeding. residual oozing was noted during explant of the pump and as a result, a 6 french angioseal device was placed. patient survived. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5972 | 1220648-2025-46295 | 23007404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007404&pc=OZD | 08/23/2025 | 11/10/2025 |  | 09/09/2025 |  |  | 08/23/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711589 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant placed an impella 5.5 to support a patient admitted in and awaiting/bridging to advanced therapies, like a heart transplant. the impella 5.5 was supporting when there were optical signal loss/unreliable alarms. the impella 5.5 has not been explanted as a result and the patient remained on support until successful weaning. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of the two reports. this report represents the pump and the other report represents the automated impella controller. refer to section h.10 for the related report number. | H.10 a new related automated impella controller report number was added. | Correction: d4 the investigation of the reported failure modes has been completed. the only product returned was the motor housing/cannula with the catheter cut off. optical sensor issue: data logs were reviewed and the alarms were confirmed. there were no visual abnormalities noted with the optical sensor, though the typical testing of the device could not be conducted since the catheter/patient cable were cut off and not returned. based on the review of data logs, it is apparent that the console could be responsible for the reported optical sensor issue. device history review: the device passed all the post sterile inspection checks.
5973 | 1220648-2025-46298 | 23007512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007512&pc=OZD | 08/12/2025 | 10/07/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705484 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that impella 5.5 was inserted and advanced via axillary artery but could not advance past proximal axillary artery due to calcification in artery. the impella 5.5 was removed and impella cp was inserted without difficulty. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. unable to deliver or advance: the root cause of the delivery issue was determined to be patient condition due to calcification in artery. device history review: the device passed all post sterile inspection checks.
5974 | 1220648-2025-31320 | 23007666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007666&pc=OZD | 09/25/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 09/25/2023 | 09/09/2025 |  |  | 03/31/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327005 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump¿s placement signal alarmed intermittently, with readings fluctuating between high and low. placement monitoring alarms were disabled due to the unreliability of the signal. imaging confirmed that the pump was in the correct position. pump support continued without interruption, and there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned. as per clinical, placement signal was intermittent and fluctuating and the placement monitoring was disabled with unreliability of the signal. pump was verified to be in good position. data logs were reviewed and several impella position unknown and false impella in aorta alarm are observed. the 5s is having poor fluctuating trend with spiking placement signals during entire case at p-7. previous pump used also had several impella position unknown and false impella in aorta alarm per log review. the root cause of the optical signal issue was most likely patient condition related with several impella position unknown alarms and false aorta alarm based on log and clinical review. all pertinent information available to abiomed has been submitted at this time.
5975 | 1220648-2025-31321 | 23007985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23007985&pc=PYX | 01/10/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 01/10/2023 | 09/09/2025 |  |  | 03/31/2023 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2023265853 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Placement signal lost. placement signal previously displayed wasn¿t correlating. no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the data logs were reviewed and it was determined that the placement signal not reliable (psnr) alarm was triggered due to optical sensor values being out of range. the data logs show the optical sensor values drop by 50mmhg and then begin to drift. this is indicative of sensor wear. the pump was returned and the psnr alarm was reproduced. the optical sensor cavity was found to be much larger than when it was originally measured to be during calibration, suggesting sensor wear leading to a breach in the cavity of the sensor. the pump was then placed in the uson to pressurize the sensor. the pressure would equalize over time confirmed that there was a breach in the cavity of the sensor. the root cause of the optical signal issue was determined to be sensor wear. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5976 | 1220648-2025-46299 | 23008018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008018&pc=OZD | 08/11/2025 | 10/21/2025 |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681794 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock received an impella cp device for mechanical circulatory support. subsequently, the patient was moved to a different facility. an echocardiogram (echo) performed at the bedside revealed that the impella was directed towards the papillary muscle. the physician made an effort to reposition and torque the impella towards the apex, but this attempt was unsuccessful. before the placement of the impella, the patient underwent percutaneous coronary intervention, which resulted in complications. the patient continued to code and expired within three hours of being admitted to the unit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported cardiac arrest and positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the cardiac arrest and positioning issue could not be determined.
5977 | 1220648-2025-46294 | 23008031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008031&pc=OZD | 08/23/2025 |  |  | 09/09/2025 |  |  | 08/23/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075485 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 with automated impella controller for support of a patient admitted in and awaiting/bridging to advanced therapies, like a heart transplant. while on support, there were optical signal loss/unreliable alarms. the patient remained on support until successful weaning and there was no patient harm. | [5 empty] | Information was received that the serial number was reported incorrectly on the initial report and due to our system having limitations a new report number 1220648-2025-47064 was created to represent the automated impella controller (aic) with the correct serial number. any additional information received regarding the aic will be reported under the new aic report number 1220648-2025-47064. therefore, this report is being downgraded to not reportable due to system limitations. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of the two reports. this report represents the automated impella controller and the other report is for the pump. refer to section h.10 for the related report number.
5978 | 1220648-2025-46300 | 23008151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008151&pc=OZD | 08/09/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669961 |  | 0048-0045 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a left ventricle thrombus and stroke during impella cp support. the pump was explanted shortly afterwards. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
5979 | 1220648-2025-46301 | 23008195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008195&pc=OZD | 08/10/2025 | 11/06/2025 |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026699447 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 had intermittent runs of ventricular tachycardia when coughing or moving but resolved after. it was further reported that a purge system and purge pressure blocked alarm occurred. tissue plasminogen activator was administered. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of the reported failure mode has been completed. no product returned. high purge pressure (hpp): data log review shows a ~4 hour rise in pp, then sustained hpp for ~18 hours before it gradual resolves in the next two days. there was no motor current spike prior to the pp rise. after the pp rise, there were fluctuations in motor current (mc) and pump flow. there is also a long cassette change that occurred during the pp rise that could have contributed to the continued rise but was not the sole cause. the cause of the high purge pressure was biomaterial buildup since there was a gradual rise in pp with no mc spike. tachycardia: the cause of the injuries was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5980 | 1220648-2025-46302 | 23008289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008289&pc=OZD | 08/10/2025 |  |  | 09/09/2025 |  |  | 08/10/2025 | 09/09/2025 |  |  | 01/31/2026 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457844 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 49 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced access site complication and suction events, prior to the impella placement, an intra-aortic balloon pump was successfully attempted due to a pump failure. the decision was made to switch to an impella cp, which was successfully placed without incident in the right groin. the patient received 8000 u heparin intravenous-procedurally. activated clotting time was 252 seconds at the end of the case. the patient was transferred to the unit on impella mcs. when on the unit, the support level was dropped from p-9 to p-7 to break the suction. the patient was administered 250 milliliters of albumin bolus and was responsive to the fluid bolus. suction events stopped and was resolved. additionally, the patient experienced groin site oozing. the dressing was removed, and pressure was held over the insertion site until oozing stopped. a large, firm hematoma was also noted in the right groin (notedly not present in the catheterization lab at time of departure). labs revealed a drop in hemoglobin and the patient received one unit of packed red blood cells. distal pulses were confirmed present by doppler. the patient continued on mcs for an additional five days before being successfully weaned and the device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5981 | 1220648-2025-31322 | 23008312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008312&pc=PYX | 07/24/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 07/24/2024 | 09/09/2025 |  |  | 01/31/2026 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025513849 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal failed and placement signal not reliable alarm occurred. there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. on the second day of support, it was noticed that the placement signal had been drifting for several hours and mean flows had reached 5 l/min. then, the psnr alarm triggered, placement signal and flow calculations disappeared from the aic. at the time of the complication, various hemodynamic measurements had worsened since case start, including co, ci, and papi. the pump was not removed due to the failure mode and was weaned a day later. product was not returned for investigation. data log were reviewed: for roughly 9 hours leading up to the psnr alarm, the ps was trending downward which was then causing the mean flow calculation to drift upward to 5 l/min. additionally, lt logs show cvp starting out noisy, but then stabilized during the time that ps was trending down. all the while, motor current remained stable. when the psnr alarm triggered, ps, mean flows, and cvp calculations were all disabled. the mc remained stable as well as all optical sensor parameters. in looking at the rt logs, it was confirmed that cvp remained stable based on the raw optical sensor pressure reading. the root cause of the placement signal issue was determined to be sensor drift. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5982 | 1220648-2025-46289 | 23008316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008316&pc=OZD | 08/11/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675556 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant in japan had a cp pump and automated impella controller (aic) support being initiated to support the patient in need of pci for the ami/cgs patient. the pump stopped and was unable to be restarted successfully. in troubleshooting they also wished to replace the aic but there was no backup/not in patient use console available. the pump was explanted and pci performed without hemodynamic support of impella. patient survived. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of the 2 report, where this report represents the pump and the other report is for aic.
5983 | 1220648-2025-46303 | 23008334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008334&pc=OZD | 08/09/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639809 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 121 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had hemolysis concerns, and arrythmia. the pump had continuous suction with flows of zero the patient was anticoagulated and there was no evidence of clot. the pump was exchanged for a 5.5. the 5.5 was successfully placed and the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5984 | 1220648-2025-31323 | 23008350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008350&pc=PYX | 10/22/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 10/22/2024 | 09/09/2025 |  |  | 01/31/2026 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025461505 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a rp flex support ongoing for a patient admitted in post coronary artery bypass graft. the pump lost placement signal after 3 days of support. the pump remained on for support despite the loss/malfunction. no report of patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. data logs were reviewed and it was confirmed that the placement signal not reliable alarm triggered on 10/21. at this time alarm 1054 triggers and pap and cvp rail indicating an optical signal issue. as for the 5s parameters, snr drops to zero, contrast drops but the amplitude increases, lamp was already at 100%, so it remains there, light drops initially but becomes variable, and gain increases. the psnr alarm remains active through the rest of the case. the behavior of the 5s parameters show characteristics of both an air gap and fiber break, without a clear distinguishing factor that would suggest one over the other. clinical details do not provide any clear evidence for the cause of the complication and product was not returned for investigation. due to limited clinical details and data logs not showing a clear optical failure pattern, the root cause of the optical signal issues could not be determined. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5985 | 1220648-2025-46304 | 23008508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008508&pc=OZD | 08/12/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026705470 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on bipella support developed bleeding and hematoma at the impella 5.5 right axillary access site. two units of blood products were administered. it was further reported that while the patient was on impella 5.5 and ecmo support, a high purge system blocked alarm occurred. tissue plasminogen activator was administered. the patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint #(b)(4). investigation summary: no product returned. access site adverse event: clinical reported patient has bleeding at subclavian/ r axillary cutdown. it was noticed after transfer to icu but limited information on what caused bleeding was provided. the cause of this issue was not determined as insufficient clinical details were provided. high purge pressure: clinical reported purge system blocked, tpa protocol administered. the logs show a motor current spike with speed deviation and shift in pump flows to become saturated. after ingestion purge pressure became elevated and purge flow dropped triggering alarms. the purge pressure was elevated for 6 hours before returning to normal values. the cause of the issue was determined to be biomaterial ingestion as evidenced by log pattern and tpa resolution. device history review: the pump sn: (b)(6) passed all post sterile inspection checks
5986 | 1220648-2025-46305 | 23008580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008580&pc=OZD | 08/11/2025 |  |  | 09/09/2025 |  |  | 08/11/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654660 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 80 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that prior to impella 5.5 support, the patient presented to the hospital with decompensated heart failure and cardiogenic shock. the patient was started on inotropes, and an intra-aortic balloon pump was placed. the decision was made six days later to escalate the patient to an impella 5.5 to rehabilitate the patient and bridge the patient to a heart transplant or a durable left ventricular assist device (lvad). it was reported that during impella 5.5 support, the patient experienced heparin induced thrombocytopenia (hit) confirmed by serotonin release assay. heparin was discontinued and the patient was started on argatroban. the patient also received plasmapheresis for hit. the patient had plasmapheresis a total of three times performed every other day. due to hit, the patient was rejected as a lvad candidate. the patient is now on a do not resuscitate order and is in palliative care. the patient remains on impella 5.5 support. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that prior to impella 5.5 support, the patient presented to the hospital with decompensated heart failure and cardiogenic shock. the patient was started on inotropes, and an intra-aortic balloon pump was placed. the decision was made six days later to escalate the patient to an impella 5.5 to rehabilitate the patient and bridge the patient to a heart transplant or a durable left ventricular assist device (lvad). it was reported that during impella 5.5 support, the patient experienced heparin induced thrombocytopenia (hit) confirmed by serotonin release assay. heparin was discontinued and the patient was started on argatroban. the patient also received plasmapheresis for hit. the patient had plasmapheresis a total of three times performed every other day. due to hit, the patient was rejected as a lvad candidate. the patient is now on a do not resuscitate order and is in palliative care. the patient was weaned and explanted to enter home hospice. |  |  |  |  | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | No product was returned for investigation. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. h6 codes have been updated to reflect investigation closure. b5: narrative updated to include device was explanted and the patient entered home hospice. the revised b5 provides a complete event narrative. d6b: explant date added.
5987 | 1220648-2025-31324 | 23008995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23008995&pc=PYX | 11/03/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 11/24/2024 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED INC. |  |  | US |  | 2025546004 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp flex support and there were placement signal not reliable alarms. support continued. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the product was not returned. in the data logs, the placement signal is noisy throughout the whole case. during the time of the placement signal alarms over three hours, they are in suction conditions. the cvp is negative, signal not reliable slightly decreases, contrast slightly decreases, and gain, light and lamp are stable. once the patient is out of the suction conditions, the placement signal alarms resolve. there are no hard failures seen. the root cause of the optical signal issue is most likely due to the patients condition. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5988 | 1220648-2025-31325 | 23009034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009034&pc=OZD | 05/14/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 05/14/2024 | 09/09/2025 |  |  | 12/31/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450013 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the placement signal stopped functioning. despite this, pump support continued without interruption, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the impella 5.5 pump was not returned. upon review of data logs placement signal not reliable observed. please refer to manufacturer report number 1220648-2024-12317 for an investigation into the console. the cause of the optical signal issue was most likely sensor silicone wear per log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5989 | 1220648-2025-31326 | 23009501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009501&pc=OZD | 04/16/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 04/16/2024 | 09/09/2025 |  |  | 10/31/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418733 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, a ¿placement signal not reliable¿ alarm occurred. the alarm was disabled, pump support continued without interruption, and the patient experienced no harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the impella 5.5 pump was not returned. upon review of data logs placement signal not reliable observed. please refer to manufacturer report number 1220648-2024-12317 for an investigation into the console. the cause of the optical signal issue was most likely sensor silicone wear per log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5990 | 1220648-2025-31327 | 23009543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009543&pc=OZD | 09/12/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 07/25/2025 | 09/09/2025 |  |  | 06/30/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357307 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal was not reliable. the flow numbers remained consistent, and patient continued to be supported on impella. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. according to the clinical, shortly after beginning impella cp support placement signal numbers were being recorded in the 200¿s. however, flow remained consistent and the patient continued to be supported hemodynamically with no other issues. no product was returned for a visual inspection. data log analysis could not be performed since the only ltlog that was returned lasted ten seconds. the cause of the optical signal issue was not determined because no significant product or logs were returned and not enough clinical information was given. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5991 | 1220648-2025-31328 | 23009688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009688&pc=OZD | 10/19/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 07/25/2025 | 09/09/2025 |  |  | 02/28/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333468 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient who was just placed on impella support experienced a placement signal was lost on the automated impella controller (aic). no impella problems were noted, motor current remained pulsatile. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. pump was not returned for investigation. as per the clinical details, placement signal was lost after the shockwave device used. the cause of the optical signal issue was most likely damaged sensor due to the shockwave device interaction, based on given clinical details. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5992 | 1220648-2025-31329 | 23009743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009743&pc=OZD | 11/25/2023 | 09/09/2025 |  | 09/09/2025 |  |  | 07/25/2025 | 09/09/2025 |  |  | 05/31/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366002 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal low alarm was present. impella was pulled back 2 cm. alarm persisted. flows are optimal and position is satisfactory, so alarm was disabled. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. pump was not returned for investigation. data logs were investigated. the 5s parameters were in specification during entire case. after pump was reduced to p0, the placement signal was observed to decrease to -40. at that time, "impella in aorta" and "impella in ventricle" alarms were observed. the placement signal shot up to 3000 and returned multiple times. after 1 hour the pump was restarted without any abnormalities and placement signal and optical parameters were in specification. without the product, the failure cannot be further investigated. therefore, the root cause of the optical issue was not determined since product was not returned for investigation. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5993 | 1220648-2025-31330 | 23009848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009848&pc=OZD | 01/31/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 01/31/2024 | 09/09/2025 |  |  | 11/30/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024425054 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the aortic placement signal failed, most likely due to pump damage during insertion. despite this, pump support continued without interruption, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned for evaluation. there was no visible damage to the pump, 5s parameters were taken and all parameters were within range, pump passed laser light continuity test. data logs were reviewed and no issues with 5s parameters were observed. clinical details note that patient was given multiple units of blood during the first day or so of support. the root cause of the optical signal issues, specifically, the placement signal unreliability, is most likely due to patient condition. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5994 | 1220648-2025-31331 | 23009875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009875&pc=OZD | 02/19/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 02/19/2024 | 09/09/2025 |  |  | 11/30/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422137 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal not reliable alarm occurred. despite this, pump support continued without interruption, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned. data logs were reviewed, and placement signal lost with placement signal not reliable alarm observed after 20 minutes of pump plug connect and immediately after pump insertion/start. no other abnormalities were found. the cause of the optical signal issue was most likely damaged optical fiber line per log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5995 | 1220648-2025-31332 | 23009888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23009888&pc=OZD | 03/06/2024 | 09/09/2025 |  | 09/09/2025 |  |  | 03/06/2024 | 09/09/2025 |  |  | 05/31/2025 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024343266 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the placement signal drifted, triggering a false suction alarm and preventing operation above performance level p4. the sensor was re-zeroed, which resolved the issue and allowed appropriate support at a higher performance level. pump support continued without interruption, and the patient remained unharmed. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned. data logs confirm that the pump's placement signal was drifting downwards. there are placement signal low alarms at the start and end of the case. there are suction and impella position unknown alarms at the time of both drift events. the clinical states that the pump was confirmed to be in good position. upon investigation of the pump, no catheter kinks were observed. the pump passed the laser continuity test. the root cause of the optical signal drift issue was most likely a damaged sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5996 | 1220648-2025-46308 | 23012104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23012104&pc=OZD | 08/12/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628425 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was placed successfully, however during placement optimization the impella pulled out of the ventricle to the ascending aorta. the physician attempted to re-advance the impella and passed through the aortic valve, but the pump bent on itself. right radial access was needed to insert snare to straighten the impella. the impella was removed, and a new cp was placed. patient survived. | [5 empty] | Product complaint # (b)(4). investigation summary: pump not returned. clinical states pump was placed successfully but when optimizing the position it was pulled out of the ventricle. it was readvanced & able to pass the aortic valve but the cannula was bent on its self. data review: data logs showed impella position in aorta alarm. device in wrong position: the cause of the positioning issue was most an unintentional use error as the pump was pulled out of the ventricle while repositioning of the pump. product damage cannula kink: the cause of the cannula kink was most likely an unintentional use error as the pump folded in on itself when readvancing into the lv. due to attempted to readvance the pump wirelessly through aortic valve. device history lot: device lot #: 1898304. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5997 | 1220648-2025-46309 | 23012113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23012113&pc=OZD | 08/12/2025 | 10/30/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716932 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 75 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was removed due to concern for left ventricle thrombus on echo. patient survived. | [5 empty] | The investigation for the reported thrombosis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of thethrombosis could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5998 | 1220648-2025-46310 | 23012114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23012114&pc=OZD | 08/12/2025 | 10/24/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693984 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had an intra-aortic balloon pump (iabp) for one week prior to impella. impella cp was placed to support higher-risk percutaneous coronary intervention for a chronic total occlusion (cto). the patient was sent to intensive care unit, and a small hematoma was noted to the left insertion site. two units of blood products were administered. manual pressure applied. site was ecchymotic, soft. the dressing was redressed and angle matched. the patient had positive palpable pedal pulse. it was further reported that the patient expired, and the duration of cp support was for five hours. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma could not be determined.
5999 | 1220648-2025-46311 | 23012118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23012118&pc=OZD | 08/12/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  | 02/28/2026 |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474990 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered positioning issues. after the impella 5.5 was placed in the left axillary artery, the pump was pulled back and multiple unsuccessful attempts were made to recross the aortic value. due to the patient's anatomy, the surgeon was unable to readvance pump and the pump was removed. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary delivery issue: the cause of delivery issue was most likely patient condition related since the pump was unable to be repositioned due to the patient's anatomy. device history lot device lot: 1803689. device history batch subcomponent lot: n/a. device history review device (sn: (b)(6) passed all post sterile inspection checks.
6000 | 1220648-2025-46313 | 23012121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23012121&pc=OZD | 08/12/2025 | 09/09/2025 |  | 09/09/2025 |  |  | 08/12/2025 | 09/09/2025 |  |  |  |  | 09/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693726 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during insertion of the impella 5.5 the device failed to deliver post successful shockwave of left subclavian vessel. after several attempts and catheter kink, the decision was made to abort case and place impella cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: complaint device not returned for evaluation. delivery issue: the root cause of the deliver issue was determined to be patient condition as the patient had pad (peripheral artery disease) vessels condition unable to pass impella through vasculature led to kink the catheter due to attempted to advance impella. product damage (catheter kink): the cause of the catheter kink most likely was patient condition related due to pad vessels condition unable to pass impella through vasculature. device history review: the pump passed all post sterile inspection checks.
6001 | 1220648-2025-46231 | 22990855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22990855&pc=OZD | 08/09/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693741 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented for a left ventricle assisted device work up and was placed on an impella 5.5. during support, it was reported that there was a loss of placement signal. the pump was flowing at five liters per minute and the placement signal was not available. pulsatile motor current was noted. the placement signals was disabled. the position was confirmed under echocardiogram (echo). the physician was aware that the impella was measuring deep on the cho. the patient continues on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The explant date was added to this report. no product or event logs returned. the clinical states that there were issues with the placement and lv signal. positioning was confirmed but the echo was showing the pump was positioned slightly deep. it is unknown if those issues resolved during support. due to a lack of product or data logs returned, and insufficient clinical details, the root cause of the optical sensor issue cannot be determined.
6002 | 1220648-2025-46232 | 22991302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22991302&pc=OZD | 08/10/2025 |  |  | 09/08/2025 |  |  | 08/10/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708101 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 69 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported during impella cp support that the patient experienced restlessness and altered mental status. a stroke code was called and computed tomography of the head did not reveal any bleeding. the patient was reintubated, and it was noted there were no impella issues at the time of the event. the following day, the patient was weaned to p2. the patient was noted to be within goals of the care team after weaning the impella, so the medical team decided to explant the device. the patient was successfully weaned and explanted. | [5 empty] | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary ==> the pump was not returned for investigation. major bleed: clinical details report that the patient was bleeding from their gums and the triple lumen access site on (b)(6) 2025. the patient was on heparin at the time. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. access site adverse event: clinical details report that once back in the icu on day one of support, the patient's impella access site was oozing. it was reported that the patient's ptt was >200 at the time. the insertion angle was adjusted several times and the ptt was brought back under control, and the groin site was noted to be intact and dry the next day. product was not returned for investigation. since several troubleshooting methods were reported, the cause of the minor bleed could not be determined. stroke: it was reported that the patient had altered mental state overnight and code stroke was called on (b)(6) 2025. ct of the head was negative for bleed. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1947073 device history batch ==> subcomponent: . device history review ==> pump sn (b)(6) passed all post-sterile inspection checks.
6003 | 1220648-2025-46233 | 22991449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22991449&pc=OZD | 08/09/2025 | 10/24/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708103 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 64 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp support experienced right foot mottling. thee decision was made to remove the impella cp and replace with another impella cp. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
6004 | 1220648-2025-46234 | 22991603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22991603&pc=OZD | 08/09/2025 | 10/24/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636093 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 64 KG | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 for mechanical circulatory support. the implant procedure was aborted due the patient's anatomy. it was noted that the impella 5.5 could not cross at the innominate due to vessel size. the pump was explanted and replaced with an impella cp. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issue has been completed. the root cause of the delivery issue was determined to be patient condition as the patient had large vessel size preventing successful delivery of impella.
6005 | 1220648-2025-46236 | 22992172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22992172&pc=OZD | 08/08/2025 |  |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710659 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that during the impella cp insertion, there were difficulties obtaining optimal position of the impella in the ventricle. it was reported that the impella was inserted, removed, and re-wired four times in order to place the pump in the correct position as the pump was getting caught in the papillary muscle. eventually, the pump was placed in the correct position and support was initiated. on the second day of support, it was reported that the patient¿s laboratory samples were all hemolyzed, and the patient was experiencing dark red urine. an echocardiogram was performed to verify the pump was placed. later that day, the pump was weaned and removed for a planned mitral clip procedure and was not removed early for the complication. the patient was placed on an intra-aortic balloon pump after the impella cp was removed for support during the procedure. the mitral clip procedure failed, and the patient required intubation. the patient was then placed on an impella 5.5 for mechanical support and to be stabilized for surgery. the patient remained on impella 5.5 support until stable for surgery. the patient successfully had the surgery and was subsequently successfully weaned and the impella 5.5 was explanted. | [5 empty] | Investigation summary: product was not returned for review. data logs showed there were positioning alarms triggered in early of the case but resolved quickly. no motor current spikes were observed outside of p-level changes. mechanical interaction with blood/hemolysis: the cause of the hemolysis was unable to be determined as limited clinical details were provided, and no product was returned for review. device in wrong position: the cause of the positioning issue was unable to be determined as limited clinical details were provided, and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | H6 medical device problem code a01 code was not reported in the initial report.
6006 | 1220648-2025-46237 | 22993440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22993440&pc=OZD | 08/09/2025 |  |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685129 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient vomited and due to the movement, access site bleeding occurred. the dressing was changed and manual pressure was held for 15 minutes. the patient also started to code. the patient had ventricular fibrillation prior which required the patient to be shocked three times and to receive 15 minutes of cardiopulmonary resuscitation. the team also observed signs of hemolysis and hematuria that prompted pump repositioning discussions and diuresis. additionally, optical signal issues were noted. none of these harms or injuries prompted further intervention and the pump remained on for six days til wean and explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Correction g4, h6. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematuria: the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. access site adverse event/minor bleed: it is unclear based on the details provided which of the several potential factors, was the cause. for these reasons, the cause of the access site adverse event could not be determined. optical sensor issue: data logs were reviewed, and the alarms were confirmed. since clinical details report that the pump was properly positioned without any other details, data logs were inconclusive, and product wasn't returned for investigation, the cause of the optical signal issue could not be determined. device history review: the device passed all the post sterile inspection checks.
6007 | 1220648-2025-46239 | 22993716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22993716&pc=OZD | 08/09/2025 | 09/15/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1675422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the console fell to the floor in prep for patient use. the automated impella controller was being moved from storage for patient use. afterwards the display was "grainy". instructions for use will be followed and a different/backup console was used for patient support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: added code g02014. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was returned for analysis data analysis: console logs from the reported day of event did not contain any information relevant to the reported issue. device analysis: console was inspected and there were no signs of mechanical damage or impact on the aic housing. console was powered on, and the screen operated within specification, as there were no defects or distortion on the aic screen. reported issue was not confirmed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6008 | 1220648-2025-46238 | 22994405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22994405&pc=OZD | 08/09/2025 | 10/22/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694915 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and a hematoma at the incision site was still noted. the team is aware that it is starting to cause tingling and decrease in pulses in the right arm. the plan is to start an anticoagulation tomorrow. per the registered nurse, a clot was removed from the jackson-pratt drain and drain put out 15 milliliters, so the drain will be left in for now. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported limb ischemia, thrombosis, and hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia, thrombosis, and hematoma could not be determined.
6009 | 1220648-2025-46241 | 22994635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22994635&pc=OZD | 08/09/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666083 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an impella cp was placed to support a patient admitted in with acute myocardial infarction complicated by cardiogenic shock. while on support, the impella was alarming "position unknown" and p-levels were decreased. echocardiogram was completed and the doctor pulled the impella back. the correct position was verified. a company representative confirmed that anticoagulation remained off due to a stroke that occurred prior. there was a plan for a a repeat computed tomography of the patient¿s brain. at this time, the patient is still on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6b, explant date, has been added. | The investigation of the reported failure mode has been completed. positioning issue: data logs review showed that there were position unknown alarms and suction alarms triggered during the case but resolved quickly. product was not returned for review. the cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the device passed all post sterile inspection checks.
6010 | 1220648-2025-46240 | 22994654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22994654&pc=OZD | 08/09/2025 |  |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702730 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during impella cp support the patient became hypotensive after sedation was given. neurovascular changes were noted with no distal flow. the cardiologist and cardiovascular surgeon were consulting and discussing likely clamping down around the sheath after hypotensive episode and medication titrations. the patient will be closely monitored for any changes to extremity. if the patients vascular status does not improve with weaning medications and more continual sedation then the plan will be for external femoral-femoral bypass only. it was noted that the left foot was pale and unable to doppler pulses. runoff fluoroscopy was performed for a second time with no distal flow. neurovascular status of the left leg with active warming and serial doppler checks will be monitored. precedex drips were added. two units of packed red blood cells were given and left foot neurovascular assessments continued. the patient now has better color and cap refill with intermittent doppler pulses. doppler patient pulse was present at times. the impella cp was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure modes has been completed. the device was not received for investigation. ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
6011 | 1220648-2025-46242 | 22994780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22994780&pc=OZD | 08/11/2025 | 10/30/2025 |  | 09/08/2025 |  |  | 08/11/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700903 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 61 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient on impella cp experienced a small hematoma both distal and proximal to insertion site, as well as oozing being present upon arrival from transport. distal hematoma was massaged and manual pressure which seemed to resolved. additionally, manual pressure held on proximal hematoma; femstop placed at lower pressure to better manage proximal hematoma since patient was supra-therapeutic on heparin. bedside echocardiogram showed optimal placement. femstop was slowly taken off to assess for bleeding; the physician resutured blue suture pad; small hematoma distal to site redeveloped. manual pressure applied to proximal site; patient seemed to vagal at increased pressure and become hemodynamically unstable. compressions started for less than 2 minutes before epi push was given and stopped immediately once blood pressure became extremely hypertensive. 2 packed red blood cell¿s and 1 l ng given stat along with several small doses of epi during the next hour; levo was also started. patient was hemodynamically stable; site was not oozing - dressing changed; hematomas present and stable; femstop in place with positive dopplerable pt pulses. flows within expected range. the impella was successfully weaned and explanted. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported major bleed/ hematoma and hypertension has been completed. no product was returned for investigation. the root cause of the major bleed/ hematoma and hypertension was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
6012 | 1220648-2025-46243 | 22994794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22994794&pc=OZD | 08/09/2025 | 10/06/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584627 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the physician came to the bedside and discussed concerns for malpositioned device and hemolysis. the medical doctor repositioned the device. waveforms and mean arterial pressure flow improved. the impella cp was able to be successfully weaned and explanted. the patient was stable post explant. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis and positioning issue has been completed. the impella device was not returned for evaluation. data logs show fluctuations in the placement signal, pump flow, and motor current while running at a steady p-level throughout the case run. no motor current spike was reported, and the data log did not confirm any position-related issues. the root cause of the hemolysis was most likely improper device positioning, based on the clinical details. the cause of the improper positioning issue was not determined since sufficient clinical information was not provided.
6013 | 1220648-2025-46244 | 22994910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22994910&pc=OZD | 08/11/2025 | 10/28/2025 |  | 09/08/2025 |  |  | 08/11/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693833 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support the patient along with extracorporeal membrane oxygenation. during the support there was suction that was not resolving and the device was repositioned. later, it was reported that suction alarms started again, flows dropped, and waveforms looked dampened. blood products were administered and suction alarms were disabled. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D.7a revised as an incorrect selection was entered on the initial report that was submitted. e.1 added initial reporter address line 1 and initial reporter phone number as they were inadvertently omitted from the initial report that was submitted. h.6 medical device problem code was submitted as it was not reported on the initial report that was submitted. investigation summary: data review: data logs confirmed suction when the pump was started on 2025/08/09 with low placement signal and remained until 2025/08/15 then resolved for the rest of the case. product was not returned for review. pump flow issue: suction resolved by (b)(6) 2025 per data log review. placement signal was low until (b)(6) 2025 when suction resolved and right ventricle (rv) was reported to be improving. the cause of the suction was most likely a patient condition related since the patient had severe rv failure. positioning issue: (b)(6) 2025 inlet was pointed toward mitral, but still in free space. the cause of positioning issue was patient condition related because transthoracic echocardiography done and showed left ventricle cavity very small. device history review: the pump passed all post sterile inspection checks.
6014 | 1220648-2025-46246 | 22994957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22994957&pc=OZD | 08/10/2025 | 10/07/2025 |  | 09/08/2025 |  |  | 08/10/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684830 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 97 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, suction alarm occurred with decoupling of placement signal and left ventricle (lv) signal. the pump was adjusted multiple times and inlet appeared to be free floating in lv but could not break the suction. the suction persisted even at p-2. at times, the lv tracing was not appearing even at p-9. decision made to increase pressor support and wean impella off. impella inspected for clot ingestion but none was noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6015 | 1220648-2025-46245 | 22995002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22995002&pc=OZD | 08/08/2025 |  |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640374 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 91 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a patient admitted in with acute myocardial infarction/cardiogenic shock. while on support, the patient had some heavy bleeding from the impella site that required blood products to be given. the patient was sleeping when the bleeding started, most likely from moving in the bed. the intensive care unit staff was able to manage the bleeding by holding occlusive pressure at the site. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of major bleed was most likely use issue related since the bleeding started when the patient was sleeping and most likely moved. the device passed all post sterile inspection checks.
6016 | 1220648-2025-46247 | 22995010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22995010&pc=PYX | 08/10/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/10/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025588011 |  | 000000000001000323 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 54 KG |  | Black Or African American | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos) was initially implanted with an impella 5.5 followed by an impella rp flex for bipella mechanical circulatory support. the patient would experience complications and subsequently expire. patient had removal of mediastinal mass, mitral and tricuspid valves replacement and maze (modified atrial fibrillation surgery with electrophysiological zero ablation) on two days prior to the impella 5.5 device implant and was unable to come off pump. the patient was, therefore, placed on central extracorporeal membrane oxygenation (ecmo). the patient underwent chest washout and closure. impella 5.5 was placed and ecmo decannulated. the patient began to have suction events, which led to increased vasopressor requirement and moderate right ventricular dysfunction. the impella rp flex was implanted via left subclavian vein for bipella support. following, it was noted the patient was severely coagulopathic. numerous blood products were given to correct. then impella rp flex flows decreased, mean pulmonary artery pressure and motor current increased from baseline. the patient was noted to be well supported with mixed venous of 68 with plans to start heparin in 6 hours if bleeding was stable. (note: it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient.) over the night, the patient continued to bleed requiring additional products and ultimately was unable to flow right to left despite the rp flex support. the patient would soon thereafter expire. the physician commented and suspected the lungs were not allowing flow back into the left atrium due to pulmonary embolism. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding for which blood product was administered. it was reported that the device had a high purge pressure alarm and that the motor current had gradually increased after receiving multiple blood products. the tissue plasminogen activator (tpa) was going to be used as purge fluid; however, the high purge pressure resolved on its own and the tpa remained on standby. additional information noted that the patient expired. |  |  |  |  | Section b5 was corrected, updated to include complete event details. medical safety review: medical safety officer reviewed the event and noted there is not enough evidence to exclude the impella rp flex as a contributing factor in the demise outcome. the physician suspected that the impella rp flex threw pulmonary embolism and prevented flow from right to left as they were unable to flow on the impella 5.5. the bleeding event associated with both devices; suction event and actions occurring with the impella 5.5 is a serious injury and should be reported. h6 coding: health effect: clinical and impact coding were updated, repopulated accordingly. related medical device reports: this impella rp flex report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. high purge pressure (hpp): the data logs show that during priming the purge pressure was high but then it went back down to be within specification. about 42 minutes in here was a spike in motor current followed by a trend in purge pressure. it rose from about 400 mmhg to 600 mmhg and triggered hpp alarms about 2 hours in and then returned to 600mmhg. hpp was triggered again after by another motor current (mc) spike about 7 hours into the case but it resolved after an hour. during this whole time the mc was trending up to almost 1000ma with spikes the entire time and slightly lower flows. based on the data log analysis the root cause of the high purge pressure is most likely due to ingested biological material. embolism/major bleed: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
6017 | 1220648-2025-31297 | 22995127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22995127&pc=PYX | 10/08/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 10/08/2024 | 09/08/2025 |  |  | 02/28/2026 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025484906 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, a ¿placement signal unreliable¿ alarm occurred. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or logs were returned for evaluation. a failure mode of "placement signal issues" was added for investigation to include potential issues with the sensor. the root cause of the placement signal issues cannot be definitively determined as no product or logs were returned for evaluation and limited clinical details were provided. the root cause of the optical signal issues cannot be definitively determined as no product or logs were returned for evaluation and limited clinical details were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6018 | 1220648-2025-31298 | 22996245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22996245&pc=PYX | 11/07/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 11/07/2024 | 09/08/2025 |  |  | 07/31/2024 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560597 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump¿s flow readings and waveforms were lost. despite this, the patient remained hemodynamically stable and continued to receive effective circulatory support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the failure mode was changed to "placement signal issue" as issue was in relation to sensor. pump was returned and upon inspection blood was found on the sensor face. there were no other abnormalities or damages to sensor or sensor area. the data logs from the field do not show any placement signal alarms. the motor current was stable while placement signal drifted up and flow drifted down before flow calculation loss. the pump was tested in the flow loop and sensor drift reproduced. the root cause of the placement signal issue was determined to be sensor drift of the sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6019 | 1220648-2025-31299 | 22996889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22996889&pc=PYX | 01/26/2023 | 09/08/2025 |  | 09/08/2025 |  |  | 01/26/2023 | 09/08/2025 |  |  | 05/31/2024 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023213702 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs were returned and confirm the unreliable placement signal on the first day of support. the dp sensor was out of range and slowly returned to normal range over the course of 5 hours. the displayed flow returns with the placement signal. the root cause of the temporary inaccurate placement signal was unable to be determined as no product was returned for analysis. all pertinent information available to abiomed has been submitted at this time.
6020 | 1220648-2025-46249 | 22997895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22997895&pc=PYX | 08/09/2025 | 02/27/2026 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026666673 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Increase in Pressure | Device Sensing Problem | [5 empty] | Heart Failure/Congestive Heart Failure | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex for mechanical circulatory support. while on impella rp flex support, the placement signal not reliable alarm occurred. motor current was within range. the alarm was disabled. additionally, there was a high purge pressure alarm and tissue plasminogen activator was administered. purge pressure had since been lowered. it was further reported that the family decided on comfort measures with withdrawing care. the patient expired. | [5 empty] | H6 health effect - clinical code was corrected form e2403 to e0611. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Corrections: d4 catalog number updated h6 component code changed from g04105 to g07001 h6 type of investigation removed b13 the investigation for the cardiac arrest, purge pressure high, and placement signal issue has been completed. the cause of the cardiac arrest was unable to be determined due to insufficient clinical details. due to a lack of sufficient data logs or product returned and an inconclusive data log review, the cause of the purge pressure high issues cannot be established. without sufficient details on the time of transport or return product to confirm any damage of the dp sensor, the cause of the placement signal issue cannot be confirmed.
6021 | 1220648-2025-31300 | 22997992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22997992&pc=PYX | 09/01/2023 | 09/08/2025 |  | 09/08/2025 |  |  | 09/01/2023 | 09/08/2025 |  |  | 06/30/2025 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024363980 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. according to the clinical, two hours after beginning impella rp support, frequent suction and placement signal not reliable alarms were recorded. no attempts were made to resolve the issue and the patient remained on support for another five days until the pump was removed. no product was returned for a visual inspection. data log analysis confirmed the presence of suction and placement signal not reliable alarms. the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6022 | 1220648-2025-31301 | 22998286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998286&pc=OZD | 12/20/2024 | 10/27/2025 |  | 09/08/2025 |  |  | 12/20/2024 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527310 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, there was a sudden loss of both the placement and left ventricular signals. however, the patient remained hemodynamically stable and continued to receive effective pump support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump and data log were not provided for investigation. according to the clinical information, the placement signal was recovered after properly connecting the patient cable to the console, which indicates that an air gap caused the issue. the cause of the optical signal failure was most likely air gap from between the console and the patient cable plug or within the middle of the red handle ferrule, based on given clinical details. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D6b: corrected the explant date. | G3 ¿date received by manufacturer¿ corrected.
6023 | 1220648-2025-46252 | 22998299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998299&pc=OZD | 08/09/2025 | 10/22/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  | 02/28/2026 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474973 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Inflammation | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an impella 5.5 placed for a patient with acute myocardial infarction/cardiogenic shock. the patient had orthopedic surgery post a fall. the pump was placed in combination with extracorporeal membrane oxygenation (ecmo). the patient's ecmo sites were bleeding, as was the liver. as such there was no anticoagulation/heparin in the pump or systemic. on day two of support, the impella 5.5 limb became swollen and they evaluated for a deep vein thrombosis. the arterial flow was noted to be good. on day seven, the pump was successfully weaned and explanted. the patient survived. | [5 empty] | The investigation for the reported inflammation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the inflammation could not be determined. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6024 | 1220648-2025-46254 | 22998366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998366&pc=OZD | 08/11/2025 | 12/31/2025 |  | 09/08/2025 |  |  | 08/11/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - |  |  | US |  | 2026708749 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Difficult to Remove | Malposition of Device | Deformation Due to Compressive Stress | Patient Device Interaction Problem |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received the device was discarded. | The user facility reported the patient had a cp placed to support the patient admitted in with a cardiac history and new acute myocardial infarction with cardiogenic shock. the pump was supporting but the team was repositioning the pump (there was elevated pfhg) the pump came out of the left ventricle into the aorta. the team then observed the pump to prolapse/kink upon itself. the team had to place a balloon into the aorta to assist with the pump removal. the pump was removed and replaced. |  |  |  |  | Additional information b5, h6. correction h6. | Reportability status has been updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6025 | 1220648-2025-31302 | 22998390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998390&pc=OZD | 11/03/2023 | 09/08/2025 |  | 09/08/2025 |  |  | 11/03/2023 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the automated impella controller (aic) was having on-going placement signal not reliable issues. echo was ordered and completed to confirm placement. staff continued support with the aic and impella. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. data analysis: the imc logs confirm that there was a placement signal unreliable alarm i.e. [(b)(6) 2023 07:24:01 imcgui: event set: #130 from 9 (1)]. see (b)(6)_imc_logs.pdf. the alarm was preceded by numerous messages of invalid optical bench data samples due to a low optical snr. i.e. [(b)(6) 2023 07:22:51 ossiopsens set invalid_bad_snr_sample_mask]. see (b)(6)_imc_logs.pdf. device analysis: the console was inspected in the lab. after optical connector was inspected and cleaned, snr test was conducted (section 13 of the functional test (0042-7316)). it was observed that the snr was passing the spec >2000. see (b)(6)_obt.png. then the fringe pattern of the optical bench was inspected both with an oscilloscope when a pump attached to the console and without one. it was observed that there was an anomaly with the pattern consistent across the two test setups. see (b)(6)_no_pump_fringe_ptrn. png and (b)(6)_pump_fringe_ptrn.png. conclusion: the root cause for the unreliable placement signal alarms was a low signal not reliable due to a defective optical bench. this failure mode of optical signal issue is not directly related to the reported pump complaint investigated in 23-19573-1 of pump stoppage. repair task: field service instructed to replace the defective optical bench, (0042-1012). return defective subcomponent to engineering. perform sections 23 of the preventative maintenance procedure. (0042-7309) perform full functional check (0042-7316). a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6026 | 1220648-2025-46253 | 22998407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998407&pc=OZD | 08/10/2025 | 10/30/2025 |  | 09/08/2025 |  |  | 08/10/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695450 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered positioning issues during impella cp support. the echocardiogram (echo) measured pump at 5cm, and the slack was pulled back and left at 4.1cm. the heart was slightly rotated and the impella was angled toward the mitral valve (mv). thus, an attempt was made to rotate the pump under echo but the team was unable to so. a few hours later, another echo was obtained and the pump was pulled back another 0.5 cm and they attempted to angle it away from the mv, however, the issue persisted and they were unable to do so. patient survived. | [5 empty] | The investigation for the reported positioning issue has been completed. no device was returned for review. the root cause of positioning issue was most likely patient condition related since the heart of patient was rotated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6027 | 1220648-2025-46255 | 22998416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998416&pc=OZD | 08/11/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/11/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651074 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 99 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was taken to the cath lab and on angio, multivessel occlusions were found. right coronary artery occluded but was able to be opened by stenting. a decision was then made to place impella in the right femoral artery. vital signs slightly improved but respiratory remained critical. the patient had a total blockage (occlusion) of the circumflex artery and a chronic total occlusion (cto) of the left anterior descending artery (lad). the patient began coding and compressions were initiated. during compressions the cp was pulled out of the left ventricle. md decided due to severe multi-occlusions, to not replace impella. cpr continued but the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections: b1, e4. the investigation of the reported failure modes has been completed. no product was returned. the cause of positioning issue was most likely use issue related since during compressions the cp was pulled out of the left ventricle by the physician. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
6028 | 1220648-2025-46256 | 22998424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998424&pc=OZD | 08/11/2025 | 10/30/2025 |  | 09/08/2025 |  |  | 08/11/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634521 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support coded with ventricular tachycardia after line was placed in the neck. no known intervention. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported tachycardia and cardiac arrest has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the tachycardia and cardiac arrest could not be determined.
6029 | 1220648-2025-46257 | 22998450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998450&pc=OZD | 08/21/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/21/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1544031 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered an automated impella controller (aic) battery issue. there was a decreased battery on aic while ambulating, which alarmed to plug in immediately. the aic was only on battery power for 10 min. the next day, the aic had critical low battery alarms when unplugged and the following days continued to have battery issues. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. this report represents the aic (controller). refer to section h10 for related report number 1220648-2025-46250, which represents the pump. | Correction e4 the investigation of the reported failure mode has been completed. data analysis: the console logs did not show any battery level low (23) alarms. in the logs, however, multiple battery level low (23) alarms were observed, along with several false impella position alarms. across the case, multiple battery level low and ac power disconnected alarms were seen. battery data embedded in the lt log revealed incorrect value (a single sample) of the battery temperature was observed being sent from battery 1 and incorrect value (a single sample) of the battery full capacity was observed being sent from battery device analysis: the console was returned to field service for further investigation. the reported complaint could not be reproduced, even after the console was power cycled and operated with the pump and purge cassette for a couple of hours. communication between the power battery manager (pbm) and the batteries was monitored, revealing no issues. additionally, a visual inspection of the internal circuitry showed no abnormalities. root cause: the root cause of the "battery level low" alarm was most likely an erroneous trigger of the invalid data sample in the smbus communication between the pbm and the batteries. however, the specific component responsible was not determined. device history review: the device passed all the post sterile inspection checks.
6030 | 1220648-2025-46260 | 22998457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998457&pc=OZD | 08/27/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/27/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1733469 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered an automated impella controller (aic) battery issue. the complainant encountered a battery failure and controller error shut down. no further details were provided. | [5 empty] | Patient information unknown. | The investigation was completed and the device was returned. data analysis: the data logs are consistent with the reported complaint. the battery failure alarm (transport connection blown/missing #360) was triggered immediately after the aic was booted on (b)(6) 2025. examining the data logs for the previous case, which began on (b)(6) 2025, with pump sn: (b)(6), reveals that the pump was connected and disconnected three times. after the third connection, a controller error (opm communication stopped #1043) alarm was triggered, followed by the pump stopping due to impella stopped (mechanical/electrical issue #115) alarms. additionally, the absence of snr samples affected the console¿s ability to detect the pump disconnection. this necessitated a hard shutdown by the user, which involved turning off the battery transporter switch underneath the aic. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are lastly reported as +16384 values and then no more samples were recorded. the log entry process sampling data from optical bench: sent stop acquisition cmd ack indicates the sent stop acquisition cmdflag was set when entering ossi_sampling_stopped_state which eventually led to a false communication stopped condition. this was entered into jama with defect/ bug. device analysis: the console ic10644 was returned to fs for further investigation. during this process, the failure mode was successfully reproduced by starting the aic with keeping the battery transporter switch in the off state. while testing the batteries, no issues were observed; both batteries charged according to specifications, and no cells were found to be imbalanced, as shown in batttest.png. additionally, an inspection of the analog output from the ccd on the optical bench revealed no distortions, regardless of whether the aic was connected to the pump or disconnected. finally, the "5s" parameters¿particularly when tested with the test pump and test cavity 2¿showed an snr value below the limits specified in the pm procedure. however, when tested with test cavity 1, the snr values were within the acceptable range. root cause: no issues were identified with the aic related to the reported failure mode of the aic-battery issue. the root cause of the controller error (opm communication stopped #1043) was due to an aic software issue.
6031 | 1220648-2025-46259 | 22998460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998460&pc=OZD | 08/21/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/21/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025608485 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered minor bleeding and optical signal issue during impella 5.5 support. some oozing was noted at the site, which was resolved by redressing it. additionally, a pump in aorta alarm was encountered, however, it was verified that the impella was in good position. the complainant continued having impella pump in aorta alarms throughout support, motor current was stable and position was verified via echocardiogram. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. this file represent the pump. refer to section h10 for the related report number 1220648-2025-46251, which represents the aic (controller). | Explant date provided, d6b updated. correction: e4.
6032 | 1220648-2025-31303 | 22998470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998470&pc=OZD | 12/04/2023 | 09/08/2025 |  | 09/08/2025 |  |  | 12/04/2023 | 09/08/2025 |  |  | 09/30/2025 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024402969 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the ¿placement signal not reliable¿ alarm occurred since insertion of the impella. motor current pulsatile, and echo confirm device is optimally placed. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the pump was not returned for investigation. data logs were not provided for this case run. the doctor reported placement signal not reliable on the second day of pump use. the pump was utilized for a total of 8 days. the placement signal was calculated based on optical signal parameters; hence, the reported failure mode of placement signal was changed to optical signal issue as investigated. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6033 | 1220648-2025-46248 | 22998645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998645&pc=OZD | 08/09/2025 |  |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  | 08/31/2025 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024388132 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female | 80 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that during impella cp support, the patient developed a ventricular arrythmia which required a pacer and medical therapy escalation. the impella ran well and there was no alarms. the patient was weaned and the impella cp was explanted. it was noted that the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6034 | 1220648-2025-31304 | 22998667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998667&pc=OZD | 01/29/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 01/29/2024 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 190529 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support due to acute myocardial infarction. while on support there was placement signal unreliable alarms. placement was verified and support continued. the automated impella controller (aic) has been sent in for investigation. no known patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. console logs from the reported day of event are consistent with complaint and show ¿placement signal not reliable" alarm. the console was tested in the lab. and the issue was eventually reproduced, while analog output of optical bench ccd was monitored. the images captured on oscilloscope showed distortion of the ¿fringe¿ pattern, which, with diminished values of signal not reliable, hindered the position detection algorithm. unlike previously observed occurrences of similar issues (oscillating contrast), they could not be resolved by short mechanical vibrations applied to the bench or lamp assembly, but only by power-cycling optical bench. analysis of the data associated with one of prior complaints involving the same console (b)(4) revealed the same signatures of optical signals, therefore leading to the opening of second investigation ((b)(4)) against the console (originally, (b)(4), was entered against the pump). the issue addressed by (b)(4) had the same failure mode and root cause as this investigation. the cause of the aic issue was a defective optical bench. repair: replace optical bench assembly (0042-3015 or equivalent) and perform functional check. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. please refer to dhr checklist for indications of: "prior complaint involving optical bench from same lot 1395640: case-(b)(4): optical placement signal not reliable and loss of communication to optical bench, due to optical bench failure (increased current draw)." and " (b)(4): prior occurrence of psnr, originally against pump (b)(4): purge disk not detected due to flag damage." all pertinent information available to abiomed has been submitted at this time.
6035 | 1220648-2025-31305 | 22998754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998754&pc=OZD | 04/11/2024 | 10/27/2025 |  | 09/08/2025 |  |  | 04/11/2024 | 09/08/2025 |  |  | 12/31/2023 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023163844 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that while 59-year-old male patient was on impella cp support, placement signal alarm sounded. the position of the device was checked with echo. inlet free in the ventricle, outlet in the aorta. the physician checked the position again when the patient was placed back on his back. the console was turned off. the patient was stable, and the device was not removed as a result. the patient was successfully weaned, and the device explanted the next day without issue. | [5 empty] | D6a: corrected implant date. d6b: corrected explant date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. no product was returned for evaluation. clinical details noted that pump position was check via echo and found to be in proper position. data logs were reviewed and rt log at time of "impella in aorta" alarm shows a motor current spike followed by a loss in pulsatility in motor current and pump flow. additionally, a motor speed deviation is observed. the 5s parameters at the time of alarm remained stable and within range. the root cause of the optical signal issues was most likely due to biomaterial ingestion. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | G3 ¿date received by manufacturer¿ corrected.
6036 | 1220648-2025-31306 | 22998832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998832&pc=OZD | 05/09/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 07/02/2024 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721787 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a 5.5 pump placed to support a surgical patient in need of support escalation from the cp. the 5.5 was placed but lost placement signal. the 5.5 stayed on for support without any report of harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. console logs from the reported day of event are consistent with complaint and show three alarms ¿placement signal not reliable¿ (#130, due to low snr ratio) after 3 days of operation. logs also revealed multiple events 1045 (loss of opm data due to crc error) throughout the case, but no corresponding alarms because of short duration of the triggering condition. console ic10520 was later sent to danvers for an unrelated complaint (b)(4) but tested and analyzed in connection with both investigations and failure modes. it¿s been concluded that ¿psnr¿ alarms (#130) were unrelated to the console, and separate investigation (b)(4) refers to possible relation to the pump (501233) that was used. events 1045, unrelated to psnr condition, although did not result in any alarms, are related to momentary loss of opm data. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the controller error (but not psnr alarm) was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. please refer to dhr checklist for indication of: "24-36962-2: purge disk not detected". this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. repair: perform functional check of the console.
6037 | 1220648-2025-31307 | 22998880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22998880&pc=OZD | 05/30/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 05/30/2024 | 09/08/2025 |  |  | 01/31/2026 |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED INC. |  |  | US |  | 2025453573 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient had impella cp support due to needing angioplasty, shockwave, and stent to left main, left circumflex coronary artery. after the shockwave, there was a placement signal not reliable alarm. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. no pump was returned for investigation. clinical states that placement signal not reliable (psnr) alarms occurred after shockwave use. data logs show that about 1 hour into the case the ps spikes with drop in snr, gain, and contrast. lamp increases. there are psnr alarms. the ps does not return. the root cause of the optical signal issues was most likely a damaged sensor due to shockwave use. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6038 | 1220648-2025-31308 | 22999038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999038&pc=OZD | 06/26/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 06/26/2024 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1076289 |  | 1000100 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella support when the automated impella controller (aic) had placement signal issues reported. echocardiogram was performed to check placement of pump. motor current remained pulsatile and flows were appropriate for level of support. support continued. no serious patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. device analysis: the console (b)(6) was sent back to field service for further investigation. the reported complaint could not be reproduced. visible light was detected from the optical pump connector in the console, and "5s" parameter testing confirmed that the optical system was within specification. please see the attached test_cavity.jpg and connector.jpg. root cause: the root cause for the reported placement signal not reliable (opm invalid signal, optical pump connected]) - alarm issued and controller error (defective optical sensor lamp) was most likely due to the defecting optical bench or the lamp assembly. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. please refer to dhr checklist for indication of: "complaint: (b)(4): controller error (defective optical sensor lamp) and lamp assembly was changed." corrective actions: before returning the console to the customer, the following actions are instructed to perform: replace the optical bench (0042-1012). perform a full functional check on the console and optical system (0042-7316). all pertinent information available to abiomed has been submitted at this time.
6039 | 1220648-2025-31309 | 22999274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999274&pc=OZD | 08/11/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 08/11/2024 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519451 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal was disabled due to a false alarm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the failure mode "placement signal issue" was changed to "optical signal issue" because the cp pump does not have a differential pressure sensor, "placement signal issue" is the failure mode used for devices with that sensor. no product was returned for investigation. the data logs show multiple "impella position in aorta" alarms before the placement signal monitoring was disabled. no mc spikes observed in the logs. it was not mentioned in the clinical details how proper positioning was confirmed. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6040 | 1220648-2025-31310 | 22999312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999312&pc=OZD | 10/18/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 10/18/2024 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507062 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported the patient was on impella cp support. while on support, alarm "placement signal not reliable" triggered. no report of patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the failure mode was changed from placement to optical signal issue because the pump in question is a cp and only has an optical sensor. data logs were reviewed, and the psnr alarm triggered as reported on the 3rd day of support. at this time, signal not reliable drops to zero, contrast increases in both value and amplitude, lamp maxes out, gain increases, and light drops. these are all signatures of a fiber optic break. returned product was investigated and when connected to the console, 5s parameters were out of specification and a controller error triggered. additionally, the pump had a kink in the catheter just distal to the kink protector and had a light leak from that same location during laser continuity testing. based on data logs and returned product, the root cause of the optical signal issues was determined to be a damaged fiber line. during data log review, it was also noted that a controller error, alarm 1045, triggered roughly an hour after the placement signal not reliable alarm triggered. after discussion with the durables team, it was determined that this was unrelated to the reported complication but should be investigated. refer to complaint # (b)(4) for the investigation into that issue. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6041 | 1220648-2025-46261 | 22999386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999386&pc=OZD | 08/09/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691689 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 79 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had an impella cp pump placed to support the patient admitted in with acute myocardial infarction with cardiogenic shock. while on support, vascular surgery was called to check on the groin site as it was still oozing. care was withdrawn and the patient expired. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. minor bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6042 | 1220648-2025-46263 | 22999397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999397&pc=OZD | 08/09/2025 | 10/24/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680098 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 64 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella cp on (b)(6) 2025. the decision was made to escalate patient due to right leg mottling and lactate increasing. the patient¿s leg was pallor in color. patient survived. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6043 | 1220648-2025-31312 | 22999471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999471&pc=OZD | 02/18/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 02/18/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607595 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on impella 5.5 experienced a placement signal not reliable alarm, that was resolved after purge cassette was changed. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. device implanted (b)(6) 2025 without issue. placement signal not reliable (psnr) alarm first occurs after 4 days on support, resolves with purge cassette change. 3 days later, alarm returns and continues on and off the remainder of the case. pump position confirmed via echo on (b)(6) 2025. pump not returned. data log review confirms psnr alarms occurring intermittently after day 4 then throughout the case. every purge cassette and/or bag change resolves the alarm and returns snr values to nominal, but issue quickly returns every time. the cause of the optical signal issue was not determined since product was not returned and lack of clinical details. fm changed to optical signal issue from placement signal issue as 5.5 pumps utilize an optical sensor. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6044 | 1220648-2025-46262 | 22999490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999490&pc=OZD | 08/29/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/30/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1684102 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 128 KG | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that they encountered automated impella controller (aic) issues concerning an inaudible alarm. it was reported that there was no audio from the aic although there were active alarms. a suction alarm and placement signal low alarm were encountered with no audio coming from either alarms. it was advised to address the alarm and then change to backup aic if necessary. no further mention of the issue was noted, the patient is currently stable and remains on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown. | Additional information was received and provided (same) explant date for the impella cp and subsequently implanted impella 5.5 with a survival outcome at explant as patient expired (approximately 15 days after the aic event occurrence). the case has been referred for medical safety review accordingly. however, this follow up 1 report is being submitted as an update to the patient's status/outcome as initially reported. note: the type of reportable event remains unchanged at this time.
6045 | 1220648-2025-46265 | 22999530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999530&pc=OZD | 08/09/2025 |  |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708100 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient went into ventricular arrest and was converted four times. lidocaine and amio drips were started. electrocardiogram and angiograms were completed and based on results, the physician decided to place an impella cp for unloading. the following day in the morning labs hemolyzed. impella position in aorta alarm triggered, and the impella cp was repositioned with phone support, pulled back as it was at 7cm on echocardiogram and hitting the wall. after repositioning to 4.5cm improved waveforms were demonstrated and blood pressure was improved. of additional note, swan was coiled in the right ventricle before going to the pulmonary artery. however, the physician did not want to adjust at this time. central venous pressure read 6 without suction alarms. there were ongoing placement signal low alarms that were managed. placement signal alarms were turned back on, and right side looked good on echocardiogram. lasix was given and if suction occurs the physician noted volume will be given to the patient. urine remained clear and yellow and groin soft/stable, and there were good pulses. the following day, the patient was extubated in the evening. suction alarms were triggered, and albumin was given. lidocaine was discontinued. a hematoma was noted as the site; however, it was managed with manual pressure without further intervention. the patient continued support. two days later the patient began decompensating. impella explant planned for today was cancelled. the patient was planned for catheterization and coronary artery bypass graft surgery the next day. impella was set at performance level p-6. the following day, the patient was successfully weaned, and the device was explanted with a survived outcome. | [5 empty] | Complaint coding specifically for h6 medical device problem codes were updated to better reflect the reported event and therefore fully repopulated. note: h6, health effect - clinica/impact, component and investigation type/findings and conclusions coding remain unchanged as initially reported. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6046 | 1220648-2025-31311 | 22999548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999548&pc=PYX | 08/29/2024 | 09/08/2025 |  | 09/08/2025 |  |  | 08/29/2024 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529688 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal unreliable alarm appeared during impella rp flex support. the staff was advised to disable the audio and alarm to continue using the implanted pump. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the pump was not returned for investigation. data logs show several 1051 placement signal not reliable (psnr) alarm after a downward trend in the placement signal. the cvp 300 was active during the event, and the placement signal was returned to the normal range after 7 hours. no issues were noted with motor current and optical signal parameters. failure indicates dp sensor drift. the doctor was concerned about the negative placement signal value and psnr without a reported motor current trend. the reported failure mode of the optical signal issue was changed to a placement signal issue because the optical signal 5s parameters were within specification range and placement signal values derived from both the optical and dp sensors. dp sensor drift was confirmed from the data log review. the severity remained unchanged at 1, as the device was not removed due to the failure mode. the cause of the placement signal issue was sensor drift, based on data log review. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6047 | 1220648-2025-46264 | 22999565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22999565&pc=OZD | 08/09/2025 | 09/08/2025 |  | 09/08/2025 |  |  | 08/09/2025 | 09/08/2025 |  |  |  |  | 09/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691667 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a cp pump for the support of the patient admitted in with cardiogenic shock (cgs) and being bridged to a heartmate3. on day 13 of pump support, the pump came out of left ventricle (lv) position. the pump had "dislocated" and returned to the cath lab for pump repositioning. the pump was successfully weaned and explanted and patient was successfully placed on the left ventricular assist device (lvad). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. device in wrong position (positioning issue): no product was returned. the cause of the positioning issue was unable to be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
6048 | 1220648-2025-31266 | 22989089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989089&pc=OZD | 07/31/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 07/31/2023 | 09/07/2025 |  |  | 02/28/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311136 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump triggered a low placement signal alarm. the pump¿s position was verified to be correct via echocardiography, and the decision was made to disable placement monitoring. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump and logs were not returned for evaluation. the root cause of the placement signal issue is most likely due to the patients condition. all pertinent information available to abiomed has been submitted at this time.
6049 | 1220648-2025-31267 | 22989092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989092&pc=OZD | 09/18/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 09/18/2023 | 09/07/2025 |  |  | 04/30/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024339452 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump triggered a "placement signal not reliable" alert on impella connect. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for evaluation. upon reviewing the log values at the time the alarm occurred, an increase in lamp intensity, a decrease in gain and signal-to-noise ratio, and a "barcoding" pattern in the contrast were observed. this leads to the conclusion that there was a fiber break due to torque in the optical sensor at the time of failure. the root cause of the optical signal issue was a damaged optical fiber line. the location of the optical fiber break could not be determined as no product was returned. all pertinent information available to abiomed has been submitted at this time.
6050 | 1220648-2025-31268 | 22989097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989097&pc=OZD | 11/20/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 11/20/2023 | 09/07/2025 |  |  | 09/30/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403878 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, intermittent "position unknown" alarms occurred. the alarms were attributed to the patient's low pulsatility and, despite recurring, resolved immediately without the need for additional intervention. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data log shows several impella position unknown alarms during the entire case run, with no suction alarms reported. no relation was established between the positioning-related alarms and the optical signal parameters. the root cause of the optical signal issue was most likely patient condition related, based on data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6051 | 1220648-2025-31269 | 22989269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989269&pc=OZD | 01/10/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 01/10/2024 | 09/07/2025 |  |  | 10/31/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422960 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. approximately two weeks into support, the pump indicated an optical signal loss. despite this, the pump remained in use to continue supporting the patient. there was no harm reported to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for review. data logs confirm alarm which eventually resolves. the cause of the optical signal issue was most likely a damaged optical fiber but the exact cause of the damage could not be determined without product return and lack of clinical details. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6052 | 1220648-2025-31270 | 22989270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989270&pc=OZD | 01/26/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 01/26/2024 | 09/07/2025 |  |  | 09/30/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395892 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, a 'placement signal unreliable' alarm was triggered. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the impella 5.5 s2 was returned for evaluation. no visible damage or issues with the pump. this confirms that there was no damage to the optical fiber resulting in loss of light. data logs were reviewed. the root cause of the optical signal failure is unable to be determined as the failure was unable to be reproduced. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6053 | 1220648-2025-31271 | 22989277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989277&pc=OZD | 02/10/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 02/10/2024 | 09/07/2025 |  |  | 08/31/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024386026 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump triggered a 'placement signal not reliable' alarm. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for investigation. data logs were reviewed. the parameters were within the specification. per the clinical information proved and data log review, the root cause of the optical signal issue was not determined since product was not returned for investigation and data logs were inconclusive of the root cause. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6054 | 1220648-2025-31272 | 22989279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989279&pc=OZD | 03/15/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 03/15/2024 | 09/07/2025 |  |  | 08/31/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390730 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, a 'placement signal unreliable' alarm occurred. the alarm was disabled, support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for investigation. data logs were reviewed. the root cause of the optical signal drift issue was most likely sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6055 | 1220648-2025-31273 | 22989307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989307&pc=OZD | 03/11/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 03/11/2024 | 09/07/2025 |  |  | 10/31/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409541 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump experienced intermittent placement signal issues. the alarm was disabled, support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. impella 5.5 returned for investigation. visually inspected the pump and no catheter kinks or other damages were found. optical bench parameters are observed to be in normal range and light continuity test passed without any issues. no optical placement signal issue observed or reproduced. the optical sensor pressure test done at various set point parameters and no major abnormalities were seen. data logs were reviewed, and several placements signal low alarms were observed. the optical parameters were all found to be stable in normal range. the imc logs were reviewed and no known abnormalities relative to the failure was observed. no other issues were found. the cause of the optical signal issue was patient condition related per log and field investigation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6056 | 1220648-2025-31274 | 22989314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989314&pc=OZD | 05/09/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 05/09/2024 | 09/07/2025 |  |  | 01/31/2026 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025457868 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump lost placement signal. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for investigation. data logs were reviewed but were inconclusive. the root cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6057 | 1220648-2025-31275 | 22989316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989316&pc=OZD | 05/18/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 05/18/2024 | 09/07/2025 |  |  | 12/31/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450106 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump lost placement signal. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and multiple placement signal low alarms were noted. no issues with optical parameters were observed at time of alarms. clinical details noted that pump position was checked and found to be in proper position, with no changed in motor current. the root cause of the optical signal issue is most likely patient condition. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6058 | 1220648-2025-31276 | 22989321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989321&pc=OZD | 07/04/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 07/04/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502066 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the patient got out of bed and moved to a chair. at that time, the placement signal and left ventricular signals drastically increased. an echocardiogram was performed and revealed that the pump was shallow. the physician repositioned the impella; however, the placement signal did not normalize. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was returned for investigation. the pump passed the light test and electrical testing. the failure mode couldn't be reproduced. the pump passed the sensor pressure test. data logs showed that signal not reliable, light and lamp were stable. upon review of pump and the data logs, no problem found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6059 | 1220648-2025-31277 | 22989323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989323&pc=OZD | 08/12/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 08/12/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501637 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the pump generated a 'placement signal not reliable' alarm. the audio alarm was disabled. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data log shows that the placement signal trended downward, reaching negative values before becoming unreliable. the cause of the optical signal issue was most likely silicone wear on the sensor, as indicated by the observed drift in the placement signal, which was confirmed by the trends in the optical signal parameters in the return data log. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6060 | 1220648-2025-31278 | 22989328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989328&pc=OZD | 09/21/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 09/21/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515218 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the pump suddenly displayed 'placement signal not reliable' alarms. pump position was verified. the audio alarm was disabled. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. neither data logs nor product were returned for investigation. clinical details provide limited details that could explain why the placement signal was lost. due to lack of product return or data logs, the root cause of the optical signal issues could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6061 | 1220648-2025-31279 | 22989330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989330&pc=OZD | 09/26/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 09/26/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515201 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the pump provided support for 48 days before the optical signal was lost. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data logs were reviewed. there were no major changes noted in any of the optical signal parameters. the cause of the optical signal issue was most likely sensor silicone wear, based data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6062 | 1220648-2025-31280 | 22989332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989332&pc=OZD | 02/11/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 02/11/2024 | 09/07/2025 |  |  | 03/31/2024 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023207854 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. the complainant reported that the implanted impella cp pump lost signal on second day of support. the optical signal loss did not cause any harm or injury. the device remained in use without causing any harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump and logs were not returned for investigation. upon review of the data on the device it is apparent that the console is the potential cause of the issue. no problem was found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6063 | 1220648-2025-31281 | 22989336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989336&pc=OZD | 04/29/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 04/29/2024 | 09/07/2025 |  |  | 10/31/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444322 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump lost signal during support. the optical signal loss did not cause any harm or injury. the device remained in use without causing any harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs were returned. the root cause of the optical signal issue was not determined because no product or data logs were returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6064 | 1220648-2025-31282 | 22989338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989338&pc=OZD | 05/03/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 05/03/2024 | 09/07/2025 |  |  | 06/30/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024378542 |  | 0048-0044 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the impella cp, initially set at p2, was increased to p7; however, there was no corresponding increase in patient output. echocardiography demonstrated good myocardial motion. despite proper device placement, an alarm repeatedly indicated ¿impella in aorta,¿ raising concerns that the optical sensor may have been damaged. there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the impella cp pump was returned. visually inspected and no kinks were found. biomaterial was observed on cannula external to optical fiber/sensor face. light continuity test passed and optical bench 5s parameters were in normal range. no other issues seen. pump was connected to console and ran in bucket of water, no ps issue alarm observed. optical signal issue did not reproduce. the data logs were not returned. the root cause of the optical signal issue was not determined since data logs were not returned and due to inconclusive evidence per return product analysis. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6065 | 1220648-2025-31283 | 22989339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989339&pc=OZD | 09/22/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 09/22/2023 | 09/07/2025 |  |  | 03/31/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326205 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal issue during impella support. the position waveform was not displayed; however, there was no problem with the impella operation. the device remained in use without causing any harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product and data logs were not returned for review. the root cause of the placement signal issue was unable to be determined as limited clinical details were provided and no product was returned for review. all pertinent information available to abiomed has been submitted at this time.
6066 | 1220648-2025-31284 | 22989343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989343&pc=OZD | 10/10/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 10/10/2024 | 09/07/2025 |  |  | 01/31/2026 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025462179 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the implanted impella cp pump lost its optical signal during support after 25 days of use. the signal loss did not result in any harm or injury to the patient. the clinical team was aware that the pump had been used beyond its indicated duration. the following day, the patient was successfully bridged to a left ventricular assist device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation, and the data log was not provided. the provided clinical image shows that the pump outlet cage was cut down, and a larger purge gap was observed. according to the clinical details, a placement signal issue was reported before end of the case run. the pump was not removed during this placement signal issue. the root cause of the optical signal issue issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6067 | 1220648-2025-31285 | 22989344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989344&pc=OZD | 11/05/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 11/05/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506012 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, an alarm was triggered due to a low minimum value of the position detection signal,. however, device positioning was confirmed to be appropriate via fluoroscopy and ultrasound. the device remained in use without causing any harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data logs show that the parameters were stable throughout the case. based on data log and clinical review, the root cause of the optical signal issue is most likely due to the patients condition. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6068 | 1220648-2025-31286 | 22989354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989354&pc=OZD | 11/08/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 10/08/2023 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137667 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the console generated a positioning signal alarm. the physician confirmed via echocardiography that pump function was not affected. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console and logs were returned. the failures were unable to be replicated during investigation. the cause of the aic issue was most likely a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6069 | 1220648-2025-31287 | 22989357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989357&pc=OZD | 11/02/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 11/02/2023 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301039 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella displayed 'placement signal not reliable' alarms. the console and the pump were not replaced. engineers confirmed that the issue was related to the automated impella controller, not the pump. no harm occurred to the patient as a result of this incident. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console and logs were returned. the logs confirm that the placement signal not reliable alarm was issued on the complaint date, three times. the consoles core as well as the cladding of the optical connector fiber that was originally part of ic5320 was contaminated even after rigorous cleaning. the cladding also appears to have regressed into the outer buffer. the root cause of the placement signal issues was a defective front panel connector causing lower signal not reliable and contrast. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6070 | 1220648-2025-31288 | 22989358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989358&pc=OZD | 02/26/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 02/26/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1439042 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, a placement signal loss was observed and triggered an alarm on the console. despite the alarm, motor current and pump function remained normal. the patient stayed hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console and logs were returned and reviewed. real time logs confirmed that all signals received from optical bench. the root cause for the failure could not be determined due to difficulty to reproduce and is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of reoccurrence. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6071 | 1220648-2025-31289 | 22989361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989361&pc=OZD | 09/03/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 09/03/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1457200 |  | 0042-0010 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the optical sensor was not working. the patient stayed hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console was not returned. the data logs were reviewed and showed multiple instances across case where alarm was triggered. the root cause of the optical sensor failure alarms were due to a low signal not reliable optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6072 | 1220648-2025-31290 | 22989363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989363&pc=OZD | 02/04/2025 | 09/07/2025 |  | 09/07/2025 |  |  | 02/04/2025 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1576134 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the console displayed a 'placement signal not reliable' alarm. the intervention steps taken are unknown. no harm was reported to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console and data logs were returned and reviewed. the log shows that the placement signal dropped to a negative value and gradually became more negative. this console was tested and was unable to reproduce the reported failure mode by running the optical test pump. there was no issue with the analog output, confirming that there was no problem with the optical bench. the root cause for abnormal placement signal was not determined due to the inability to reproduce. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6073 | 1220648-2025-31291 | 22989372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989372&pc=OZD | 09/20/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 09/20/2023 | 09/07/2025 |  |  | 04/30/2025 |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366087 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, an optical signal issue during impella support. it was confirmed that the motor waveform was functioning normally. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned. the optical bench parameters were verified over the console and all the parameters were in normal range. no other abnormalities were observed. data logs were reviewed, and the placement signal low alarm was observed. the root cause of the optical signal issue was patient condition related based on log, clinical review and field investigation. all pertinent information available to abiomed has been submitted at this time.
6074 | 1220648-2025-31292 | 22989374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989374&pc=OZD | 07/16/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 07/16/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521569 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the impella displayed 'placement signal unreliable' alarms. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs were returned. the 5s parameters alone are not sufficient to conclude the failure reason. therefore, the root cause of the optical signal issue was not determined since product was not returned for investigation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6075 | 1220648-2025-31293 | 22989377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989377&pc=OZD | 12/26/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 12/26/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565517 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. a shockwave was given to the patient and the waveforms went away on the console. flows and motor current still worked as normal. the pump was still supporting the patient. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned product was inspected and the 5s parameters were taken which showed very low signal not reliable. the data logs show 5s parameters characteristic of a sensor damage pattern. the root cause of the optical signal issue was due to a damaged sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6076 | 1220648-2025-31294 | 22989378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989378&pc=OZD | 06/14/2023 | 09/07/2025 |  | 09/07/2025 |  |  | 06/14/2023 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1677314 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal alarm occurred. proper positioning was verified by fluoroscopy. no harm was done to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs confirm multiple instances of suction, impella position unknown and placement signal low. these alarms could likely have been triggered due to patient condition. the console was inspected in the lab. there was no observable anomaly when a known good pump was run with the console. the root cause could not be determined due to failure not being reproducible and therefore not reproducible. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6077 | 1220648-2025-31295 | 22989381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989381&pc=OZD | 10/22/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 10/22/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025511273 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us medical center had placed a 5.5 pump to support a transfer patient with cardiogenic shock (cgs). the pump replaced the cp used along with extracorporeal membrane oxygenation (ecmo) for the transferred patient. the 5.5 was supporting when the optical signal was lost. the pump signal loss did not prompt any intervention or explant. the patient remained on the pump for support with no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. no product was returned for evaluation. data logs were reviewed and controller error alarms were noted with placement signal not reliable alarms. at time of controller error alarms, all optical parameters spike out of range and then immediately return. clinical details noted there was a loss in placement signal and controller error for less than 30 seconds. pump continued to flow and no change in hemodynamics were noted per staff. upon review of the data logs, definitive automated impella controller (aic) failure type is suspected. the root cause of the optical signal issue cannot be definitively determined as no product was returned to confirm no issues with the pump. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6078 | 1220648-2025-31296 | 22989382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22989382&pc=OZD | 03/21/2024 | 09/07/2025 |  | 09/07/2025 |  |  | 03/21/2024 | 09/07/2025 |  |  |  |  | 09/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103727 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. the patient stayed hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs showed that the alarm occurred. the product was returned and the failure mode was able to be reproduced. the root cause of the console issue was a software issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6079 | 1220648-2025-46170 | 22987226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987226&pc=PYX | 08/06/2025 | 10/30/2025 |  | 09/06/2025 |  |  | 08/06/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026651089 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device and impella rp flex for mechanical circulatory support. it was reported on the first day of impella rp flex support that the purge pressure was found to be higher than expected, and there were no alarms that had indicated the high purge pressure. the following day, there were alarms for high purge pressure noted. the medical team was advised to follow hospital protocols, and were advised to change purge tubing, check for kinks in tubing, and verify the patient¿s activated clotting time was in therapeutic range. the medical team followed hospital protocol to add tissue plasminogen activator to the purge solution. the following morning, the high purge pressure alarms resolved, and the purge solution was switched back to bicarbonate solution. additionally, it was reported on (b)(6) 2025 that the patient lacked pulses in their left leg, where both impella cp and rp flex were inserted, and the patient¿s right arm. it was noted that the patient had previous deep vein thrombosis and previous surgical complications which most likely contributed to the limb ischemia. the vascular team was consulted, but no further information was provided on any interventions that were performed for limb ischemia. on (b)(6) 2025 it was reported that the impella rp flex placement signal results were inaccurate as it was reporting flows higher than expected at p-4 with a negative pulmonary artery pressure. there was no change in motor current and no change in the patient condition. the patient remained on support until heart function recovered. the impella cp was successfully weaned and explanted, and the following day the impella rp flex was successfully weaned and explanted. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: rp flex sn (b)(6) returned for investigation. high purge pressure: clinical reported a "high purge pressure" alarm on the console a few hours after case start. the pump was returned for investigation and tested for high purge pressure. the issue did not reproduce, maintaining a purge pressure of ~400 mmhg which is within product specifications. no abnormalities were observed on the product. data log review shows a stable placement signal prior to starting the pump, indicating the pump was properly primed. a bag change was started at 21:39, but it was not finished until four minutes later at 21:43. per ifu, the bag change should be performed within 90 seconds. about 7 minutes after the bag change was completed, purge pressure begins to rise and spikes up, triggering "high purge pressure" as well as "purge system blocked" alarms. the motor current also becomes dampened at this time, which indicates biomaterial ingestion. tissue plasminogen activator (tpa) was added to the purge fluid, and the issue was resolved. the cause of the high purge pressure was most likely biomaterial ingestion that occurred due to a long bag change, as seen by the dampened motor current and timing of the bag change at the time of the alarms, which was resolved by tpa. the high purge pressure also did not reproduce in testing. placement signal issues: clinical reported that the console was showing higher flows than normal at p-4, and placement signal was negative with no change in motor current. the product was returned for investigation and no damage to the pump was observed. data logs confirm the negative placement signal and higher flow readings that occurred on the fourth day. motor current remains pulsatile and consistent, and p-level isn't changed at this time. the differential pressure (dp) fluctuated down to negative, but it does not trigger a placement signal not reliable alarm. in physical product investigation, the electrical resistance readings of the dp sensor line was below spec, reading at 108 ohms. the resistance from sensor v+/v- should fall between 3900 and 5000 ohms. there was no biomaterial or damage observed on or around the sensor. there was no blood observed in the blue handle and no holes were found in the catheter. the cause of the placement signal issue was most likely and electrical short in the dp voltage line, as the resistance readings were below specification, but the cause of the short was not determined. device history lot: device lot number: 1899487. device history batch: dp sensor subcomponent lot: 2025532117. device history review: pump sn (b)(6) passed all post sterile inspection criteria. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. there are two associated devices for the reported event. the first pump (impella rp flex) is addressed under manufacturer report number. the second pump (impella cp) is addressed under manufacturer report number 1220648-2025-46177. | Medical safety review of the limb ischemia event for association given bipella support (impella cp via left femoral artery and impella rp flex via left femoral vein) was completed and noted the report of limb ischemia goes against the impella cp as it could be pressing against/occluding flow to the artery which is associated with limb ischemia. coding update: as the limb ischemia is unrelated to the rp flex, section h6 (health effect-clinical code) has been updated, corrected accordingly. note: b5 (event description), b1 (adverse event/product problem), b2 (outcomes attributed), and h1 (type of reportable event) and h6 (health effect-impact, medical device problem, component coding) remain unchanged. device status: additional information noted the impella rp flex device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related report numbers: this impella rp flex report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the limb ischemia.
6080 | 1220648-2025-46178 | 22987368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987368&pc=OZD | 08/07/2025 | 10/16/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672888 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 66 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Puncture/Hole | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support developed a right neck hematoma. anticoagulant was held and one unit of packed red blood cells was given. the patient was taken for neck exploration. the surgeon was attempting to get control of the bleeding of graft and punctured the graft with needle multiple times. noted blood by the red plug. the patient was explanted following impella damage during neck exploration. veno-arterial extracorporeal membrane oxygenation continued. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. | B.7 added information that was inadvertently omitted from the initial report that was submitted. d.4 added as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number, zip code and zip code extension as they were inadvertently omitted from the initial report that was submitted. g.1 manufacturer site fax number and manufacturer site phone number should not have been reflected on the initial report that was submitted as they were already reflected in the manufacturer contact fax number and manufacturer contact phone number fields. h.6 codes f1905 was reported incorrectly on the initial report that was submitted and a new code was added. the investigation was completed. the impella 5.5 was returned for investigation. data log review was not required for this case run. bleed/hematoma: the cause of the bleed/hematoma was not determined since sufficient clinical information was not provided. product damage (hole in catheter): the returned product was visually damage-free. the red handle was opened, and blood was found inside. further inspection of the catheter revealed multiple suture holes at the 20 centimeter mark. the cause of the product damage was related to unintended use error during suturing. device history review: the pump with passed all post-sterile inspection checks.
6081 | 1220648-2025-31247 | 22987736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987736&pc=OZD | 03/19/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 04/30/2025 | 09/06/2025 |  |  | 07/12/2022 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1670339 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day 12 of impella support, a placement signal 'not reliable alarm' occurred on the automated impella controller. the audio was disabled. the team monitored motor current. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | The cause of the placement signal not reliable alarm is likely due to a failing bulb in the optical bench lamp assembly.
6082 | 1220648-2025-31248 | 22987757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987757&pc=OZD | 03/31/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 03/31/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103512 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the aortic placement signal did not correlate with arterial line pressures. motor current and flow remained within expected limits for the set p-level. the investigating engineer attributed the issue to the automated impella controller (aic). no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs confirmed that the console had significantly low signal not reliable. the console was returned and while running a pump, there was again low signal not reliable throughout the case. the root cause of the low signal low reliable while running pumps was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6083 | 1220648-2025-31249 | 22987762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987762&pc=PYX | 05/05/2025 | 10/24/2025 |  | 09/06/2025 |  |  | 05/05/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025570426 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, the placement signal displayed inaccurate values, and flow measurements were likely unreliable. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the fm optical signal issue was changed to placement signal issue as data log and product review indicate it was the dp sensor causing the issue. the data logs show that the placement signal starts to gradually drift down while flows improve again. it remains relatively stable throughout the case and no placement signal not reliable alarms were triggered. the returned product the placement signal drifted down while the flows drifted up. based on the data logs and returned product testing, the root cause of the placement signal issue is most likely due to dp sensor drift. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date added.
6084 | 1220648-2025-31250 | 22987766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987766&pc=OZD | 05/24/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 05/26/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577208 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant in japan had placed an impella cp pump to support a patient. the pump was placed but optical signal loss occurred perhaps due to damage during application of pressure at the site dressing. the pump remained on despite the placement signal loss, and there has been no harm or injury. the pump was successfully weaned and explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. after cp placement and before pt transfer to the intensive care unit, a dressing film was applied while maintaining the puncture angle, but during this process, the placement signal (ps) received by the optical sensor disappeared. it was possible to continue to manage the condition by changing the angle and moving the catheter back and forth and fixing it at a point where the waveform did not disappear. data log review showed low ps alarms but never triggering the placement signal not reliable alarm (psnr). ps briefly spikes to 3000mmhg before quickly recovering. issue happens intermittently over the first day of support. during ps railing, the signal-to-noise ratio goes to zero and contrast decreases. gain and lamp both remain constant. ps railing occurs less frequently as case goes on. the pump was returned and inspected. the pump passed uson and laser continuity testing. there was biomaterial found wrapped around the impellor blades but it was determined it had no impact on the optical signal. the cause of the optical signal issue was not determined since returned product showed no abnormalities related to the optical signal, data log review was inconclusive, and insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6085 | 1220648-2025-31251 | 22987769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987769&pc=PYX | 05/27/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 05/27/2025 | 09/06/2025 |  |  | 02/28/2026 |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478693 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella rp flex and cp for bipella support. the right pump lost optical signal. the team considered the use of a non-abiomed wire at insertion may have caused the malfunction. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. optical signal issue: clinical suspected damage to the optical sensor causing loss of placement signal and flows. no problem was found as the root cause of the optical signal issue as investigation revealed sensor drift caused loss of placement signal and no issues with optical sensor during support placement signal issue: the data logs show a drift down in placement signal with a corresponding drift up of pump flows for about an hour before placement signal went out of range and flow calculation was lost. the flow calculation and placement signal were lost for about 8 hours before recovering. central venous pressure was also out of range during this time. the pump's 5s parameters were stable and within normal range confirming no issues with the optical sensor. the pattern observed is characteristic of sensor drift of the dp sensor. the root cause of the placement signal issue was determined to be due to a sensor drift as evidenced by the data log analysis indicating sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6086 | 1220648-2025-31252 | 22987773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987773&pc=OZD | 05/29/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 05/29/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669218 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an implanted impella cp pump was lost during patient support. support continued with the implanted pump until planned escalation later that day. the pump was weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. clinical details report that the placement signal went out within hours of case start. neither the pump nor console were replaced due to the complication. data logs were reviewed and at roughly 07:30, contrast began to oscillate between +/-3200 and all other 5s parameters were affected. roughly 7 hours later, contrast and all other 5s parameters returned to acceptable values, and the placement signal was returned to calculating physiological values. the intermittent nature of the issue indicates that there was no failure of the optical sensor/fiber line in the pump. product was not returned for investigation. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6087 | 1220648-2025-31253 | 22987777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987777&pc=OZD | 05/28/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 05/28/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 84 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had an impella cp implant and the optical sensor was damaged upon insertion. troubleshooting and appropriate steps were followed to no success. the physician was notified that the cp would still be able to provide support but would not be able to get a placement signal. the doctor acknowledged this and continued insertion and use of the impella cp. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. clinical details report that the placement signal not reliable (psnr) alarm triggered out of box. the pump was still used to support a percutaneous coronary intervention procedure. no further details were provided. data logs were reviewed, and from the first rt log, all 5s parameters were out of specification. of note, gain was ~2.44 and light was ~1.9. these are potentially indications of damage to the optical fiber line, however, without returned product, other failures cannot be definitively ruled out. the root cause of the optical signal issue could not be determined since product was not returned for investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6088 | 1220648-2025-31254 | 22987779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987779&pc=PYX | 05/27/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 05/27/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025550997 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp flex placed to support a patient. the pump was placed, however, there was an optical signal loss that did not prompt an explant or replacement. pump remained on for support despite malfunction. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. rp differential pressure sensor went out two hours after transferring to the intensive care unit. flows based entirely off motor current and p level. sensor was zeroed by physician and sensor pulmonary artery (pa) waveform returned and automated impella controller displayed flow rates again. the sensor would continue to drift from monitored pa catheter values over the next 24 hours and was zeroed again. device required zeroing approximately every 24-48 hours for the duration of use. data log review confirmed placement signal drift, trending flows, and placement signal not reliable alarms. the pump was returned and inspected. no visual abnormalities were noted and pump passed electrical testing. the pump was flow loop tested for 48 hours and the drift was able to be recreated. drift reproduced over the course of the first 2 hours where placement signal went from 0 to -70mmhg. placement signal then recovered over the next 26 hours from -70 to 0mmhg. the cause of the placement signal issue was sensor drift as the data logs showed drift and the drift was able to be recreated during investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6089 | 1220648-2025-31255 | 22987796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987796&pc=OZD | 05/29/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 05/29/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669833 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 73 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that post placement of an impella 5.5, the placement signal was noted to be approximately 30 pts higher than patient's mean arterial pressure from arterial line and blood pressure cuff. all other parameters including waveforms, flows, and motor current appeared within normal range. the team monitored. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. according to the clinical notes, the placement signal was reported to be higher than the a-line reading. the data log shows a placement signal of 100-150 during the case run, with no abnormalities noted on the 5s optical signal parameters. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6090 | 1220648-2025-31256 | 22987823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987823&pc=OZD | 06/02/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 06/02/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657542 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced placement signal issues. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data log was not provided or retrieved from remote server. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6091 | 1220648-2025-46179 | 22987837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987837&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697173 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during the insertion of the impella, a lot of force was applied at the distal end of the long sheath due to aorta stenosis. after the impella was delivered and support was initiated, flows would not go past 1 l/min. the decision was made to explant the impella cp and replace it with a new device. the patient did not experience any complications with the new device and was successfully weaned and explanted after heart function recovered. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: impella cp sn (b)(6) returned for evaluation. difficulty advancing pump: the return pump was visually inspected and no physical abnormalities or related defect found. as per clinical patient had aortic stenosis causing difficulty advancing pump. the cause of the difficulty advancing pump issue was patient condition related due to stenosis. low pump flow: returned cp complaint pump sn (b)(6) visually inspected. no cannula kink, no biomaterial or broken blade observed. reproducing low flow test performed and pump run for more than 10 minutes at p9. no low pump flow observed and flows were within specified range. data logs were reviewed and suction alarms occurred. but no motor current spike, positioning issue or placement signal issues seen. the cause of low flow issue cannot be determined since the issue did not reproduce with return device, no conclusive log evidence and insufficient clinical information. device history lot: device lot: 1943231 device history review: the pump sn (b)(6) passed all post sterile inspection checks.
6092 | 1220648-2025-46180 | 22987845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987845&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707246 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 91 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp and escalated to an impella 5.5 device for continuation in mechanical circulatory support (mcs). the following day after start of the impella 5.5, the patient had a right m1 occlusion involving the majority of the right temporal lobe, frontal and parietal mid complexities with perfusion abnormalities. the physician noted the finding as grim. support continued with plan to wean as necessary and assess functions and effects from stroke. the patient would continue on support for an additional four days before expiring. care was withdrawn given the poor prognosis. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for impella 5.5 with smartassist system: section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | Product complaint # (b)(4). investigation summary : stroke : in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot : device lot : 1942023. device history batch : sub-component lot : n/a. device history review : the pump s/n: (b)(6) passed all the post sterile inspection checks.
6093 | 1220648-2025-46181 | 22987849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987849&pc=OZD | 08/06/2025 | 11/03/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696954 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 109 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device and the decision was made to immediately escalate to impella 5.5 increased support. the patient experienced hemolysis and subsequently expired. the patient was admitted to the hospital post cardiac arrest, ventricular tachycardia (vtach) and ventricular fibrillation (vfib), requiring multiple rounds of defibrillation and over one hour of resuscitation efforts. return of spontaneous circulation was obtained and patient was then taken catheterization lab with successful intervention to the left anterior descending artery with single drug-eluting stent. post procedure, while in the unit, the patient began to deteriorate hemodynamically requiring increased doses of several vasopressors. decision was made to bring patient back to lab to establish mechanical circulatory support (mcs) with impella cp. immediately after the patient was escalated to the impella 5.5 device for increased hemodynamic support. two days later the patient was noted to be hemodynamically stable; oxygenation significantly improved. continuous renal replacement therapy was started. another two days later, hemolysis was suspected. the patient was on concurrent extracorporeal membrane oxygenation (ecmo) support with impella 5.5 at support level p-3. hemoglobin was 160 along with lactate dehydrogenase greater than 1000. possible configuration to veno-arterial-venous or veno-veno ecmo was noted. after the next two days, the patient began experiencing north-south syndrome with regained pulsatility. the patient was taken to catheterization lab to reconfigure for vav ecmo with venous oxygenated return in the right internal jugular. pulse oxygenation saturation improved from 80s to 94%. however, two days later there was suspicion of spinal ischemia as patient was not moving extremities, and the patient would expire later this day. care was withdrawn | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for impella 5.5 with smartassist system: section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section use of echocardiography for positioning of the impella catheter - ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section impella catheter too far into the left ventricle - ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hemolysis, clinical details noted that after hospital transfer, patient had increased lactate dehydrogenase and hemoglobin. the pump position unable to be confirmed at that time. patient also on continuous renal replacement therapy and extracorporeal membrane oxygenation (ecmo) support. two days later, pump position was confirmed. no details provided if hemolysis resolved during pump run. data logs show no persistent suction or positioning alarms around time of reported hemolysis. no motor current spikes indicating ingestion were observed. the cause of the hemolysis could not be definitely determined as limited clinical details were provided, and no product was returned for evaluation. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6094 | 1220648-2025-46182 | 22987862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987862&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701882 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support (mcs) and, post implant same day, experienced runs of ventricular tachycardia that caused reduced flows with a mean arterial pressure (map) of 40. the impella flows were suboptimal at 2.7 l/pm and a 500 ml was ordered in an effort to improves flows. the impella was at performance level p-2 with suction occurring with attempt to increase flows. the patient remained acidotic and impella cp device flows never improved. a radial art line was being placed. impella mcs continued for an additional three days. the patient was successfully weaned off ecmo and impella cp support in the operating room and cannulas and impella cp device were surgically removed with a survived patient outcome. | [5 empty] | The impella device was not received from the customer. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical reported impella at p-2 with suction occurring with attempted higher flows. the day prior also reported patient had v-tach runs that caused reduced and suboptimal flows. patient also supported on extracorporeal membrane oxygenation (ecmo) with 400ml of blood loss since ecmo started. the logs confirm suction alarms. clinical does not report any positioning issues. the cause of the issue was likely patient condition related as evidenced by concurrent ecmo use and suction at attempt of higher flows with ventricular tachycardia causing reduced flows. regarding the tachycardia, in order to make a root cause determination, essential clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6: investigation: type, findings and conclusion codes and h6: component code were updated accordingly based on the completed investigation.
6095 | 1220648-2025-46183 | 22987883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987883&pc=OZD | 05/01/2024 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  | 12/31/2025 |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456708 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Cardiogenic Shock | Thrombosis/Thrombus | [10 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that experienced deterioration of cardiogenic shock with an "unknown (undetermined)" relationship to the impella device used, occlusive impella inflow and limb ischemia with a possible relationship to the impella device used. the patient would ultimately expire. the incapacitated patient presented with st-segment elevation myocardial infarction post in-field tpa via lifeflight and was rushed into emergent left heart catheterization. percutaneous coronary intervention (pci) was completed to the right coronary artery with one drug-eluting stent. prior to pci, an intra-aortic balloon pump was inserted, and the patient was started on dopamine and levophed. during the pci, the patient went into sustained ventricular tachycardia (vt) requiring cardioversion 200j, and then subsequently went into atrial fibrillation with rapid ventricular response (afib rvr) and was given adenosine and atropine. at this time, the was on levophed 10 and iabp 1:1. same day in the afternoon, the patient went into cardiac arrest and required cardiopulmonary resuscitation x2 and defibrillation. the patient continued to deteriorate with worsening cardiogenic shock overnight. the patient was not deemed a candidate for oht, extracorporeal membrane oxygenation or other therapies. however, the patient was placed on continuous renal replacement therapy (crrt). status was changed to do not resuscitate. the following day in the afternoon, the iabp was removed and impella cp was inserted via the right groin. however, impella sheath clotted post-implant, so the patient received heparin and left and right crossover sheath. the next day, lactate remained elevated. concern from impella clot occlusion for right lower extremity ischemia. the patient continued to worsen, was changed to comfort care and subsequently expired this day. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the following events and device malfunctions were investigated and noted the following: low flow: the data log review showed brief suction alarm with drop in placement signal pulsatility on first day of support that resolved within one hour. there were no further suction alarms or low flow observed. motor current is stable. the cause of the low flow was not determined due to no product return and lack of clinical details. ischemia/thrombosis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. cardiogenic shock: the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella introducer. | Patient-specific information section a4: weight and a6: race is unknown. this is one of two medical device reports associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella sheath introducer. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for the impella cp with smartassist system: section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6096 | 1220648-2025-46185 | 22987889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987889&pc=OZD | 08/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677839 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, the user accidentally stepped on the cord and the pump stopped. the issue could not be troubleshot despite efforts and restarting attempts. the team had to replace the 5.5 (and console). no harm or injury noted from the pump stop. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | H6: added code g02004. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary impella 5.5 returned for evaluation. pump stop: clinical details report that the patient stepped on the pump cable, and it unplugged from the console and stopped. the pump couldn't be restarted with various consoles. data logs were reviewed, and all pump metrics were stable through (b)(6) 2025, when a sudden pump stop alarm #115 triggered with no mc spike. attempts to restart the pump failed with more alarm #115's and mc spikes above 32,000. returned product was investigated, and all motor cable phases had open lines. the pump was unable to be restarted, triggering alarm #115 and #13 pump stops. there were no catheter kinks, so the red pump plug was opened, and necking/fracture of the motor cable wires was confirmed on the catheter side. based on the clinical details provided that the patient stepped on the cable and returned product had open motor cable lines with clear stretching and fractures, the cause of the pump stop was determined to be an open electrical line due to an unintended use error. device history review pump sn (b)(6) passed all post-sterile inspection checks.
6097 | 1220648-2025-31257 | 22987925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987925&pc=OZD | 06/05/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 06/05/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622089 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced optical signal failure issues. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs were returned. the root cause of the optical signal issue was not determined as no product or data logs were returned for analysis. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6098 | 1220648-2025-31258 | 22987948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987948&pc=PYX | 06/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 06/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025602829 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. shortly after implantation, a placement signal issue occurred. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the fm was updated from optical signal issue to placement signal issue since data log review and returned product review revealed an issue with the dp sensor, not the optical sensor. the returned product was inspected. the pump had a confirmed electrical short with low resistance readings. the data logs from the case show a drift down in placement signal coinciding with an upward drift in pump flow before placement signal went out of range and flow calculation was disabled. the root cause of the placement signal issue was most likely an electrical short due to the low sensor resistance readings but the cause of the short is unknown. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6099 | 1220648-2025-31259 | 22987970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987970&pc=OZD | 01/21/2023 | 09/06/2025 |  | 09/06/2025 |  |  | 04/25/2023 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155705 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller exhibited an unreliable placement signal. as a result of this ongoing issue, the alarms were intermittently disabled. patient support continued using the implanted pump, and no harm was caused to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console logs from the reported day of event are consistent with complaint. initial testing did not reveal any anomalies; however the issue was successfully reproduced after 3 days during long-term testing. data analysis and troubleshooting performed with regards to other known complaints with same symptoms indicated that the issue could be caused either by underpowered optical bench or an undetermined issue with the optical bench itself. by elimination, we¿ve determined that the issue was caused by defective lamp assembly of the optical bench. the root cause of the issue was a defective component. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6100 | 1220648-2025-46186 | 22987990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987990&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536897 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was inserted on a patient without issue with heparin onboard. the impella was ramped up in auto. however it was reported that after about 10 minutes the pump stopped. attempts were made to restart however they were unsuccessful. decision was made to remove the pump and replace with a new cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. | The investigation of the reported failure mode has been completed. no product was returned for evaluation. pump stop: data logs were reviewed and no pump stop alarms were observed over 1 hour long term log available. real time logs showed fluctuating motor current and motor speed but again no pump stop alarms. the cause of the pump stop could not be definitively determined as insufficient data logs were provided and no product was returned for evaluation. device history review: the device passed all the post sterile inspection checks.
6101 | 1220648-2025-31260 | 22988003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988003&pc=OZD | 04/15/2023 | 09/06/2025 |  | 09/06/2025 |  |  | 05/23/2023 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123074 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a automated impella controller exhibited sensor failure. the sensor did not prompt any intervention. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console logs from the reported day of event are consistent with complaint and show two occurrences of ¿placement signal not reliable¿ alarms. the returned console was tested, the issue was reproduced. it¿s been determined that the issue was most likely caused by defective lamp assembly, but in abundance of caution, entire optical bench would be replaced. the root cause of the issue was a defective component. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6102 | 1220648-2025-46187 | 22988005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988005&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing for over 42 days for the patient who had been admitted in with cardiogenic shock (cgs) and cardiomyopathy. during support, the user accidentally stepped on the cord resulting in pump stop. the team tried to get the original console to restart. the console needed to be replaced. no harm or injury noted from the console stop. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: console logs from the reported day of event are partially consistent with complaint and show that 42 days and 16 hours after case start, pump motor current suddenly dropped to 0 and impella stopped alarm (#115, due to mechanical or electrical issues) presented. device analysis: console was inspected and tested in the lab, and there was no damage to the optical pump connector, which might be expected if plug is forcefully pulled out. the console was able to detect a test pump and run it, with valid values of motor current and placement signal. inspection of the console revealed broken purge disk detect flag (unrelated to complaint), and significant contamination of purge pressure sensor area, presumably due to prior purge fluid leak. device history review: the device passed all post sterile inspection checks.
6103 | 1220648-2025-46188 | 22988016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988016&pc=OZD | 08/07/2025 | 10/21/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615235 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella 5.5 pump placed to support the patient admitted in for cardiomyopathy. while on support, the impella exit site noted to have purulent drainage. the surgeon discussed plan and decided to explant the impella due to an infection at the site. the impella was removed and surgical site explored. the white blood cell count and lactic acid levels were normal, no signs of bacteremia or sepsis at that time. blood cultures and wound cultures were sent; results were pending | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported major bleed and infection has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed and infection could not be determined.
6104 | 1220648-2025-31261 | 22988018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988018&pc=OZD | 05/03/2023 | 09/06/2025 |  | 09/06/2025 |  |  | 05/03/2023 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1682677 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller displayed an optical sensor error/failure. the console was returned for repair. no harm was reported to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs showed that the calculated placement would suddenly spike at the beginning as if the pump connector was not fully seated in the blue pump receptacle. the console was returned and reviewed. the failure mode was not reproduced before and after cleaning the fiber parameters were within specification. ran optical pump simulator for few hours but there were no alarms related to an abnormal placement signal. the placement signal issues could not be reproduced during testing, the root cause was unable to determine. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6105 | 1220648-2025-31262 | 22988022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988022&pc=OZD | 05/01/2023 | 09/06/2025 |  | 09/06/2025 |  |  | 06/12/2023 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller exhibited sensor failure. the sensor did not prompt any intervention. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console was tested in the engineering lab and the error unable to reproduce. the fringe pattern show that the optical bench was working as expected. the root cause of the issue was not determined since issue couldn't be reproduced. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6106 | 1220648-2025-31263 | 22988037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988037&pc=OZD | 06/06/2023 | 09/06/2025 |  | 09/06/2025 |  |  | 06/06/2023 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024325850 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. the automated impella controller displayed a 'placement signal not reliable' alarm, which resulted in the loss of placement and left ventricle signal. the intervention steps taken are unknown. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump and data logs were returned. the data logs showed the console triggered a placement signal unreliable error alarm that did not resolve until the pump was unplugged. the parameters for the console showed error. the cause of the low signal not reliable was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6107 | 1220648-2025-31264 | 22988040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988040&pc=OZD | 06/12/2023 | 09/06/2025 |  | 09/06/2025 |  |  | 06/12/2023 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1649323 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, a 'placement signal unreliable' alarm occurred. proper positioning was verified by fluoroscopy, and the physician was able to complete the atherectomy. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs confirm numerous optical-signal-not-reliable events. the failure mode was reproduced in the lab. optical bench test was taken before and after cleaning of the optical bench blue receptacle connector. the first results showed low signal not reliable which was consistent with the reported complaint. the root cause was an air gap caused by a dirty optical receptacle. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6108 | 1220648-2025-46189 | 22988043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988043&pc=OZD | 08/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026712242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, bleeding was noted from groin site and manual pressure was held. heparin was stopped. furthermore, upon impella removal, it was noted that the 6 french angioseal did not deploy properly and was removed. the doctor initially held pressure for 10 minutes and then the staff held an additional 15 minutes. hemostasis was achieved and groin site was dressed by the staff. the doctor examined groin site again after the dressing was applied. the groin site was clean, dry and intact. it was soft with no hematoma noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. access site bleeding: the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
6109 | 1220648-2025-46190 | 22988052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988052&pc=OZD | 08/07/2025 | 10/21/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026655213 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a patient on impella cp experienced some oozing at right groin, in addition to a oozing at a previous venous site. dressing was taken down and angle readjusted and sutured. the impella was successfully weaned and explanted without issue. perclose x3 utilized to close groin. patient is stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleeding has been completed. the root cause of access site bleeding was most likely use issue related since hemostasis was achieved by adding suture and readjusting the angle.
6110 | 1220648-2025-46192 | 22988062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988062&pc=PYX | 08/09/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026659326 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Crack | Increase in Pressure | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that a patient was on bipella support and experienced bleeding around rp flex repo sheath insertion site, 2 unites of blood were given in additional to a suture place at bedside to controlled the bleeding. additionally, there was a purge system blocked alarm that cleared on its own after 1.5-2hrs. impella flows were minimal 1.2l/min @ p-6, decision was made to remove rp flex. after removal of device it was noted the blue impella plug and purge side arm was under the patient¿s foot/heel against footboard of the bed. the side arm appeared kinked to some extent, the patient was sedated however, legs moved some which would account for the purge pressure intermittent issues. | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp and rp flex device for bipella mechanical circulatory support (mcs). patient with chest pain for 2 days and presented to the emergency room with late presentation of inferno lateral myocardial infarction with hypotension. lactate was normal. the patient was sent to catheterization lab for left and right heart catheterization and underwent percutaneous coronary intervention (pci). the patient was still hypotensive after successful thrombectomy with pnumbra and stenting of the mid-proximal right coronary artery. the impella cp device was implanted, and the patient was sent to the unit on mcs. the following day after start of the impella cp support, the patient experienced hematuria and multiple hemolyzed blood samples collected from the patient. physician pulled back impella cp with echocardiogram at bedside. based on a hemodynamic marker for the right ventricular function and tricuspid annular plane systolic excursion measurement, the physician elected to proceed with impella rp flex support. the patient now on bipella support. prior to leaving the catheterization lab sites were dry. upon arrival at the unit, some bleeding was noted. staff felt blankets weighing down the impella catheter and propped back up. the patient was bleeding ¿profusely¿ from groin site, which was managed by placing additional purse string suture around site. two units of blood were given. the team was unsure if the impella cp (right femoral artery) or impella rp flex (right femoral vein) site was bleeding. bipella support continued. (additional information subsequently received noted the bleeding was from the impella rp flex access site). two days later, purge system blocked occurred with the impella rp flex. the physician spoke with medical affairs for tissue plasminogen activator protocol. alarm cleared on its own, however, after 1.5-2hours. impella flows were minimal 1.2l/min at p-6 performance level. the patient remained stable despite minimal right-sided support and the impella rp flex was explanted. after removal of the impella rp flex, it was noted the blue impella plug, and purge side arm were under the patient¿s foot/heel against footboard of the bed. the side arm appeared kinked to some extent. the patient was sedated; however, the patient's legs move some which the team noted accounted for the purge pressure intermittent issues. the swan-ganz catheter and sheath from right internal jugular were also found to be pulled back several centimeters. team was advised about risk of clot ingestion with device and needed to decrease impella flow to p-2 performance level during catheter manipulation to decrease chance of ingestion of biomaterials into device. impella cp support continued. however, the following day the patient would expire. care was withdrawn. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Section b5 was corrected to include complete details surrounding the event, which includes additional information that was received regarding the bleeding event and added b6: tests/lab data including dates. medical safety review: the patient on impella cp and had issues with hemolysis. the impella cp cannot be dissociated with the outcome. the patient was on bipella support. there was access site bleeding requiring transfusion; which access site was bleeding was initially undetermined and then found to be the impella rp flex. following, the patient was able to manage with minimal rp flex support given the purge system blocked issue; therefore, the rp flex was explanted. this malfunction and bleeding event is a serious injury and should be reported. the impella cp was present at the time of death and is the more likely contributor. however, the patient had significant arrhythmias and in this case is what led to a significant decline in the patient's function and the family decided to withdraw support which appears to have led to along with the patient's underlying condition the demise outcome. the medical safety officer reviewed the event and noted the same in that impella cp device cannot be dissociated from the demise outcome. and the impella rp flex is a serious injury type of reportable event. device status: additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related medical device reports: this impella rp flex report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome. | The investigation of the reported failure mode has been completed. the pump was returned for analysis. high purge pressure - data log review confirmed purge system blocked occurred after a purge pressure rise over 4 minutes with no motor current (mc) spike. the returned pump had biomaterial on the impellor blades and in the purge gap, however high purge pressure was unable to be recreated. the cause of the high purge pressure was biomaterial build up due to the purge pressure rising over 4 minutes with no mc spike and the biomaterial found on the returned pump. product damage - after removal of device it was noted the blue impella plug and purge side arm were under the patient¿s foot/heel against footboard of the bed. the side arm appeared kinked to some extent. the returned pump confirmed a crack on the purge side arm protective covering. the cause of the material crack on the sidearm retainer was unintended use of the device being set under the patient causing the crack. patient code: major bleed - the cause of the bleeding was not determined due to insufficient clinical details. device history lot- device lot: 1901681. device history batch- subcomponent lot: n/a. device history review- pump s/n (b)(6) passed all post sterile inspection checks.
6111 | 1220648-2025-46193 | 22988067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988067&pc=OZD | 08/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706530 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 64 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in non-st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the patient had absent distal pulses and mottling of color in the leg the impella was placed in. an external femoral to femoral bypass was performed and the patient¿s leg color improved. however, the patient continued to deteriorate hemodynamically and the medical team determined the patient needed to be escalated to an impella 5.5 for increased hemodynamic support. the patient was successfully escalated to an impella 5.5. it was reported that during impella 5.5 insertion into the right subclavian artery, the guidewire would not advance and there was resistance at the at innominate artery. contrast-enhanced echocardiography was performed and confirmed the artery was occluded. impella 5.5 insertion was then transitioned to the left subclavian artery, and the impella 5.5 was inserted without further complications. upon removal of the impella cp a thrombectomy was performed, and the impella cp access site was surgically closed. it was reported 4 days after impella 5.5 support was initiated, the patient was noted to have right sided weakness and forced gaze. computed tomography revealed a large stroke to left middle cerebral artery. two days later, the family decided to withdraw care of the patient. the patient¿s care was withdrawn and the patient expired. this complaint involves 2 pumps: pump 1: impella cp with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). this mdr specifically addresses impella cp with serial number (b)(6), which is the subject of this report. a separate mdr will be submitted for the other device associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety review was completed and the following was noted: the impella cp supported the patient for one day and experienced limb ischemia and thrombus before being taken to the operating room for escalation of impella 5.5 with a survived outcome. this is a serious injury and should be reported. while on the impella 5.5 the patient experienced a middle cerebral artery stroke and expired two days later. there is not enough information to dissociate the impella 5.5 from the outcome. however, of note, the patient presented with non-st-segment elevation myocardial infarction and hemodynamically deteriorated becoming tachycardic and diaphoretic. after the embolic stroke there was no neurological intervention; the neurological prognosis was discussed with family, and ultimately, family decided to withdraw all care leading to the patient's demise. the medical safety officer reviewed the event and noted the same in that impella 5.5 device cannot be dissociated from the demise outcome. and the impella cp is a serious injury type of reportable event.
6112 | 1220648-2025-46194 | 22988071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988071&pc=OZD | 08/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694623 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 64 KG | Hispanic | White | Death | Yes | Yes | Difficult to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in non-st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the patient had absent distal pulses and mottling of color in the leg the impella was placed in. an external femoral to femoral bypass was performed and the patient¿s leg color improved. however, the patient continued to deteriorate hemodynamically and the medical team determined the patient needed to be escalated to an impella 5.5 for increased hemodynamic support. the patient was successfully escalated to an impella 5.5. it was reported that during impella 5.5 insertion into the right subclavian artery; the guidewire would not advance and there was resistance at the at innominate artery. contrast-enhanced echocardiography was performed and confirmed the artery was occluded. impella 5.5 insertion was then transitioned to the left subclavian artery, and the impella 5.5 was inserted without further complications. upon removal of the impella cp a thrombectomy was performed, and the impella cp access site was surgically closed. it was reported 4 days after impella 5.5 support was initiated, the patient was noted to have right sided weakness and forced gaze. computed tomography revealed a large stroke to left middle cerebral artery. two days later, the family decided to withdraw care of the patient. the patient¿s care was withdrawn and the patient expired. this complaint involves 2 pumps: pump 1: impella cp with serial number: (b)(6), pump 2: impella 5.5 with serial number: (b)(6). this mdr specifically addresses impella 5.5 with serial number: (b)(6), which is the subject of this report. a separate mdr will be submitted for the other device associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.1 added selection as it was inadvertently omitted from the initial report that was submitted. b.2 added selection that was inadvertently omitted from the initial report that was submitted. h.6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation findings. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ investigation summary: difficult to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had resistance at the innominate artery preventing successful delivery of impella. stroke : in order to make a cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | Additional information was received and noted there was no neurological intervention after the stroke was discovered, which was confirmed to be embolic. the family quickly decided to withdraw all life sustaining measures. medical safety review: the impella cp supported the patient for one day and experienced limb ischemia and thrombus before being taken to the operating room for escalation of impella 5.5 with a survived outcome. this is a serious injury and should be reported. while on the impella 5.5 the patient experienced a middle cerebral artery stroke and expired two days later. there is not enough information to dissociate the impella 5.5 from the outcome. however, of note, the patient presented with non-st-segment elevation myocardial infarction and hemodynamically deteriorated becoming tachycardic and diaphoretic. after the embolic stroke there was no neurological intervention; the neurological prognosis was discussed with family, and ultimately, family decided to withdraw all care leading to the patient's demise. the medical safety officer reviewed the event and noted the same in that impella 5.5 device cannot be dissociated from the demise outcome. and the impella cp is a serious injury type of reportable event. device status: the additional information also included the device status which noted the pump was unfortunately not saved. should the device or any new information be received, a supplemental mdr will be filed.
6113 | 1220648-2025-46196 | 22988072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988072&pc=OZD | 08/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668975 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 123 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.5 device for mechanical circulatory support. during support, patient experienced hemolysis. the physician stated that the impella caused hemolysis which lead to creatine elevation to 3.5 and lactate dehydrogenase (ldh) elevation to over 1,000. the pump ran without alarms and volume/position status were troubleshot by the physician. the physician also repositioned the impella overnight and one unit of packed red blood cells (prbc¿s) was given to the patient for hemoglobin drop. the impella was then removed from graft and no bleeding was noted at surgical site. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ | The investigation was completed and the product was not returned for investigation. hemolysis: clinical reported suspected hemolysis based on the patient's elevated creatine levels and ldh. data log reviewed showed a suction alarm on the first day of the case, where motor current, flow level, and motor speed briefly dropped. there were also alarms for impella in aorta and ventricle at the beginning of the case that were resolved. the placement signal was trending downward before and at the time of the suction alarm, and clinical reported troubleshooting the position and volume conditions, but did not specify any positioning issues or if any volume was administered. the motor current and motor speed are consistent with one another, but systolic motor current decreased in pulsatility around the same time the suction alarms occurred. there are also no motor current spikes occurring at any point in the case. the cause of hemolysis was not determined due to insufficient clinical details and no product returned. the device history review was completed and passed all post sterile inspection criteria.
6114 | 1220648-2025-46195 | 22988076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988076&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555629 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 40 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the first impella cp pump stopped and was immediately replaced with an impella cp pump 2; however, this pump had suction and required intervention. the patient would ultimately expire. the patient presented in pulseless electrical activity during cardiac resynchronization therapy defibrillation, received cardiopulmonary resuscitation (cpr) and venoarterial extracorporeal membrane oxygenation (va ecmo) with return of spontaneous circulation achieved post cannulation (this occurring three days prior to the impella implant). there was a noted pulmonary hemorrhage from the cpr; neuro remained intact. however, a computed tomography scan of the head showed a 5mm subdural hematoma. while recovering in the intensive care unit, it was noted the patient had recurrent supraventricular tachycardia/ventricular tachycardia while on ecmo. the decision was made place the patient on impella mechanical circulatory support (mcs) and protek duo with decannulation of ecmo. of note, an impella 5.5 was considered with protek duo. however, due to the nature of the patient's size and vessels, the decision was made to place impella cp axillary which was successfully completed. approximately 13 days after start of the impella mcs, the pump suddenly failed/stopped. the patient was emergently taken cardiovascular lab to replace the impella cp pump 1, which was removed from axillary site without an issue. the site was then closed. a new impella cp pump 2 was placed. however, the next day suction alarm was triggered, and consequently, the patient was given blood and albumin. subsequently, the next day, the patient expired. | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. after 12 days of impella cp support, the pump had sudden pump failure without rise in motor current until the event. the pump as a result was immediately removed and a new impella cp was used instead. no patient harm was reported. |  |  |  |  | Medical safety review was completed and noted the following: the event describes a pump stop/failure with impella cp pump 1 and suction events requiring intervention with impella cp pump 2, which contributed to the serious injury. however, there is not enough information to exclude impella cp pump 1 as an associated factor in the demise outcome. the immediate issue is the pump failure and is what led to or contributed more to the demise outcome. medical safety officer reviewed the event and noted the same in that the impella cp pump 1 cannot be dissociated for the demise outcome and impella cp pump 2 event resulted in a serious injury and should be reported. change in reportability: after review of the case by the medical safety officer, it was determined the impella cp pump 1 is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect - clinical and impact codes) have been updated, corrected accordingly. additionally, section b5 (event description) was updated to include complete event details. related medical device reports: this impella cp pump 1 report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp pump 2. device status: additional information noted the impella device pump 1 was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿. | Product complaint # (b)(4). investigation summary: pump (b)(6) and purge tube returned + data stick returned separately on 9/2/2025. pump stop - after 12 days on support, pump stop occurred. data log review confirmed pump ran for 12 days, with pump stop occurring with mc spike. the returned product recreated the pump stop upon attempted restart and imaging under the keyence shown a large impella purge gap, indicative of bearing wear. the cause of the pump stop was material wear related to the ball bearings. this cp pump failed on day 12 of support, beyond the intended design life of a pp of 8 days per design trace matrix (b)(4). device history lot: device lot: 1833664. device history batch: subcomponent lot: n/a. device history review : q1) any qns/mrrs/reworks associated with the complaint device: no. q2) were there any failures of the device to meet the requirements for manufacturing testing, qualification, and inspection?: no. q3) were there any other product malfunction complaints in this device lot related to this failure mode?: no. phr summary: s/n (b)(6) passed all post sterile inspection checks.
6115 | 1220648-2025-46197 | 22988081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988081&pc=OZD | 08/07/2025 | 09/11/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1283654 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 122 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, there was a purge disc not detected on day 83 of the support. the team had to replace the aic and support proceeded without harm or injury. | The patient received a heart transplant and survived. |  |  |  |  | Additional information was provided in d6 and h6. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B5 updated with the patient outcome. d6a and d6b should have been left blank as the console is external device. h5 was erroneously selected; yes on the initial manufacturer device report. h6 component code was erroneously entered on the initial manufacturer device report. h8 was erroneously selected on the initial manufacturer device report.
6116 | 1220648-2025-46198 | 22988083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988083&pc=OZD | 08/07/2025 | 10/10/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700759 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Break | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had gastrointestinal bleeding (gi bleed) and product damage during impella 5.5 support. during pump implant, the physician had to surgically cutdown to close as a clot was noted at the profunda. afterwards, tegaderm dressing was placed on axillary site as well as 3 point fixation with number side up on red plug, and the patient was then returned to intensive critical unit (icu) for continued management. during support, gi bleed was continuously noted with bloody bowel movements. as a result, bivalirudin and anticoagulants were stopped and several units of blood and platelets were administered to the patient. moreover, the patient ripped the impella catheter apart into 2 pieces at the red plug. the initial plan was to replace the pump, however, the patient deteriorated, despite inotrope/pressors, and did not want to be reintubated or proceed with care. discussions were made since patient continued to endure active severe gi bleed, had undergone multiple previous failed weans and ultimately the decision was made to withdraw care per patient request. the patient subsequently expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the access site bleeding and product damage has been completed. the available logs from the case do not show any electrical pump stop alarms. the root cause of the product damage was determined to be use related on the evidence of the patient intentionally damaging device and refusing care. the root cause of the access site bleeding was not determined due to insufficient clinical details.
6117 | 1220648-2025-46199 | 22988084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988084&pc=OZD | 08/08/2025 | 10/28/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711535 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 62 KG | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. there was a crack at the impella 5.5 red plug and consistent drip of purge fluid noted. purge pressure and purge flow maintained. troubleshooting was done but unsuccessful. a couple days later, impella 5.5 was removed due to the leak by the red plug. no patient harm was reported due to the issue. | [5 empty] | The investigation for the reported product damage (pump crack) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the product damage (pump crack) could not be determined. h6 med dev prob code a0404 added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6118 | 1220648-2025-46200 | 22988086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988086&pc=PYX | 08/08/2025 | 10/24/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - |  |  | US |  | 2025575277 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the patient received blood product. a ct scan showed a small retroperitoneal hemorrhage and heparin was held. patient is stable post issue. | [5 empty] | Patient ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported retroperitoneal hemorrhage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the retroperitoneal hemorrhage could not be determined.
6119 | 1220648-2025-46191 | 22988087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988087&pc=OZD | 08/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026702670 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented for high-risk percutaneous coronary intervention with impella cp support. it was noted that the sheath was inserted into right femoral artery. wire access obtained into left ventricle. however, while attempting to insert impella the pump would not pass thru lower aorta. despite a couple of attempts it seemed as if it was hanging up and would not pass. pump pulled out and intervention done thru 14fr impella sheath. post procedure, sheath was removed and great hemostasis achieved using 2 perclose. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Device was not returned for investigation. the cause of the deliver issue was determined to be patient condition as the patient had stenosis in the lower aorta preventing successful delivery of impella. the pump s/n: (b)(6) passed all the post sterile inspection checks.
6120 | 1220648-2025-46201 | 22988088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988088&pc=OZD | 08/08/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026684232 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was causing ventricular tachycardia. the impella was run shallow for the duration of the pci procedure. impella was then weaned and explanted. patient is stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint #(b)(4). investigation summary: no product was returned. tachycardia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the root cause of the positioning issue was unable to be determined as insufficient clinical details were provided. device history lot: device lot: 1937629. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
6121 | 1220648-2025-46202 | 22988090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988090&pc=OZD | 08/08/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698563 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had a cat scan and was positive for cerebrovascular accident (cva). additionally, the patient had loss of distal pulses. the impella was removed, support was for 26 hours. successfully weaned. patient not requiring any inotropic or pressor support and is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. no product was returned for investigation. ischemia/ stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. a device history review was completed and passed all post sterile inspection checks.
6122 | 1220648-2025-46203 | 22988092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988092&pc=OZD | 08/08/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708057 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. post impella implant left leg without dopplerable pulse and decision made to utilize reperfusion catheter. post reperfusion catheter placement left leg showed to have laminar flow distally. additionally, the operating room staff reported continuous suction. impella was said to be in correct position. then again it was reported of multiple suction events noted over despite administering multiple blood products and albumin. it was also noted that 5500 cc clots were removed from chest cavity due to leaking jp drain. the patient remained critically ill, presenting very vasoplegic. prognosis was grim. ultimately the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Pump was not returned for investigation. clinical symptoms(ischemia): in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: data logs show several suction alarms until noon on 8-august. the placement signal was fluctuating and decreasing trend observed during suction events. the motor current also showed fluctuations while running at a steady p-level. according to clinical reports, the patient was on ecmo support with the pump in use, and multiple blood products were administered on 9-august to address the suction. by 10-august, the patient was tolerating p2 levels without suction after the blood products were given. the cause of the pump suction issue was most likely patient condition based on data log review and provided clinical details. the pump sn (b)(6) passed all post sterile inspection checks.
6123 | 1220648-2025-46204 | 22988094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988094&pc=OZD | 08/08/2025 | 10/02/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620590 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 52 KG | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was supported with impella cp for approximately an hour, but it was difficult maintaining adequate flows. troubleshooting was done but unsuccessful. the impella was exchanged for another cp. | [5 empty] | Patient race is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. data logs confirmed low pump flow when pump started & remained to the rest of the case. per data log review, no mc spikes observed. product was not returned for review. the cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
6124 | 1220648-2025-46205 | 22988095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988095&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026686484 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 96 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. after 39 hours of impella 5.5 support the device was exchanged due to high purge pressure alarm. tissue plasminogen activator (tpa) was unsuccessful in resolving the alarm. no patient harm reported. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the pump was not returned. data log review showed purge pressure (pp) high alarms occurring after a bag change. 1 hour leading up to the bag change, the purge pressure had dropped from 600 to 200mmhg and the purge flow had dropped to 0 l/min. it was noted that the purge flow ticks had flatlined at this time as well. the cause of the purge pressure high was not determined as no product was returned, the data log review was inconclusive, and insufficient clinical details were provided. device history review: the pump passed all post sterile inspection checks.
6125 | 1220648-2025-46206 | 22988096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988096&pc=OZD | 08/08/2025 |  |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636273 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 106 KG | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during impella 5.5 support, there were alarms for placement signal low and that pressures displayed where significantly lower than pressures obtained from a non-invasive blood pressure (nibp) cuff. an echocardiogram was performed to confirm placement of the impella device. placement monitoring was then disabled. there was no harm to patient reported. the medical staff was not able to use placement signal monitoring for the duration of impella support and monitored motor current and patient¿s pressure through nibp. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | Investigation summary: the pump (cut) was returned. optical signal issue: clinical details report placement signal (ps) low alarms, suction alarms, and inaccurate ps readings. placement signal monitoring was disabled, and the placement signal not reliable (psnr) alarm triggered a few days later. data logs were not available for investigation. the pump cannula was the only thing returned. since the pump was cut, standard testing for optical signal issue could not be conducted. there were no visual abnormalities or defects noted on the cannula or motor housing. since insufficient product and data logs were available for investigation, the cause of the reported optical signal issue could not be determined. device history review: the pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. h6 codes and section d9 have been updated to reflect product return.
6126 | 1220648-2025-46208 | 22988099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988099&pc=OZD | 08/08/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026652092 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | Thrombosis/Thrombus | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp device was implanted for a patient admitted in with acute myocardial infarction/cardiogenic shock and ventricular fibrillation. the patient was also placed on extracorporeal membrane oxygenation (ecmo). cardiac function temporarily recovered with the impella cp and ecmo support, however, subsequently, pulse pressure disappeared again. due to right heart failure, an oozing rupture was detected during percutaneous coronary intervention (pci) of the remaining branch, and drainage was performed. it was noted that because it was ¿merely¿ oozing, surgical treatment was not performed, and the patient was monitored. however, it was determined that the cardiac function recovery would be prolonged; therefore, a thoracotomy was performed, and hemostasis was performed with plan to upgrade the patient to an impella 5.5 device. additional information noted that there was no relationship between the event and the device. thoracotomy revealed no obvious active oozing. it was determined that a blood clot due to bleeding during chest compression had likely adhered to the pericardium and was interfering with cardiac contraction. hemostasis was not performed, and only clot removal was performed. physician¿s comments stated oozing rupture was suspected after cp insertion. no connection to the device was noted. | [5 empty] | Given the information at hand including the additional information with the physician¿s comments, there is not enough information to exclude the impella cp device from the event. patient-specific information: patient details other than race and ethnicity is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # ==> (b)(4). investigation summary ==> major bleed: the root cause of major bleed could not be determined as insufficient clinical details were provided. thrombosis: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1906637. device history batch ==> subcomponent lot: n/a. device history review ==> device (sn: (b)(6)) passed all post sterile inspection checks.
6127 | 1220648-2025-46207 | 22988101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988101&pc=OZD | 08/08/2025 |  |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692840 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and placed on extracorporeal membrane oxygenation therapy. the following day after start of impella mcs, constant suction alarms from impella device support level p-2 were encountered. the risk of adverse events as a result of unresolved suction alarms was discussed with the care team and the decision was made to maintain impella support at p-2. an echocardiogram was completed and showed the impella position as appropriate and left ventricle was decompressed. the patient was noted to be in stable condition following the event. | [5 empty] | Product complaint # (b)(4). investigation summary: pump suction: clinical details report that the pump was in continuous suction while at p-2 midday 8/aug/2025. the pump was confirmed to be in good position, but the patient was on ecmo and there was low volume in the ventricle. no additional volume was added. data logs were reviewed and the report was confirmed. the case had 125 continuous suction #73 triggers, and 44 continuous suction #14 triggers, indicating low mcmax values, or flows during systole. additionally, there were 14 triggers of position unknown #33, indicative of low mc modulation and low ps pulsatility. the alarms, specifically suction #73 were persistent throughout the case. a closer look at the lt log on 8/aug showed pump flows were averaging 0.47 l/min (spec: 1.1-2.1 l.min). there were no signs of improper positioning or ingestion events in data logs. this all aligns with the clinical details provided that the the patient was on ecmo and had low volume in the lv. product wasn't returned for investigation. based on the low mc values/flow calculations and alarms in data logs aligning with the clinical details provided, the cause of the pump suction was determined to be patient condition related. device history lot: device lot: 1936081. device history batch : subcomponent: na. device history review: pump sn (b)(6) passed all post-sterile inspection checks. | A4: weight is unknown. device status: the impella device was not received from the customer as it was discarded after use. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6128 | 1220648-2025-46209 | 22988102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988102&pc=OZD | 08/08/2025 |  |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693554 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and the follow, and approximately 12 days into the support, the patient reported the placement signal was unreliable. it was further noted an echocardiogram was completed and showed the pump's position as appropriate. the outcome of the patient as a result of the event is unknown. however, the patient continued on impella mcs for an additional 7 days before being successfully weaned and the device explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and noted post event, the pump was successfully and regularly used and then weaned without any further problems. the patient remained stable after explantation. regarding the status of the device it was not returned as it discarded by the hospital staff. per the update, the staff was unaware a complaint intake would be created upon explantation of the device. should any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. optical signal issue - after 12 days on support, placement signal not reliable (psnr) alarms occurred. echocardiogram confirmed proper positioning. data log review confirmed psnr alarms occurring with placement signal intermittently railing to 3000mmhg. raw optical sensor rails to -3200mmhg and does not form a coherent aortic waveform. signal not reliable and motor current oscillate together in no known pattern. cross functional review could not identify any known pattern in the data logs, and the issue is thought to be unrelated to the automated impella controller. the cause of the optical signal issue could not be determined due to no product returned, inconclusive data log review, and insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6129 | 1220648-2025-46211 | 22988103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988103&pc=OZD | 08/08/2025 | 11/18/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701628 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and the following day, there was bleeding at the impella access site. the nurse noted the bleeding did not appear to be coming from the arterial. it appeared to be instead coming from the skin tract. perclose sutures were already in place and tightened around the impella catheter. quikclot was placed at the access site as well and the bleeding subsequently resolved. the patient was administered two units of red blood cells as a result of the event. post the event the patient was noted to be in stable condition. impella support continued for an additional eight days before being weaned and explanted with a survived patient outcome. the patient continued on extracorporeal membrane oxygenation support thereafter. | [5 empty] | Section a4: weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the major bleed, the cause was most likely use issue related since hemostasis was achieved by tightening the perclose sutures. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6130 | 1220648-2025-46210 | 22988104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988104&pc=OZD | 08/07/2025 | 10/14/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703516 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support the patient was receiving albumin and volume for suction due to low central venous pressure. suction resolved slowly. the following day it was reported that heparin with held due to gastrointestinal bleed. later it was reported that care was withdrawn and the patient expired. | [5 empty] | B.2 added selection as it was inadvertently omitted from the initial report that was submitted. b.3 revised date of event as it was entered incorrectly on the initial report that was submitted. b.7 information was added as it was inadvertently omitted from the initial report that was submitted. d.7a revised as it was entered incorrectly on the initial report that was submitted. e.1 added address line 1 as it was inadvertently omitted from the initial report that was submitted. h.6 revised medical device problem code since the initial report was submitted. the investigation has been completed. pump suction: no product was returned. clinical information reported the patient received albumin and volume for suction overnight due to low central venous pressure and suction slowly resolved. the data logs confirmed suction alarms and showed no other related log abnormalities related to issue such as positioning alarms or ingestion signatures. the cause of the issue was likely patient condition related as evidenced by issue resolving after volume administration and no data log abnormalities. major bleed: clinical information reported heparin was on hold due to a gastrointestinal bleed. the cause of the clinical symptoms could not be determined as insufficient clinical details were provided. the device history review was completed and the pump passed all post sterile inspection checks. | A4 is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
6131 | 1220648-2025-46213 | 22988105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988105&pc=OZD | 06/10/2021 | 10/29/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  | 02/28/2023 |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022051953 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Study | Health Professional | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Syncope/Fainting | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complaints team received three stemi-dtu study notifications reporting the following: possible relationship between the impella device and procedure to the patient's ventricular tachycardia. post procedure, patient was having frequent runs of ventricular tachycardia. transthoracic echocardiography was ordered and 500cc of normal saline bolus was administered, probable relationship between the impella device and procedure to the patient's impella site oozing. the impella site was saturated with blood and actively bleeding. the nurse held pressure, and decision was made to bring patient back to catheterization lab for removal. probable relationship between the impella device and procedure to the patient's vasovagal episode. prior to the impella removal, the patient had a vasovagal event and briefly required levophed for blood pressure support. the patient also received an additional 1 l of intravenous fluids as well as narcan and flumazenil to reverse the sedation. based on the report the above noted events occurred on the same day. the outcome of the patient as a result of these events is unnoted. | [5 empty] | Patient-specific information including implant and explant dates is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hemorrhage/bleeding and vasovagal reaction, the root cause of the injury was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. pertaining to the arrhythmia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6132 | 1220648-2025-46215 | 22988106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988106&pc=OZD | 08/08/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661023 |  | 000000000001000100 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and approximately nine days after start of the support, placement signal issue occurred. there were no changes in the patient's hemodynamics or motor current. the care team informed that pump performance is not affected and was walked through on disabling audio. there was no harm to the patient as a result of this event. | [5 empty] | The following corrections were identified: additional information was received via medsun reporting program - user facility medwatch report (b)(4) and was not attached to the follow up 1 report submitted on 22-oct-2025. (e4 was updated and user facility medwatch report has been attached to this follow-up 2 report accordingly). details of the additional information received were the following: "the display of the impella device stopped showing left ventricle and aorta tracing. fiberoptic cable malfunction with device." | Patient-specific information section a5: ethnicity is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint #(b)(4). g.2 medwatch report was received. additional information was received from the field rep and noted the optical sensor error was transient and no longer an issue at the time of impella explant. it was surmised that perhaps the optical sensor may have been "stunned" but not necessarily damaged intra-operatively as the error was transient. investigation: the impella pump was not returned; therefore, an evaluation/analysis of the device was not possible. however, the event logs were received and analyzed. the investigation was completed and noted a device history review confirmed the pump passed all the post sterile inspection checks. regarding the optical sensor issue data logs were reviewed, and the placement signal not reliable alarm was confirmed. at this time, signal not reliable drops to zero, lamp maxes out at 100%, light decreases, gain decreases, and contrast oscillates between +/-3,000. the real time logs were closely examined during the time of the placement signal not reliable event, and the oscillating contrast trend was seen. the pump remained in use during placement signal not reliable issue. optical signal parameters and placement signal was seen returned after 36 hours of case run. upon review of data logs, it is apparent that the console is the potential cause of the issue. h6: investigation type, findings and conclusion codes and component codes were updated accordingly based on the completed investigation. a request for an automated impella controller (aic) complaint was made, and a complaint file was created and processed accordingly. refer to the related manufacturer report number as noted in section h10 for the report on the automated impella controller.
6133 | 1220648-2025-46214 | 22988107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988107&pc=OZD | 08/08/2025 | 10/21/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709026 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. a decision was made to place impella cp for support via axillary access due to patients heavily calcified lower extremity access sites. multiple attempts to advance impella cp into the left ventricle (lv) but were all unsuccessful due to patient¿s anatomy. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported delivery issue has been completed. the root cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella.
6134 | 1220648-2025-46216 | 22988108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988108&pc=OZD | 08/08/2025 |  |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636921 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, continuous suctioning issues were noted. during start of procedure, continuous suction was experienced and held throughout duration of procedure and position was confirmed several times prior to transport to intensive critical unit (icu). additionally, after patient was settled in icu, pump experienced continuous suction and position of the impella was confirmed again. however, the patient was in a critical condition and family decided to withdraw care. patient subsequently expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B: added product malfunction h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: pump suction: the pump was not returned for investigation. it was implanted at noon on 8-august and used for 15 hours. data logs were reviewed, and the pump was noted to be in suction condition throughout the case. during the first 7 hours of the case, a slowly improving pump flow trend was observed while the pump operated at a steady p-level, with noted fluctuations/trending placement signals. a motor current spikes were reported around 7 hours into the case, accompanied by a drop in pump flow and trending placement signals until the end of the case. clinically, the iabp site was bleeding during the time of suction, and blood was administered during the case. the doctor confirmed good pump position and ruled out rv issues. the cause of the pump suction was most likely related to the patient condition, based on data log review and the provided clinical details. device history review: the pump passed all post sterile inspection checks.
6135 | 1220648-2025-46217 | 22988109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988109&pc=OZD | 08/06/2025 | 10/21/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025541877 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 78 YR |  | Female | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had a cp pump placed for support placed for the patient with a cardiac history and known aortic valve replacement . the pump was delivered across the valve with efforts, and while on the support the pump site developed a hematoma that did require 2 units of red blood cells and continued monitoring. the pump support continued despite the access site injury for 5 days and went into atrial fibrillation with rapid ventricular response and was cardioverted x3. that day the team made decision to explant and replace the cp with an intra-aortic balloon pump. patient survived the event. | [5 empty] | The investigation for the reported hematoma and arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma and arrhythmia could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6136 | 1220648-2025-46218 | 22988110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988110&pc=OZD | 08/07/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653699 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 147 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as an escalation from the impella cp and decannulated from venoarterial extracorporeal membrane oxygenation and after approximately 47 days of support, the patient was taken to the operating room with a plan to change from right axillary impella 5.5 to left axillary impella 5.5 due to persistent infection of the right axillary surgical site. the surgeon turned down the impella to p-2 and saw patient was pulsatile not requiring pressors, heart failure was called and decided to remove the right axillary impella 5.5 and then c-shock called to decide on replacing the impella 5.5 in left axillary or leaving patient unsupported. no further information was noted other than the device explant date two days later. the outcome of the patient is unknown. | [5 empty] | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | B.7 revised information as it was previously entered incorrectly on the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the product was received and the investigation was completed. device history review was completed and the pump passed all post sterile inspection criteria. clinical details reported on the 16th day, an infection was observed at the right axillary impella site and cultures came back with gram negative bacilli. the returned product had no kinks or abnormalities observed on the pump. biomaterial was observed on the outside of the cannula, but it likely occurred during explant and it's unlikely that it's related to the patient injury. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and component code were updated accordingly based on the completed investigation.
6137 | 1220648-2025-46219 | 22988112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988112&pc=OZD | 08/08/2025 | 10/28/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708992 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the first impella cp pump stopped and was immediately replaced with an impella cp pump 2; however, this pump had suction and required intervention. the patient would ultimately expire. the patient presented in pulseless electrical activity during cardiac resynchronization therapy defibrillation, received cardiopulmonary resuscitation (cpr) and venoarterial extracorporeal membrane oxygenation (va ecmo) with return of spontaneous circulation achieved post cannulation (this occurring three days prior to the impella implant). there was a noted pulmonary hemorrhage from the cpr; neuro remained intact. however, a computed tomography scan of the head showed a 5mm subdural hematoma. while recovering in the intensive care unit, it was noted the patient had recurrent supraventricular tachycardia/ventricular tachycardia while on ecmo. the decision was made place the patient on impella mechanical circulatory support (mcs) and protek duo with decannulation of ecmo. of note, an impella 5.5 was considered with protek duo. however, due to the nature of the patient's size and vessels, the decision was made to place impella cp axillary which was successfully completed. approximately 13 days after start of the impella mcs, the pump suddenly failed/stopped. the patient was emergently taken cardiovascular lab to replace the impella cp pump 1, which was removed from axillary site without an issue. the site was then closed. a new impella cp pump 2 was placed. however, the next day suction alarm was triggered, and consequently, the patient was given blood and albumin. subsequently, the next day, the patient expired. | The complainant reported that the patient encountered suction during impella cp support. due to a suction alarm, the patient was given blood an albumin as a result. however, after 2 days of support, the patient expired. no further information was provided. |  |  |  |  | Medical safety review was completed and noted the following: the event describes a pump stop/failure with impella cp pump 1 and suction events requiring intervention with impella cp pump 2, which contributed to the serious injury. however, there is not enough information to exclude impella cp pump 1 as an associated factor in the demise outcome. the immediate issue is the pump failure and is what led to or contributed more to the demise outcome. medical safety officer reviewed the event and noted the same in that the impella cp pump 1 cannot be dissociated for the demise outcome and impella cp pump 2 event resulted in a serious injury and should be reported. change in reportability: after review of the case by the medical safety officer, it was determined the impella cp pump 2 is a serious injury type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect - clinical and impact codes) have been updated, corrected accordingly. additionally, section b5 (event description) was updated to include complete event details. related medical device reports: this impella cp pump 2 report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp pump 1, which is associated with the demise outcome. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B7 additional information was added that was inadvertently omitted from the initial report that was submitted. d7a revised selection as it was entered incorrectly on the initial report that was submitted. e1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. revised initial reporter phone number as it was entered incorrectly on the initial report that was submitted. investigation summary: the product was not returned for investigation. hemodynamic instability: the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. pump suction: clinical details report suction alarms on (b)(6) 2025. blood and albumin were administered, and there were no further concerns, however, the patient expired on support the following day. data logs were reviewed, and the suction alarms were confirmed. until mid-afternoon on (b)(6) 2025, pump flows were within specification and all other pump metrics were stable. however, around 16:00, mcmax values started to trend down while the pump was running at p-8. this resulted in increased dpmax values, ultimately higher than the ps at mcmax, indicating continuous suction. for this reason, the low pulsatility algorithm triggered. these alarms triggered on and off for the remainder of support. this aligns with the clinical details provided: blood volume and albumin were administered, reportedly resolving suction, but the patient expired on support the following day, explaining intermittent alarms throughout the remainder of support. based on the clinical details provided and review of suction alarms in data logs, the cause of the pump suction was determined to be patient condition related. device history review: the pump passed all post-sterile inspection checks.
6138 | 1220648-2025-46220 | 22988118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988118&pc=OZD | 08/01/2025 | 10/07/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571820 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 76 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Malposition of Device | [5 empty] | Fever | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Purge system blocked alarm. md notified and discussed tpa protocol and to reach out to abiomed medical director. plan was to take 5.5 out already this morning so he made decision just to leave purge system as is and proceed with already planned or time this morning to explant device. fever: (b)(6) 2025: patient spiked a fever overnight; however, wbc down today. care team suspects patient has an infection and that it might have been the cause of her valve issue. patient started on cefepine for suspected infection. patient received concentrated albumin today and 1/ 10 pack of platelets. (b)(6) 2025: patient started on vancomycin yesterday with the cefepine for antibiotic coverage. positioning issues: (b)(6) 2025: echo performed today to check position. position shallow at 2.5, so dr. (b)(6) advanced position. new cm marking 44 with new lv measurement of 5.7. (b)(6) 2025: echo performed this am, impella position showed it had advanced to 6.4 from 5.7; however, icu team did not want to reposition at this time. thrombocytopenia: (b)(6) 2025: patient remains on crrt as she is a long-term dialysis patient. (b)(6) 2025: patient remains on crrt with goal to pull 100 ml/hr. (b)(6) 2025: crrt circuit has clotted 3-4 times in last 24 hours. still no anticoagulation due to low platelet count and concern for hit, although panel was negative when sent a few days ago. the us complainant had an impella 5.5 pump placed to support the patient for an upcoming mitral valve surgery. while on support, the patient spiked a fever overnight; however their white blood cell count was down. the care team suspected an infection and the patient was started on cefepine and vancomycin. an echocardiogram was performed and found the pump was shallow at 2.5cm. the pump was advanced; the new cm marking was 44 with a new left ventricle measurement of 5.7cm. another echocardiogram showed the pump advanced to 6.4cm from 5.7 cm but the icu team did not want to reposition at that time. the patient remained on continuous renal replacement therapy as they were a long term dialysis patient. there was a concern for heparin induced thrombocytopenia although the panel was negative a few days prior. there was a purge system blocked alarm. the doctor was notified and discussed tissue-type plasminogen activator protocol and an abiomed medical director was contacted. the device was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue and high purge pressure(hpp) has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. data log review showed purge pressures rise over 12 hours leading up to hpp alarms with decreases in purge flows. no mc spikes noted outside of p-level changes. the root cause of the pump positioning issues was not determined due to the pump not being returned and insufficient clinical details. the cause of the high purge pressure was biomaterial buildup due to the 12 hours of purge pressure trending up in parallel with purge flows trending down.
6139 | 1220648-2025-46221 | 22988120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988120&pc=OZD | 08/08/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691878 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Intraoperative Cardiac Injury | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was supported by cp pump and extracorporeal membrane oxygenation (ecmo) but, the team made decision to explant and escalate to the 5.5 in the operating room on day 9 of the support. the explant was being performed when the team observed a tamponade/lv damage that required surgical repair and red blood cells (rbc) infusion. the patient later expired on the 5.5 pump. the cause of the effusion and cp causing/contributing to it was not known. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Product complaint # (b)(4). investigation summary : heart injury/hemorrhage : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot : device lot : 1926327. device history batch : sub-component lot : n/a. device history review : the pump s/n: (b)(6) passed all the post sterile inspection checks.
6140 | 1220648-2025-46222 | 22988122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988122&pc=OZD | 08/07/2025 | 10/17/2025 |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Consumer | User facility |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during use in the cath lab the console had a battery failure alarm. they switched the power switch on underside of console to ¿on¿. attempted switch off for 60 seconds, and then back to on and then restarted without resolution of alarm. the console was exchanged. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1, e.3 and e.4 added information that was inadvertently omitted from the initial report that was submitted. the investigation has been completed. data analysis: the console logs confirmed the reported failure. there was a battery failure and battery 1 comm. failure alarm issue due to loss of communication to the battery 1. further investigation into the automated impella controller (aic) logs revealed that the aic had been unused for months prior to the reported issue. additionally, the aic was booted several times, but the alarms continued to persist. device analysis: the console was tested on a lab bench, where the failure mode was successfully reproduced. upon arrival, battery 1 was found to be dead, which confirmed the aic was kept unplugged for months. upon replacing the original batteries with known good ones, the battery 1 was still unable to charge. it was also determined that the original power battery manager (pbm) had an issue with charger 1. however, when the original pbm was replaced with a known good one, both batteries charged successfully. additionally, no battery-related alarms or issues were observed when running the aic with a known good pump and purge cassette for an hour. root cause: the battery failure and battery 1 comm. failure alarms was due to the combined effect of a defective battery 1 and a defective pbm. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6141 | 1220648-2025-46223 | 22988128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988128&pc=OZD | 08/08/2025 | 10/22/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695420 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after 38 hours of support the patient exhibited a cerebral bleed. additional information noted care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
6142 | 1220648-2025-46224 | 22988203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988203&pc=OZD | 08/08/2025 | 09/06/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695396 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Distributor |  |  |  | 52 YR |  | Female | 86 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support had ventricular fibrillation. the patient was shocked, and an amino bolus was given. additionally, the patient developed a gastrointestinal (gi) bleed and anticoagulant was held. the patient continued on support, but was unable to advance therapies, due to the bleeding. additional information noted care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure modes has been completed. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. ventricular fibrillation : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
6143 | 1220648-2025-46225 | 22988327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988327&pc=OZD | 08/03/2024 | 09/15/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508833 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 40 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Hematoma | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient placed on impella 5.5 after undergoing emergent aortic resection and mediastinal re-exploration for persistent aortic root bleeding. the patient had worsening atrial flutter and went into pulseless electrical activity (pea) arrest due to suspected bleeding, requiring hemoperitoneum evacuation. patient nonexpanding omental hematoma of the transverse colon. the patient continued to have intermittent atrial fibrillation rapid ventricular response. the patient was transitioned to do not resuscitate (dnr) and comfort care, with impella, crrt (continuous renal replacement therapy), and pacemaker being turned off. the patient ultimately expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D2 date of death has been updated. | The investigation for the access site bleeding, arrhythmia, and hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed, hematoma, and arrhythmia and could not be determined.
6144 | 1220648-2025-31265 | 22988461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988461&pc=OZD | 02/24/2023 | 09/06/2025 |  | 09/06/2025 |  |  | 01/17/2024 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller was indicating intermittent false position alarms during patient support. there was no patient harm, and support continued with the implanted pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs were returned and the console was tested. unable to reproduce the reported symptom, the fringe pattern looks good, which confirms no issue with the optical bench. no other issue was found with the console hardware. the root cause of the degraded optical signals was most likely the contaminated front panel optical connector and the noisy signal from the optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6145 | 1220648-2025-46226 | 22988531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988531&pc=OZD | 08/08/2025 |  |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656328 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 60 YR |  | Male | 64 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had fever, signs of sepsis and suspected gastrointestinal bleeding (gi bleed). after showing some signs of sepsis, including fever, the patient was administered antibiotics. the antibiotics resolved the fever. additionally, the patient experienced a drop in hemoglobin and platelets with no major signs of blood loss from all indwelling sites. as a result, a gastrointestinal bleed was suspected and the patient was administered one unit of blood. the patient survived the procedure. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. sepsis: the root cause of the sepsis was unable to be determined due to insufficient clinical details. major bleed: the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6146 | 1220648-2025-46227 | 22988768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988768&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/07/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700141 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 54 YR |  | Male | 100 KG | Non Hispanic | White | Death | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Tachycardia | Hematuria | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp and rp flex device for bipella mechanical circulatory support (mcs). patient with chest pain for 2 days and presented to the emergency room with late presentation of inferno lateral myocardial infarction with hypotension. lactate was normal. the patient was sent to catheterization lab for left and right heart catheterization and underwent percutaneous coronary intervention (pci). the patient was still hypotensive after successful thrombectomy with pnumbra and stenting of the mid-proximal right coronary artery. the impella cp device was implanted, and the patient was sent to the unit on mcs. the following day after start of the impella cp support, the patient experienced hematuria and multiple hemolyzed blood samples collected from the patient. physician pulled back impella cp with echocardiogram at bedside. based on a hemodynamic marker for the right ventricular function and tricuspid annular plane systolic excursion measurement, the physician elected to proceed with impella rp flex support. the patient now on bipella support. prior to leaving the catheterization lab sites were dry. upon arrival at the unit, some bleeding was noted. staff felt blankets weighing down the impella catheter and propped back up. the patient was bleeding ¿profusely¿ from groin site, which was managed by placing additional purse string suture around site. two units of blood were given. the team was unsure if the impella cp (right femoral artery) or impella rp flex (right femoral vein) site was bleeding. bipella support continued. (additional information subsequently received noted the bleeding was from the impella rp flex access site). two days later, purge system blocked occurred with the impella rp flex. the physician spoke with medical affairs for tissue plasminogen activator protocol. alarm cleared on its own, however, after 1.5-2hours. impella flows were minimal 1.2l/min at p-6 performance level. the patient remained stable despite minimal right-sided support and the impella rp flex was explanted. after removal of the impella rp flex, it was noted the blue impella plug, and purge side arm were under the patient¿s foot/heel against footboard of the bed. the side arm appeared kinked to some extent. the patient was sedated; however, the patient's legs move some which the team noted accounted for the purge pressure intermittent issues. the swan-ganz catheter and sheath from right internal jugular were also found to be pulled back several centimeters. team was advised about risk of clot ingestion with device and needed to decrease impella flow to p-2 performance level during catheter manipulation to decrease chance of ingestion of biomaterials into device. impella cp support continued. however, the following day the patient would expire. care was withdrawn. | The user facility reported that the patient on impella cp support was having hematuria and hemolysis which was resolved by repositioning the impella cp. additionally, impella was repositioned in the middle of the night due to ventricular tachycardia. it was further reported that the family did not wish to escalate care any further and the patient expired. |  |  |  |  | Section b5 was corrected to include complete details surrounding the event, which includes additional information that was received regarding the bleeding event and added b6: tests/lab data including dates. additionally, device problem coding update was not represented in the initial medwatch report. therefore repopulation, update to h6: health effect: clinical and impact coding and h6: medical device problem coding was completed. medical safety review: the patient on impella cp and had issues with hemolysis. the impella cp cannot be dissociated with the outcome. the patient was on bipella support. there was access site bleeding requiring transfusion; which access site was bleeding was initially undetermined and then found to be the impella rp flex. following, the patient was able to manage with minimal rp flex support given the purge system blocked issue; therefore, the rp flex was explanted. this malfunction and bleeding event is a serious injury and should be reported. the impella cp was present at the time of death and is the more likely contributor. however, the patient had significant arrhythmias and in this case is what led to a significant decline in the patient's function and the family decided to withdraw support which appears to have led to along with the patient's underlying condition the demise outcome. the medical safety officer reviewed the event and noted the same in that impella cp device cannot be dissociated from the demise outcome. and the impella rp flex is a serious injury type of reportable event. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Corrections: b1, e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: data review: data logs showed pump ran without issue during support. positioning issue: the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. vt: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks.
6147 | 1220648-2025-46228 | 22988840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988840&pc=OZD | 08/08/2025 | 10/02/2025 |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701028 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Gastrointestinal Hemorrhage | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient escalated from an impella cp to an impella 5.5 pump for support due to cardiomyopathy. while on support, gastrointestinal bleeding was noted so heparin was turned off. additionally, a low platelet count of twelve was noted. the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombocytopenia and access site bleeding has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia and access site bleeding could not be determined.
6148 | 1220648-2025-46230 | 22988959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22988959&pc=OZD | 08/07/2025 |  |  | 09/06/2025 |  |  | 08/08/2025 | 09/06/2025 |  |  |  |  | 09/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691868 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 42 YR |  | Male | 97 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Tachycardia | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella 5.5. the patient had been on two previous impella cps where complications occurred, requiring escalation of therapy. the patient had a hypotensive episode during ventricular access, requiring a dose of epinephrine. after pump insertion to the right axillary pocket, the pocket was surgically closed. the patient was bleeding from the mouth, had a small hematoma, blood in stool, and had swelling with a pressure dressing in place. the patient required blood transfusion and anticoagulation was held. the patient had multiple episodes of ventricular tachycardia (vt) with suction alarms, resolved with repositioning, amiodarone, levo, vaso, and bivalirudin initiation. the patient had additional respiratory issues requiring tracheostomy, peg tube, bronchoscopy with lavage, and antibiotics for possible atelectasis vs pneumonia, mucus plugs, bloody sputum, and clots. the patient ultimately expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product returned. hypotension/ hematoma/major bleed: clinical details noted patient was hypotensive during pump insertion. small hematoma was observed at impella access site. patient also had bleeding from mouth and required blood products. the cause of the injury was not determined due to insufficient clinical details. tachycardia/infection: clinical details noted that patient had episodes of vt during support. patient also being treated for pneumonia. . the root cause of the injury was unable to be determined due to insufficient clinical details. suction: clinical details noted that patient had multiple suction events overnight throughout support when patient was being turned and would resolved on its own or when returned to supine position. pump position multiple times confirmed to be free floating in lv and free of all structures after suction events. data logs show intermittent suction alarms throughout support. the cause of the suction was user related as the suction alarms were occurring when patient was being turned and would resolve when returned to supine position per clinical details. device history review: the device passed all the post sterile inspection checks.
6149 | 1220648-2025-31224 | 22974803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974803&pc=OZD | 05/16/2025 | 10/24/2025 |  | 09/05/2025 |  |  | 05/16/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645596 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. the pump was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data log review shows a high number of "placement signal low" alarms that are constant throughout this day, but they also appear prior to and after this day. it is unclear if and when the patient was coughing outside of the complaint date. there were also a few impella in ventricle alarms on the day of the complaint. the optical 5s parameters showed no abnormalities during the case. no product was returned. the root cause of the optical signal issue was not determined since insufficient clinical details were provided and product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date updated.
6150 | 1220648-2025-31225 | 22974828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974828&pc=OZD | 05/22/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 05/22/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661751 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the pump¿s optical sensor failed, resulting in loss of placement signal. despite this, the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. impella 5.5 pump was not returned. data logs were returned and show optical signal issue with several alarms. the failure trend indicate the optical fiber line is likely damaged due to excessive torqueing. the root cause of the optical signal issue was most likely a damaged optical fiber line based on 5s failure pattern observed in data logs but the location of damage is unknown without return product. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6151 | 1220648-2025-31226 | 22974836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974836&pc=OZD | 05/22/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 05/22/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655474 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. the pump triggered an "unreliable placement signal" alarm. despite this, the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data log review shows that the placement signal went out of range and contrast increased in amplitude and oscillated intermittently. no product was returned. the root cause of the optical signal issue was most likely due to a damaged fiber line since the data logs show 5s parameter changes that are characteristic of a fiber break but the location/origin of the damaged fiber line is unknown since product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6152 | 1220648-2025-31227 | 22974848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974848&pc=OZD | 05/24/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 05/24/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570741 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. the pump triggered a ¿placement signal low¿ alarm, although proper positioning was verified. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs show stable 5s parameters throughout support. placement signal low alarms were observed. the root cause of the optical signal issue was not determined as no product was returned for analysis and no data log abnormalities in relation to reported issue were observed. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6153 | 1220648-2025-31228 | 22974867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974867&pc=OZD | 05/25/2025 | 10/24/2025 |  | 09/05/2025 |  |  | 05/25/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630322 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the ¿placement signal not reliable¿ alarm was active. the placement signal was disabled due to the alarm. despite this, the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs are not available and product was not returned. the root cause of the optical signal issue was not determined since product/data logs were not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date updated.
6154 | 1220648-2025-46151 | 22976395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22976395&pc=OZD | 08/06/2025 |  |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464426 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during the support, the nurse called describing a brief low battery alarm on the automated impella controller (aic). they unplugged and the battery dropped to 50% when plugged back in and returned then to 100%. currently, the aic is reading 100%. they have not been able to recreate the alarm by unplugging. the team made decision to replace the aic. no patient harm was reported due to the issue. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
6155 | 1220648-2025-46107 | 22976520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22976520&pc=OZD | 08/05/2025 |  |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132457 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 120 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while on impella support there was a placement signal and the left ventricle signal dashed out with a yellow alarm ¿controller error¿ on the automated impella controller. motor current and flows were unchanged. there was no patient harm. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6156 | 1220648-2025-46154 | 22976885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22976885&pc=OZD | 08/06/2025 | 10/28/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655529 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 112 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a paitent was escalated from a impella cp to an impella 5.5 as an escalation of care. however, it was reported that during the impella 5.5 insertion, there was a miscommunication between the surgeon and perfusion, resulting in the perfusionist starting the impella 5.5 while the guide wire was still inside of it. the console immediately had a pump failure warning. once the issue was discovered, the wire was removed. the surgeon was advised of the option of changing the pump to a new one. the surgeon decided to remove the wire and proceed with the current impella that was in position. the impella started appropriately. there was no patient harm. | [5 empty] | Investigation summary: temporary pump stop: clinical details reported the impella 5.5 was started while the guide wire was still inside of it. guidewire was removed and the pump was started. the logs aligned with clinical details and showed the impella stopped motor current high alarms at case start which resolved. the cause of the issue was due to an unintended use related error. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. entering cardiac output (impella cp with smart assist). section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6157 | 1220648-2025-46155 | 22976981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22976981&pc=OZD | 08/06/2025 |  |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660143 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was implanted with an impella cp and bleeding around the impella access site was noted. multiple troubleshooting interventions were completed, including manual pressure, forward tension, angle matching, afterload reduction, heparin and integrillin boluses, brilanta, blood products, and vascular surgery was to be consulted. the patient was transported to the cardiovascular intensive care unit with manual pressure being applied to the access site. the femoral artery was found to be perforated, and required placement of a covered stent by the vascular team and impella cp removal. the damage occurred at the time of impella access in the catheterization lab, post sheath exchange. the patient required three units of blood products. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: vascular injury: the root cause of vascular injury cannot be determined since the product was not returned, and insufficient clinical details were provided. access site bleeding: the root cause of access site bleeding could not be determined since the product was not returned, and insufficient clinical details were provided. device history review: device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
6158 | 1220648-2025-31229 | 22977133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977133&pc=OZD | 05/27/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 05/27/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626499 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction complicated by cardiogenic shock. during shockwave, the pump lost the optical signal. support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs were returned for evaluation. the root cause of the optical signal issue could not be definitively determined as no product or data logs were returned for evaluation and limited clinical details were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6159 | 1220648-2025-31230 | 22977220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977220&pc=OZD | 06/04/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 06/04/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2026654940 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the nurse reported a placement signal low alarm, which was subsequently disabled. additionally, it was noted that flow could not be increased above p4 without suction events. the device was repositioned, after which the performance level could be increased without suction, and waveforms improved. the device was successfully weaned and explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs were reviewed and the reported abnormalities were confirmed. based on clinical details provided and data log review, the root cause of the optical signal issue was determined to most likely be a use issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6160 | 1220648-2025-46156 | 22977229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977229&pc=OZD | 08/06/2025 |  |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694756 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 64 KG |  |  | Malfunction | No | Yes | Crack | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was explanted due to high purge pressure system blocked alarm. a new impella 5.5 was implanted. additional information was received. following a complex coronary artery bypass graft and valve procedure, the patient was taken to the intensive care unit. therapeutic intravenous heparin was added back into the treatment regimen within 12 hours. within 24 hours, a purge flow low, purge pressure high alarm persisted. the problem was not solved with standard procedure. tissue plasminogen activator (tpa) was added appropriately to the purge solution with no change to purge flow and pressure for eight hours. within 12 hours, a leak was noticed at the purge sidearm. the pump was exchanged that day. no patient harm was reported. | [5 empty] | Impella 5.5 sn (b)(6) + purge cassette returned for evaluation. high purge pressure: returned product was investigated and there were no relevant abnormalities. there were no kinks in the purge cassette line or catheter. there was no biomaterial noted in the outflow/inflow or cannula. the pump was connected to a console and run in a bucket. initially purge pressure was in the 900's but trended down gradually into the 600's after 3 hours and remained steady for the rest of the run. based on the pattern noted in data logs and the clinical report that tpa was administered in conjunction with high purge pressure resolving gradually while running in a bucket of water, the cause of the high purge pressure was determined to be biomaterial buildup. purge sidearm crack: data logs were reviewed, and there were no signs of a leak, only the reported high purge pressure that persisted through the end of the case. returned product was visually inspected and there were no visual abnormalities noted at first. the pump was then connected to a console and purged while running for over 24 hours. at this point, there were signs of fluid/moisture leakage from the sidearm reservoir membrane into the radial seal. additionally, it was noted that there was an abnormality at the edge of the membrane that presented like a small tear or void. the sidearm was then cut off of the pump and pressurized to 950 mmhg for a period of time to see if a more severe leak could be produced, however, there was only a small increase of moisture noted around the membrane perimeter. based on the evidence of moisture around the perimeter of the sidearm reservoir membrane, the cause of the purge leak was determined to be due to insufficient seal of the purge sidearm reservoir. phr: pump sn (b)(6) passed all post sterile inspection checks. | A.5 and a.6 are unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6161 | 1220648-2025-46158 | 22977374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977374&pc=OZD | 08/06/2025 | 09/29/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with was implanted with an impella 5.5 as a bridge to transplant for mechanical circulatory support. the complainant reported that impella automated impella controller (aic) after unplugging from wall power the battery immediately dropped to 50% and then popped back up to 100% once plugged back in. the aic was exchanged. that aic also had similar issue , and was exchanged with 3rd aic without any issue. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data log review confirmed the reported failure mode. the console was returned to field service for further investigation. the reported complaint could not be reproduced, even after the console was power cycled and operated with the pump and purge cassette for a couple of hours. communication between the power battery manager (pbm) and the batteries was monitored, revealing no issues. additionally, a visual inspection of the internal circuitry showed no abnormalities. root cause: the root cause of the "battery level low" alarm was most likely an erroneous trigger of the invalid data sample in the smbus communication between the pbm and the batteries. however, the specific component responsible was not determined. product history review summary: there were 1 other complaint discovered when reviewing the product history of console ic9997, which is related to the reported issue.
6162 | 1220648-2025-46157 | 22977388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977388&pc=OZD | 08/01/2025 | 09/30/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679411 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp pump was placed for a patient admitted in with planned high risk surgery for the ventricular septal defect. while in the coronary care unit, persistent blood leakage was observed from the groove next to the shaft of the repositioning sheath (sterile sleeve side). due to the persistent blood leakage, the outside of the blue nose was ligated with silk suture, resulting in symptom improvement. a small amount of blood leakage continued thereafter. the pump was explanted several days later and the patient survived. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6163 | 1220648-2025-46160 | 22977617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977617&pc=OZD | 08/06/2025 | 10/21/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699483 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 119 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered femoral site bleeding. as a result, the pump was removed. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the reported major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed could not be determined.
6164 | 1220648-2025-46159 | 22977647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977647&pc=OZD | 08/06/2025 |  |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636342 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 97 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in with multiple cardiac and systemic comorbidities and placed on the impella cp pump for support the pump was supporting when the pump came out of the appropriate left ventricle position. this occurred as the patient was being transferred from the catheterization lab table to the bed. the pump migrated to the aorta and could not be re-delivered. the team made the decision to explant and replace the pump. a new impella cp was placed and there was no reported patient harm. | [5 empty] | The investigation was completed. positioning issues: no product was returned. impella migrated into aorta. the cause of the positioning issue was determined to be an unintentional use error as the pump most likely moved out of position when the patient transferred from cl table to bed. the device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6165 | 1220648-2025-46152 | 22977913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22977913&pc=OZD | 08/06/2025 | 10/07/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186351 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was placed on impella cp support with increasing chest pain and need for coronary artery bypass graft. while on support, the automated impella controller (aic) alarmed for a placement signal not reliable alarm. the placement signal was high and not correlating with the cuff pressure. the signal went away but recurred, aligning more with the blood pressure cuff. the pump was not explanted and the aic was not replaced, as the physician was aware and was not concerned. the position of the pump was confirmed via echocardiogram. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this file represents the automated impella controller. refer to section h.10 for the related report number that represents the pump. | D.9 the device was returned and the date was added. h.6 codes were updated as the device was returned. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
6166 | 1220648-2025-46162 | 22978614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22978614&pc=OZD | 08/06/2025 | 10/10/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 | [6 empty] | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA OPTICAL CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371242 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Automated impella controller (aic) started with a system self-check failure. fwcheck showed powermod1 was failing. there was no patient involvement. | [5 empty] | The investigation for the reported ¿system self check failure has been completed. the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to this communication issue, the aic console rebooted automatically (as designed) to attempt another power on after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | D2a revised as it was entered incorrectly on the initial report that was submitted. e1 added initial reporter address line 1 as it was omitted from the initial report that was submitted. postal code was added. it was incorrectly placed in the zip code field on the initial report that was submitted. g1 manufacturer site postal code was added as it was incorrectly reported on the initial report that was submitted in the manufacturer site zip code field. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed as the product was returned for evaluation. data analysis: the console logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to this communication issue, the automated impella controller console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the console was sent back for further investigation. the system self check failure screen was reproduced upon boot up during testing. visual inspection confirmed pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. product history review summary: there were no issues related to the complaint discovered when reviewing the product history for the console. | H6 health effect - impact code f2601- problem identified before clinical use/exposure has been corrected to only f26- no health consequences or impact. | The device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6167 | 1220648-2025-46153 | 22978767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22978767&pc=OZD | 08/07/2025 | 10/07/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656118 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Pain | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was placed on impella cp support with increasing chest pain and need for coronary artery bypass graft. the patient had access site bleeding requiring blood transfusion. the patient became bradycardic and almost required intubation. the patient was having small runs of ventricular tachycardia and chest pain requiring medications. the pump was later explanted, without the need for vascular surgery, as the site was closed with percloses, pressure was held, and hemostasis was achieved. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this file represents the pump. refer to section h.10 for the related report number which represents the automated impella controller. | B.3 updated date of event as it was reported incorrectly on the initial report that was submitted. g.3 the date received by manufacturer for the pump was reported incorrectly on the initial report and should of been reflected as 07-aug-2025. the investigation for the reported access site bleeding and tachycardia/bradycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleeding and tachycardia/bradycardia could not be determined. the device history review was completed and the device passed all post sterile inspection checks.
6168 | 1220648-2025-46119 | 22978772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22978772&pc=OZD | 08/06/2025 | 10/17/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682854 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, patient experienced hemolysis. pump p-level was reduced in response to the event. access site was slightly oozing, pressure was held and dressing changed. | An echocardiogram was ordered and the pump was repositioned. there were some suction alarms. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: hemolysis: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ hemmorhage/bleeding. impella cp ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | B.1 added selection as it was inadvertently omitted from the initial report that was submitted. b.5 information was added that was inadvertently omitted from the initial report that was submitted. d.4 model number and d.7a revised as they were previously entered incorrectly on the initial report that was submitted. h.6 code e0506 and f23 were reported incorrectly on the initial report that was submitted. a new impact code and medical device problem codes have been added. the investigation has been completed. the pump was not returned for investigation. the case started on ic2895 and pump ran for about eight hours before being transferred to ic11184. there was only one instance of a suction alarm reported at the end, just before the pump was transferred to ic11184. on ic11184, the pump was used for another 33 hours, with several suction alarms reported during the first 13 hours while running on p6 and p7. the placement signal was also trending downward during the suction events. no abnormalities were reported regarding positioning issues or motor current spikes throughout the entire case run. mechanical interaction with blood(hemolysis): the cause of the hemolysis was not determined since the product was not returned for investigation, and data logs/clinical details were not sufficient to determine the cause of hemolysis. pump suction: the cause of the pump suction was most likely related to the patient condition, based on data log review, which indicated that the patient was not tolerating the high p-level. a device history review was completed and the pump passed all post sterile inspection checks.
6169 | 1220648-2025-46124 | 22978787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22978787&pc=OZD | 08/06/2025 | 09/30/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639787 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support of the impella cp, 14 fr sheath was removed and swapped with the repositioning sheath. site with minor oozing was noted. site was dressed, and positive distal pulse was noted. patient was out in holding area for recovery and became restless at affected distal extremity and bleeding/oozing with hematoma formation was noted both proximal and distal to access site of impella repositioning sheath. pressure was held for 30 minutes without any improvement. the surgeon suggested weaning the impella while prepping for explant. impella was weaned without hemodynamic compromise and then explanted | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. major bleed: no issues were found with the returned product. the cause of major bleed was most likely use issue related since the patient became restless at affected distal extremity and bleeding/oozing with hematoma formation noted both proximal and distal to access site of impella repositioning sheath. device history review: the device passed all post sterile inspection checks.
6170 | 1220648-2025-46161 | 22978793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22978793&pc=OZD | 08/07/2025 |  |  | 09/05/2025 |  |  | 08/07/2025 | 09/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1133364 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with was implanted with an impella 5.5 as a bridge to transplant for mechanical circulatory support. the complainant reported that an impella automated impella controller (aic) after unplugging from the wall power the battery immediately dropped to 50% and then popped back up to 100% once plugged back in. the aic was exchanged. that aic also had similar issue, and was exchanged with a third aic without any issue. there was no harm to the patient. | [5 empty] | The automated impella controller (aic) was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. this report represents the second aic. refer to section h.10 for the related report number which represents the first aic. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: the logs were consistent with the reported issue. device analysis: the console was returned to field service for further investigation. the reported complaint could not be reproduced, even after the console was power cycled and operated with the pump and purge cassette for a couple of hours. communication between the power battery manager (pbm) and the batteries was monitored, revealing no issues. additionally, a visual inspection of the internal circuitry showed no abnormalities. root cause: the root cause of the battery percentage went from 100% to 50% was most likely an erroneous trigger of the invalid data sample in the smbus communication between the pbm and the batteries. however, the specific component responsible was not determined. device history review summary : there were 2 other complaint discovered when reviewing the product history of console (b)(4), which are not related to the reported issue.
6171 | 1220648-2025-46164 | 22978825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22978825&pc=OZD | 08/06/2025 | 10/22/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655245 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 56 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported a patient on impella 5.5 support was undergoing a coronary artery bypass grafting procedure. during procedure, the patient coded, with rhythm going into ventricular tachycardia (vt)/ ventricular fibrillation (vf), 2 shocks before return of spontaneous circulation (rosc) was achieved. the patient continued to have runs of bleeding time overnight until dobutamine was held. the pump was explanted via surgical wound closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in impella ifu: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported ventricular fibrillation and tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ventricular fibrillation and tachycardia could not be determined.
6172 | 1220648-2025-31231 | 22978827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22978827&pc=PYX | 05/24/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 05/24/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED INC. |  |  | US |  | 2025584363 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, the patient experienced an abrupt increase in placement signals accompanied by a drop in flows. despite this, the device was successfully weaned and explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for investigation. data log review confirmed the abrupt increase in placement signal and drop in flows. the optical parameters showed no abnormalities. the root cause of the placement signal issue was most likely sensor drift since the data logs showed drifting ps and pump flow along with stable mc and optical parameters. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6173 | 1220648-2025-46163 | 22979056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979056&pc=OZD | 08/06/2025 |  |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552511 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of an impella cp, the implant was aborted due to veno-arterial extracorporeal membrane oxygenation placement and contralateral access unusable due to the patient¿s vascular anatomy. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | A.2, a.4, a.5, and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6174 | 1220648-2025-46166 | 22979069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979069&pc=OZD | 08/06/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705245 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. upon placement, placement signal was noted to be significantly higher than patient¿s arterial line pressure. the patient continued on support until the device was successfully weaned. there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. no damages or abnormalities were noted with the pump. due to a lack of imc logs provided to be able to observe sensor calibration, the optical sensor issue cannot be determined. device history review: the device passed all the post sterile inspection. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6175 | 1220648-2025-46167 | 22979099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979099&pc=OZD | 08/07/2025 |  |  | 09/05/2025 |  |  | 08/07/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701180 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female | 121 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs). three vessel coronary artery bypass graft surgery was performed with the impella 5.5 device centrally placed and false placement alarms (impella in aorta) triggered. placement signal monitoring was disabled. a transesophageal echocardiogram showed the impella 5.5 device across the aortic valve at 3.5 cm. the patient was continued on impella mcs. approximately four days later, suction events were encountered during runs of ventricular tachycardia (treatment unnoted). a couple of days later impella in aorta alarm triggered again, and the care team reduced the impella flow and confirmed the positioning. heart function was noted recovering well and no pressers present. next step was noted as weaning the patient from the impella device which was successfully completed and the device explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint #: (b)(4). complaint product experience coding was updated based the investigation team recommendation to better align with the reported details. therefore, section h6 medical device problem coding was repopulated to reflect the update.
6176 | 1220648-2025-46168 | 22979162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979162&pc=OZD | 08/07/2025 |  |  | 09/05/2025 |  |  | 08/07/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643365 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 101 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock underwent percutaneous coronary intervention for a chronic totally occluded right coronary artery. four drug-eluting stents were implanted. throughout the morning the patient continued to decompensate, and a decision was made to take the patient back to the catheterization lab for venoarterial extracorporeal membrane oxygenation (va ecmo) with an impella cp for mechanical circulatory support (mcs), the following day the impella was toward the mitral valve overnight with lactate dehydrogenase of 3290. the impella was repositioned and explanted later this day with a survived patient outcome. the patient continued on va ecmo therapy. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows the impella cp with smartassist system: section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section use of echocardiography for positioning of the impella catheter - ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. section impella catheter too far into the left ventricle - ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿. section suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | Product complaint # (b)(4). investigation summary- product and data logs not returned. suspected hemolysis on the 2nd day of support. pump was confirmed in improper position contacting the mitral apparatus. the patient had small lv and primary rv failure. the pump was repositioned and did not resolve the hemolysis until the pump was explanted. data & product were not returned for review. mechanical interaction with blood/hemolysis: the cause of the hemolysis was unable to do determined due to no product or data logs returned. device history lot- device lot #: 1898427. device history batch- subcomponent lot#: n/a. device history review= pump sn: (b)(6) passed all post sterile inspection checks.
6177 | 1220648-2025-31232 | 22979303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979303&pc=OZD | 07/04/2023 | 09/05/2025 |  | 09/05/2025 |  |  | 07/04/2023 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024327961 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during set up the console gave a "optical sensor system error" alarm. the console was exchanged for a different one and the problem was resolved. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data log review confirmed the alarm. the console was tested on an engineering lab bench. the alarm was reproduced by creating the air gap between the white plug from the impella catheter and the blue catheter plug in the console. the root cause of the ¿placement signal not reliable¿ - alarm from the aic was most likely due to the air gap. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6178 | 1220648-2025-46165 | 22979312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979312&pc=OZD | 08/05/2025 | 01/25/2026 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623559 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an impella cp was placed to support a patient admitted in with cardiogenic shock complicating acute myocardial infarction. the pump was supporting when there was a visualized clot on the day after implant. the day of implant there was no clot observed so the thrombosis was new. the team did note the activated clotting time to be in the 180's. the team made the decision to transfer to the tertiary center and they explanted the pump successfully and escalated to an impella 5.5. no embolization was noted. the patient survived | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | Additional information was received regarding an update to the patient's outcome. after successful escalation to the impella 5.5, the patient was eventually, successfully transitioned to a left ventricular assist device. h6 investigation: type, findings and conclusion codes and h6 component code previously submitted remain unchanged. update to patient outcome received and there is no impact to the investigation completed.
6179 | 1220648-2025-31233 | 22979333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979333&pc=OZD | 02/01/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 02/01/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED INC. |  |  | US |  | 1142003 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the console experienced signal not reliable alarm. the audio was disabled. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs from console confirm that the placement signal not reliable alarm was issued. the console was investigated in the lab. the front optical connector was found to be clean. a known good pump and purge line was run with the console for more than an hour with no observable abnormality. the root cause for the oscillating contrast could not be determined due to not being to reproduce the issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6180 | 1220648-2025-31234 | 22979519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979519&pc=OZD | 05/24/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 05/24/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548998 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the automated impella controller displayed a 'placement signal not reliable' alarm. the intervention steps taken are unknown. no harm was reported to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console and logs were returned and reviewed. console logs from the reported day of event are consistent with complaint and show that after 2 weeks of hemodynamic support, the console presented with placement signal not reliable alarm, which kept self-resolving and repeating. the console was tested, but the issue was not reproduced, and optical bench illumination was stable. the root cause of the console issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6181 | 1220648-2025-46169 | 22979577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979577&pc=OZD | 08/06/2025 | 10/14/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646685 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced hemolysis. the position was checked on imaging and the pump was noted to be in good position. the impella cp was successfully weaned and explanted. | [5 empty] | A.4 and a.5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported mechanical interaction with blood has been completed. there was no defect found with return product. data logs reviewed and few position alarms seen which resolved immediately on (b)(6) 2025. placement signal low was seen for some time during support. no other mc spike or other abnormalities seen. the cause of mechanical interaction with blood was not determined as no defect found with return product, inconclusive log evidence and due to insufficient clinical information. d9 device return date added. h6 type of investigation code b01 has been added.
6182 | 1220648-2025-46118 | 22979612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979612&pc=OZD | 08/06/2025 | 10/28/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026701730 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Pulmonary Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. patient was taken for computed tomography (ct) for chest, abdomen and head, which was positive for large bilateral pulmonary emboli. previous echo showed left ventricular ejection fraction (lvef) of 50% with severely reduced right ventricle (rv) function. decision was made to remove impella and treat with lytics. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported pulmonary embolism has been completed. the product was not returned for investigation. the root cause of the injury was not determined due to insufficient clinical details.
6183 | 1220648-2025-31235 | 22979708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22979708&pc=OZD | 07/25/2025 | 10/27/2025 |  | 09/05/2025 |  |  | 07/25/2025 | 09/05/2025 |  |  | 12/31/2025 |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025447204 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the automated impella controller displayed a 'placement signal not reliable' alarm. the intervention steps taken are unknown. no harm was reported to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. data logs show the placement signal was not reliable, along with an oscillating trend in contrast and zero signal not reliable values. this trend returned to normal after 8 days of pump use. upon review of data logs, no problem was found with the pump and it is apparent the console was the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D6b: should be blank, the explant date is unknown. | G3 ¿date received by manufacturer¿ corrected.
6184 | 1220648-2025-46171 | 22980475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22980475&pc=OZD | 08/08/2025 |  |  | 09/05/2025 |  |  | 08/08/2025 | 09/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701757 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a placement signal not reliable alarm. the alarm was encountered when pump flows were at p4. the nurse was walked through how to disable the audio and a plan was made to conduct an echocardiogram. no further details regarding the issue were mentioned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: added code g0301204 to the component code section. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary" optical sensor issue: the pump was not returned for investigation. data logs show that the placement signal drifted to -50 for about one hour before it lost to 3000 on 8-august (18th day of pump use), without significantly impacting the 5s optical signal parameters. there was an active psnr alarm, and the placement signal was not recovered during the rest of the case run. the cause of the optical sensor issue was determined to be sensor wear, based on data log review. device history review: the pump passed all post sterile inspection checks.
6185 | 1220648-2025-46172 | 22980777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22980777&pc=OZD | 08/08/2025 |  |  | 09/05/2025 |  |  | 08/08/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546550 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the decision was made to place the impella cp via the right femoral artery pre-percutaneous coronary intervention. during impella support, the patient was in atrial fibrillation and had a few runs of ventricular tachycardia. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6186 | 1220648-2025-46173 | 22981373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22981373&pc=OZD | 08/08/2025 |  |  | 09/05/2025 |  |  | 08/08/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694449 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. during insertion of the impella cp, the impella was primed and loaded onto the wire. however, they were unable to advance past the sheath. buddy wire was attempted but the vessel was not tortuous and appeared stenotic at inflow. single acces with a 23 french sheath was attempted but they were still unable to advance with the motor still in the sheath. impella/wire was removed. multiple attempts were made with no success and they were unable to advance past the sheath. the impella and wire were removed. intravascular ultrasound was inserted which showed 70% stenosis extending the entire abdominal aorta to common. the impella was aborted. no patient harm was reported due to the issue. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. delivery issue: the root cause of the deliver issue was determined to be patient condition as the patient had stenosis and difficult anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
6187 | 1220648-2025-46150 | 22982318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22982318&pc=OZD | 08/06/2025 |  |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702600 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. while on support, the right groin site was having bleeding and the site was redressed. in addition to adding quick clot and the dressing became saturated within 10-15 minutes. the site was redressed two times. the nurse took the patient for a computed tomography (ct) scan, and with the movement of the patient, bleeding became worse. the site was redressed two more times since return from ct scan. it was noted that the bleeding seemed to be slowing down. additionally, the partial thromboplastin time increased overnight to 132 and treated with two units of fresh frozen plasma and protamine. blood cultures came back positive for gram positive cocci in pairs and chains resulting in enterococcus faecalis. a consult for antibiotic treatment was made. currently on vancomycin & zosyn. there was increasing requirements of levophed, and methylene blue was administered. ultimately, care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the bleeding from the right groin is unclear from the provided clinical details. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. infection: the root cause of the infection was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A.4 and a.5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6188 | 1220648-2025-46174 | 22982516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22982516&pc=OZD | 08/07/2025 | 10/21/2025 |  | 09/05/2025 |  |  | 08/07/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707308 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 103 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a decision was made to upgrade a patient on intra-aortic balloon pump to the impella 5.5. however, it was reported that the impella 5.5 was unsuccessful in multiple attempts to pass from the subclavian artery into the aorta due to 90 degree angled curve. the decision was made to place an impella cp. the impella 5.5 was removed and the impella cp was advanced all the way into the left ventricle. | [5 empty] | The investigation for the reported delivery issue has been completed. the root cause of the deliver issue was determined to be patient condition as the patient had small aortic arch and difficult anatomy preventing successful delivery of impella. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6189 | 1220648-2025-31236 | 22983559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22983559&pc=OZD | 06/07/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 06/07/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2025466986 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, there were intermittent optical signal alarms. despite this, the pump remained operational throughout the support period until weaning and explantation. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs were returned and it indicates that the placement signal was intermittently abnormal. this console was tested after arriving in fs. unable to reproduce the reported failure mode symptom by running the test optical pump for 8 hours. no problem was found with the console hardware. the root cause for the intermittent placement signal (ps) spike was not determined as there was no explicit evidence that caused the event. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6190 | 1220648-2025-31237 | 22983661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22983661&pc=OZD | 05/14/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 06/28/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095486 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the automated impella controller experienced placement signal issues. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console was returned for investigation and confirmed alarm. the logs show a placement signal not reliable alarm. the root cause of the placement signal not reliable alarm and oscillating contrast was the defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6191 | 1220648-2025-31238 | 22983857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22983857&pc=PYX | 06/18/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 06/18/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095635 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the automated impella controller displayed a 'placement signal not reliable' alarm and disappeared. no harm was reported to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console and logs were returned and reviewed. console logs from the reported day of event are consistent with complaint and show that 5 minutes after pump was started, optical placement signal became unreliable. after 15 minutes it self-resolved, but after another minute became invalid for the duration of the case. the console was tested and issue was reproduced. light was intermittently absent from the optical pump connector. the root cause of the issue was a defective component. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6192 | 1220648-2025-46176 | 22986871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22986871&pc=OZD | 08/07/2025 | 10/21/2025 |  | 09/05/2025 |  |  | 08/07/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708790 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced ventricular tachycardia during support. the patient was treated by shock/defibrillation. additionally, the patient experienced limb ischemia treated by a distal perfusion catheter line placement. additional information noted the patient expired on support. | [5 empty] | Patient race is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ventricular tachycardia and limb ischemia could not be determined.
6193 | 1220648-2025-46175 | 22986879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22986879&pc=OZD | 08/06/2025 | 09/05/2025 |  | 09/05/2025 |  |  | 08/06/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690221 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Hospitalization | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5 device via the right axillary artery for additional mechanical circulatory support. two days after start of the additional mcs, the patient underwent redo coronary artery bypass graft surgery (left internal mammary artery had an issue distally; it was detached and reattached). cabg surgery was uneventful and successful per the notes. however, the patient's left lower extremity had pulse but was cyanotic; fasciotomy remained open with woundvac. however, below the knee amputation was needed and eventually completed. impella site was clean dry and intact. on day 7 of impella mcs, the patient experienced a hypotensive event over the night correlating with rhythm change requiring impella performance level to be increased. the patient remained intubated and sedated, vasopressin was added, and the patient remained on continuous renal replacement therapy. the patient continued on impella mcs. four days after the hypotensive event, the impella was having intermittent suction alarms and the impella device was repositioned which aided in decreasing suction alarms. the patient received one unit of packed red blood cells. the next day an attempt was made to wean the impella. through the night, the patient had increased shortness of breath followed by ventricular tachycardia episodes where the patient lost pulsatility. patient had to be intubated, decision then to increase impella support level to p-7. hypotension caused the team to add pressors and send for a cta overnight which revealed a left internal jugular deep vein thrombosis and r lobe pulmonary embolism. interventional radiologist stated the clot burden was not high enough for thrombectomy. echocardiogram at bedside was completed to reevaluate the right heart function. blood cultures were sent, and the patient remained on crrt. then next day an attempt to wean the patient was impella was made again; however, there was increased ectopy/ventricular tachycardia runs. amiodarone and lidocaine restarted. the patient continued impella support. ventricular tachycardia episodes noted minimal since restarting amiodarone, lidocaine. days following the patient had bleeding from the below the knee amputation and received two units of packed red blood cells overnight. (it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient). no ventricular tachycardia was noted overnight and the next day there was no arrhythmic event noted. lidocaine drip discontinued and amiodarone switched to by mouth. impella support continued, and slow weaning restarted. the impella support level was down to p-3 and the following day thereafter the impella 5.5 device was explanted with a survived patient outcome. | [5 empty] | Patient-specific information section a4: weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. investigation of the events(s) and product malfunction(s) noted the following: ventricular tachycardia: clinical details noted that patient was experiencing runs of vtach while on impella support. in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. bleeding: clinical details noted that patient required multiple units of blood volume throughout support. no additional details provided for reason of additional volume being provided. the cause of the injury was not determined due to insufficient clinical details. low flow: clinical details noted that patient was having intermittent suction alarms. pump was repositioned the following day and aided in decreasing suction alarms. additionally, patient was on crrt with volume removal. data logs were reviewed and lt log showed low flows and suction alarms at time of reported issue. no motor current spike or trend in placement signal was observed. the cause of the low flows was likely due to a use related issue as clinical details noted that repositioning the pump decreased the suction alarms. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6194 | 1220648-2025-31239 | 22986910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22986910&pc=OZD | 07/01/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 07/01/2024 | 09/05/2025 |  |  | 11/30/2025 |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428311 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, there was a "placement signal not reliable" alarm. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no pump returned. data logs were reviewed and the report was confirmed. the root cause of the optical signal issue was most likely due to sensor wear of the optical sensor as evidenced by the log pattern and duration of support. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6195 | 1220648-2025-31240 | 22986942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22986942&pc=OZD | 07/01/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 07/01/2024 | 09/05/2025 |  |  | 04/30/2025 |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332832 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced fiber optic sensor artifact, with readings above scale. the issue would temporarily resolve but then recur. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned product was inspected and there were no physical abnormalities. the pump passed light continuity test and the 5s parameters were within specification. upon review of the data logs and returned product, it is apparent that the console is the potential cause of the issue. no problem was found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6196 | 1220648-2025-31241 | 22986964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22986964&pc=PYX | 08/19/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 08/19/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529616 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. shortly after implantation, a ¿placement signal not reliable¿ alarm occurred. the alarm audio was disabled, and support was continued. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned pump was inspected and there were no physical abnormalities. the 5s parameters were within specification. the data logs confirm placement signal not reliable alarms against the optical sensor. the root cause of the optical signal issue was most likely patient condition. the logs also showed placement signal trending in downwards before going out of range. placement signal then trends upwards with a concurrent downward trend in pump flow. the pump was tested in the lab and sensor drift reproduced. the root cause of the placement signal issue was determined to be a sensor drift of the sensor as evidenced by identified log pattern and returned product analysis. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6197 | 1220648-2025-31242 | 22986994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22986994&pc=PYX | 09/01/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 09/01/2024 | 09/05/2025 |  |  | 02/28/2026 |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478754 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, there was a placement signal alarm from the pump. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no data logs were available. the product was returned and evaluated. upon plugging into a lab console, a "placement signal not reliable" alarm was triggered. there were no visual abnormalities with the dp sensor, but dp sensor drift was observed in the lab flow loop. the optical parameters were also evaluated with no issues found. the root cause of the placement signal issue was most likely sensor drift since it was reproduced in the lab. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6198 | 1220648-2025-31243 | 22987001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987001&pc=OZD | 04/30/2025 | 10/24/2025 |  | 09/05/2025 |  |  | 04/30/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084304 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. following implantation, placement signal not reliable alarm was active on the console. pump placement was verified, and the alarm was silenced. support therapy continued without interruption. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of the event are consistent with the complaint. the logs show multiple instances of the suction alarms along with placement signal low and multiple position alarms throughout the case. the console ic1167 was returned to field service for further investigation, as the reported complaint could not be reproduced. during the inspection, visible light was observed coming out from the external optical pump connector. the root cause for the placement signal not reliable alarm was likely due to the intermittent failure of the optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date updated.
6199 | 1220648-2025-31244 | 22987090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987090&pc=OZD | 06/28/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 06/28/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2025479673 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, a placement signal unreliable alarm occurred. the pump remained operational throughout support and the pump was weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console logs from the reported day of event are consistent with complaint and show that after 2 weeks of support, a placement signal not reliable alarm was triggered, and remained unresolved until the pump was stopped and disconnected several hours later. the returned console was tested and issue was reproduced. the root cause of the issue was a defective component. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6200 | 1220648-2025-31245 | 22987149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987149&pc=OZD | 10/18/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 10/18/2024 | 09/05/2025 |  |  | 12/31/2025 |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456428 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the optical signal loss was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was inspected and there were no physical abnormalities. the pump passed the laser continuity test and ran without reproducing the placement signal not reliable alarm. the data logs were reviewed and there was low signal not reliable during support. the root cause of the optical signal issue was not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6201 | 1220648-2025-31246 | 22987152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987152&pc=OZD | 11/18/2024 | 09/05/2025 |  | 09/05/2025 |  |  | 11/18/2024 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569058 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported using a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. during the procedure, a ¿placement signal unreliable¿ alarm occurred intermittently, accompanied by a pulsatile motor current, while flow readings remained correct. no patient harm occurred. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned. the pump passed laser continuity testing. there were no kinks or biomaterial observed. the 5s parameters were in spec. the alarm did not reproduce when plugged into console. no abnormalities were found with the pump. log review showed upon pump start-up. this issue was also found with multiple pumps on the same console prior to this complaint occurrence date. the root cause of the optical signal issue was unable to be reproduced on the returned pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6202 | 1220648-2025-46177 | 22987224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22987224&pc=OZD | 08/07/2025 |  |  | 09/05/2025 |  |  | 08/07/2025 | 09/05/2025 |  |  |  |  | 09/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628484 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device and impella rp flex for mechanical circulatory support. it was reported on the first day of impella rp flex support that the purge pressure was found to be higher than expected, and there were no alarms that had indicated the high purge pressure. the following day, there were alarms for high purge pressure noted. the medical team was advised to follow hospital protocols, and were advised to change purge tubing, check for kinks in tubing, and verify the patient¿s activated clotting time was in therapeutic range. the medical team followed hospital protocol to add tissue plasminogen activator to the purge solution. the following morning, the high purge pressure alarms resolved, and the purge solution was switched back to bicarbonate solution. additionally, it was reported on (b)(6) 2025 that the patient lacked pulses in their left leg, where both impella cp and rp flex were inserted, and the patient¿s right arm. it was noted that the patient had previous deep vein thrombosis and previous surgical complications which most likely contributed to the limb ischemia. the vascular team was consulted, but no further information was provided on any interventions that were performed for limb ischemia. on (b)(6) 2025 it was reported that the impella rp flex placement signal results were inaccurate as it was reporting flows higher than expected at p-4 with a negative pulmonary artery pressure. there was no change in motor current and no change in the patient condition. the patient remained on support until heart function recovered. the impella cp was successfully weaned and explanted, and the following day the impella rp flex was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. there are two associated devices for the reported event. the first pump (impella rp flex) is addressed under manufacturer report number 1220648-2025-46170. the second pump (impella cp) is addressed under manufacturer report number. | Medical safety officer review of the limb ischemia event for association given bipella support (impella cp via left femoral artery and impella rp flex via left femoral vein) was completed and noted the report of limb ischemia goes against the impella cp as it could be pressing against/occluding flow to the artery which is associated with limb ischemia. related report numbers: this impella cp report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex where changes were made to remove the limb ischemia event accordingly.
6203 | 22962113 | 22962113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22962113&pc=OZD | 06/10/2025 | 08/19/2025 | 08/19/2025 | 09/04/2025 | 08/19/2025 | 09/04/2025 |  | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | OZD | [5 empty] | ABIOMED INC |  |  | US |  | 1959009 |  |  |  | No |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella purge cassette was changed, per policy. new purge cassette alarmed for low purge pressure. purge cassette was exchanged for a new one without issue.
6204 | 1220648-2025-45793 | 22962468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22962468&pc=OZD | 07/13/2025 | 08/29/2025 |  | 09/04/2025 |  |  | 07/13/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692773 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | Mitral Valve Insufficiency/ Regurgitation | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple complications while implanted with an impella 5.5. after two days of support, the patient endured ventricular tachycardia. the pump support level was lowered in response to the condition. two days later, the patient was diagnosed with severe mitral valve regurgitation. the staff did not express any concerns and support continued uninterrupted. upon eventual planned explant, a thrombus was found on the catheter. shortly after, the patient lost radial pulses and experienced numbness/tingling in their extremity. the patient was taken to the operating room and a thrombectomy was performed treat the condition. it was noted that anticoagulants were only given intermittently during support and leading up to the event. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The root cause of the reported ventricular tachycardia was unable to be determined due to insufficient clinical details. the root cause of the mitral valve incompetence cannot be determined as insufficient clinical details were provided.
6205 | 1220648-2025-46121 | 22962590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22962590&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708837 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female | 55 KG |  |  | Injury | Yes | Yes | Failure to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp for mechanical circulatory support. it was reported that the impella sheath was unable to be advanced during implantation. both the short and long sheaths were attempted and were unsuccessful. shockwave intravascular lithotripsy utilizing shockwave to the left femoral artery was also unsuccessful. the impella sheath would not advance. the impella cp implant was aborted. however, upon explant, the perclose device failed to achieve hemostasis at the impella access site. the patient was then brought to the operating room for left femoral artery surgical closure. | [5 empty] | A.5 and a.6 are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Investigation summary: delivery issue: the root cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy along with calcified and tortuous vessels preventing successful delivery of impella. access site bleeding: the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. device history review: device sn: (b)(6) passed all post sterile inspection checks.
6206 | 1220648-2025-31213 | 22962833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22962833&pc=OZD | 08/08/2023 | 09/04/2025 |  | 09/04/2025 |  |  | 08/08/2023 | 09/04/2025 |  |  | 05/31/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352219 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient with acute myocardial infarction and cardiogenic shock. during use, the pump position was confirmed; however, the device exhibited an ¿impella in aorta¿ alarm. despite the alarm, the impella continued to provide effective hemodynamic support, and its function was not otherwise affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. per the given clinical information provided, the root cause of the optical signal issue was not determined since product was not returned and data logs were not returned as well. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6207 | 1220648-2025-46122 | 22963536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22963536&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589776 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was being implanted with an impella 5.5 for mechanical circulatory support after decompensating following surgery for four coronary artery bypass grafts. it was noted pre-operatively that the patient had calcium build up in their arteries. it was reported that the impella 5.5 was unable to pass through the calcium during delivery. multiple attempts were made with different wires and a buddy wire, but all were unsuccessful in delivering the impella 5.5 to the left ventricle. the decision was made to abort the impella 5.5 due to patient anatomy and instead placed an impella cp. the impella cp was successfully placed, and it was reported that during support with the impella cp, the patient experienced runs of torsades that was addressed with medication. it was also reported that the pump had placement signal low alarms, which resolved with repositioning the impella cp. on the final day of support, the patient coded for forty minutes, however, it was noted there were no issues with the impella cp during that time. the patient¿s family then decided to withdraw care. care was withdrawn and the patient expired. this complaint involves two impella pumps. this report represents the impella 5.5. another report will represent the impella cp with serial number: (b)(6). | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The medical safety officer reviewed the event and the impella cp device cannot be dissociated from the demise outcome. the event describes a delivery issue with the impella 5.5 due to the patient anatomy in which the case was aborted; switch was made to the impella cp which was successfully placed. the delivery issue is a product malfunction and should be reported. additionally, regarding the impella cp, the event describes runs of torsades and the patient was cardioverted out, addressed with medication. the next day the patient coded again with return of spontaneous circulation. however, the patient would eventually expire as the family withdrew care. there is not enough information to dissociate the impella cp present at the time of the patient's demise. however, the patient's underlying condition likely contributed most to the demise outcome, and the family decided to withdraw support which appears to have led to the patient's demise. investigation: the impella device was not returned as it was discarded, and therefore an analysis of the device was not possible. device history record review was completed and revealed the device passed all post sterile inspection checks. the delivery issue was investigated, and the cause was determined to be patient condition due to calcium in the vessels. the pump was unable to be delivered into the left ventricle using multiple wires and buddy wire. related medical device reports: this impella 5.5 report is one of two devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome.
6208 | 1220648-2025-46126 | 22963840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22963840&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694660 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 78 KG | Hispanic | White | Injury | Yes | No | False Alarm | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp support it was reported that the patient experienced no flow to the right leg, above where the impella cp was inserted. it was also reported the same day, that there was an alarm for impella in aorta. impella position was confirmed with imaging. as a result of the lack of flow down the leg, the impella cp was weaned and removed. | [5 empty] | The impella device was not received from the customer as the customer reported it has been discarded. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). b5 is being updated to include related device information not included in the initial report. the revised b5 provides a complete event narrative. h6. health effect: clinical code, hemorrhage/bleeding was added for the surgical closure for the impella access site. h6. medical device problem code false alarm was previously reported and has been removed; malposition code has been added. upon internal review, it was determined that the alarm was a true alarm that was caused by malposition of the device. the device was subsequently repositioned.
6209 | 1220648-2025-46127 | 22963944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22963944&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656091 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient with heart failure and in cardiogenic shock was escalated to an impella 5.5 device from an impella cp for increased mechanical circulatory support. the patient had been previously supported on the impella cp for 8 days. during impella 5.5 support, it was reported that there were alarms for high purge pressure and purge system blocked. the medical team made the decision to explant the impella 5.5 and replace it with an intra-aortic balloon pump (iabp) due to the purge pressure issues. the patient was successfully placed onto the iabp, however later the same day the patient¿s care was withdrawn and the patient expired. | [5 empty] | Given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. product status: the impella device was not received from the customer as the customer reported it has been discarded. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D.3 manufacturer fax was inadvertently omitted from the initial manufacturer report number 220648-2025-46127. d.9 updated as the pump is available for investigation. e.1 initial reporter address line 2 was inadvertently omitted from the initial manufacturer report number 220648-2025-46127. g.1 manufacturer contact fax number was inadvertently omitted from the initial manufacturer report number 220648-2025-46127. h.6 investigation codes were updated as the product had been returned for investigation. h.10 updated the additional manufacturer narrative as the product is available for investigation. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the impella 5.5 was returned for evaluation. the clinical details state that the purge pressure rose above 1100mmhg and the purge flow rate fell below one milliliter per hour, even after replacing the purge system. high purge pressure and purge system blocked alarms were triggered. to avoid failure the pump was removed as it was not properly unloading. the event logs showed that during the whole case there was variability in the motor current, pump flows and placement signal. the purge pressure was stable for about five hours until it began to gradually increase over the course of seven hours, then there is a sharp increase over the course of two hours leading to high purge pressure and purge system blocked alarms. there are no motor current spikes before this. there are no purge cassette or bag changes before. the purge pressure continued to be high for about 28 hours, and the it slowly trended down over the remainder of the case. there is also no anticoagulation used throughout the case. the returned product had some biomaterial in the purge gap, and when it was connected to the console purge pressure was slightly elevated but it was able to return to normal 500mmhg levels after purging the pump and soaking the gap in warm water. there was no blood ingress in the line. there were no other abnormalities observed. the root cause of the high purge pressure and purge system blocked is due to biological material buildup based on data log and returned product analysis. device history review: the impella passed all post sterile inspection checks. | H.11 the first follow up rmanufacturer device report number 1220648-2025-46127 stated incorrectly that section h.10 was updated when it should of stated h.11 was updated.
6210 | 1220648-2025-46128 | 22963993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22963993&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634564 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported upon initiation of support, the impella alarmed for suction and there were no flows. the medical team decided to explant the pump and replace it with a new impella cp. the patient was successfully implanted with the new impella cp and remains on support with information available at time of this report. | [5 empty] | The impella device was not received from the customer as the customer reported it has been discarded. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: low flow occurred when pump started. troubleshooting was unsuccessful. pump was removed & replaced by a new one that worked without issue. low pump flow: data review: data logs confirmed low pump flow when pump started that triggered auto suction response & suction alarms. no motor current or placement signal spikes observed during support. this event remained throughout the rest of the case. product was not returned for review. the cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. the pump passed all post sterile inspection checks.
6211 | 1220648-2025-46129 | 22964051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964051&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700770 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Backflow | Difficult to Advance | [5 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that while inserting the impella, there was difficulty in obtaining ventricular access under fluoroscopy and transesophageal echocardiogram and a 6fr pig/advantage glide wire, 5fr straight taper glidecath/straight-stiff glidewire, and 5fr al-1/straight-stiff glidewire were attempted unsuccessfully. ventricular access was successfully obtained using angled tapered glidecath and straight tip stiff glidewire which was then exchanged for 6fr pigtail for improved positioning in the left ventricle. support was initiated, and the patient was sent back to the intensive care unit following the implant of impella 5.5. it was noted that the patient was hypertensive and there were occasional retrograde flow alarms. on the second day of support, the patient experienced a small ischemic stroke in the parietal lobe. the patient remained stable, and it was noted that the stroke resolved and the patient recovered. the patient was successfully bridged to a heart transplant, the impella was removed during transplant surgery. | [5 empty] | The impella device was not received from the customer as the customer reported it has been discarded. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint #: (b)(4). investigation summary: there was no product returned. delivery: it was stated that there was difficulty getting lv access and it was eventually obtained by using an angled tapered guide catheter and straight tip wire. due to lack of clinical details provided, the root cause of the delivery issue cannot fully be determined. retrograde flow: on the 5th, it was stated that there were retrograde flow alarms. the data logs show the retrograde flow alarms were triggered at the beginning of the case but shortly resolved. the motor current was not pulsatile, with a slight variation in flows and ps at p-4. the motor current and flows were within specification. this alarm was said to be triggered as the patient became hypertensive. due to lack of product returned, the root cause of the retrograde flow cannot be determined. stroke: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1955452. device history batch: subcomponent lot: n/a. device history review: impella 595556a passed all post sterile inspection checks.
6212 | 1220648-2025-46130 | 22964100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964100&pc=OZD | 08/06/2025 |  |  | 09/04/2025 |  |  | 08/06/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687532 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Fever | Ventricular Fibrillation | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported on the third day of support that the patient experienced a stroke. the stroke was confirmed on computed tomography and the patient was noted to have right side deficits present. additionally, it was noted that the previous day it was determined the patient was at risk for stroke from impella placement due to calcifications in the aorta found prior to surgery. as a result of the stroke, anticoagulation therapy was withheld. on the 7th day of support, it was reported the impella was too deep and alarmed for suction. the patient experienced ventricular tachycardia that self-resolved. the p level was lowered from p8 to p6. the physician came to evaluate the impella position, however no further information was provided on device position or if adjustments were made. the following day it was noted that the patient was febrile, no information was provided on any interventions that were performed due to the fever. the patient was successfully weaned and explanted after heart function recovered and the explant was not performed earlier than planned as a result of the complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6213 | 1220648-2025-46062 | 22964366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964366&pc=OZD | 08/02/2025 | 10/24/2025 |  | 09/04/2025 |  |  | 08/02/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694563 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during impella 5.5 support, intermittent placement signal loss and position in ventricle were noted to happen when wedging or turning the patient followed by resolution without any intervention. the patient successfully received a heart transplant and the impella 5.5 was explanted. the patient was stable. | [5 empty] | The investigation for the reported access site bleeding and optical signal issue has been completed. no product was returned. clinical reported placement signal low and pump in ventricle which triggered during patient movement and resolved after. clinical noted patient also on extracorporeal membrane oxygenation. the logs confirm occurrence of both alarms and their resolutions and show no identifiable optical signal failure patterns. the root cause of the optical signal issue was likely use-related as issue occurred during patient movement and resolved after. the root cause of the access site bleeding was not determined as limited clinical information was provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6214 | 1220648-2025-46131 | 22964544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964544&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674186 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | Distributor |  |  | 86 YR |  | Male | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp for mechanical circulatory support. during insertion, it was reported that the physician was unable to deliver the impella cp due to vascular complications of the patient. the impella cp was removed, and a different longer introducer was used to successfully deliver a replacement impella cp. the patient remained on the replacement impella cp until recovered and was successfully weaned and explanted. | [5 empty] | Investigation summary: the investigation has been completed. unable to deliver or advance : the cause of the delivery issue was determined to be patient condition as the patient had tortuosity of vessel preventing successful delivery of the impella. device history review: the pump passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6215 | 1220648-2025-46132 | 22964649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964649&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706513 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the following day, an erroneous low aortic placement signal occurred triggering suction and placement signal low alarms. the impella's position was confirmed, and motor current and flows were appropriate. support continued. however, the device was subsequently explanted as the patient had an escalation in therapy to an impella 5.5 with a survived patient outcome. there was no harm to the patient as a result of the event. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary: the pump was not returned for investigation. data logs show that several placement signal was low alarms and suction alarms occurred throughout the case until noon on 08-05, when the placement signal suddenly jumped around 90/50 from lowest 35/0, with resulting in resolution of the suction and ps low alarms. the placement signal gradually trended downward until a sudden resolution occurred. a spike in motor current was reported in the middle of the case run, without impacting the placement signal. no trend was noted in the 5s optical signal parameters throughout the entire case run. optical sensor issue: the cause of the optical sensor issue was most likely related to the patient's condition, based on data log review. pump suction: the cause of the pump suction issue was most likely related to the patient's condition, based on data log review. device history lot: device batch: 1943509. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks.
6216 | 1220648-2025-46112 | 22964798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964798&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688186 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a purge side arm leak issue. it was reported that the purge side arm leaked during setup. as a result, the pump would not complete prime and the implant was aborted. | [5 empty] | Product complaint # (b)(4). the impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | A.4 added weight of the patient as it was inadvertently omitted from the initial report that was submitted. b.7 added information that was inadvertently omitted from the initial report that was submitted. d.3 no number should of been added behind the manufacturer name that was submitted on the initial report that was submitted. d.7a revised as an incorrect selection was made on the initial report that was submitted. e.1 information was added that was inadvertently omitted from the initial report that was submitted. e.4 added selection as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturing site name and manufacturer site phone were reported on the initial report that was submitted and should of been left blank. g.2 consumer was selected incorrectly on the initial report that was submitted. company representation was added. purge leak-pump: the pump and purge cassette were returned for investigation. no visual damage was noted on the returned product. the pump was then primed, and a purge leak was observed from the sidearm filter hole. the filter was further examined, no damage was seen at air hole and weld line. data log review was not required for this case. as per the clinical details, purge side arm leak was observed during setup. the cause of the purge leak was sidearm filter damage where purge fluid leaked out for the filter vent hole. device history review: the pump passed all post sterile inspection checks.
6217 | 1220648-2025-46106 | 22964805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964805&pc=OZD | 08/05/2025 | 10/20/2025 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026673832 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a impella cp placed for the support of the patient with coronary artery disease. the cp supported for just 18 hours when the pump was explanted and escalated to the 5.5 pump. the cp access site had developed a hematoma and pseudoaneurysm that prompted blood products to be infused. the patient survived the harm and was supported by the 5.5 until care was withdrawn after just under 10 days. | [5 empty] | Product complaint # (b)(4). the impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma could not be determined.
6218 | 1220648-2025-46133 | 22964834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964834&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698615 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced access site bleeding requiring intervention. the impella sheath was inserted into an already existing 7 french sheath that showed signs of hematoma and contrast extravasation on an earlier angiogram. the physician hoped the 13 french repositioning sheath would occlude the vessel and stop the bleeding. once the sheath was in place, it only made the bleeding worse. the physician did not think the site would improve with manual pressure. the perclose was placed before the peel-away was placed and was sinched down with no improvement to the bleeding. contralateral access was gained to move the impella cp to the left femoral artery. however, the bifurcation was high, and the stick was questionable. a hematoma formed around the fresh 6 french sheath. surgical options were discussed and, however, quickly dismissed. reevaluation of the patient showed an improved left ventricular end-diastolic pressure of 11 mmhg, lactic acid of 1.2, and improved chest pain symptoms. the physician elected to remove the impella pump and close the access site with two perclose and one angioseal. the patient remained on levophed was sent to the unit for recovery and management. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history review confirmed the pump passed all the post sterile inspection checks. regarding the major bleed, the cause of the injury is most likely use related since the impella sheath was inserted into an already existing 7 french sheath that showed signs of a hematoma and contrast extravasation on an angiogram. once the sheath was in place it only made the bleeding worse. h6 investigation: investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation.
6219 | 1220648-2025-46123 | 22964841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22964841&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026682025 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 81 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Pericardial Effusion | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella cp on (b)(6) 2025, there were large amounts of oozing from left femoral artery impella site w repo sheath in. pressure heldx15min. redressed site w forward tension sutures and different angle match. no hematoma noted. patient hemoglobin dropped from 7.9- 6.3. during intubation of the patient in icu patient was also believed to have bleeding ulcer and gi bleed. the patient received 2p red blood cells. afterwards, the patient became hypotensive and laminar. the bedside echocardiogram showed a large pericardial effusion. 1l removed and family arrived and wished to withdraw care, and the patient expired. | [5 empty] | Correction: d3, d4, e4, g1. the investigation of the reported failure mode has been completed. no product was received for evaluation. pericardial effusion/hypotension: the cause of the injury was not determined due to insufficient clinical details. major bleed: the root cause of the injury was unable to be determined to be use issue because the complication was managed by angle matching and forward tension sutures. device history review: the device passed all post sterile inspection checks. | Product complaint # (b)(4). the impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: access site bleeding. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6220 | 1220648-2025-46135 | 22965505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22965505&pc=OZD | 06/30/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665230 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | Distributor |  |  | 89 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. it was reported that during impella cp support, a leadless pacemaker communication failure occurred with a competitors (atrial) and (ventricular) pacemakers. it was reported that atrial-pace and ventricular-pace pacing were possible, but dual pacing was not possible. since atrial and ventricular only pacing was still available, no actions were taken and the patient remained on impella support until the patient was recovered and was able to be weaned and explanted. after the impella was explanted, the pacemakers were able to dual pace again. there was no harm reported to the patient for the issue. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Investigation summary: device interaction: clinical details report that the patient had atrial and ventricular pacemakers implanted prior to impella support. the impella was inserted on 25/jun/2025, and it was reported that during a ventricular septal perforation/defect procedure on 30/jun/2025, the pacemakers lost their ability to communicate/dual pace. the complication resolved when the impella was explanted the following day. data logs were reviewed and there were no abnormalities. there were many shifts in motor current (mc) and several mc spikes throughout support, all of which aligned with changes in performance (p)-level or the pump being restarted. there were 26 triggers of impella position in aorta alarms. some of these were around the time of the surgical procedure, and some were around the time of weaning. it cannot be determined the relationship between the reported issue and the alarms. product was not returned for investigation. since clinical details provided were limited, data log review was inconclusive, and product wasn't returned for investigation, the cause of the device interaction could not be determined. device history lot: device lot: 1914157. device history review: pump sn (b)(6) passed all post-sterile inspection checks.
6221 | 1220648-2025-31214 | 22965604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22965604&pc=OZD | 03/26/2024 | 09/04/2025 |  | 09/04/2025 |  |  | 03/26/2024 | 09/04/2025 |  |  | 11/30/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432437 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported using a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. upon connection of the impella to the console, an ¿optical sensor issue¿ notification was displayed. disconnecting and reconnecting the same console was unsuccessful, as the same alarm reoccurred. the physician proceeded with the procedure without impella sensor function. no harm was reported to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for investigation. data logs confirm the controller error optical bench alarm immediately upon connecting the pump. checking the pumps used on the console before and after confirm the issue to be pump related. the root cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6222 | 1220648-2025-46136 | 22965773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22965773&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  | 08/31/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393082 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during placement of impella 5.5, the doctor had a difficult time advancing the wire and catheter due to anatomy. a 035 rosen j wire and catheter were eventually able to be advanced to the left ventricle. however, the doctor elected to attempt initial placement over this wire after having such difficulty earlier. the impella 5.5 was placed through the graft but would not advance beyond the subclavian artery so the impella 5.5 was removed. angiogram was attempted to be ran to evaluate anatomy and did not provide useful information. the impella 5.5 was removed with plans to attempt an impella cp. during wire exchange, the catheter could not be readvanced. therefore, efforts were aborted. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. delivery issue: the root cause of the deliver issue was determined to be patient condition as the patient had a difficult anatomy along with stenosis preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
6223 | 1220648-2025-46008 | 22965914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22965914&pc=OZD | 07/14/2025 |  |  | 09/04/2025 |  |  | 07/14/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601481 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Stroke/CVA | Fever | Hemolysis | Low Blood Pressure/ Hypotension | Hematuria | Pericardial Effusion | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hematuria, hemolysis, and suction alarms. it was reported that a patient arrived post aortic valve replacement and aortic repair. an impella 5.5 was placed due to low ejection fraction after transesophageal echocardiogram. dark red urine was noted. the impella positioned was confirmed with echocardiogram and central venous pressure was assessed. some suction alarms were noted and one unit of blood was given. the patient was noted to be on heated high flow, mild fever, and cultures were sent. hemolysis was noted due to over diuresis. diuretics were off. the patient also went hypotensive while working with physical therapy, albumin was administered and vasopressin drip was started. additionally, a computed tomography (ct) scan head showed intraparenchymal hemorrhage. the plan was to change the purge to bicarbonate. the patient was moving all extremities and following commands. the plan was to volume resuscitate with colloids and blood products and monitor bleeding. it was also noted that to repeat ct to assess brain bleed. echocardiogram was pending to assess hemopericardium. new finding of pericardial effusion from ct scan on (b)(6) 2025. non-occlusive was noted. possible systemic heparin start. the patient was successfully weaned from the impella 5.5 and explanted. | [5 empty] | The impella device was not received from the customer, but the event logs were received. therefore, the investigation of the device is not possible, but event logs evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | No product was returned for investigation. data logs were reviewed. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of suction issue was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the stroke was unable to be determined due to insufficient clinical details. the cause of the hypotension cannot be determined as insufficient clinical details were provided. the device passed all post sterile inspection checks.
6224 | 1220648-2025-31215 | 22965928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22965928&pc=OZD | 04/09/2024 | 09/04/2025 |  | 09/04/2025 |  |  | 04/09/2024 | 09/04/2025 |  |  | 02/28/2026 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480775 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the optical placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and 5s parameters show signal not reliable. 5s parameters are consistent with a broken sensor face. clinical details noted that failure of optical sensor with "placement signal not reliable" occurred after first wave of shockwave was administered. additionally, support from the impella was not affected with motor current remaining pulsatile and pump was checked with imaging to be in proper position. the root cause of the optical signal issue was due to a broken sensor face due to use of shockwave device. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6225 | 1220648-2025-31216 | 22966483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22966483&pc=OZD | 04/15/2024 | 09/04/2025 |  | 09/04/2025 |  |  | 04/15/2024 | 09/04/2025 |  |  | 11/30/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445238 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported using a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. during support, a ¿placement signal unreliable¿ alarm repeatedly occurred. pump position was verified by fluoroscopy and found to be normal. the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was evaluated and the 5s parameters were within normal range. the alarm did not reproduce in the lab. the pump passed the light continuity test and there were no visual abnormalities. the data logs confirm placement signal unreliable alarms and decreasing signal not reliable. the root cause of the optical signal issue was not determined as issue could not be reproduced. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6226 | 1220648-2025-46109 | 22966719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22966719&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hematuria | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a patient admitted in for st-elevation myocardial infarction. although the impella was placed without difficulty, it was noted that wire access was gained with some thrombosis seen. furthermore, suction alarms occurred, which were treated with fluid and a reduction in the device¿s p-level. once the device was at p4, the impella shifted back into the aorta and repositioning did not occur for approximately an hour. the patient¿s urine was reported as dark red with labs pending. no subsequent lab results were noted and the impella was explanted successfully. | [5 empty] | B1: added malfunction per information in the complaint record. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. data review: data logs confirmed pump was in improper position corresponded with the time positioning issue & hemolysis was reported per clinical description that triggered alarms impella position in aorta & suction. the positioning issue & suction alarms were then resolved. no other issues were found on the logs. about 6 hours later, pump was removed. hemolysis: product was not returned for review. based on clinical description & data analysis, the cause of hemolysis was most likely due to pump was not in the proper position. suction issue: the cause of suction issue was most likely due to pump was not in the proper position. positioning issue: the root cause of positioning issue was unable to be determined as limited clinical details were provided, and no product was returned for review. thrombus: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: pump passed all post sterile inspection checks. | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
6227 | 1220648-2025-31217 | 22967164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967164&pc=OZD | 04/26/2024 | 09/04/2025 |  | 09/04/2025 |  |  | 04/26/2024 | 09/04/2025 |  |  | 12/31/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451958 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the optical placement signal was lost following the use of shockwave therapy. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the impella pump was returned for investigation. visually inspected and found no physical damages or abnormalities. optical bench 5s parameters were verified. light continuity test passed without any issues. no other abnormalities were found. as per clinical, once shockwave use was complete and being removed, placement signal waveforms and numbers went away on the aic with placement signal not reliable alarm appearance. data logs were not returned for investigation. the cause of the optical signal issue was most likely a damaged sensor due to shockwave interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6228 | 1220648-2025-46137 | 22967255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967255&pc=DQX | 08/04/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018"PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  | 8853995 |  | 0052-3005 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. when introducing the impella cp, the 0.018 guidewire was shaped and advanced into the pigtail, but it did not come out from the tip of the pigtail and encountered strong resistance. upon checking the tip of the guidewire, a bend that had not been made during shaping was found, which was thought to be the cause. therefore, a new wire was used, which passed through without issue and the impella cp was successfully placed. there was no harm to the patient. | [5 empty] | D.4 the serial number is unknown and therefore the serial number portion was left out of the primary udi number. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The product was returned for investigation, and the guidewire was bend near the tip of the wire. however, the pump was not returned to determine if the pump related to the cause of resistance, and it's unclear in clinical details if the kink was present prior to advancing it through the pigtail. the cause of the guidewire kink could not be determined due to insufficient clinical details. 0.018 guidewire passed all post sterile inspection criteria.
6229 | 1220648-2025-31218 | 22967292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967292&pc=OZD | 08/29/2024 | 09/04/2025 |  | 09/04/2025 |  |  | 08/29/2024 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532452 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the use of shockwave therapy in the patient¿s left main coronary artery interfered with the optical placement sensor. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data log shows the placement signal not reliable alarm along and all 5s parameter was down which indicate hard failure of the sensor. as per the clinical report, the placement signal was lost immediately after pump insertion while the shockwave was being used. the cause of the optical signal issue was most likely damaged sensor due to the shockwave device interaction, based on given clinical details and data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6230 | 1220648-2025-46138 | 22967491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967491&pc=OZD | 08/05/2025 | 02/24/2026 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692699 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainantreported that an impella 5.5 was placed in a cardiomyopathy patient. the pump was supporting while the purge pressure was rising. the team troubleshot by adding tissue plaminogen activator to the purge solution. this was done on second day of the support, but the pump remained on for a total of nine days. the pump was then explanted as the patient was successfully bridged to a left ventricular assist device . | [5 empty] | A.2 - a.5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d.9, h.6 and h.11 were updated to reflect the new information received that the product was returned. | Investigation summary: the investigation has been completed. the device was returned for investigation. the clinical details stated that there was high purge pressure that was experienced on the sixth and tissue plasminogen activator (tpa) was started which reduced the purge pressure. another high purge pressure event occurred on the 12th so tpa was added again. there were no data logs returned. the pump was returned cut and borescope images were taken and there was no biomaterial seen around the impeller or purge gap. without sufficient product or logs returned, the root cause of the high purge pressure cannot be established. device history review: the device passed all the post sterile inspection checks.
6231 | 1220648-2025-31219 | 22967493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967493&pc=OZD | 09/03/2024 | 09/04/2025 |  | 09/04/2025 |  |  | 09/03/2024 | 09/04/2025 |  |  | 09/30/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400875 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the optical sensor malfunctioned; however, the pump continued to operate until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned however the cannula was severed off from the rest of the catheter making it insufficient for a thorough product evaluation. the 5s parameters could not be obtained and the light leak test could not be performed. there was biomaterial beneath the impeller. no data logs from the case were available for analysis. the root cause of the optical signal issue was not determined as insufficient product was returned for analysis, no logs were available and insufficient clinical information was provided a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6232 | 1220648-2025-31220 | 22967532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967532&pc=OZD | 09/19/2024 | 09/04/2025 |  | 09/04/2025 |  |  | 09/19/2024 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532423 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, optical signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned product was inspected. the 5s parameters were obtained which showed signal not reliable and contrast below normal parameter range. the pump passed the light continuity test. the data logs confirm placement signal not reliable alarm and show placement signal going out of range abruptly. there were no motor current spikes prior and the 5s parameters align with the pattern of a sensor break with a decrease in signal not reliable and contrast and an increase in lamp and gain. this aligns with clinical mention of shockwave use during support. the root cause of the optical signal issue was determined to be due to a damaged sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6233 | 1220648-2025-46139 | 22967563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967563&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710556 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient's groin site developed some oozing during impella cp support. the site was managed and one unit of blood was given to treat the bleed. support continued uninterrupted following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6234 | 1220648-2025-46140 | 22967852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22967852&pc=OZD | 08/05/2025 | 10/03/2025 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678482 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 62 KG | Non Hispanic |  | Malfunction | No | Yes | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered a pump stop during impella 5.5 support. the pump was implanted and it was started at p2 without issue. when increasing to p4, saturated flows were noted with spike in motor current. as a result, the device was turned off and the patient was placed on veno-arterial extracorporeal membrane oxygenation. the patient survived. | [5 empty] | Investigation summary: the impella 5.5 and purge cassette were returned for investigation. field logs were not available. saturated flow / pump stop: clinical details noted that after pump implant, the p-level was increased to p-4 and they got saturated flows with a motor current spike. the motor current was still pulsatile but then the pump stopped working. upon removal, there was a large piece of tissue noted in the outflow of the cannula that looked like layers of arterial wall per the surgeon. the pump was returned and the outflow cage and cannula were free of any biomaterial. no visual abnormalities were observed. the pump was successfully run with mean flows of 5.7 liters per minute at p-9 and 4.5 liters per minute at p-6, which are high for their respective p-levels. there were also fluctuations in motor current and pump flow while running the pump in the lab at p-9 indicating potential biomaterial shifting. the motor was then torn down and there was evidence of light biomaterial with some sections stained with purge fluid. the cause of the saturated flow was biomaterial based on the provided clinical details about the biomaterial/motor current spike and since the pump reproduced the high flows with evidence of biomaterial on the returned product. the cause of the pump stop was biomaterial based on the provided clinical details about the biomaterial/motor current spike and evidence of biomaterial on the returned product. the device history review was completed and the pump passed all post-sterile inspection checks. | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
6235 | 1220648-2025-46141 | 22968378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22968378&pc=PYX | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025587985 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on bipella support was taken back for a chest washout on (b)(6) 2025 in the morning due to the coronary artery bypass graft procedure. post washout, the impella rp flex alarmed purge system blocked. no kinks in the purge line were observed but purge cassette was changed out as troubleshooting technique. this briefly corrected the alarm for about five minutes. however, the purge pressure continued to climb, and the alarmed purge system blocked. the physician did not want to proceed with tissue plasminogen activator due to the patients post operative bleeding complication. impella flows on the impella rp flex were not affected at this time. the physician was made aware that the pump could stop if the alarm does not resolve. throughout the evening, the purge system slowly started correcting itself, ultimately coming back to 14 milliliters per hours on the purge flow rate. however, the care team started to notice that patients' urine output was dropping off and urine color was red (fruit punch) colored. placement signal unreliable alarm also occurred with loss of placement signal in the impella rp flex. the next day, the team discussed the patient conditions, and it was felt that the impella rp flex most likely was experiencing some form of clot burden and was leading to hemolysis. after discussion, a decision was made by the heart team to remove the impella rp flex and place the patient on peripheral veno-arterial extracorporeal membrane oxygenation and continued with impella 5.5 support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6236 | 1220648-2025-46142 | 22968449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22968449&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635956 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female | 94 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was reported to be waiting for a heart transplant. the patient was on impella 5.5 support for around three months when impella stopped. the impella was restarted multiple times but each time the pump would stop again. the impella was exchanged. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A.4. added weight of the patient as it was inadvertently omitted from the initial report that was submitted. d.9. the device was returned and the date was added. e.1. added zip code extension as it was inadvertently omitted from the initial report that was submitted. the pump was returned for investigation. pump stop: clinical details report that the pump stopped and it was unable to be restarted. data logs were reviewed and the pump stop alarms were confirmed. on (b)(6) 2025, data logs showed a sudden trigger of pump stop with a motor current spike to 1200 ma and speed deviation. the pump was able to be restarted briefly, but there were many more pump stop alarms triggered. returned product was investigated and the pump failed electrical testing, with an open line in motor phase 1. there was also a kink in the catheter just distal to the red pump plug/kink protector. finally, it was noted that there was biomaterial in the outflow cage and cannula. the pump was attempted to be restarted in a bucket of water, but alarm for pump stops reproduced. the red handle was then opened and there was no necking of the motor cables inside. a cut was made in the catheter at the location of the kink, and continuity was checked to either side in all three motor cable phases. to both sides, all lines were continuous, indicating that the source of the open line noted in initial electrical testing was isolated to the location of the kink. based on the fact that the first alarm triggered in the field and returned product had a kink in the catheter with an open motor cable line, it was determined that the biomaterial was unrelated to the reported pump stop. the cause of the pump stop was determined to be an open electrical line. device history review: the pump passed all post-sterile inspection checks.
6237 | 1220648-2025-46143 | 22968475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22968475&pc=OZD | 08/05/2025 | 10/14/2025 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678150 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis. echo showed the impella in a shallow position, but the team did not reposition the device. due to the patient¿s religion, blood products were not administered. it was further reported that the family withdrew care and the patient expired. | [5 empty] | The investigation for the reported hemolysis and positioning issues has been completed. no product was returned. pt has had several echos that showed questionable pump position. labs confirmed the hemolysis. pump confirmed to be shallow. after reviewing logs, no suction alarms were observed. position in ventricle and aorta at the case start resolved. motor spike on 8/3 with all parameters immediately resolved. placement signal is variable. the cause of hemolysis was most likely use issue since per clinical hemolysis was confirmed by labs and the pump was found to be shallow. the cause of the positioning issue was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6238 | 1220648-2025-46144 | 22968509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22968509&pc=OZD | 11/27/2024 | 10/22/2025 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025545619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification from abiomed's long-term outcome and quality indicator (loqi) impella registry was received regarding a patient that endured refractory biventricular shock with a "remote" relationship to the impella device used. the patient experienced hypotension and bradycardia that was not responsive to atropine, aminophylline, and phenylephrine pushes. a transvenous pacemaker placement (tvp) was placed and the patient was intubated. the procedure was complicated by refractory shock requiring impella cp support. the patient did not recover and expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hypotension bradycardia and ventricular fibrillation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hypotension bradycardia and ventricular fibrillation could not be determined.
6239 | 1220648-2025-46146 | 22968645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22968645&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723019 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 67 KG |  |  | Malfunction | No | Yes | Break | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the insertion of the initial impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful as it was unable to cross the iliac artery due to calcium and the use of a short sheath (13cm). the physician made the decision to shockwave the common femoral and illiac arteries and utilize a 25cm peel-away sheath. however, upon visual inspection of the pump damage was observed. a new impella cp device was decided, which was successfully implanted, and the percutaneous coronary intervention was completed without any issues. there was no harm to the patient as a result of the event. | [5 empty] | Patient-specific information section a5: ethnicity and a6: race are unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6240 | 1220648-2025-46147 | 22968691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22968691&pc=OZD | 08/05/2025 |  |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1157452 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Consumer | User facility |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during prepping for impella implant, the aic had a battery failure alarm. the alarm resolved by resetting the aic. another aic was used for the case. there was no harm to the patient. | [5 empty] | Correction: d3, d4, e4. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6241 | 1220648-2025-46148 | 22968711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22968711&pc=OZD | 08/05/2025 | 10/29/2025 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697780 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 91 KG | Non Hispanic |  | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant report that while a patient was on impella cp support labs showed signs of hemolysis. the doctor adjusted the impella to a shallow position and dropped the p-level. the following day the pump was explanted and the patient survived the surgery. | [5 empty] | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hemolysis and positioning issue has been completed. product was not returned for review. data logs showed pump ran without issue during support. there were positioning alarms triggered in early of the case but resolved quickly. no mc or ps spikes observed on the logs. the root cause of hemolysis and positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review.
6242 | 1220648-2025-46149 | 22970779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22970779&pc=OZD | 08/05/2025 | 09/04/2025 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723019 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 67 KG |  |  | Malfunction | No | Yes | Break | Failure to Advance | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of the initial impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful as it was unable to cross the iliac artery due to calcium and the use of a short sheath (13cm). the physician made the decision to shockwave the common femoral and illiac arteries and utilize a 25cm peel-away sheath. however, upon visual inspection of the pump damage was observed. a new impella cp device was decided, which was successfully implanted, and the percutaneous coronary intervention was completed without any issues. there was no harm to the patient as a result of the event. | [5 empty] | Patient-specific information section a5: ethnicity and a6: race are unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary- the impella cp pump was not returned. data logs are not applicable for reported failures. delivery issues(unable to deliver or advance): as per clinical, it was reported the pump did not cross iliac artery due to calcified vessels at femoral. the cause of delivery issues was patient related since pump was unable to pass vasculature due to calcified vessel. product damage(catheter kink): as per clinical, it was reported the there was catheter kink damaged due to resistance at femoral artery. patient had calcified vessel condition. the cause of product damage(catheter kink) was due to patient condition related. device history lot- device lot: 1931536. device history batch- subcomponent lot: n/a. device history review- the pump sn (b)(6) passed all the incoming post sterile inspection checks.
6243 | 1220648-2025-46145 | 22973910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22973910&pc=OZD | 08/05/2025 | 09/04/2025 |  | 09/04/2025 |  |  | 08/05/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026674495 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male | 86 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported that during support of the impella cp that there was concern for gi bleed overnight and hemorrhagic shock. gi was consulted but turned down for surgery and recommended medical management. the patient continued gi bleeding noted overnight with the need to transfuse packed red blood cells (prbcs).care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Product complaint # (b)(4). major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided.
6244 | 1220648-2025-31221 | 22974664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974664&pc=OZD | 03/30/2025 | 10/24/2025 |  | 09/04/2025 |  |  | 03/30/2025 | 09/04/2025 |  |  | 01/13/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618596 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was unreliable. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. the data log review shows intermittent decreases in signal not reliable and contrast around the time of the complaint. the root cause of the optical signal issue could not be determined since the pump was not returned and insufficient data logs were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date updated.
6245 | 1220648-2025-31222 | 22974739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974739&pc=OZD | 05/14/2025 | 10/24/2025 |  | 09/04/2025 |  |  | 05/14/2025 | 09/04/2025 |  |  | 11/30/2025 |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431171 |  | 1000361 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues; however, flow rates, motor current, and purge values remained within expected ranges. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. lack of data logs or product returned, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date updated.
6246 | 1220648-2025-31223 | 22974799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22974799&pc=OZD | 05/03/2025 | 09/04/2025 |  | 09/04/2025 |  |  | 05/03/2025 | 09/04/2025 |  |  |  |  | 09/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653265 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. the pump was successfully weaned and explanted. no patient harm was reported, and the console was not replaced. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and showed increased placement signal with decreased signal not reliable and contrast at time of "placement signal not reliable" alarms. upon review of the data logs, it is apparent that the console is the potential cause of the issue. the root cause of the optical signal issue could not be definitively determined as no product was returned for evaluation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6247 | 1220648-2025-31198 | 22951081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22951081&pc=OZD | 06/13/2025 | 10/20/2025 |  | 09/03/2025 |  |  | 08/06/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026664383 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with an impella 5.5 was reported to have lost aortic and left ventricle placement signals. motor current pulsatile and flow were consistent. the catheter plug was pushed to ensure it was fully engaged with aic and the placement signal returned. it was further reported that care was withdrawn and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. there was no product returned. optical signal issue. the clinical states that "around the same time as the tubing snapping, device lost aorta and left ventricle placement signals. motor current pulsatile and flow consistent. customer service center called and relayed that purge system and optical sensor are unrelated. after some time, chronic care management pushed catheter plug to ensure it was fully engaged with automated impella controller (aic). placement signal returned immediately. the data logs show that there were placement signal not reliable (psnr) alarms where contrast decreases, gain increases, lamp increases, signal-to-noise ratio (snr) decreases, and light was stable. this shortly resolves. based on the data logs and clinical details provided, the root cause of the optical signal issue is most likely due to a small air gap between the aic and pump plug. purge leak - cassette. the clinical states that "purge tubing snapped near yellow leur during transport from or to unit immediately post-implant. rn quickly replaced cassette and tubing to resolve purge system open alarm." the data logs show that around that time there were purge system open alarms with purge flow increasing. based on the clinical description, the root cause of the purge leak - cassette was most likely due to use as the purge tubing snapped during transport. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6248 | 1220648-2025-31199 | 22951156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22951156&pc=PYX | 06/13/2025 | 09/03/2025 |  | 09/03/2025 |  |  | 08/12/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026619125 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that patient had a rp flex placed for the support of the patient with a aortic valve replacement, mitral valve replacement, and cardiac arrest. the patient was placed on extracorporeal membrane oxygenation as well. the pump lost placement signal (ps) but, remains on despite the issue. patient survived. | [5 empty] | The investigation for the optical signal issue has been completed. the device was not returned for evaluation. however, the data logs were reviewed but did not show any issues with the optical sensor but rather the dp sensor. clinical details state that there was a sudden change in the pa placement signal (ps) as well as impella flows. there were multiple motor current (mc) spikes and suction alarms prior to the large increase in ps and drop in flows. optical parameters were stable throughout the entire case. there were no psnr (placement signal not reliable) alarms. the cause of the placement signal issue was most likely biomaterial since there were mc spikes prior to the drop in flows and increase in ps. the cause of the low pump flow was most likely biomaterial ingestion since there were mc spikes prior to the drop in flows. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
6249 | 1220648-2025-46023 | 22951304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22951304&pc=OZD | 08/07/2025 | 09/03/2025 |  | 09/03/2025 |  |  | 08/07/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692534 |  | 0048-0014 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 48 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient had pericardial effusion. after the impella cp was inserted, extensive bloody pericardial effusion was found and the application of a pigtail catheter for relief of the pericardial effusion. sustained stabilization was not successful. the patient required resuscitation. removal of the impella cp, emergency thoracotomy in the cardiac catheter laboratory and surgical over suturing of a perforation of the left ventricle - most likely induced by the impella guide wire. return of spontaneous circulation and slow stabilization of the patient under initially high-dose catecholamines and after massive blood transfusion and administration of multiple coagulation products. the patient is stable post explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this file represents the pump. refer to section h.10 for the manufacturer device report number which represents the guidewire. | Additional information was received and noted the following: per the physician, it was a normal implant of the impella without any resistance except the guidewire was very deep. also, there was a possible disease for the ventricle to perforate this fast, but the physician could not tell what it was. the patient outcome was good. note: h6 investigation as previously reported remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The investigation of the reported failure mode has been completed. pericardial effusion: the product was not returned for investigation. the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
6250 | 1220648-2025-45991 | 22951419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22951419&pc=DQX | 08/07/2025 | 09/03/2025 |  | 09/03/2025 |  |  | 08/11/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018"PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  | 9187767 |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 48 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient had pericardial effusion. after the impella cp was inserted, extensive bloody pericardial effusion was found and the application of a pigtail catheter for relief of the pericardial effusion. sustained stabilization was not successful. the patient required resuscitation. removal of the impella cp, emergency thoracotomy in the cardiac catheter laboratory and surgical over suturing of a perforation of the left ventricle - most likely induced by the impella guide wire. return of spontaneous circulation and slow stabilization of the patient under initially high-dose catecholamines and after massive blood transfusion and administration of multiple coagulation products. the patient is stable post explant. | [5 empty] | D.4 the serial number is unknown and therefore the serial number portion was left out of the primary udi number. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this file represents the guidewire refer to section h.10 for the related manufacturer device report number which represents the pump. | Additional information was received and noted the following: per the physician, it was a normal implant of the impella without any resistance except the guidewire was very deep. also, there was a possible disease for the ventricle to perforate this fast, but the physician could not tell what it was. the patient outcome was good. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. cardiac perforation/mechanical interaction with tissue: clinical reported that the left ventricle had a perforation after the impella insertion, and suspected that it was from the guidewire. the cause of the perforation was not determined due to insufficient clinical details and no product returned. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
6251 | 1220648-2025-31200 | 22951880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22951880&pc=OZD | 07/16/2025 | 09/03/2025 |  | 09/03/2025 |  |  | 07/02/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654140 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | Material Split, Cut or Torn |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient admitted with cardiogenic shock (cgs) and cardiomyopathy. during support, the pump experienced optical signal issues, although position was confirmed. after approximately two months of support, the pump was found to have migrated out of position due to sleeve damage, allowing it to cross beyond the valve and into the left ventricle. the device was explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs show stable placement signal and stable 5s parameters. there were no motor current spikes. limited clinical information was provided. the root cause of the optical signal issue was not determined as no product was returned and no data log abnormalities were observed. the root cause of the positioning issue was most likely use related as issue occurred during patient ambulation. the root cause of the product damage (sterile sleeve damage) was not determined as no product or images were returned for analysis a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6252 | 1220648-2025-46098 | 22952189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22952189&pc=OZD | 08/05/2025 |  |  | 09/03/2025 |  |  | 08/05/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707503 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support (mcs) experienced gastrointestinal bleeding requiring intervention and would subsequently expire. the patient was admitted for heart failure exacerbation and work up for advanced therapies. the patient went into cardiac arrest requiring cardiopulmonary resuscitation and return of spontaneous circulation was achieved. the patient was upgraded to the critical care unit in cardiogenic shock and transported to the operating room for innominate veno-arterial extracorporeal membrane oxygenation (va ecmo) and impella 5.5 placement as a bridge to advanced therapies. the innominate artery was cut down. a 10 millimeter graft was attached and tunneled with a plan to cut and sew the bifurcating side arm to be used for arterial cannula and impella 5.5. heparin was given. activated clotting time was 326. oozing at graft site and hypotension were noted. the bleeding was controlled surgically. cellsaver product was given. the decision was made to cannulate va ecmo first due to the patient¿s hemodynamic instability and impella 5.5 to follow, which were completed as decided. the patient was sent to the unit on support. two days later, in the morning, lactate dehydrogenase trended up. an echocardiogram was completed. the impella 5.5 appeared to be in stable position. later that afternoon, the patient was acutely decompensating with metabolic acidosis and lactate 11. active gastrointestinal bleed was noted. the patient was reintubated with increased pressors and inotropes and planned for exploratory laparoscopy. diffuse bowel necrosis was noted, and the patient was subsequently no longer a candidate for advance therapies. the family made the decision to withdraw care and thereafter the patient expired. additionally, of note, the aortic sensor read consistently 20 mmhg lower than arterial line after the device was started. a discussion with the surgeon was had and the surgeon opted to recalibrate the aortic signal. no abnormalities noted to aortic sensor tracing during pump prime and no excessive force was applied during prep and insertion. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned. optical sensor issue: clinical details note that the ao consistently read 20 mmhg lower than the arterial line after the device started. ao signal was recalibrated. data log review showed ps at ~50 mmhg after pump startup. shortly after, a 22 mmhg adjustment was made to the ps. there were no abnormalities with the optical parameters during the case. based on imc logs, there were no abnormalities with optical sensor calibration (factory g0=15764 and calibrated g0=15750). the cause of the optical sensor issue was not determined since product was not returned and insufficient clinical details were provided. major bleed: clinical details note that there was an active gi bleed on 08-aug-2025. heparin had been administered during pump implant. data logs are not applicable. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. device history lot : device lot: 1942024. device history batch : subcomponent lot: n/a. device history review : the pump sn (b)(6) passed all post-sterile inspection checks. the device was not returned; therefore, an evaluation/analysis of the device was not possible. event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. | The impella 5.5 cannot be dissociated from the patient's outcome. it cannot be determined if the use of heparin for implant of the impella 5.5 device impacted/exacerbated the outcome of the bleed for the patient. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6253 | 1220648-2025-46099 | 22952333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22952333&pc=OZD | 08/05/2025 |  |  | 09/03/2025 |  |  | 08/05/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690216 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs). it was reported that after implantation and on initiation of support, the cannula was noted to be severely kinked near the motor housing, compromising forward flow. attempts by the physician to troubleshoot were unsuccessful in resolving the issue. as a result, the impella cp was removed and the patient sent to intensive care unit without mcs. the patient was noted to be in critical condition, and over all outcome is not known. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the impella cp was returned for evaluation. product damage (cannula kink): a kink in the cannula was observed in the returned product. the cause of product damage (cannula kink) was most likely use issue related since cannula was noted to be severely kinked near the motor housing, compromising forward flow. low pump flow: cannula kink was noted. the cause of low pump flow was most likely material kink since the cannula was found to be kinked. device history review: the device passed all post sterile inspection checks. | H6 medical device problem code a140508 added to capture "compromised forward flow" in initial report.
6254 | 1220648-2025-31201 | 22952649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22952649&pc=OZD | 05/08/2024 | 09/03/2025 |  | 09/03/2025 |  |  | 05/08/2024 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161466 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an aic controller experienced an optical sensor failure. there was no patient involvement associated with the noted issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs revealed that at the time of the placement signal low alarm, the optical parameters were within specification. the console was returned for further investigation. the optical system was tested and all optical parameters were confirmed to be within specification. the reported complaint was unable to be reproduced. no issues were found with the console. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6255 | 1220648-2025-46102 | 22953001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22953001&pc=OZD | 07/22/2025 | 09/02/2025 |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  | 03/31/2017 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276512 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during impella 5.5 support, the patient became dizzy and alerted medical staff. the patient then went into cardiac arrest and required cardiopulmonary resuscitation to achieve a return of spontaneous circulation. it was noted that the automated impella controller did not alarm for suction alarms which occurred just prior to the patient¿s cardiac arrest. two days after this event, the patient expired. it was noted this was likely due to septic shock. there was no information available on the source of sepsis. this complaint involves one impella 5.5 pump (serial number (b)(6)) and one automated impella controller (serial number (b)(6)). this mdr specifically addresses automated impella controller, which is the subject of this report. a separate mdr was submitted for the impella 5.5 associated with this complaint. | [5 empty] | Investigation summary ==> data analysis: review of the imc logs confirms the occurrence of a yellow suction alarm at the date and time of the complaint. suction alarms, as well as ¿impella position unknown¿ alarms are prevalent throughout the case the occasional occurrence of ¿impella position in aorta¿ and ¿impella position in ventricle¿ alarms is also noted the rt logs at the time of the complaint show no remarkable variation in the optical 5s parameters conferred with members of the software department who stated that they were unaware of any open bugs or defects that might result in suction alarms with no audible alert, and that after their own review of the data logs they did not see anything that might explain how such an event could have occurred . device analysis: per section 7 of the preventative maintenance procedure the console was placed on a bench and powered on; the usual ¿beeps¿ were all noted during the boot sequence. a purge cassette, purge disc, and impella 5.5 pump were connected to the console, and pump flow was started at p-level 9. this resulted in a yellow ¿suction¿ alarm appearing on the console display; the expected audible alarm was also generated (suctionalarmaudible.mov). the aic hardware was inspected for any loose connectors or cables and no issues were found. root cause: the occurrence of a suction alarm(s) is confirmed in the imc logs, but the ability to trigger this alarm with no audible alert could not be reproduced in the lab. the root cause of the ¿suction alarm w/o acoustic alarm¿ complaint is therefore undetermined. | A.4 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. refer to section h.10 for the related report number representing the impella 5.5.
6256 | 1220648-2025-46101 | 22953252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22953252&pc=OZD | 08/04/2025 |  |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689917 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Life Threatening | Required Intervention |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced an arrhythmic event and expired. the patient was a surgical turndown and was implanted with an impella cp pre percutaneous coronary intervention (pci). the pci was completed and the patient decompensated. the patient experienced ventricular tachycardia and required multiple shocks. the patient was eventually transferred to the unit on impella mcs. the patient continued to decline despite extracorporeal membrane oxygenation efforts. care was withdrawn in accordance with the patient's wishes. the patient expired. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history review confirmed the pump passed all the post sterile inspection checks. regarding the ventricular tachycardia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6257 | 1220648-2025-31202 | 22953559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22953559&pc=OZD | 05/24/2023 | 09/03/2025 |  | 09/03/2025 |  |  | 05/24/2023 | 09/03/2025 |  |  | 11/30/2023 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. post procedure, the pump triggered a "placement signal unreliable" alarm that occurred intermittently. despite the alarm, the motor current and pump function appeared normal. the patient remained hemodynamically stable and continued to receive support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. the data logs were reviewed, and the placement signal unreliable alarm occurred throughout the case. the cause of the optical signal issue could not be determined since the complaint device not returned for investigation.
6258 | 1220648-2025-46104 | 22953877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22953877&pc=OZD | 08/04/2025 |  |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694431 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp for mechanical circulatory support. the impella cp placement was aborted due to the patient having a small vessel for insertion per the medical doctor. the decision was made to abort the procedure and pursue transfer to other facility. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | No product was returned for investigation. delivery issue: the cause of delivery issue is determined to be patient condition because the vessels are too small for insertion and unable to pass through vasculature. device with sn: (b)(6) passed all post sterile inspection checks.
6259 | 1220648-2025-31203 | 22953945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22953945&pc=OZD | 05/14/2023 | 09/03/2025 |  | 09/03/2025 |  |  | 05/14/2023 | 09/03/2025 |  |  | 12/28/2022 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283727 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the measured placement signal disappeared following a shock applied to the left main artery. despite the loss of placement signal, motor current and pump flow remained within normal limits, and patient support continued. there was no patient harm as a result of the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs and pump were not returned. the root cause of the optical signal issue was most likely sensor damage due to an interaction with shockwave device. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6260 | 1220648-2025-46100 | 22954301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22954301&pc=OZD | 08/04/2025 |  |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693771 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was transferred in on an intra-aortic balloon pump for coronary artery bypass graft surgery versus high-risk percutaneous coronary intervention (pci) work up. the decision was made to place an impella 5.5 to optimize the patient for the procedure. the impella 5.5 was successfully implanted, and the patient was transferred to the unit on impella mechanical circulatory support. approximately two days later, the patient underwent high-risk pci noted without issues. the patient was noted to be stable. the following day, the impella 5.5 was slowly weaned with plans to explant in two days. the impella 5.5 was explanted on the day planned, and the patient immediately decompensated. emergently, a new impella 5.5 was prepped and inserted into the axillary graft and advanced into the left ventricle. upon starting the new impella 5.5 pump, motor current was flat and would not increase. alarms noted "block in purge system." the purge line was immediately checked, and no obstructions were noted. the physician immediately removed this impella 5.5 and requested a new impella device. the patient was still actively decompensating. chemical pressor support was initiated in attempt to stabilize the patient. at this time, the physician decided to place an impella cp to support the patient which was successfully completed. however, the patient would expire in the procedural area. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. during pump start, the motor current was flat and block in purge alarm noted. the pump was removed in favor of another pump. data log review showed, the purge pressure rises with lowering purge flows. motor current spike was seen prior to the rise of purge pressure and pump flows also were high after motor current spike. at case end, pump flows begin to trend down but remain within 5.5 p-4 flow range of 1.9-3.3l/min. the cause of the purge issues was biomaterial ingestion due to the rise in purge pressure with purge flow decreases after a motor current spike occurred. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The medical safety officer reviewed the event and noted the following: impella 5.5 pump 1 worked as intended; no known device malfunction/deficiency. the impella was removed as planned and unexpectedly it was noted mechanical circulatory support (mcs) was still needed; therefore, a new impella 5.5 was selected for implant. however, the physician was unable to start the pump after insertion. alarms noted "block in purge system." the purge line was immediately checked, and no obstructions were noted. immediately another impella device, impella cp pump 3, was obtained and implanted without incident; however, the patient would expire approximately 30 minutes of start of impella cp mcs. impella cp pump 3, there was no report of a device malfunction/deficiency and there is no information at hand that suggests, indicates otherwise. the demise outcome is more related to the patient's condition; however, there is not enough evidence to rule out, dissociate the impella 5.5 pump 2 from the patient's outcome. there was an alarm "block in purge system" and the physician was unable to get the impella 5.5 pump 2 device to start in emergency situation. | A.4 is unknown. reason for rapid decompensation after planned removal of impella support and cause of death are unknown. however, based on the details at hand there is not enough information to dissociate the second impella 5.5 from the outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6261 | 1220648-2025-46103 | 22954374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22954374&pc=OZD | 02/07/2025 | 10/21/2025 |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580208 |  | 000000000001000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 placed in (b)(6) 2025 for a patient admitted in for a planned surgery for coronary artery bypass graft and mitral regurgitation. the patient was enrolled in an impact study and was supported for just under six days til wean and explant. the impact team in (b)(6) 2025 notified the complaints team of an adverse event during the support. bleeding from the graft site for the impella delivery was deemed a ¿definite¿ relationship and was noted to be a major bleed that required surgical intervention. the patient was stable post issue. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hemorrhage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemorrhage could not be determined.
6262 | 1220648-2025-46105 | 22954524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22954524&pc=OZD | 08/05/2025 |  |  | 09/03/2025 |  |  | 08/05/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683634 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 139 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported while on impella 5.5 support, the patient experienced bleeding at the impella access site that required intervention. it was reported that on (b)(6) 2025, a patient experienced a witnessed collapse where a bystander initiated cardiopulmonary resuscitation (cpr). emergency medical services were called, and when arrived it was reported that cpr had been in progress for 15 minutes. the patient was transported to a local hospital where they were defibrillated four times enroute due to pulseless ventricular fibrillation (vf). upon arrival to the emergency department, the patient was intubated and briefly coded with ventricular tachycardia (vt) arrest and was defibrillated a total of 26 times for vt. return of spontaneous circulation (rosc) was achieved after administration of lidocaine and amiodarone. the decision was made to emergently implant an impella cp and transfer the patient to the complainant hospital for advanced support. the patient continued to have episodes of vt during impella cp insertion which required defibrillation to successfully convert the patient to normal sinus rhythm. upon arrival at the complainant hospital, the patient was noted to be in third degree heart block. advanced cardiac life support (acls) protocol was initiated, and the patient was taken to the cardiovascular operating room for veno-arterial extracorporeal membrane oxygenation (va ecmo) placement. the patient remained receiving acls and cpr throughout cannulation of va ecmo and was defibrillated multiple times, however remained in vf. the patient received a total of one hour and six minutes of cpr until finally regaining a rosc with normal sinus rhythm. the following day, on (b)(6) 2025, the medical team decided to escalate the patient to an impella 5.5 for increased long-term support. on the second day of support, the patient experienced oozing from the axillary impella 5.5 access site. the patient received four units of packed red blood cells, and a unit of ten packs of platelets. on the third day of support, the patient was found to have herniated their brain stem with progressive signs of neurological issues and was unable to maintain pressures despite va ecmo and impella 5.5 support. the family decided to withdraw care of the patient. care was withdrawn and the patient expired. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6263 | 1220648-2025-46108 | 22954744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22954744&pc=OZD | 08/04/2025 |  |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696374 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient had access site bleeding, requiring interventional cardiology to add an additional stitch. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | No product was returned for investigation. the cause of the access site bleeding is determined to be use issue due to patient movement and the bleeding was resolved and the hemostasis was achieved by adding an additional stitch. device with sn: (b)(6) passed all post sterile inspection checks.
6264 | 1220648-2025-46110 | 22954913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22954913&pc=OZD | 08/04/2025 |  |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637031 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 117 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted via the left femoral artery with an impella cp for mechanical circulatory support. approximately two and a half hours post implant, echocardiogram at bedside was done and the impella measured six centimeters (cm) across the valve. the device was pulled back to measure 3.6 cm. it was noted there was no interference with any structures. performance level was p-9 flowing at 3.8 liters per minute. the following day, possible removal of the impella cp was discussed. it was noted that the patient was not on any inotropes or vasopressors. external femorofemoral bypass was doing well. (there is very limited information and unable to rule out impella related limb ischemia given the device and access.) additionally, and subsequently noted, the impella cp was eventually removed this day due to when turning the patient, the pump dove in and the pigtail wrapped around the mitral valve tendons. the outcome of the patient as a result of this event is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the positioning issue, impella in aorta alarms started after the impella dived in and got tangled around mitral valve tendon during echo. the cause of the positioning issue was an unintentional use error per clinical as the pump came out of position when turning the patient for echo. pertaining to the limb ischemia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6265 | 1220648-2025-31204 | 22955388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22955388&pc=OZD | 02/18/2023 | 09/03/2025 |  | 09/03/2025 |  |  | 02/18/2023 | 09/03/2025 |  |  | 12/31/2024 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298767 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs and pump were not returned. the root cause of the issue was not determined since product and data logs were not returned. all pertinent information available to abiomed has been submitted at this time.
6266 | 1220648-2025-31205 | 22955535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22955535&pc=OZD | 08/28/2023 | 10/27/2025 |  | 09/03/2025 |  |  | 08/28/2023 | 09/03/2025 |  |  | 05/31/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351218 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs shows the continuous placement signal low alarm during case run. all the optical 5s parameter were within specification except the beginning of the case where only signal not reliable values were reported low only during pump was set to p9. the cause of the optical signal issue was most likely patient condition related based on data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | G3 ¿date received by manufacturer¿ corrected. | D6a: corrected the implant date. d6b: corrected the explant date.
6267 | 1220648-2025-31206 | 22956127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22956127&pc=OZD | 06/17/2025 | 11/07/2025 |  | 09/03/2025 |  |  | 06/18/2025 | 09/03/2025 |  |  | 12/31/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447462 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia Stroke | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with an impella cp developed an ischemic stroke following the next day of procedure. it is unclear of the etiology of the stroke, but the patient did suffer a 9mm midline shift and herniated. the patient was made comfort care. care was withdrawn. | The complainant also reported that during implantation of an impella cp, the optical sensor failed and a yellow ¿optical sensor unreliable¿ alarm populated. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for evaluation. optical signal issue the clinical expresses that aic displayed a yellow "optical sensor unreliable¿ alarm. a review of the imc logs indicated a " g0=0" which is an indication that the pump wasn't calibrated correctly and that resulted in the aic displaying the optical sensor unreliable alarm. therefore, the root cause of the optical signal issue was most likely the eeprom corruption which resulted in the component reliability issue. stroke/ cva in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated. h6 updated. d10 concomitant medical products and therapy dates updated.
6268 | 1220648-2025-31207 | 22956199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22956199&pc=OZD | 09/20/2023 | 09/03/2025 |  | 09/03/2025 |  |  | 09/20/2023 | 09/03/2025 |  |  | 06/30/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374870 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the pump operated for approximately one hour before displaying a ¿placement signal unreliable¿ alarm. the pump continued to provide support to the patient. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs showed that the alarm occurred about an hour into support. these parameters are indicative of a broken sensor face. the product was returned and the 5s parameters were similar to the values in the field. light was found leaking from the sensor face. the sensor face was removed and there was built up dextrose on it. the dextrose was washed off and the sensor face was found to be fractured. the root cause of the optical signal issues is most likely due to a damaged sensor face. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6269 | 1220648-2025-46114 | 22956495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22956495&pc=OZD | 08/03/2025 | 09/29/2025 |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564637 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Fever | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while a patient was on impella cp support, the patient became febrile and received antibiotics. later, it was reported that the patient coded twice which required with two shocks per the nursing staff. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported cardiac arrest and fever has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the cardiac arrest and fever could not be determined.
6270 | 1220648-2025-46111 | 22956572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22956572&pc=OZD | 08/04/2025 | 01/29/2026 |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1701427 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 140 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered automated impella controller (aic) battery issues. the battery was not charging when plugged in and so the staff walked through an aic to aic transfer. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6271 | 1220648-2025-31208 | 22956693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22956693&pc=OZD | 10/07/2023 | 09/03/2025 |  | 09/03/2025 |  |  | 10/07/2023 | 09/03/2025 |  |  | 06/30/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369181 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for investigation. data logs show that the placement signal raised an unreliable alarm. this optical signal does not align with any known trend of sensor failure. the doctor confirmed the proper pump position during the placement signal's unreliable alarm. therefore, the root cause of the optical signal issue was not determined since sufficient clinical information was not provided and product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6272 | 1220648-2025-31209 | 22956978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22956978&pc=OZD | 11/19/2023 | 09/03/2025 |  | 09/03/2025 |  |  | 11/19/2023 | 09/03/2025 |  |  | 09/30/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395372 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, continuous placement signal low alarms were observed. the physician confirmed correct device positioning using echocardiography. volume status and positioning were reviewed again. despite the alarms, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs were not returned as well. without the pump data or pump, the complaint cannot be investigated. therefore, the root cause of the issue was not determined since product and data logs were not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6273 | 1220648-2025-31210 | 22957158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22957158&pc=OZD | 12/07/2023 | 09/03/2025 |  | 09/03/2025 |  |  | 12/07/2023 | 09/03/2025 |  |  | 08/31/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024394237 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient admitted for a high risk percutaneous coronary intervention (hrpci). the pump lost placement signal but, remained on without any harm or injury for the hrpci for 0.25 hours. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. pump was returned and investigated. the "placement signal not reliable" alarm cound not be reproduced. data logs were not returned for investigation. without the data logs the failure cannot be further investigated. therefore, the root cause of the optical signal issue was not determined since the failure cannot be reproduced and data logs were unavailable.
6274 | 1220648-2025-31211 | 22957382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22957382&pc=OZD | 02/27/2024 | 09/03/2025 |  | 09/03/2025 |  |  | 02/27/2024 | 09/03/2025 |  |  | 11/30/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434010 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the optical placement signal was lost. despite this, the patient remained hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. impella cp pump was returned for investigation. visually inspected the pump and optical bench 5s parameters were verified. no other abnormalities were found. data logs were reviewed and observed the 5s parameters signal not reliable decreasing. the return product 5s verification and log review indicate failure relative to damaged sensor trend likely due to shockwave interaction. placement signal unreliable alarms observed during this time. as per clinical placement signal lost during shockwave use. the root cause of the optical signal issue was a damaged sensor due to shockwave interaction during support.. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6275 | 1220648-2025-31212 | 22957682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22957682&pc=OZD | 03/04/2024 | 09/03/2025 |  | 09/03/2025 |  |  | 03/04/2024 | 09/03/2025 |  |  | 05/31/2025 |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024367306 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp. while on support, the aortic placement signal was lost. support continued until the impella was weaned. clinical representative confirmed that the pump was not weaned due to the loss of the placement signal. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. data logs show that the placement signal goes out of spec around the same time the clinical mentions shockwave. data logs show that snr drops to 0, contrast and gain drop, and lamp increases. there was a kink in the catheter about half a cm below the motor housing. the pump passed the laser continuity test. the returned 5s parameters showed extremely low snr and lamp at 100%. contrast was low. gain and light are within normal ranges. the cavity length is 32768 nm. no biomaterial was observed. the root cause of the optical signal issue was most likely a damaged sensor due to interaction with shockwave. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer.
6276 | 1220648-2025-46115 | 22959758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22959758&pc=OZD | 08/04/2025 | 10/13/2025 |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689354 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced gastrointestinal bleeding and limb ischemia requiring intervention. however, the patient would subsequently expire. the patient was admitted for possible left ventricular lv pseudo aneurysm and complaint of chest pain anterior st-segment elevation myocardial infarction. veno arterial extracorporeal membrane oxygenation (va ecmo) was cannulated at bedside after the decision to implant the impella 5.5 device was delayed as the patient developed a left ventricle clot and was noted allergy to heparin. the left ventricle clot dissipated and thereafter the decision was made to place the impella 5.5 via the right axillary. the patient was sent to the unit on impella and ecmo support with the plan to remain on these therapies until heart transplant. approximately 12 days later there were concerns regarding the patient¿s lungs as the patient was failing wean trials of oxygenator. computed tomography (ct) of the lungs was completed and nothing significant showed up. however, the patient did go back for another ct as the patient was not improving and was still having shortness of breath. hemoptysis was noted. anticoagulants remained on hold due to bloody secretions. the next couple of days it was noted that the patient had gastrointestinal bleeding. the patient received one unit of packed red blood cells. anticoagulation remained off. the next day, the patient¿s feet were mottled, and the patient was actively bleeding; source unknown. three endoscopic clips were placed for the gastrointestinal bleed. ecmo was decannulated the following day. now day 26 of support, the patient had a hemoglobin of 6.7 g/dl and received one unit of red blood cells. the patient remain intubated this time (noted: the patient was intubated/extubated 5-6 times already). the impella was repositioned due to suction alarms and no issues with suction was noted post repositioning. the patient would continue impella mcs and lungs were noted still not doing well. eventually, the patient would expire 8 days later. care was withdrawn. | [5 empty] | It cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. additionally, given the impella related event of limb ischemia there is not enough information to dissociate the device from the demise outcome. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported vascular injury and ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular injury and ischemia could not be determined.
6277 | 1220648-2025-46116 | 22959789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22959789&pc=OZD | 08/04/2025 |  |  | 09/03/2025 |  |  | 08/04/2025 | 09/03/2025 |  |  |  |  | 09/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686393 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and approximately 10 days after starting the impella support, the patient was having some purge issues. two bags of tissue plasminogen activator were administered, and the issue resolved. however, seven days later, the patient was noted resting this morning and purge flows dropped again. the physician therefore performed several rapid power level changes from performance level p-8 to p-0 back and forth to try to break up anything on the motor. update four days thereafter noted purge issues resolved and seem to be better. the purge cassette was changed. the patient continued impella mcs for an additional seven days before receiving a heart transplant. post transplant the patient arrived back in the cardiovascular intensive care unit now with an intra-aortic balloon pump and vasopressin support. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Initial report section g4: pma/510k number erroneously reflected an additional digit (number 2) at the end and has been corrected accordingly. | B.1 adverse event was incorrectly selected on the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. h.6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the purge pressure high: after 10 days on support, high purge pressure alarms occurred. tissue plasminogen activator was given and alarms resolved. data log review showed 24-hour build-up of purge pressure prior to high purge pressure alarms. no long bag changes were noted. no motor current spikes outside of p-level changes. the root cause of the purge pressure high event was biomaterial build-up; an adverse event related to patient condition.
6278 | 1220648-2025-45844 | 22929851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929851&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026684983 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 105 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced a subacute hemorrhagic stroke. the exact timing of the stroke was unknown. the patient was unable to move their right side as a result of the stroke. the patient's family made the decision to withdraw care as a result of the patient's condition. the patient passed away following explant. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The device was not returned for investigation. the root cause of the stroke/cerebrovascular accident was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6279 | 1220648-2025-46013 | 22929857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929857&pc=OZD | 07/07/2024 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025488409 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Necrosis | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced limb ischemia which prompted a fem-fem to maintain adequate distal flow. the patient required amputation of the leg above the knee. the patient survived. | [5 empty] | The device was not returned for investigation. the cause of the ischemia was not determined due to insufficient clinical details. the pump passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6280 | 1220648-2025-45854 | 22929859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929859&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708061 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon explant of an impella cp pump a perclose suture failed, resulting in access site bleeding. a covered stent was placed and two units of blood were transfused to achieve hemostasis. | [5 empty] | The device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6281 | 1220648-2025-45966 | 22929860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929860&pc=OZD | 07/30/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller generated multiple alarms. there was no adverse impact to patient management and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data analysis: logs from the console associated with pump showed that the placement signal started drifting downwards before pump was transferred a different console, but trending wasn¿t severe enough to trigger alarm 1048. device analysis: not necessary for this complaint, as the reported issue was concluded to have been caused by the pump device history review: not applicable. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6282 | 1220648-2025-46024 | 22929865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929865&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688367 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 61 YR |  | Male | 110 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Patient-Device Incompatibility | Priming Problem |  |  |  |  | Atrial Fibrillation | Atrial Flutter | Fever | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Renal Failure | Hypovolemia | Respiratory Failure | Hematuria | Hemodynamic instability |  |  | Death | Required Intervention | Hospitalization | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (mcs) and would experience complications. impella cp mcs was provided for a total of three days before being escalated to the impella 5.5 device which continued for an additional 18 days. the patient would experience complications and ultimately expire. during the impella cp device explant and the impella 5.5 implant, the patient¿s estimated blood loss was approximately 2.2 liters, receiving 5 units of prbcs, 2 units of platelets, 3 units of frozen fresh plasma, 500 mls. of albumin, and 29,000 units of heparin. the patient was eventually transferred to the unit on impella 5.5 support. on arrival to the critical care unit, there was no pulsatility to the arterial line and the patient¿s urine was tinged pink. a transthoracic echocardiogram (tte) was performed, confirming normal positioning with a depth of 5.3 cm, with severe global hypokinesis and underfilled left ventricle. the patient was given an additional 500 mls. of albumin, with all labs being repeated, and anticoagulation being held until reassessment. it was noted that the patient became febrile yesterday morning, prompting a bronchoscopy, blood cultures, and intravenous antibiotics. the bronchoscopy revealed a foreign body in the lung, suggesting aspiration. the following day, overnight, the patient received 250 mls. of albumin and was transfused one unit of packed red blood cells (prbc) with urine clearing. later this evening it was noted the patient¿s condition remained unchanged/stable. however, an air in the purge alarm was triggered on this day and required a cassette change. three cassettes were used to resolve the issue. an automated impella controller (aic) error occurred as well and was resolved prompting the team to place a backup aic outside the patient¿s room. over the next couple of days intermittent suction was encountered requiring lowering of support level to p-5 then to p-4. over the next three days, the patient was administered 3 units of prbc. a bedside bronchoscopy was performed due to a poor looking chest x-ray. the patient had bloody secretions and was febrile over the night with a temperature max of 102 degrees fahrenheit. blood cultures were sent, and the patient was already on broad spectrum of antibiotics. the patient remained intubated, febrile with a temperature maximum of 102 degrees fahrenheit and positive for cough and gag reflex. the next day, the patient failed spontaneous breathing trial (sbt) due to tachypnea. a chest x-ray indicated acute respiratory distress syndrome (ards). ventilator settings were adjusted. a chest computed tomography (ct) was completed, the following day and indicative of bilateral pneumonia with ards. a beside bronchoscopy performed this morning due to increased oxygen requirements that showed diffuse alveolar hemorrhage in the bilateral lobes. the patient remained intubated and febrile with a temperature maximum of 102 degrees fahrenheit. the patient was hypotensive requiring an increase in vasopressor support and was noted not a candidate for a tracheostomy at this time due to status. due to worsening lung function, the patient was started on continuous cisatracurium. diuretics to be increased to assist with interstitial pulmonary fluid removal. it was additionally noted the patient was reported to have atrial fibrillation (af) with rapid ventricular response (rvr) and the following day the patient had atrial fibrillation and flutter with unsuccessful cardioversion. the patient converted to normal sinus rhythm the next morning. now day 11continuation of support, the patient had ongoing endotracheal tub and oral bleeding that continued through the next day. the patient was given one unit of prbc for low hemoglobin levels at 7.1. heparin remained suspended. an echocardiogram was planned later in the day with hopes of explanting the impella 5.5 as soon as possible due to the patient¿s ongoing inability to tolerate anticoagulation and left ventricle ejection fraction (lvef) at 40%. the next day, the pump was increased to support level p-8; however, a suction alarm occurred, and support level was decreased to p-6. additionally, 250 mls of albumin were given before returning the patient to support level p-8. the following day suction alarms were again encountered, a total of four. however, these self-resolved. heparin was noted to have restarted this day. renal function remained elevated, and the patient was started on continuous renal replacement therapy (crrt) the next day. subsequently, the patient was tachypneic and became acutely hypotensive during an exchange of intravenous tubing. multiple suction alarms were encountered, and support level was briefly dropped to p-5. eventually, the patient was placed in palliative care and comfort care initiated. the patient would ultimately expire. | [5 empty] | Based on the details at hand, there is not enough information to dissociate the impella 5.5 device from the demise outcome. this report on the impella 5.5 is one of three medical device reports associated with the patient's implant experience. refer to the related manufacturer report numbers as noted in section h10. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6. investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. fever/infection/artial fibrillation/atrial flutter: fever/infection/artial fibrillation/atrial flutter: the root cause of the injury was unable to be determined due to insufficient clinical details. respiratory failure/hematuria/hypovolaemia/hemodynamic instability/hemorrhage/major bleed/renal failure/hypotension: the cause of the injury was not determined due to insufficient clinical details. air in purge system/purge system failure: the clinical details state that there was an "air in purge" alarms and 3 purge cassettes had to be used to resolve the alarm. the data logs show that there were air in purge, purge pressure low, purge system open, purge line click-on not detected, and controller errors alarms on 8/1. the purge pressure can be seen to decrease to about 0mmhg with purge flow increasing to about 30ml/h during this time. the alarms resolve on and off. there are no imc logs to confirm the changing of the cassettes or the key presses on the console. due to a lack of product returned, imc logs and sufficient clinical information on what was occurring at the time of changing the cassettes, the root cause of the air in purge system, cannot be established. patient anatomy or condition prior to therapy: the clinical details the patient had many conditions while on cp support before being escalated to the 5.5. due to the lack of sufficient clinical details, the root cause of the patient anatomy or condition prior to therapy cannot be established. pump suction: the clinical details state that the pump had suction throughout support that was typically resolved with lowering the p-level. additionally, the patient was also receiving multiple blood products and was hypotensive throughout support but then also had atrial fibrillation. the data logs show that there were multiple instances of suction triggering throughout support that seemed to have resolved by lowering of the p-level. an rt log was taken during one of the suction events and the ps was spiking with decreases in motor current and pump flows. based on these values and the differential pressure, the alarms seemed to be accurately triggering. the ps did not appear to be an aortic pressure tracing during this time. based on the data logs and clinical details provided, the root cause of the suction issue is most likely due to the patient¿s condition. | The patient was initially supported with an impella cp for three days and experienced complications and is a serious injury. the patient was escalated to an impella 5.5 and supported for approximately 18 days and experienced complications including but not limited to significant bleeding requiring transfusion 5 units. there is not enough evidence to exclude the impella 5.5 as associated factor to the outcome. however, it should be noted the patient presented in critical condition: diaphoretic, anterior/inferior stemi with a 100% lm occlusion. the patient was unable to fully recover from the cardiac event. regarding the aic, patient expired for reasons unrelated to the aic. there was no interruption in support as the aic does not appear to have been exchanged, or issue impacted the therapy. aic device remained in service after the controller failure and defective purge cassette encounter. the medical safety officer reviewed the event and noted the same in that impella 5.5 device cannot be dissociated from the demise outcome. and the impella cp and aic is a serious injury and product malfunction type of reportable event respectively. related medical device reports: this impella 5.5 device report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device reports on the impella cp and automated impella controller.
6283 | 1220648-2025-46033 | 22929886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929886&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708848 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | Restricted Flow rate | Malposition of Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received that the indication for use was cardiomyopathy. | It was reported that a patient implanted with an impella cp experienced suction and impella in aorta alarms. the alarm was disabled. later, the patient was successfully weaned from the pump and survived. |  |  |  |  | B5 additional information was received and added. e1 added initial reporter phone number as it was omitted from the previously submitted reports. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was not returned for investigation. the data logs confirm the complaint condition. the cause of the false alarm issue was most likely related to biomaterial ingestion during the case. the cause of the pump suction issue was most likely related to biomaterial ingestion during the case. the pump passed all post sterile inspection checks.
6284 | 1220648-2025-45995 | 22929923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929923&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026646463 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. upon removal, the patient became aphasic and stroke alert was called. a computed tomography confirmed an embolic stroke and the patient was taken for emergent thrombectomy. a thrombectomy was performed successfully and the patient returned to intensive care unit for recovery. | It was reported that a patient implanted with an impella cp experienced hematuria. in response, the pump was explanted. at the time of explant a cerebrovascular accident occurred and the patient became aphasic. the patient survived. |  |  |  |  | B.5. additional information was received. d.3. manufacturer name submitted on the initial report should not have contained numbers behind it. e.1. initial reporter phone was added as it was inadvertently omitted from the initial report that was submitted. h.6. code e0509 was reported incorrectly on the initial report submitted and a new health effect - clinical code was added. health effect - impact code f1902 was inadvertently omitted from the initial report submitted and was added. h.11. added instructions for use for the health effect - clinical code added. no product was returned for investigation. the root cause of the reported hemolysis was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6285 | 1220648-2025-45841 | 22929934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929934&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026673865 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following successful connection and priming of an impella cp pump, the automated impella controller (aic) displayed a purge pressure low alarm. the set up was verified to be correct and both the cassette and aic were swapped, but the issue remained. the pump was subsequently replaced and the issue resolved. no patient harm was reported. | [5 empty] | The pump was returned for investigation. a crack and a white discoloration stress mark at the reservoir-to-filter joint. a low purge pressure trend was verified on the data log. the cause of the purge pressure low issue was a damaged sidearm at the reservoir-to-filter joint, which was most likely due to unintended use error. the pump passed all post sterile inspection checks. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6286 | 1220648-2025-45963 | 22929949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929949&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026691903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Sepsis | Hematuria | Unspecified Kidney or Urinary Problem | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp exhibited hematuria. the impella was repositioned and continuous renal replacement therapy was started. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H.6 health effect code (e1311) was incorrectly entered on the initial manufacturer device report number. a new code was added. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ investigation summary: hematuria: clinical details noted the patient had red urine and decreased output. the cause of the injury was not determined due to insufficient clinical details. renal injury: clinical details noted that patient was placed on continuous renal replacement therapy at the end of impella support. the cause of the injury was not determined due to insufficient clinical details. sepsis: clinical details noted the patient was in septic shock due to community acquired pneumonia. in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hemolysis: clinical details noted that hemolysis was identified by red urine. the pump was identified to be deep in the left ventricle and was repositioned under echocardiogram. data logs were reviewed and long-term logs showed intermittent suction alarms during support. no positioning alarms were noted at time of reported hemolysis. no motor current spikes occurred prior to reported hemolysis. the cause of the hemolysis could not be definitively determined as no definitive issues were found on returned data logs and clinical details did not note if hemolysis resolved after repositioning. device history review: the pump passed all post-sterile inspection checks.
6287 | 1220648-2025-45996 | 22929963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929963&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025555903 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 121 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella 5.5 the patient experienced optical signal loss and suction. the patient's family decided to withdraw care and the patient expired. | [5 empty] | The cause of the optical signal issue could not be definitively determined as no hard failure of the optical sensor was observed, no product was returned for evaluation, and limited clinical details were provided. the cause of the low pump flow was likely patient condition related as low flows were occurring throughout support while patient was also on ecmo support, and no motor current spikes were observed to indicate ingestion. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6288 | 1220648-2025-45998 | 22929965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929965&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699638 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Sepsis | Cardiogenic Shock | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced multiple complications coinciding with impella cp support. following implant, the patient developed hemolysis. the p level was reduced from nine to seven and the condition began to resolve. the following day, the patient suffered from septic and cardiogenic shock. the cause of the condition was not determined. the patient was started on antibiotics to treat the sepsis. as the patient's underlying condition did not improve, the family opted to withdraw care. the patient passed away following explant. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6289 | 1220648-2025-45843 | 22929988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929988&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026651228 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 82 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp developed ischemia in their left lower extremity, as a result of the loss in flow to the extremity, the decision was made to remove the pump. | [5 empty] | D.3 manufacturer name should not have contained numbers behind it on the initial report that was submitted. manufacturer fax number was added as it was inadvertently omitted from the initial report that was submitted. e.1 initial reporter phone number and zip code extension were added as they were inadvertently omitted from the initial report that was submitted. e.4 selection was added as it was inadvertently omitted from the initial report that was submitted. g.1 . manufacturer contact fax number was added as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks. | A5 is unknown. the impella device was not discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6290 | 1220648-2025-45938 | 22929994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929994&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following impella cp implant, the patient noticed increased mottling from the groin to knee in the right leg and thigh to bilateral flanks in the left. the staff was unable to doppler a pedal signal and the patient felt cold to the touch. the physician opted against intervening and the patient continued to deteriorate from their underlying condition. the patient was placed on comfort care and passed away the following morning. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6291 | 1220648-2025-46016 | 22929995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929995&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026664641 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing and hemolysis. the p level of the pump was reduced, fluids were transfused. additionally, the perclose was tightened and manual pressure was held. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The device was not returned for investigation. data logs showed pump ran without issue during support. the cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of access site bleeding was most likely use related since the bleeding was resolved with tightening sutures. the pump passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6292 | 1220648-2025-45955 | 22929998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22929998&pc=PYX | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  | 02/28/2026 |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025566425 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  |  | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp experienced multiple pump stops due to a suspected thrombus. the pump was explanted and the patient survived. | [5 empty] | The customer's impella rp was returned for investigation. no data logs were returned for case. the returned pump was inspected and biomaterial was found on and beneath the impeller. the pump passed electrical testing. the pump was run in the lab and pump stop was not reproduced. the cause of the issue was unable to be determined as no data logs were returned and issue was not reproduced. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6293 | 1220648-2025-46002 | 22930000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930000&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025570461 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 72 KG |  |  | Injury | Yes | No | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. an echocardiogram showed the impella was positioned deep, so it was repositioned. the next day the patient was successfully escalated to the 5.5. | [5 empty] | The cause of positioning issues was not determined due to insufficient clinical details. the root cause of the hemolysis was most likely due to incorrect positioning since hemostasis was achieved by repositioning the pump per clinical. no issues were found in the logs after reviewing. the pump passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6294 | 1220648-2025-46036 | 22930003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930003&pc=PYX | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026661890 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 93 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered placement signal issues and that the patient encountered an acute ischemic infarction and a parenchymal hemorrhage while implanted with an impella rp flex. the placement signal waveform was bouncing around and dropping off the screen while flows were fluctuating as the placement signal was showing flows as high as 6.0 before the automated impella controller (aic) would alarm placement signal not reliable. the placement signal went out and flow calculation disabled was displayed since approximately midnight. additionally, a cta (computed tomography angiography) of the patient's brain was performed and followed by mri of the brain. both showed a large volume acute ischemic infarction of right mca (middle cerebral artery), pca (posterior cerebral artery), and left sca (superior cerebellar artery) territories with a small area of parenchymal hemorrhage of the anterior right frontal lobe. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6295 | 1220648-2025-46060 | 22930005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930005&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708755 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced symptoms of hemolysis. the pump was explanted from what was noted to be a smaller left ventricle, which may have contributed to the hemolysis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella cp and purge cassette returned for evaluation. visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue and passed hemolysis test. the cause of hemolysis was unable to be determined as limited clinical details were provided and no issues were found on the pump. the pump passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6296 | 1220648-2025-45999 | 22930006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930006&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710508 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery. the implant was aborted as a result of the complication. | [5 empty] | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. delivery issue: the root cause of the delivery issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of the impella. the device history review was completed and the pump passed all the post sterile inspection checks. | A2, a4, a5, and a6 are unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6297 | 1220648-2025-46007 | 22930028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930028&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026675392 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male | 163 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp developed significant access site bleeding. the repositioning sheath angle was adjusted, tension sutures were re-done, and a mattress suture was added, but the bleed persisted. the patient was then weaned from the pump and survived. | [5 empty] | The root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. the device passed all post sterile inspection checks. | A2 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6298 | 1220648-2025-46032 | 22930030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930030&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025604732 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 103 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Biocompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was attempted using a right axillary approach. after multiple attempts, that device implant was aborted. no patient harm was reported. | [5 empty] | The device was not returned for investigation. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass beyond innominate artery because of patient anatomy and insertion was aborted. the device with passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6299 | 1220648-2025-46039 | 22930035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930035&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708755 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that there were concerns of hemolysis for a patient implanted with an impella cp. the pump was explanted from what was noted to be a smaller left ventricle, which may have contributed to the hemolysis. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6300 | 1220648-2025-46030 | 22930052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930052&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708151 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the implanted impella cp was explanted due to multiple repositioning alarms of "impella in the aorta". the patient survived. | [5 empty] | The root cause of the positioning issue was unable to be determined as insufficient clinical details were provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6301 | 1220648-2025-46059 | 22930056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930056&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026696916 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the placement signal of an implanted impella cp pump became erroneous and then dashed out during patient support. later, the patient was successfully weaned from the pump and survived. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. the cause of the optical signal issue could not be definitively determined as limited clinical details were provided and no abnormalities were found on returned product. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6302 | 1220648-2025-46015 | 22930130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930130&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708767 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp with systemic heparin experienced dark, tarry stool indicative of a possible gastrointestinal bleed. the patient was transfused one unit of packed red blood cells. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks. (b)(6). | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6303 | 1220648-2025-45994 | 22930136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930136&pc=OZD | 07/30/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026656015 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an automated impella controller (aic) generated multiple error alarms while supporting an impella 5.5 pump. the aic was swapped for another aic. no patient harm was reported. | [5 empty] | Product complaint # (b)(4). investigation summary ==> 5.5 pump, sn:(b)(6) was returned. optical signal issue: clinical details note that the ao position waveform suddenly deviated from the a-line and showed abnormal values. replacing the console did not resolve the issue. psnr alarm was noted and the ps then did not display values. pump was not removed due to this issue. data log review showed a drop of ~50 mmhg over three hours on 26-jul-2025 before a more gradual decrease in ps reaching almost -50 mmhg on 28-jul-2025 and then going out of range. data logs are characteristic of optical sensor wear. the pump was returned and evaluated. there were no abnormalities with the 5s parameters and the pump passed laser continuity test. the pump did not produce a psnr alarm when plugged into console. the cause of the optical signal issue was sensor wear since the ps downward trend matches the known data log pattern and also occurred after 18 days of support, which is within the intended design life of 60 days per design traceability matrix 0550-9900. phr summary: the pump sn (b)(6) passed all post-sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6304 | 1220648-2025-46018 | 22930140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930140&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026683084 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thromboembolism | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced suction and low flow events. a thrombus was identified. it is unclear how the thrombus was treated. impella support continues. | [5 empty] | The customer's device was returned for evaluation but data logs were not recovered. no abnormalities were noted on the pump. the pump was run in the lab at p-9 for 10 minutes and the low flows were unable to be reproduced. the cause of the low pump flows could not be determined due to no defect found with the returned device, data logs not being recovered, and insufficient clinical details. the cause of the pump suction could not be determined due to no defect found with the returned device, data logs not being recovered, and insufficient clinical details. | D6b. explant date provided. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6305 | 1220648-2025-45997 | 22930146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930146&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026646866 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced bleeding within the iliopsoas muscle. the pump was removed and a transcatheter arterial embolization was performed on the lumbar and iliac arteries. the patient survived. | [5 empty] | The product was not returned for investigation. the root cause of bleeding was determined to be patient condition because the patient was already experiencing multiple organ failure due to circulatory failure, liver function was in the four digits, and coagulation abnormalities were present. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6306 | 1220648-2025-46019 | 22930147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930147&pc=OZD | 07/25/2025 |  |  | 09/02/2025 |  |  | 07/25/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026664373 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation with rapid ventricular response, and also a leak from the luer connection. it is unclear how the atrial fibrillation was treated. the purge cassette was exchanged, but the purge leak was not resolved. a crack was noticed on the female connection side. a steri-strip was placed on the female port and a 6 inch extension line was added to resolve the leak. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6307 | 1220648-2025-46035 | 22930153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930153&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025586657 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues when attempting implantation of an impella 5.5. the physician pushed the catheter to where it folded up on itself with the wire still in the cannula. the wire and catheter were tangled and the decision was made to remove the device. a second pump was successfully implanted. no patient harm was reported. | [5 empty] | No product was returned for investigation. the cause of the delivery issue was determined to be use issue as the physician tangled wire and catheter together that the implant was aborted preventing successful delivery of impella per clinical. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6308 | 1220648-2025-46061 | 22930181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930181&pc=OZD | 07/25/2025 |  |  | 09/02/2025 |  |  | 07/25/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026664373 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support went into atrial fibrillation with rapid ventricular response. no intervention was reported. it was further reported that there was leak from the luer connection. the purge cassette was exchanged, but the purge leak was not resolved. a crack was noticed on the female connection side so a steri-strip was placed on the female port and a 6¿ extension line was added. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The data logs were reviewed and confirm the complaint condition. the returned product revealed that there was a large crack in the sidearm luer connection along with white discoloration and smaller cracks. the root cause of the purge leak - pump is most likely due to a crack in the purge sidearm luer to reservoir likely due to stress from use. the root cause of the arrythmia was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6309 | 1220648-2025-46020 | 22930198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22930198&pc=PYX | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690546 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 87 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with both an implella 5.5 and an impella rp flex lost the optical signal from the impella rp flex. no interruption to patient management was reported and impella support continues. | [5 empty] | The pump was not returned for investigation. the cause of the placement signal issue could not be determined since product was not returned and sufficient clinical information was not provided. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6310 | 1220648-2025-45911 | 22931014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931014&pc=OZD | 08/05/2025 |  |  | 09/02/2025 |  |  | 08/05/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026685009 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a "defective impella catheter" alarm appeared upon connecting an impella cp pump to the automatic impella controller during preparation for patient support. the pump was replaced to resolve the issue. no patient harm was reported. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6311 | 1220648-2025-45953 | 22931016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931016&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1360190 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Operating System Version or Upgrade Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced a system self check failure and software issues. the device was rebooted multiple times, but the issue persisted. there was no noted patient involvement. the device was removed from service in response to the failure. | [5 empty] | The root cause of the ¿system self check failure¿ was determined to be power manager corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | The reported event did not occur during patient management. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6312 | 1220648-2025-46065 | 22931029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931029&pc=OZD | 07/31/2025 | 08/31/2025 |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026680816 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 63 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Difficult or Delayed Positioning | Restricted Flow rate | Decrease in Pressure | Increase in Pressure | Patient-Device Incompatibility | Difficult to Advance | Adverse Event Without Identified Device or Use Problem | Arrhythmia | Hematoma | Low Blood Pressure/ Hypotension | Unspecified Infection | Ischemia | Necrosis | Pain | Renal Failure | Hemodynamic instability |  |  |  |  | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with difficult anatomy experienced multiple events when implanted with an impella 5.5: difficulty delivering and placing the pump, hematoma, access site infection requiring antibiotics, limb ischemia, hypotension and hemodynamic instability requiring blood products, low pump flow, pump suction, low purge pressure, ventricular tachycardia requiring resuscitation, renal failure and acute tubular necrosis requiring continuous renal replacement therapy and continuous veno-venous hemodialysis. the patient eventually expired. | [5 empty] | Correction: b1, h6 medical device problem code 141102. the investigation of the reported failure mode has been completed. no product was received for evaluation. difficult to deliver or advance/patient anatomy or condition prior to therapy: the cause of difficulty delivering the pump or patient anatomy prior to therapy was due to patient condition related. access site adverse event/hematoma: the cause of access site adverse event was not determined due to insufficient clinical information. pump suction: data logs were there was no motor current (mc), placement signal (ps) and pump flow issues. no other issues seen during this time. the cause of pump suction was not determined due to product not being returned and inconclusive log/clinical evidence. low pump flows / high purge pressure: data logs were reviewed and on (b)(6) first there was mc variation/slight downward with pump flow reducing low at p-8 and pp increase. later after few hours flow rises becoming saturated at same p-8 with some pp variation. the mc had upward trend/variation with rise in pump flow, no p-level change. purge pressure slightly increased with small drop in purge flows. the cause of low pump flow/rise in purge pressure was due to biomaterial buildup based on log and clinical review ischemia, tachycardia, infection: the root cause of the injury was unable to be determined due to insufficient clinical details. pain, necrosis, renal failure, hypotension, temporary impairment, hemodynamic instability) the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6313 | 1220648-2025-45907 | 22931033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931033&pc=OZD | 08/02/2025 |  |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703911 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male | 90 KG |  |  | Death | Yes | Yes | Inaccurate Flow Rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 displayed higher than expected left ventricular end-diastolic diameter values and high flow rates. no intervention was performed. impella support continued for another seven days, at which point care was withdrawn due to the patient's prolonged postcardiotomy syndrome. the patient subsequently passed away. | [5 empty] | No product was returned for evaluation. the cause of the optical signal issue could not be determined as limited clinical details were provided and no product was returned for analysis. data logs at time of reported issue were observed and no hard failures of the optical sensor were observed and all optical parameters are stable and within range at time of reported issue. no alarms are noted at time of reported issue. the cause of the saturated flows could not be determined as no product was returned and limited clinical details were provided. data logs were reviewed and showed increased pump flows with slight trend up in motor current. no active alarms at that time. the device passed all post-sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6314 | 1220648-2025-46070 | 22931044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931044&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1116352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that following patient transfer, an automated impella controller (aic) was unable to detect a newly placed purge cassette. the purge cassette was swapped to troubleshoot the issue, but the aic once again failed to detect the purge cassette. the aic was subsequently swapped and the issue resolved. support continued following the console replacement. no patient harm was reported due to the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
6315 | 1220648-2025-45912 | 22931047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931047&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026665184 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that biomaterial was found around the outlet of an impella 5.5 pump following planned explant. no intervention was required in regards to the biomaterial. | [5 empty] | The returned product was investigated. biomaterial was found on the inflow cage. dried blood was seen around the blue suture hub. no other abnormalities were seen. the complete pump was not returned, so further investigation could not be performed. thrombosis: clinical details mention that biological material was found around the outlet area. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | A4, a5, and a6 are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6316 | 1220648-2025-46056 | 22931050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931050&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698601 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp required treatment with amio due to arrhythmia. later, the patient was weaned from the pump and survivied. | [5 empty] | The root cause of the tachycardia was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6317 | 1220648-2025-45956 | 22931076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931076&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1478158 |  | 0042-0000-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following a scheduled procedure, the flow rate of an implanted impella cp pump could not be increased. it was noted that the pump had been placed in surgical mode during the procedure and the cannula had been clamped with a blocking forcep to address a lack of increase in base pressure. upon release and cancellation of surgical mode, the auxiliary flow rate failed to increase and suction alarms appeared. despite repositioning the pump, the issue persisted and the decision was made to remove the device. an intra-aortic balloon pump was placed and the patient remained supported via ecmo. | [5 empty] | A2, a3, a4, and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6318 | 1220648-2025-46052 | 22931081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931081&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026671043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 68 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced atrial fibrillation with rapid ventricular response. an amiodarone drip and second bolus were administered but were ineffective in slowing the heart rate. the patient was successfully cardioverted. after the patient was transferred to another facility, doppler pulses in the right leg were lost. an arterial ultrasound was performed and the patient was treated with an external fem bypass. later, care was withdrawn and the patient expired. | [5 empty] | The customer's device was not returned for investigation. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6319 | 1220648-2025-45867 | 22931085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931085&pc=OZD | 08/05/2025 |  |  | 09/02/2025 |  |  | 08/05/2025 | 09/02/2025 |  |  | 12/31/2025 |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025453952 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Difficult or Delayed Positioning | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp implant, the repositioning sheath could not be sufficiently advanced and hemostasis could not be achieved at the access site. the tissue was noted to be deep and doughy at the site. after multiple failed attempts to reposition the sheath, vascular surgery was consulted to perform a vascular cutdown and surgically repair a tear at the arteriotomy. blood products were given to help treat the patient. the pump was able to successfully support the planned procedure and the patient was successfully weaned from the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6320 | 1220648-2025-46043 | 22931091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931091&pc=OZD | 08/02/2025 |  |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707235 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Pain | Sepsis | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced oozing and hematoma at the access site. a pressure dressing was applied and systemic heparin was withheld. some pain was noted at insertion site. a nitro drip was discontinued. later, the ao placement signal waveform appeared incorrect with elevated left ventricle waveform. the pump was pulled back to resolve the issue. impella support continues. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the hemhorage, pain, and fever was not determined due to insufficient clinical details. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the cause of the bleeding was not determined due to insufficient clinical details. the data logs show very noisy placement signal and show continuous 'impella position unknown' alarms including on the occurrence date. clinical confirmed that the issue resolved after pump was repositioned. the cause of the optical signal issue was due to improper positioning as evidenced by issue resolution after pump pulled back the device passed all post sterile inspection checks. | The customer's device was not returned for evaluation. the cause of the reported hemorrhage, sepsis, and bleeding was not determined due to insufficient clinical details. the device's data logs confirme the optical signal issue. the cause of the optical signal issue was due to improper positioning as evidenced by issue resolution after pump pulled back.
6321 | 1220648-2025-45866 | 22931122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931122&pc=OZD | 11/10/2022 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  | 06/30/2024 |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2023229689 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp developed hemolysis. the pump was explanted the following day and the condition resolved. | [5 empty] | The product was not returned for review. the cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the pump passed all post sterile inspection checks. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6322 | 1220648-2025-45903 | 22931127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931127&pc=OZD | 08/09/2025 |  |  | 09/02/2025 |  |  | 08/09/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026643588 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced bleeding complications. it was documented that the patient was bleeding from their nose and puncture site roughly five days into support. multiple blood transfusions were performed and anticoagulation was monitored to manage the condition. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The customer's device was not returned for evaluation. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6323 | 1220648-2025-45863 | 22931128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931128&pc=OZD | 08/02/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1748660 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that upon connecting an impella 5.5 pump to the automated impella controller, a placement signal not reliable message appeared. the controller was exchanged for a new one to resolve the issue. no patient harm was reported. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6324 | 1220648-2025-46029 | 22931131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931131&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1478158 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The product was used in surgical mode during a vsp operation. to address the problem of a failure to increase base pressure, the pump position was changed several times during the operation, and the shaft and cannula were clamped with a blocking forceps. after treatment, the pump was returned to its proper position, surgical mode was deactivated, and the support level was increased. however, the support flow rate did not increase, and a suction alert occurred at approximately 0.1. further adjustments to the pump position did not result in an adequate flow rate or abatement of suction. therefore, the cp was removed and replaced with an iabp. the patient returned to the room with ecmo and iabp support. | The complainant reported that following a scheduled procedure, the flow rate of an implanted impella cp pump could not be increased. it was noted that the pump had been placed in surgical mode during the procedure and the cannula had been clamped with a blocking forcep to address a lack of increase in base pressure. upon release and cancellation of surgical mode, the auxiliary flow rate failed to increase and suction alarms appeared. despite repositioning the pump, the issue persisted and the decision was made to remove the device. an intra-aortic balloon pump was placed and the patient remained supported via ecmo. |  |  |  |  | B5: additional event information. h6: updated codes according to updated even information. | Correction: d3. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: console logs from the reported day of event are consistent with complaint device analysis: console was tested, but no evidence was found that would implicate the console as the cause of reported flow issues, and suction alarms. during testing the console operated within specification, and no unexpected reboots or restarts were observed. device history review: the device passed all post sterile inspection checks. | A2, a3, a4, and a6 are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6325 | 1220648-2025-46053 | 22931134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931134&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Malposition of Device | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A literature review identified the following title: impella-supported off-pump coronary artery bypass grafting (cabg) in a patient with severe mitral regurgitation: a case report in which an impella cp implanted preoperatively deepened its position during off-pump coronary artery bypass grafting decannulation. this change in position caused the impella to interfere with the posterior mitral valve leaflet and led to severe mitral regurgitation and difficulty with intraoperative circulatory control. the patient is stable. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6326 | 1220648-2025-46066 | 22931136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931136&pc=OZD | 08/02/2025 |  |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026700160 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during the support a large left ventricle thrombus was seen on imaging. the impella 5.5 was explanted to prevent thrombus ingestion into the pump. | [5 empty] | Correction g1: the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6327 | 1220648-2025-46054 | 22931151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931151&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026667568 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 44 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp the pigtail was unable to cross a calcium nodule. the impella was removed and the calcium was stented. the impella was re-inserted but it was still unable to pass through the previous illiac stent. the decision was made to remove the impella and proceed with the intervention without support. | [5 empty] | No product was returned for investigation. the root cause of delivery issue is determined to be patient condition due to calcification in the vessels. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6328 | 1220648-2025-45874 | 22931159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931159&pc=OZD | 08/06/2025 |  |  | 09/02/2025 |  |  | 08/06/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699464 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the anticontamination sleeve of an implanted impella cp pump had a break at the distal end of the catheter lock. no patient harm was reported. impella support continues. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Clinical details state that the anticontamination sleeve was broken at the distal end of the catheter lock during the case. data logs are not applicable and product was not returned. the cause of the sterile sleeve damage was not determined since product was not returned and insufficient clinical details were provided.
6329 | 1220648-2025-45865 | 22931160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931160&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026697519 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 63 YR |  | Male | 75 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received long-term outcome and quality indicator (loqi) impella registry information. it reported "upon ae #1 in subject (b)(6) of postoperative blood loss anemia with relationship to impella listed as possibly. loqi states: subject with baseline h/h of 16.8 g/dl/ 51.8% on (b)(6) 2025. impella assisted procedure performed on (b)(6) 2025 and impella remained in place until (b)(6) 2025. h/h started to trend down on (b)(6) 2025 with a nadir or 6.7 g/dl/18.8% on (b)(6) 2025. subject noted to have ¿oozing/bleeding¿ from groin sites and ij site until impella was removed. subject received 1u prbcs on (b)(6) 2025, 2u prbcs and 1u platelets on (b)(6) 2025, and 1u prbcs on (b)(6) 2025. subject was discharged from index hospitalization on (b)(6) with h/h of 8.1 g/dl/24.4%." | The complainant reported that a patient implanted with an impella cp experienced persistent bleeding issues from multiple access sites. heparin was halted, medications were adjusted, and blood products were given throughout support to manage the ongoing bleeds. later, the patient was successfully weaned from the pump and survived. |  |  |  |  | B.3 event date was updated due to new information received. b.5 new information was received. d.3 manufacturer name should have not contained numbers behind it on the initial report that was submitted. manufacturer fax was inadvertently omitted from the initial report submitted and was added. e.1 zip code extension was added as it was inadvertently omitted from the initial report that was submitted. e.4 selection was made as this field should not have been left blank on the initial report submitted. g.1 manufacturer contact fax number was added as it was inadvertently omitted from the initial report that was submitted. g.2 a new selection was added due to new information received. h1.11 instructions for use were inadvertently omitted from the initial report submitted and were added. the customer's device was not returned for evaluation. the cause of the injury is most likely patient related as the bleeding occurred from all access sites. the device history review was completed and the pump passed all the post sterile inspection checks. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6330 | 1220648-2025-45872 | 22931163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931163&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/05/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026679425 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp suffered a cerebral hemorrhage. the pump was surgically removed in response to the condition. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6331 | 1220648-2025-45952 | 22931169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931169&pc=OZD | 08/08/2025 |  |  | 09/02/2025 |  |  | 08/08/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026683206 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male | 90 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed a hematoma at their access site during impella 5.5 support. the patient was taken to the operating room where an evacuation of the hematoma was performed. support continued uninterrupted following the intervention. | [5 empty] | The pump was not returned for investigation. the pump passed all the post sterile inspection checks. | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6332 | 1220648-2025-45871 | 22931176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931176&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/05/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1087352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the battery of an automated impella controller failed to charge during use. the console was swapped as a result of the issue. there was no patient injury. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6333 | 1220648-2025-46004 | 22931187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931187&pc=OZD | 08/22/2025 |  |  | 09/02/2025 |  |  | 08/22/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1289307 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complaint reported that automated impella controller (aic) was turned on and while going through its boot up stage it would not fully complete the initiation cycle and eventually gave the screen and warning of failed internal system checks, please replace console immediately. the console was removed and a second aic was initiated without incident. there was no patient involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. a.2 should of been left blank on the initial report submitted as no patient was involved. b.5 the description should of said no patient harm was noted instead of no patient involved on the initial report that was submitted. d.3 the manufacturer name should not have contained numbers behind it on the initial report that was submitted. d.3 manufacturer fax was omitted from the initial report submitted and was added. d.9 date the device was returned was added. e.1 zip code extension was added as it was inadvertently omitted from the initial report that was submitted. h.4 device manufacture date was reported incorrectly on the initial report submitted and was updated. h.6 and h.11. updated to reflect the device being returned. | H.6 revised health effect - impact code as it was previously entered incorrectly on the initial report that was submitted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6334 | 1220648-2025-45864 | 22931190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931190&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708904 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 78 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient inadvertently pulled an implanted impella cp pump out of the left ventricle during support. the staff was able to readvance the pump, however following proper placement, persistent suction alarms began to appear. both the volume and positioning were verified, but the suction remained. the pump was subsequently removed as a result of the noted issue. no patient harm was reported. | [5 empty] | The customer's device was returned and visually inspected. biomaterial was observed wrapped around the impeller. no broken blades or cannula kinks were observed. during testing, low flows were not reproduced and flow was within specified limits. the device's data logs were reviewed. suction alarms followed by motor current spikes and sudden drops in impella flow were observed. no placement signal issues were seen. the cause of the low pump flow issue was most likely biomaterial ingestion. the cause of the positioning issue most likely was use related due to improper technique since the patient pulled out the impella. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6335 | 1220648-2025-45870 | 22931195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931195&pc=OZD | 08/05/2025 |  |  | 09/02/2025 |  |  | 08/05/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026698734 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient implanted with an impella cp developed a gastrointestinal bleed. systemic heparin was withheld to manage the condition. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: d3, e4. | B.7 added information as it was inadvertently omitted from the initial report that was submitted. d.3 manufacturer name should not have contained numbers behind it on the initial report that was submitted. manufacturer fax number was added as it was inadvertently omitted from the initial report that was submitted. e.1 initial reporter phone number and zip code extension were added as they were inadvertently omitted from the initial report that was submitted. e.4 added as information was inadvertently omitted from the initial report that was submitted. g.1 added manufacturer contact fax number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed and the product was not returned for investigation. retroperitoneal hemorrhage: as per the clinical, heparin was stopped due to gastrointestinal bleed. the cause of the retroperitoneal hemorrhage issue was not determined due to insufficient clinical information and returned product investigation. the device history review was completed and the pump passed all post sterile inspection checks.
6336 | 1220648-2025-45886 | 22931203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931203&pc=OZD | 08/06/2025 |  |  | 09/02/2025 |  |  | 08/06/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698048 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had no distal flow in their impella leg following impella cp implant. the position of the pump was verified and the doctor declined to intervene. it was additionally noted that the flow rate of the pump steadily decreased in auto mode during support. while initially only able to achieve 2.8lpm, the flow rate decreased quickly to a rate of 1.4lpm. volume was given to the patient in an attempt to increase the flow rate but the issue persisted. after four hours of support, the patient continued to deteriorate from their underlying conditions and passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Product complaint # (b)(4). investigation summary : limb ischemia: clinical reported no flow noted down the leg, and no decision to insert any perfusion catheter. the product was not returned for investigation. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. low pump flow: clinical reported not being able to achieve good flows since the start of the case, reaching 2.3 l/min on auto. position was reported to be good, and patient received 2 l of ns to improve flows, but it did not resolve. the product was not returned for investigation. data log review showed no motor current spikes, but had low pulsatility and a consistent downward trend in placement signal throughout the case. suction alarms occurred throughout the case, and placement signal low alarm occurred once placement signal dropped below 50 mmhg. flow steadily decreased over time, even at high p-levels. the cause of the low pump flow is most likely patient condition, as clinical reported the patient receiving volume support, and placement signal was trending downward with low pulsatility throughout the case, indicating a decline in the patients condition. device history lot : device lot number: 1946909. device history batch : n/a. device history review : pump sn (b)(6) passed all post sterile inspection criteria.
6337 | 1220648-2025-45951 | 22931260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931260&pc=OZD | 08/09/2025 |  |  | 09/02/2025 |  |  | 08/09/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026689073 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the placement signal of an impella cp pump experienced intermittent issues during patient support. following implant, it was noted that the placement signal disappeared on support level p-9. a visual inspection reveal no noted issues with the setup or connections that could have caused the failure. support continued for another two days, during which the placement signal was absent or intermittent on different p levels. the issue had no impact on the functionality of the pump. due to the patient's underlying condition, the family decided to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The cause of the optical signal issue was due to a damaged sensor as evidenced by returned product analysis and log pattern. the pump passed all post sterile inspection checks. | A2, a4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6338 | 1220648-2025-46055 | 22931514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931514&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026617833 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 98 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 exhibited ooze around the staples of the right axillary insertion site. packed red blood cells were transfused to obtain hemostasis. additionally, the patient converted to rapid atrial fibrillation with rapid ventricular response and was started on amiodarone. the patient then converted back into a normal sinus rhythm. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the complaint was confirmed. the device passed all post sterile inspection checks.
6339 | 1220648-2025-45922 | 22931517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931517&pc=OZD | 08/05/2025 |  |  | 09/02/2025 |  |  | 08/05/2025 | 09/02/2025 |  |  | 12/31/2025 |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025453952 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 85 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant the repositioning sheath could not be sufficiently advanced and hemostasis could not be achieved at the access site. the tissue was noted to be deep and doughy at the site. after multiple failed attempts to reposition the sheath, vascular surgery was consulted to perform a vascular cutdown and surgically repair a tear at the arteriotomy. blood products were given to help treat the patient. the pump was able to successfully support the planned procedure following the intervention. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The pump was returned for investigation. pump was returned for analysis. the visual inspection found no issues. the root cause of vascular damage was most likely patient condition related since the tissue at access site of the patient was ¿deep & doughy¿ per clinical description. the root cause of positioning issue was most likely patient condition related since the tissue at access site of the patient was ¿deep & doughy¿ per clinical description. the pump passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6340 | 1220648-2025-45896 | 22931519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931519&pc=OZD | 08/06/2025 |  |  | 09/02/2025 |  |  | 08/06/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690332 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Heart Block | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced an atrioventricular block. the patient was successfully resuscitated, but it resulted in the development of ventricular tachycardia. after cardioversion, the conditions resolved and support continued with the implanted device. the patient was later successfully weaned from the device and survived. | [5 empty] | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the heart block was not determined due to insufficient clinical details. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6341 | 1220648-2025-46003 | 22931608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22931608&pc=OZD | 08/11/2025 |  |  | 09/02/2025 |  |  | 08/11/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1133728 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered automated impella controller (aic) issues. per biomed rep, the aic was in the icu-not on a patient, and alarming a battery issue and that even when the console was plugged in, it continued to alarm. team was unable to duplicate the alarm or message since then but the aic has been plugged into appropriate hospital wall outlet since then and no alarms have been noted. no patient is involved in this event. | [5 empty] | A.1 revised to reflect none as no patient was involved. a.2 this should of been left blank on the initial report that was submitted as no patient was involved. d.3 no numbers should of been reported behind the manufacturer name on the initial report that was submitted. added manufacturer fax as it was inadvertently omitted from the initial report that was submitted. d.9 the device was returned and date received were added. e.1 revised initial reporter phone number as it was reported incorrectly on the initial report that was submitted. added zip code extension as it was inadvertently omitted from the initial report that was submitted. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6342 | 1220648-2025-46006 | 22932455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22932455&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026623474 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon explant of an impella cp closure of the groin site was complicated when the perclose sutures failed. hemostasis was achieved with cutdown surgery. the patient survived. | [5 empty] | No product was returned for investigation. it was stated that the patient had calcified and tortuous vessels. during removal vascular surgery was called for a cutdown as two perclose were attempted and failed due to the calcium. based on the clinical details, the root cause of the vascular complication is most likely due to the patients condition. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6343 | 1220648-2025-46072 | 22932497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22932497&pc=OZD | 08/01/2025 |  |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627461 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 63 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Post procedure narrative: on (b)(6) 2025 - loss wire access in lv on initial insertion requiring removing pump and rewiring ventricle. the complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the plan was to originally place the patient on bipella support: impella 5.5 and an impella rp flex. however, the patient decompensated on anesthesia induction prompting extracorporeal membrane oxygenation (ecmo). the impella 5.5 was inserted and post implant, guidewire access in the left ventricle was lost on initial insertion requiring removal of the pump and wiring the ventricle. the patient was noted to be in stable condition following the event and currently on impella mcs and ecmo support. impella rp implant was noted as possibly for a later time within the said week. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, positioning issue, during the initial insertion, guidewire access was lost in the left ventricle followed by removal of pump and rewiring the ventricle. the root cause of positioning issue is determined to be use issue because of lost wire access and later pump removed and rewired into ventricle. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and noted that the impella device was explanted with a survived patient outcome. further information per case wrap notes indicated the left ventricle recovered, and the right ventricle assist device remained for right ventricular support. d.6b: explant month, day and year were updated accordingly. e.1 zip code extension was added as it was inadvertently omitted from the initial manufacturer report that was submitted.
6344 | 1220648-2025-46051 | 22932522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22932522&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026687150 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced concerns of arrythmia and a tamponade, and concerns of renal failure. the patient was treated with amio gtt, defibrillation, and continuous renal replacement therapy. on the 9th day of the support the patient coded and expired. the relationship between the impella and cause of death is unknown. | [5 empty] | D.3 manufacturer name should not of contained numbers behind it on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. arrhythmia: the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6345 | 1220648-2025-46075 | 22932523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22932523&pc=OZD | 07/30/2025 | 10/14/2025 |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687200 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 86 KG |  | White | Injury | Yes | Yes | False Alarm | Device Sensing Problem | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced ischemia, bleeding and stroke requiring intervention. additionally, there was a false alarm pertaining to the placement signal. the patient arrived at the emergency room and coded. unable to obtain return of spontaneous circulation, the patient was cannulated for extracorporeal membrane oxygenation. impella was placed via the right femoral artery for percutaneous coronary intervention (pci). angiogram showed minimal to no flow in the right lower extremity. the physician noted that the right lower extremity was dusky prior to impella placement. plan for reperfusion catheter on ecmo circuit was noted. pci was performed to the 100% occluded left ostium. the patient remained in unstable ventricular fibrillation prior to impella placement and shocked multiple times without resolution. ecmo venous cannula on the r side, arterial on the l side and reperfusion catheter was placed. the patient was chugging on ecmo circuit entire case. 4l of crystalloid given with minimal resolution. post implant, same day, the placement signal low alarm started. however, imaging confirmed placement of the impella was seated in a "good" position. discussion was made regarding turning off the placement signal alarm. additionally, this day, a computed tomography showed ischemic bowel, and the patient was now bleeding. heparin was on hold. impella cp mcs continued. approximately four days later, the patient 3 showed signs/symptoms of a stroke. computed tomography scan noted a subacute stroke. the neuro team was concerned for left ventricle thrombus; however, no thrombus was noted on echocardiogram with and without contrast administration. plans to remove the impella cp device was made and was completed with a survived patient outcome. | [5 empty] | B.3 revised date of event as it was reported incorrectly on the initial report that was submitted. e.1 revised zip code was it was entered incorrectly on the initial report that was submitted. added zip code extension as it was inadvertently omitted from the initial report that was submitted. g.3 date of awareness on the initial report that was submitted should of reflected 30-jul-2025. h.6 added medical device problem code as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. the product was not returned for investigation. data logs were reviewed and the placement signal (ps) low alarms were confirmed. the first alarm triggered roughly one minute after the pump was started. when reviewing the long term log, optical signal metrics were all stable and within specification for the entire case. reviewing the first real time (rt) log showed that prior to starting the pump, the optical sensor was reading a pressure of approximately 90 mmhg. then, at the time that the optical sensor calibration was completed, it suddenly dropped to 0 mmhg, indicating that the applied pressure caused for improper calibration of the sensor (factory g0=16279 and calibrated g0=15971). mean pressures for the remainder of the case were around 0 mmhg, hence why the ps low alarm was triggering. false alarm/optical sensor issue: clinical details reported ps low alarms even though pump position was confirmed several times and there were no reports of the patient being extremely hypotensive during support. data logs were reviewed and the ps low alarms were confirmed. rt/impella controller logs indicate that the sensor was improperly calibrated because a 90 mmhg pressure was being applied to the sensor during the calibration process at case start. it is unclear what was being done to cause for such a high initial pressure reading, since no clinical details were provided around this time. the product was not returned for investigation. since data log review was inconclusive with limited supporting clinical details and the product was not returned for investigation, the cause of the false alarm/optical signal issue could not be determined. ischemia/stroke: it was reported that at the time of case start, the patient was showing signs of ischemia. the patient also exhibited signs of a stroke. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injuries were unable to be determined due to insufficient clinical details. major bleed: clinical details report that the patient experienced a gi bleed. heparin was on hold. no further details were provided. the product was not returned for investigation. since insufficient clinical details were provided, the cause of the major bleed could not be determined. the device history review was completed and the pump passed all post-sterile inspection checks. | Patient-specific information section a5: ethnicity and b7: medical history is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6346 | 1220648-2025-46074 | 22932539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22932539&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704159 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support prior to coronary artery bypass graft procedure. it was reported on the third day of support, that the patient experienced atrial fibrillation with rapid ventricular response. to resolve the complication the patient was cardioverted multiple times. it was noted there were no impella related issues at the time of the event. the following day, the patient experienced ventricular tachycardia and to resolve the patient was defibrillated 4 times and the impella was repositioned to 42.5 cm. on the sixth day of support, it was reported the patient experienced a stroke and as a result had a small deficit in peripheral vision. heparin was discontinued as a result of the stroke. on the 15th day of support, the patient experienced atrial fibrillation and was cardioverted to resolve. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. nstructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6347 | 1220648-2025-46073 | 22933100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22933100&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701889 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and same day post implant optical signal was lost. no further information surrounding the event was noted. the patient was successfully weaned from the device the following day and device was explanted with a survived patient outcome. | [5 empty] | D.4 revised primary udi number as it was entered incorrectly on the initial report that was submitted, e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history review confirmed the device passed all the post sterile inspection checks. the clinical details state that there was a placement signal not reliable (psnr) alarm that occurred on the 3rd. it is unknown the timeframe on which it started to occur after implanting the pump. due to a lack of product or data logs returned, the root cause of the optical sensor issue cannot be determined. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. | Patient-specific information a4 weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6348 | 1220648-2025-46001 | 22933105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22933105&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 07/31/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026622771 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 102 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding that required sutures. additionally, the patient was put on continuous renal replacement therapy. in response to abnormal flow pattern the pump was repositioned. the patient was successfully weaned from the pump and survived. | [5 empty] | B.1 added product problem as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.6 code f1901 was entered incorrectly on the initial report that was submitted and has been updated. the investigation has been completed. no product was returned. major bleed: nose bleed was observed. the cause of major bleed could not be determined as insufficient clinical details were provided. renal failure: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. positioning issue: the cause of the positioning issue was unable to be determined as insufficient clinical details were provided. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6349 | 1220648-2025-45957 | 22936863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22936863&pc=OZD | 07/31/2025 |  |  | 09/02/2025 |  |  | 08/04/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660287 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The product was used in surgical mode during a vsp operation. to address the problem of a failure to increase base pressure, the pump position was changed several times during the operation, and the shaft and cannula were clamped with a blocking forceps. after treatment, the pump was returned to its proper position, surgical mode was deactivated, and the support level was increased. however, the support flow rate did not increase, and a suction alert occurred at approximately 0.1. further adjustments to the pump position did not result in an adequate flow rate or abatement of suction. therefore, the cp was removed and replaced with an iabp. the patient returned to the room with ecmo and iabp support. | The complainant reported that following a scheduled procedure, the flow rate of an implanted impella cp pump could not be increased. it was noted that the pump had been placed in surgical mode during the procedure and the cannula had been clamped with a blocking forcep to address a lack of increase in base pressure. upon release and cancellation of surgical mode, the auxiliary flow rate failed to increase and suction alarms appeared. despite repositioning the pump, the issue persisted and the decision was made to remove the device. an intra-aortic balloon pump was placed and the patient remained supported via extracorporeal membrane oxygenation. |  |  |  |  | Product complaint # (b)(4). investigation summary: impella cp returned for evaluation. low pump flow/suction: clinical details report that a vsp procedure was done on (b)(6) 2025. during the procedure, the pump cannula had to be clamped and moved. once the procedure was completed, the clamp was removed, the pump was brought back to appropriate position, and the pump was taken out of surgical mode, but flows were low and suction alarms began to trigger. attempts to reposition the pump were made, but the issue persisted, so the pump was explanted. data logs were reviewed, and the reported issue was confirmed. after the pump came out of surgical mode at the end of the case, there were a number of suction alarm #73 triggers, indicating low mc values for the p-level. flows were never able to get higher than 0.5 l/min either. the rt log also showed mc modulation and values to be quite low. there were no indications of abnormal spikes. product was returned with a cut in the catheter, and only the motor housing side was returned. there was no damage to the impeller, but there was evidence of two separate kinks in the cannula as well as a small amount of biomaterial near the impeller. the pump was unable to be run to attempt to reproduce low flows because of the cut in the catheter. based on the report in clinical details that the cannula was clamped just prior to the low flows, and returned product had evidence of several kinks in the cannula, the cause of the low pump flows was determined to be a cannula kink. the finding of biomaterial on returned product could potentially have been a downstream effect of the cannula kink, however, unable to be related to low flows since there were no indications of biomaterial ingestion in data logs. device history lot: device lot: 1893775. device history batch: subcomponent: na. device history review: pump sn (b)(6) passed all post-sterile inspection checks. | B5 additional event information received h6 codes updated according to | A.2, a.3, and a.4 are unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. refer to section h.10 for the related report that represents the automated impella controller.
6350 | 1220648-2025-31187 | 22937051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22937051&pc=OZD | 08/19/2025 | 12/03/2025 |  | 09/02/2025 |  |  | 08/20/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1912186 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that a recurring placement signal not reliable alarm began to appear to impella 5.5 support. the alarm was disabled and the staff was advised that the signal did not impact the function of the pump. the staff was instructed to monitor the patient's hemodynamics and motor current / flow rate for any signs of pump deficiencies. support continued uninterrupted. | [5 empty] | The investigation has been completed. data analysis: 5.5 pump (b)(6) was connected on (b)(6) 2025 to (b)(6) through (b)(6) 2025 with all logs from rlfs. the imc logs had three instances of placement signal not reliable alarms during the case all on (b)(6) or (b)(6). the lt logs show that for the first couple of days, snr is low as placement signal is railing triggering the unreliable placement signal alarm. the rt logs concur at this time placement signal is switching between a nominal value and 3000 mmhg as the raw optical sensor similarly jumps to -3276.8 mmhg and snr is <4000. the optical signals are as expected for the majority of the case after the first days of the case. device analysis: the cause of the optical signal issue was not determined since product was not returned for investigation. also, it is apparent that the console could be potential cause of the issue. root cause: the cause of the intermittent unreliable placement signal could not be determined as the console was not returned for investigation. d10 concomitant medical products and therapy dates updated. | Data analysis: 5.5 pump (b)(4) was connected on 7/18 to (b)(4) through 8/28 with all logs from rlfs. the imc logs had three instances of placement signal not reliable alarms during the case all on 7/18 or 7/19. the lt logs show that for the first couple of days, snr is low as placement signal is railing triggering the unreliable placement signal alarm. the rt logs concur at this time placement signal is switching between a nominal value and 3000 mmhg as the raw optical sensor similarly jumps to -3276.8 mmhg and snr is <4000. the optical signals are as expected for the majority of the case after the first days of the case. device analysis: the cause of the optical signal issue was not determined since product was not returned for investigation. also, it is apparent that the console could be potential cause of the issue. root cause: the cause of the intermittent unreliable placement signal could not be determined as the console was not returned for investigation. h6: added codes per the results of the investigation. h10: added the results of the investigation. | Medical safety reviewed the event and noted the following: the patient was brought to the operating room for placement of an impella 5.5 device as a bridge to coronary bypass grafting (cbg). upon evaluation, the axillary artery was determined to be too small to accommodate the impella 5.5. an impella cp was implanted without issue. post-implant, hemolysis was noted, which is considered a reportable adverse event. following cabg, mitral valve replacement (mvr), and tricuspid valve replacement (tvr), the decision was made to place an impella 5.5 directly. after 24 hours of support, the impella 5.5 generated a placement signal not reliable alarm. placement signal values and waveforms were present, and the alarm did not impact hemodynamic support. the alarm was disabled. during support, the patient developed volume overload and acute kidney injury (aki), requiring initiation of continuous renal replacement therapy (crrt). on day 9 of impella 5.5 support, the patient began to decompensate and experienced a gastrointestinal (gi) bleed. the patient was suspected to be in septic shock. after continued support, it was determined that the patient was no longer a candidate for permanent lvad therapy. the care plan shifted to continued impella 5.5 support with the goal of native heart recovery. on august 18, 2025, the impella 5.5 was repositioned, resulting in improved flows. despite ongoing support, the patient expired after 41 days of impella support. the impella cp hemolysis is a reportable adverse event. impella 5.5 placement signal not reliable alarm are considered device malfunction reportable events. the renal insufficiency, infection and gi bleed are reportable patient adverse events for the impella 5.5. the patient¿s death will be conservatively reported on the impella 5.5, although it is unlikely that the device contributed to the cause of death. the patient¿s deteriorating condition and non-candidacy for permanent lvad were likely contributing factors. note: the automated impella controller will remain a reportable product malfunction accordingly. associated manufacturing report numbers are 1220648-2025-30922 (impella 5.5) and 1220648-2025-31187 (aic). | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6351 | 1220648-2025-46071 | 22937701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22937701&pc=OZD | 08/03/2025 | 10/28/2025 |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026617547 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Death | Yes | No | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella was started in auto mode, turned to p9, suction event, resolved on p6. the pump supported but there were signs of hemolysis which prompted pump repositioning, 1 unit of red blood cells, and continuous renal replacement therapy. the pump remained on despite the hemolysis (plasma free hemoglobin draw of 200) for 4 days til the family made decision to withdraw care. | [5 empty] | The investigation for the reported positioning issue and mechanical interaction with blood has been completed. no product was returned for review. data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. the root cause of the mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. | Patient race and ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6352 | 1220648-2025-46084 | 22938877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22938877&pc=OZD | 08/01/2025 | 10/17/2025 |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674736 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 47 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced limb ischemia requiring intervention. the patient presented to the emergency department was rushed to the catheterization lab and coded. a temporary pacer wire was inserted followed by the impella cp device via the right femoral artery. percutaneous coronary intervention was completed to the right coronary artery occlusion with two drug-eluting stents deployment. thereafter the physician assessed the right leg pulse, and no pulse was present on doppler. ischemic right leg was noted. the decision was made to explant the impella cp device. two perclose devices were deployed; site was dry and intact without bleeding or hematoma. follow up check on the right leg noted palpate pulse by doppler. the patient was transferred to the unit in stable condition. | [5 empty] | E.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. the investigation for the reported ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for the impella cp with smartassist system: section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6353 | 1220648-2025-46083 | 22938885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22938885&pc=OZD | 08/01/2025 | 08/30/2025 |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690227 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 55 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 as an escalation of therapy from an impella cp device who had been admitted with cardiogenic shock. it was reported that during support; purge pressure high alarms were noted. pharmacy and medical affairs spoke regarding the purge composition being changed to tissue plasminogen activator lysis (tpa). tpa was added to the purge and the purge pressure and flows improved. the device was subsequently weaned, explanted, and the patient was successfully bridged to a durable left ventricular assist device. the impella explant was not performed earlier than planned as a result of the complication. | [5 empty] | Product status: the impella device was not received from the customer; there is uncertainty what happened with the device after the patient's demise. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A.4 weight was added as it was inadvertently omitted from the initial report that was submitted. h.6 health effect - impact code listed on the initial report was incorrect and a new code was added. the impella device was not received from the customer, but the event logs were received. the investigation has been completed. the device history review was completed and the pump passed all post sterile inspection checks. purge pressure high (pph) - after six days on support, pph alarms were noted. tissue plasminogen activator (tpa) lysis protocol was executed and purge metrics began to improve. the pump was not returned. data log review showed purge pressure gradually rose on 01-aug-2025 for three hours prior to pph alarms and then resolve within the next three hours. the cause of the high purge pressure was biomaterial buildup due to data log review showing a three hour rise in pp that resolved after executing tpa lysis protocol. h.6 codes were updated to reflect investigation results.
6354 | 1220648-2025-46086 | 22938930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22938930&pc=OZD | 08/01/2025 | 08/30/2025 |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026694733 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Unspecified Infection | Hematuria | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was additionally reported that on the thirteenth day of support, it was reported that the patient's laboratory samples were hemolyzed and the patient had red urine. two days following, the signs of hemolysis cleared without any interventions noted. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported on the 11th day of support that the patient experienced a fever and hematology results indicated the patient¿s white blood count was elevated. blood cultures were drawn, and the patient¿s pulmonary artery catheter was removed. the following day, the blood cultures were reported to be positive for bacterial growth and the patient was started on antibiotics. the same day, it was noted there was a change in impella wave forms and flows, and there was concern for saturated flows. as a result, the p level of the pump was reduced to p6 from p7. later that day, after the patient was moved to the operating room to exchange the 5.5, flows and wave forms stabilized. the decision was then made to not exchange the impella 5.5 pump. the patient remained on impella 5.5 support until blood cultures returned negative growth results. the patient was then successfully bridged to a durable left ventricular assist device and the impella 5.5 was explanted. . on explant of the impella 5.5, a large thrombus was visualized at inflow of the impella. |  |  |  |  | Additional details added to b5 summary and h6 health effect - clinical code for hemolysis and hematuria. h6 codes added for investigation results. investigation summary: 16-sep-2025: pump sn (b)(6) (cut) was returned for investigation. data logs were reviewed and the report of saturated flows was confirmed on (b)(6) 2025. on this date, there was a step up in the motor current (mc) while at p-7 that resulted in saturated flows. the p-level was then dropped to p-6, but pump flows remained saturated through (b)(6) when there was a sudden step back down into the expected range. there were no suction alarms and only 1 brief position unknown alarm during this period. purge pressure remained stable. pump flows remained in specification for the rest of the case. product was returned with a cut in the catheter, so some investigation tests weren't possible. there were no relevant visual abnormalities or defects. there was no biomaterial and there was no damage to the impeller or cages. there was a minor cannula kink near the motor housing, however, it was deemed to be unrelated to the reported complications because there were no indications of suction or low pump flows throughout the case. there is potential that it was kinked during the explant/return process. saturated flows: clinical details report a change in waveforms and flows on (b)(6) 2025. the patient was brought into the or for explant on (b)(6), but at this time, flows returned to normal and imaging showed no thrombus on the inflow/outflow cages. data logs did confirm saturated flows on this date following a sudden step up in mc. however, there were no signs of the typical ingestion signature. additionally, there were no changes to purge pressure or any improper position alarms. the resolution of the saturated flows aligned with the date of the patient going into the or, and flows returned to specification for the remainder of the case. returned product did not exhibit any abnormalities that could be related to the saturated flows. since limited clinical details were provided, there were no relevant abnormalities on returned product, and data logs were inconclusive, the cause of the saturated flows could not be determined. hemolysis: clinical details reported that the patient's urine had turned red around the time of saturated flows on (b)(6) 2025. the pump flows returned to normal when the patient was brought into the or on (b)(6) 2025, hemolysis had resolved by (b)(6) 2025, and the pump was explanted on (b)(6) 2025. no troubleshooting methods were provided beyond checking position and reducing the p-level. data logs confirmed the saturated flows, but there were no signatures of an ingestion event. additionally, there were no suction alarms or positioning alarms to suggest that these were the causes of hemolysis. returned product had no visual abnormalities that could be correlated to the hemolysis. since insufficient clinical details were provided regarding what resolved the issue, returned product did not have any relevant abnormalities, and data logs were inconclusive, the cause of the hemolysis couldn't be determined. thrombus/infection: clinical details report that the patient was placed on antibiotics due to positive blood culture test results on (b)(6) 2025. additionally, thrombus was visualized in the inflow cage upon pump explant. in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the injuries were unable to be determined due to insufficient clinical details. device history review: pump sn (b)(6) passed all post-sterile inspection checks. | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6355 | 1220648-2025-46076 | 22939054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939054&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641408 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was hypotensive with a 100% left anterior descending artery (lad) blockage. an impella cp was placed for mechanical circulatory support (mcs), and the lad was stented. post impella implant pedal pulses were present and the access site dressing was clean, dry and intact. the patient was transferred to the unit on impella mcs. the following morning, however, the patient was noted to have oozing and hemolysis complications which were described as oozing at the insertion site and rather, hematuria was present. epinephrine and lidocaine were injected into groin site after the femstop was removed. no hematoma was present. the access site was managed. the nurse fixed the angle of insertion and was stable. it was noted that the team would continue to intervene with an additional suture if needed. regarding the hematuria, the impella was repositioned at bedside under echocardiogram this day in the morning. the final position was 3.8 centimeters. plasma-free hemoglobin was ordered to further assess the hematuria with plan to diurese and rest the patient. additionally noted, systemic heparin was paused this day in the morning due to supratherapeutic anti-factor level. the following day no impella related issues were noted. the patient was successfully weaned, and the impella cp was explanted with a survived patient outcome the next day. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding hemorrhage/ bleeding, there was oozing at insertion site and the root cause is determined to be use issue because the bleeding was resolved by matching the angle. pertaining to the hematuria, the root cause of the injury was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ bleeding is being coded for the stoppage of systemic heparin due to supratherapeutic anti-factor level.
6356 | 1220648-2025-46088 | 22939068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939068&pc=PYX | 07/29/2025 | 10/10/2025 |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025571386 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 104 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp flex device for mechanical circulatory support and post implant same day the device was not streaming on impella connect. support continued uninterrupted and ended approximately eight day later. the impella rp flex was explanted at bedside with no hemodynamic changes. the patient remained on all other therapy modalities. there was no report or indication of harm to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported optical signal issue has been completed. neither product nor data logs were available for investigation. since the pump nor data logs were returned for investigation, the cause of the optical signal issue could not be determined.
6357 | 1220648-2025-46087 | 22939069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939069&pc=OZD | 08/01/2025 | 01/27/2026 |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690254 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Backflow | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support to bridge to heart transplant. it was reported on the 16th day of support, the impella had a sensor failure which then initiated an ongoing ¿preventing retrograde flow¿ alarm. it was noted that the calculations were not accurate due to the optical sensor issue. the team was instructed to monitor with daily echocardiograms and to monitor purge flow and motor current. support was not affected by the sensor issue. the patient remained stable and was successfully bridged to heart transplant. | [5 empty] | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: retrograde pump flow/optical sensor issue: clinical reported sensor failure which resulted in retrograde flow alarm. no imc logs were provided. analysis of lt and rt logs revealed sudden fluctuations in placement signals, with pump flow oscillating between 0 and 5 l/min the logs confirm alarms #1048 and #1036 placement signal not reliable alarms. 5s parameters were relatively stable with periodically low and widespread snr values while placement signal went out of range intermittently. further review of logs from the exchanged aic (ic12185) showed that retrograde flow (#426), placement signal not reliable (#1036), and placement signal not reliable (#1048) alarms continued for two days. the cause of the optical signal issue was not established as no characterized failure patterns were identified, limited clinical information was provided and no product was returned for analysis. device history lot: device lot number: 1932216. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
6358 | 1220648-2025-46085 | 22939120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939120&pc=OZD | 08/01/2025 | 10/10/2025 |  | 09/02/2025 |  |  | 08/01/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677120 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 62 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Material Split, Cut or Torn | [5 empty] | Arrhythmia | Fever | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. approximately one day after starting the support, placement signal alarms/purge pressure alarm were disabled at this time due to intra-aortic balloon pump causing interaction/interference. the patient had some ventricular tachycardia overnight; the impella support level was dropped to p-7 overnight due to elevated plasma free which seemed to resolve the issue. the patient would continue on support. four days later it was noted an electrophysiologist was consulted for the patient's ventricular tachycardia as one episode occurred overnight causing the patient's implantable cardioverter defibrillator to shock once. the patient was intravascularly dry; therefore, fluid was given to prevent further suction alarms. three days later the patient experienced one suction alarm that was self-resolved and again suction was encountered. the positioning was checked under echocardiogram that showed good depth but deflecting against the wall and causing suction. observed this with imaging on p-5 and is only currently able to maintain p-4 support level due to close proximity to heart structures. discussion made with fellow that if they can torque the pump the flip in the left ventricle then these issues will resolve. staff agreed but no one on site available and willing to reposition. the patient¿s respiratory status declined some overnight from giving too much fluid the day prior. the patient was put on bipap temporarily. the patient was now edematous and given albumin now to pull fluid back into the vasculature. plan noted to optimize breathing and fluid status. repositioning took place the following day and during positioning, the physician was trying to torque the catheter in order to attempt flipping the impella as it was deflecting against the heart wall causing suction events. with the impella 5.5 device this was found difficult as the blue button must be pressed down the entire time in order to reposition. doing this while trying to twist the catheter within the sleeve was difficult, and the physician accidentally tore the anticontamination sleeve. the decision was made to discontinue the attempt to reposition, and a tegaderm was placed over the anticontamination sleeve tear. impella support continued. however, now there was concerned about a low-grade fever may be coming from this issue. team discussed to determine best steps to resolve, and the patient ended up receiving prophylactic antibiotics. the patient seemed to be improving and plan to extubate the next morning. however, of note, heparin was paused for now due to the bloody secretions. some days later the patient had some ventricular tachycardia paired with aggressive diuresis resulting in suction alarms. the next day the patient was noted feeling good and had a run of asymptomatic ventricular tachycardia requiring one shock. the next day again, the patient had runs of ventricular tachycardia requiring shocks. the patient was asymptomatic, however. the impella 5.5 device was repositioned at the bedside and pulled back 1cm to reduce ectopy. there were no more ventricular tachycardia episodes, but the support level was unable to go above p-4. it was explained how the cannula is deflected towards the outer wall and as the flow goes up it suctions to the wall. the challenge noted was getting to cannula flipped around. the team discussed how to make this changed after the failed earlier attempts. the patient continued support waiting for left ventricular assist device procedure scheduling. latest update, day 30, noted the patient doing well and any arrhythmic events. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported low pump flow, product damage, infection and fever has been completed. data log review did not show any motor current spikes/trend indicating biomaterial interaction. the root cause of the suction low pump flow was not determined since the reason they were unable to reposition correctly is unknown and insufficient product was returned. the root cause of the product damage was use-related as clinical details indicate that the medical doctor accidentally tore the sleeve while repositioning and the returned product had a torn sleeve. the root cause of the infection and fever could not be determined.
6359 | 1220648-2025-46089 | 22939213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939213&pc=OZD | 08/02/2025 | 09/30/2025 |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664251 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 110 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. now the next, the initially implanted impella cp was explanted for possible flow limitations. the pump flow did not flow above .5lpm regardless of performance level. positioning was verified multiple times with echocardiogram, noting freely moving pigtail and measuring 3-3.8cm from annulus. the decision was made to replace the impella cp device. using same access site and the access port, 0.035 guidewire was used to remove and place new 14fr peel away sheath. the new impella cp was implanted and started without complications. repositioning sheath was sutured, and good distal pulses were noted. the patient was placed on support level p-9 for maximize unloading. after approximately five days of impella cp support, the patient was successfully weaned, and the device was explanted with a survived patient outcome. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6360 | 1220648-2025-31188 | 22939231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939231&pc=OZD | 11/20/2024 | 09/02/2025 |  | 09/02/2025 |  |  | 11/20/2024 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025514040 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during impella support, a position-related alarm occurred and position adjustment was performed. after position adjustment, a "placement signal not reliable" alarm appeared and the position waveform disappeared. despite these issues, the pump remained operational throughout support until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned product was inspected. the 5s parameters showed very low signal not reliable. the pump passed the laser continuity test. the data logs show a motor current spike coinciding with a drop in signal not reliable, contrast, gain, and a step up in lamp with placement signal spiking out of range. the placement signal not reliable alarm was triggered shortly after. this pattern is characteristic of a damaged sensor. this aligns with clinical mention of issue occurring after a position adjustment. there is no other details about what occurred in this case or any other devices used during this time period. the root cause of the optical signal issue was determined to be due to a damaged sensor but the cause of how the sensor became damaged is unknown. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6361 | 1220648-2025-46090 | 22939308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939308&pc=OZD | 08/02/2025 | 08/30/2025 |  | 09/02/2025 |  |  | 08/02/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662586 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 106 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the impella cp was placed for bailout after a noted possible dissection of the right coronary artery (rca), which led to the cardiac arrest of the patient. resuscitation efforts were eventually ceased, and time of death was called. high motor current spike after patient was declared to be deceased, and impella was stopped by the physician. the patient was transferred to the icu for postmortem care and once in the icu, the patient's heart rate spontaneously commenced, and patient began to independently breathe. (note: the time in between death declaration and heart rate commencing is unknown. however, it was estimated to be 10 minutes). the impella cp device was restarted. however, a decision made to insert new impella cp, and coronary artery bypass graft surgery was continued, completed bypassing the rca. the patient¿s chest was left open, and the patient was very volume deficient. the patient was transferred to the intensive care unit for further support. however, the patient continued to deteriorate, family decided to withdraw care, and the patient expired. | [5 empty] | Additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly. sections d.9 and h.6 investigation type, findings and conclusion codes were updated accordingly. e.1 zip code extension was added as it was inadvertently omitted from the initial manufacturer report number 1220648-2025-46090. | The product with event logs were received, and the investigation was completed. device history record review was completed, and the pump passed all post sterile inspection checks. clinical details reported pump stoppage with failed restart after the patient was transferred to the operating room, and impella support was reinitiated after the patient was pronounced dead. the returned pump had no visual abnormalities and was able to run in the lab with no pump stops. the data logs confirmed the pump was stopped by physicians for about 13 minutes before restart. the pump was restarted and was at p-1 when there was an 'impella stopped - motor current high' alarm. the motor current had risen from 200s to 500s before alarm and there were placement signal spikes and a drop in motor speed with flows dropping to 0 as pump stopped. the logs showed the pump position was in the ventricle during final minutes of support. logs showed a failed restart attempt with another 'impella stopped - motor current high' alarm. the cause of the issue was not determined as the returned pump functioned as expected and limited clinical information was provided to confirm clinical timeline and events. h6. investigation: type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. | Medical safety officer reviewed the event and noted there was no device malfunction and/or adverse event during support to the patient. high motor current spike after the patient was declared dead. impella was stopped by the physician, and the patient to icu for postmortem care. after, the patient spontaneously began to have a pulse and respirations after time of death was called. the impella was restarted. however, a decision made to insert new impella cp. device status: the impella device was not received from the customer at the time of this mdr writing, therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6362 | 1220648-2025-46077 | 22939545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22939545&pc=OZD | 08/03/2025 | 10/24/2025 |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684335 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 110 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 had a brief high purge pressure system blocked alarm. the purge cassette was replaced but the issue was not resolved. suggestion was given to the team to start tissue plasminogen activator. the purge pressure was resolved. it is unknow how the issue was resolved. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure issue has been completed. pump was not returned for investigation. data log shows approximately 5 days of gradual rise in purge pressure followed by the high purge pressure alarm. there were no motor current spikes reported during the purge issue. also, the purge trend returned back to the normal specification range with a gradual decreasing trend. the root cause of the high purge pressure issue was most likely biomaterial buildup based on data log review.
6363 | 1220648-2025-46079 | 22940872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22940872&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680752 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction/cardiogenic shock was being escalated from an impella cp to an impella 5.5. when placing the impella 5.5, there was resistance at the axillary artery. the physician was unable to make a drop from the innominate artery into the aorta. the medical doctor stated that extracorporeal membrane oxygenation cannula was in the way and the impella 5.5 was unable to drop into position due to that. additionally, the impella 5.5 noted having placement signal lost and placement signal not reliable alarm occurred. impella 5.5 insertion was aborted after multiple attempts. another impella 5.5 was inserted and this time was successful. no patient harm was reported due to the issue. | [5 empty] | Product complaint # (b)(4). investigation summary: unable to deliver/advance: clinical reported ecmo cannula was in the way and pump was unable to drop into position due to that. the cause of the issue was determined to be due to a use related issue. optical signal issue: impella noted to have lost placement signal and psnr alarm occurred after failed delivery attempt with ecmo cannula in the way. logs confirm loss psnr alarm and show immediate loss of placement signal and drop in snr. at that instant there was also a drop in light. gain increased slightly and contrast increased in amplitude. the pattern observed is characteristic of a fiber break. the cause of the issue was a fiber break as evidenced by log pattern likely due to a use-related issue experienced when attempting to deliver pump. device history lot: device lot number: 1929964. device history batch: subcomponent lot: n/a. device history review: the pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6364 | 1220648-2025-45873 | 22941009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22941009&pc=OZD | 07/22/2025 |  |  | 09/02/2025 |  |  | 07/22/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692944 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient endured frequent ventricular ectopy and ventricular tachycardia during impella 5.5 support. the positioning was verified via echocardiogram. sixteen days into support, the pump began to experience suction issues. attempts to increase support were constantly met with suction and low flow alarms and the decision was made to explant the pump. the patient remained supported via central veno-arterial exracorporeal membrane oxygenation. no patient harm was reported. | [5 empty] | The customer's device was not returned for analysis. data logs were reviewed and confirm the reported issues. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the low pump flow/suction issue was most likely related to patient condition, based on data log review and provided clinical details. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6365 | 1220648-2025-46080 | 22945698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22945698&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704152 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support while undergoing assessment for therapies such as ventricular assist device and orthotopic heart transplant. the following day, an access site hematoma was noted. a unit of blood was given to the patient overnight. however, this was unrelated to the hematoma per additional information received. the hematoma was monitored. approximately four days later, there was bleeding at the incision, and heparin was suspended. a gastroenterologist was consulted, and the next day a colonoscopy was completed at bedside (results unnoted). hemoglobin test was completed, and results appear to have prompted two units of packed red blood cells and one unit of platelets administration. heparin remained suspended. the following day, there were no impella issues per the notes, and three days later, the patient was bridged to a left ventricular assist device. the impella 5.5 was explanted with a survived patient outcome. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, peri-procedural adverse event: access site bleeding, the root cause was not determined as the product was not returned and insufficient clinical details provided. h6: codes were updated accordingly based on the completed investigation. | Additional information was received and noted the results of the colonoscopy was a gastrointestinal bleeding, and the administered blood transfusion: two units of packed red blood cells and one unit of platelets was a result of this bleed. although impella is not the direct cause of the gastrointestinal bleeding, it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the outcome of the bleed for the patient. heparin as noted in previous reporting was suspended due to bleeding. note: h6 health effect (clinica and impact), medical device product, component and investigation (type, findings and conclusion) coding remains unchanged. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. h6 bleeding is being coded to represent bleeding and the stoppage of heparin.
6366 | 1220648-2025-31189 | 22946360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22946360&pc=OZD | 07/06/2023 | 09/02/2025 |  | 09/02/2025 |  |  | 07/06/2023 | 09/02/2025 |  |  | 12/31/2024 |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315801 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the pump¿s optical sensor failed, resulting in loss of placement signal. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6367 | 1220648-2025-46091 | 22948137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948137&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689476 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 100 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and developed a hematoma and bleeding of the access site requiring intervention and introducer removal difficulty. the patient presented to the emergency department with symptomatic chest pain and became pulseless. one round of cardiopulmonary resuscitation was performed before return of spontaneous circulation. electrocardiogram revealed st-segment elevation myocardial infarction. the decision was made to place an impella before proceeding with coronary intervention. post intervention, the standard protocol for removal of peel-away and placement of repositioning sheath was applied. repositioning sheath sutured x 2 and there was no oozing or hematoma at site. after weighing the options of protec versus venoarterial extracorporeal membrane oxygenation (va ecmo), a hematoma was now noted at the site. there was no oozing from the site present. manual pressure was held for approximately 20 minutes, and improvement was noted. the decision was made to cannula for va ecmo. angle of entry supported with gauze. sterile drape placed over patient for cannulation and repositioning sheath unable to be visualized. hemoglobin/hematocrit had a drop and when cannula was visualized it was no longer supported by angle matching. a hematoma was noted, and manual pressure was held under sterile drape. after visualizing the bleed and sheath under fluoroscopy it was presumed that once anticoagulation for va cannulation was started and cannula was presumed to have been unknowingly manipulated under sterile drape. the site began oozing at the lumen, and the hematoma worsened the extraction of sheath from the artery. the decision was made to control bleeding, assess the left ventricle distention, and reevaluate impella replacement once hemostasis was achieved. the impella was successfully removed and fluoroscopy confirmed sheath placement in the artery. the physician decided not to place another pump at said time. the patient was noted to have survived the event. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history review confirmed the pump passed all the post sterile inspection checks. regarding the major bleed/ hematoma, the cause of the injury was not determined since the product was not returned and insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the impella cp device is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h.10 for the medical device report on the impella peel-away introducer.
6368 | 1220648-2025-46092 | 22948140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948140&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689476 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 100 KG |  | White | Injury | No | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and developed a hematoma and bleeding of the access site requiring intervention and introducer removal difficulty. the patient presented to the emergency department with symptomatic chest pain and became pulseless. one round of cardiopulmonary resuscitation was performed before return of spontaneous circulation. electrocardiogram revealed st-segment elevation myocardial infarction. the decision was made to place an impella before proceeding with coronary intervention. post intervention, the standard protocol for removal of peel-away and placement of repositioning sheath was applied. repositioning sheath sutured x 2 and there was no oozing or hematoma at site. after weighing the options of protec versus venoarterial extracorporeal membrane oxygenation (va ecmo), a hematoma was now noted at the site. no oozing from the site was present. manual pressure was held for approximately 20 minutes, and improvement was noted. the decision made to cannula for va ecmo. angle of entry supported with gauze. sterile drape was placed over the patient for cannulation and repositioning sheath was unable to be visualized. hemoglobin/hematocrit had a drop and when cannula was visualized it was no longer supported by angle matching. a hematoma was noted, and manual pressure was held under sterile drape. after visualizing the bleed and sheath under fluoroscopy it was presumed that once anticoagulation for va cannulation was started and cannula was presumed to have been unknowingly manipulated under sterile drape. the site began oozing at the lumen, and the hematoma worsened the extraction of sheath from the artery. the decision was made to control bleeding, assess the left ventricle distention, and reevaluate impella replacement once hemostasis was achieved. the impella was successfully removed and fluoroscopy confirmed sheath placement in artery. the physician decided not to place another pump at said time. the patient was noted to have survived the event. | [5 empty] | This report is a duplicate of manufacturer report 1220648-2025-46091. any additional information received will be reported under manufacturer report 1220648-2025-46091. | A.5 is unknown. d.4 unique device identifier is unknown. h.4 manufacturer date is unknown. the impella peel-away introducer was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. the impella peel-away introducer is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h.10 for the medical device report on the impella cp. | After further review of the complaint with the investigation team, the complaint device was corrected to the impella cp. consequently, updates to section d respective fields, g4 pma/510k and h4 device manufacturer date were updated accordingly. section h1 type of reportable event was changed to serious injury. lastly, recoding of the complaint to better align with the reported event and associated device was also completed, and consequently, updates were made to section h6 health effect-clinical, health effect-impact, medical device problem code, and component code. note: h6 investigation type, findings and conclusions codes remain unchanged as initially reported. device status: the impella cp device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6369 | 1220648-2025-46078 | 22948212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948212&pc=OZD | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695112 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the pump was supporting for seven days when the pump stopped and could not be re-started. the stop occurred after there were high motor current and flow reversal. the pump was explanted to address the issue. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A.2 - a.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of the reported failure mode has been completed. the impella device was returned for investigation. pump stop: the returned pump was inspected and an abnormally large purge gap was observed indicative of a bearing wear. no data logs were returned. the pump ran from 27-july-25 at 4;30am to 3-aug-25 at 4:45am (over 8 days of support) which is beyond the intended design life of cpc9 per design trace matrix. the cause of the issue was determined to be due a bearing wear as device was used beyond design life device history review: the device passed all the post sterile inspection checks.
6370 | 1220648-2025-46097 | 22948258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948258&pc=OZD | 08/04/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702752 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced ischemia. the patient arrived at the emergency department and experienced a cardiac arrest from ventricular arrhythmia. the patient was defibrillated and return of spontaneous circulation was achieved. the patient was taken to the catheterization lab for further treatment and percutaneous coronary intervention to the left anterior descending artery was performed with balloon angioplasty and drug-eluting stenting. balloon angioplasty was also completed to areas of previous stents. left ventricular end-diastolic pressure was 17. the patient continued with hypotension, and the decision was made to place the patient on an impella cp. the impella was inserted and started at performance level p-7. flows reached 3.2 liters per minute. due to the vessel size and ischemia, the physician exchanged the peel-away introducer for the repositioning sheath. distal flow was confirmed with the repositioning sheath in place. the site was secured and dressing applied. the patient was transferred to the unit on impella mcs, and on arrival, a ¿large¿ hematoma was noted. no intervention was required per the notes. manual pressure was held for one hour, and the hematoma was resolved with bruising noted. the next day the patient was noted to be doing well, and the following day thereafter the patient was successfully weaned and the impella cp was explanted with a survived patient outcome. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hematoma, the root cause was determined to be use issue due to patient movement because the hematoma was noted on arrival to intensive care unit. cu. hemostasis was achieved by manual pressure. pertaining to the limb ischemia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6371 | 1220648-2025-31190 | 22948260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948260&pc=OZD | 07/26/2023 | 09/02/2025 |  | 09/02/2025 |  |  | 07/26/2023 | 09/02/2025 |  |  | 04/30/2025 |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337091 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. following implantation, the device¿s sensor became unreliable. despite this issue, the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6372 | 1220648-2025-46093 | 22948338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948338&pc=PYX | 08/03/2025 |  |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026686477 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp flex for mechanical circulatory support. the impella rp flex was placed in addition to the impella cp for bipella support. the impella rp flex had a drop in flows on day four and was explanted. the patient was stable post explant. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs confirmed low pump flow when pump started and remained to the rest of the case. log also showed motor current (mc) spiked during the case. based on clinical description and data review, the cause of low flow was most likely biomaterial ingestion since the data meets the biomaterial pattern cause. device history review: the device passed all post sterile inspection checks.
6373 | 1220648-2025-31191 | 22948342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948342&pc=OZD | 01/16/2025 | 09/02/2025 |  | 09/02/2025 |  |  | 01/16/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574252 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump implanted for circulatory support. during support, the pump triggered a "placement signal not reliable" alarm. according to the nurse, the alarm and absent placement signal issues occurred intermittently. in addition, the alarm audio was disabled. despite these issues, the pump remained operational throughout support with weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned. data logs were reviewed, and the placement signal unreliability was confirmed. throughout the case, snr and contrast are quite widespread and low, lamp is variable, while gain and light are stable. finally, previous pump performance on the console was checked and there were no abnormalities with the optical signal metrics. due to limited clinical details provided on the patient as well as no product returned, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6374 | 1220648-2025-46094 | 22948420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948420&pc=OZD | 08/03/2025 | 10/30/2025 |  | 09/02/2025 |  |  | 08/03/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026709138 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 67 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during an x-ray procedure, the patient sat up, leading to the pump being pulled into the aorta. the decision was made to remove the pump and insert a new one. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation was completed and no product was returned. the root cause of the positioning issue was determined to be an unintentional use error since the patient sat up during x-ray, leading to pump being pulled into the aorta. the device history review was completed and passed all post sterile inspection checks.
6375 | 1220648-2025-31192 | 22948594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948594&pc=OZD | 05/08/2025 | 10/24/2025 |  | 09/02/2025 |  |  | 05/08/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660793 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported use of an impella cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the aortic placement signal of the impella cp was lost following shockwave therapy. the signal returned shortly thereafter, but the readings remained inaccurate. as pump flow and positioning were unaffected, support was continued with the implanted device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product and data logs were returned. based on the clinical details that the placement signal went out briefly after using shockwave, data log patterns matched with a damaged sensor, and returned product had damage to the sensor face, the root cause of the optical signal issue was determined to most likely be a damaged sensor due to a shockwave interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date updated.
6376 | 1220648-2025-31193 | 22948702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948702&pc=OZD | 05/09/2025 | 10/24/2025 |  | 09/02/2025 |  |  | 05/09/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547759 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was reported to be unreliable. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | H4 device manufacture date updated. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and at time of reported issue, "impella position unknown" alarms were triggered. pump flows were low at time of placement signal spiking. a "preventing retrograde flow" alarm was also noted. at time of placement signal spiking, all 5s parameters were affected. the root cause of the optical signal issues could not be definitively determined as no hard failures of the optical sensor were observed during the case, no product was returned for evaluation, and limited clinical details were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6377 | 1220648-2025-31194 | 22948719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948719&pc=PYX | 05/11/2025 | 09/02/2025 |  | 09/02/2025 |  |  | 05/11/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625145 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during use, the optical sensor malfunctioned. displayed values were noted to be inconsistent, including negative pressures and erratic waveforms. despite this, the pump remained operational throughout the procedure and was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. neither the pump nor data logs were available for investigation. the root cause of the condition that the optical sensor was in to result in these readings could not be determined since neither data logs nor product were available for investigation and clinical details were limited. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6378 | 1220648-2025-31195 | 22948822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948822&pc=PYX | 05/11/2025 | 09/02/2025 |  | 09/02/2025 |  |  | 05/11/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025534051 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, a placement signal not reliable alarm appeared; however, the pump remained operational throughout the procedure. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned pump was inspected and there were no physical abnormalities on initial visual inspection. the pump passed the laser continuity test and 5s parameters were within normal range. the data logs observed is characteristic of an air gap. the returned pump was tested in the lab and the same issue and alarm reproduced when an air gap was intentionally made between the console and the patient cable plug. the root cause of the optical signal issue was most likely due to an air gap between the console and the patient cable plug as evidenced by the log pattern from the field and the reproduction of the issue on the returned device a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6379 | 1220648-2025-31196 | 22948838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948838&pc=OZD | 05/12/2025 | 10/24/2025 |  | 09/02/2025 |  |  | 05/12/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637968 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump generated four ¿impella in aorta¿ alarms, all of which resolved spontaneously and appeared to be erroneous. the waveforms remained unchanged and were not consistent with impella-in-aorta patterns. echocardiography later confirmed that the impella was correctly positioned and not in the aorta. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | G4 device manufacture date updated. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and at time of "impella in aorta" alarms showed no hard failures of the optical sensor. the root cause of the optical signal issue was most likely patient condition related per data log review and clinical details as the pump was confirmed to be in proper position and "impella in aorta" alarm were triggered correctly based on the based on the readings from the optical sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6380 | 1220648-2025-31197 | 22948856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22948856&pc=OZD | 05/11/2025 | 09/02/2025 |  | 09/02/2025 |  |  | 05/11/2025 | 09/02/2025 |  |  |  |  | 09/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617627 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump generated false "impella in aorta" alarms. placement monitoring was disabled. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and showed "impella in aorta" alarms triggered in first few days of case, with "impella position unknown" alarms also being triggered. the root cause of the optical signal issue was most likely patient condition related as clinical details noted that the pump was in proper position and patient was experiencing low pulsatility, and data logs do not show any hard failures of the optical sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6381 | 1220648-2025-45898 | 22921499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22921499&pc=OZD | 08/06/2025 |  |  | 08/29/2025 |  |  | 08/06/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687691 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient developed ischemia in their impella leg during impella cp support. it was noted that the patient was on a high dose of norepinephrine and vasopressin at the time of the noted condition. the patient was escalated to impella 5.5 support in response to the condition. the patient is stable post issue. | [5 empty] | A.4, a.5 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ as noted in the impella ifu: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6382 | 1220648-2025-46014 | 22921783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22921783&pc=OZD | 07/31/2025 |  |  | 08/29/2025 |  |  | 07/31/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704480 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 54 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing at the axillary site. two units of blood products were transfused. the patient continued on support and experienced episodes of laminar flow. an echocardiogram showed the inlet was deflected to the posterior wall. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device was not returned for investigation. the pump passed all the post sterile inspection checks.
6383 | 1220648-2025-45835 | 22922061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22922061&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026705924 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 102 KG |  |  | Injury | No | Yes | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following successful impella 5.5 implant, the staff attempted to reposition the pump and it was inadvertently pulled back into the ascending aorta. the pump was removed, the left ventricle was re-wired, and then the pump was readvanced into proper positioning. it was also noted that the placement signal was reading lower than expected. despite being calibrated correctly during priming, the placement signal reading did not accurately reflect the patient's blood pressure. all other waveforms were accurate and the surgeon was advised on how to monitor the blood pressure via other means. as support continued, a small hematoma was discovered at the surgical site. no noted intervention was required to manage the condition. impella support continues. | [5 empty] | No product was returned. the cause of optical signal issue was not determined due to inconclusive log evidence and insufficient clinical details. the cause of the hematoma was not determined since there was insufficient clinical information pertained to how bleeding event resolved and no product was returned. the cause of positioning issue was due to unintended use error based on clinical information. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6384 | 1220648-2025-46011 | 22922113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22922113&pc=OZD | 07/31/2025 |  |  | 08/29/2025 |  |  | 07/31/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708752 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 73 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp, the left femoral artery was lost after perclose. the procedure was completed by accessing the right femoral artery. vascular repair was needed post explant. the patient was successfully weaned from the pump and survived. | [5 empty] | The device was not returned for investigation. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6385 | 1220648-2025-45837 | 22922688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22922688&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025612331 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 endured multiple strokes. | [5 empty] | The root cause of the stroke was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6386 | 1220648-2025-45839 | 22922869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22922869&pc=OZD | 07/29/2025 |  |  | 08/29/2025 |  |  | 07/29/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025589322 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 80 KG |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered in conjunction with impella 5.5 support. roughly four days into support, purge fluid was seen leaking from the red plug of the setup. despite the leak persisting for over a week, the doctor opted to leave the pump in for ongoing support. the motor current, purge flow, and pressure remained stable throughout support. around the same time as the reported leak, the patient had an episode of non-sustained ventricular tachycardia. the condition resolved following two boluses of amio. six days later, the patient developed a hemothorax and chest tube bleed prompting an emergency washout. support continued until the patient was successfully weaned from the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was inadvertently omitted from the initial report that was submitted. d.3 added manufacturer fax as it was inadvertently omitted from the initial report that was submitted. e.1 added initial report phone number and zip code extension as they were inadvertently omitted from the initial report that was submitted. the zip code was reported in an incorrect field on the initial report that was submitted and was added to the zip code field. e.4 selection was added as it was inadvertently omitted from the initial report that was submitted. the investigation was completed. purge leak - pump: clinical details stated that there was purge fluid leaking from the red plug. motor current and purge flow/pressure were stable. the decision was made to leave in the pump despite leak. data log review showed stable purge pressure/flow throughout the case. the cause of the purge leak was not determined since product was not returned. tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. hemothorax/chest tube bleed: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. the device history review was completed and the pump passed all post-sterile inspection checks.
6387 | 1220648-2025-45838 | 22922894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22922894&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026677858 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male | 68 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 developed a gastrointestinal bleed. heparin was temporary halted to manage the condition and the patient was transfused two units of blood. support continued uninterrupted and heparin was restarted following condition resolution. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The device was not returned for investigation. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6388 | 1220648-2025-46038 | 22923126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22923126&pc=PYX | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  | 02/28/2026 |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479492 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Low Blood Pressure/ Hypotension | Hematuria | [10 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient began to show signs of hemolysis during impella rp flex support. upon reaching support level p8, the patient's urine became dark. in response to the developing condition, the staff began to reduce the pump flows. despite the adjustment, the patient continued to exhibit hypotension and ongoing signs of hemolysis. the pump flows were further decreased and the patient's blood pressure began to normalize. as the patient's hemodynamics were stabilized, the decision was made to remove the impella rp flex pump in light of the potential risks associated with the intermittent suction alarms being encountered. the patient remained supported via extracorporeal membrane oxygenation. | [5 empty] | The customer's device was returned for evaluation. biomaterial was observed in the pump housing and wrapped around the impeller. data logs from the customer's device were reviewed and confirm the reported issue of hemolysis. the cause of the suction events and hemolysis was biomaterial ingestion. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
6389 | 1220648-2025-46010 | 22923427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22923427&pc=OZD | 07/18/2025 |  |  | 08/29/2025 |  |  | 07/18/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686385 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female | 55 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced increased ectopy and a brain bleed. the patient experienced severe headaches and vomiting which led to a head computed tomography scan, showing a hemorrhage conversion. the patient went for an emergency craniotomy that night and had an external ventricular drain placed. suction alarms occurred overnight, and albumin was given. impella 5.5 support continued. | [5 empty] | A review of the device's data logs confirmed the pump suction issues. the cause of pump suction was patient condition related based on clinical details and log analysis. the root causes of the stroke and arrythmia were unable to be determined due to insufficient clinical details. | A.6 is unknown. the impella device was not received from the customer. the event logs were returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6390 | 1220648-2025-45992 | 22923750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22923750&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 showed signs of infection and had a subcutaneous hematoma. signs of device infection required removal of the temporary pacing lead, but the patient had a severe tendency to bleed and subcutaneous hematoma. due to the high procedural risk of cardiac resynchronization therapy (crt), the implantation of the micra av system was selected as a temporary bridge, with the goal of ultimately transitioning to crt. after the implantation of the micra av, the temporary pacemaker was removed without incident, and infection and bleeding tendencies improved. subsequently, crt was implanted as long-term treatment, and the impella 5.5 was successfully removed. additional information was received that the infection was not related to the impella. | [5 empty] | Section a was left blank as patient information was not provided. section e.1 initial reporter prefix/title, first and last name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. literature citation: yamamoto, r., kataoka, n., imamura, t., nakamura, m., & kinugawa, k. (2025). temporary use of leadless pacemaker as a bridge to cardiac resynchronisation therapy: a case report of fulminant myocarditis. european heart journal - case reports, 9(7). https://doi.org/10.1093/ehjcr/ytaf316. | The customer's device was not returned for investigation. a device history review could not be performed because the device was not returned.
6391 | 1220648-2025-45961 | 22923839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22923839&pc=PYX | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026668067 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. the impella rp flex was placed after the left sided support due to right ventricle that had no native contractility. the impella rp flex was supporting when on day five of the support the placement signal was found to not be reliable. the impella rp flex was explanted and patient remained on the 5.5 support. no harm or injury noted to patient due to the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The investigation was completed. optical sensor issue: the pump was returned for investigation. no evidence of physical damage was noted on the returned pump. the pump passed the laser continuity test; however, it failed the optical bench testing due to low snr. the pump also passed the dp sensor electrical testing. data log review showed that the snr was low (around 600) from the beginning of the case run, with a gain value of approximately 2.33. the pump was utilized for more than four days, and a drift pattern was observed on optical parameters. "placement signal not reliable" alarm was triggered on the day of explant, with the snr reading of 150. based on the drift pattern, the pump was sent for overnight pressure testing on the optical sensor and returned without any conclusive evidence of drift. further testing of the optical signal carrying line also revealed no abnormalities, and outflow case teardown confirmed a damaged sensor. the cause of the optical issue was a damaged sensor; however, the root cause of the sensor damage could not be determined. the pump passed all final qualification testing during manufacture. the device history review was completed and passed all post sterile inspection checks.
6392 | 1220648-2025-45840 | 22923940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22923940&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026691727 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 developed a gastrointestinal bleed. anticoagulation was put on hold to manage the condition. support continued until the patient was successfully weaned from the pump. | [5 empty] | No product was returned for investigation. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6393 | 1220648-2025-45950 | 22923961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22923961&pc=OZD | 07/29/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026658771 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella 5.5 implant, a hematoma developed at the insertion site. heparin administration was stopped due to bleeding concerns leading up to the event. the condition began to improve after two days of support. on the third day, the pump experienced suction and the patient endured a run of ventricular tachycardia. albumin was given and the suction resolved. six days later, heparin was once again discontinued due to bleeding from the arterial line. blood products were given in response to the bleed. bleeding concerns remained for another ten days as a new hematoma developed at the access site and bleeding developed around the jp drain. heparin was restarted upon removal of the jp drain and support continued uninterrupted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned for investigation. data logs were available for investigation. the cause of the hemolysis could not be definitively determined as limited clinical details were provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the pump suction events was likely due to patient condition as clinical details noted volume was given at time of suction events and no other abnormalities were observed on pump data logs. the pump passed all post-sterile inspection checks.
6394 | 1220648-2025-45916 | 22924153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22924153&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690845 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 74 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp developed hemolysis. the pump support level was dropped from p9 to p6 and an echo was completed to verify the pump positioning. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the hemolysis was not determined since no product was returned and insufficient clinical details were provided. the cause of the access site bleeding was most likely use-related since the act was above the recommended range. the pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6395 | 1220648-2025-46058 | 22924499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22924499&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while on impella 5.5 support the patient experienced ventricular tachycardia and ventricular fibrillation needing medication and defibrillation. additionally, the patient developed a hematoma at the right axillary site. two units of blood products were administered. the patient was taken to the or for a right axilla exploration with washout and drain placement. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was not returned for investigation. the pump passed all the post sterile inspection checks.
6396 | 1220648-2025-46021 | 22925842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22925842&pc=OZD | 08/01/2025 |  |  | 08/29/2025 |  |  | 08/01/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026639413 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced suction and low pressure alarms. correct pump position was confirmed using echo. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6397 | 1220648-2025-46057 | 22925954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22925954&pc=OZD | 07/19/2025 |  |  | 08/29/2025 |  |  | 07/19/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026672007 |  | 000000000001000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 66 KG |  |  | Death | Yes | Yes | Crack | Fluid/Blood Leak | Decrease in Pressure |  |  |  |  | Atrial Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient endured atrial fibrillation during impella 5.5 support. the patient initially converted back to a normal sinus rhythm, however they endured another episode the following night. an amio drip was started to counteract the condition. roughly twelve days later, a crack was discovered in the purge sidearm. fluid was seen leaking from the location and a low purge pressure alarm appeared. no intervention was taken and support continued for an additional two days before the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | A5 and a6 are unknown. the cause of the purge leak was most likely a damaged sidearm but the location on the sidearm is unknown since product was not returned for review. the root cause of the atrial fibrillation was unable to be determined due to insufficient clinical details.
6398 | 1220648-2025-45965 | 22926013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926013&pc=PYX | 07/30/2025 | 01/30/2026 |  | 08/29/2025 |  |  | 07/31/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an automated impella controller (aic) had optical signal loss while supporting impella 5.5 use. despite the error, the aic functioned until the patient was successfully weaned from the pump. | [5 empty] | Correction: d2a, d2b, h6 medical device problem code. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6399 | 1220648-2025-45974 | 22926044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926044&pc=OZD | 07/29/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701019 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 79 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient experienced ectopy during impella 5.5 support. the patient was initially implanted with an impella cp device (serial number (b)(6), however within the first twenty minutes of support a high purge pressure alarm occurred which could not be resolved. the first impella cp pump was replaced with a second impella cp device (serial number (b)(6). after 1 day of support, the medical team decided to escalate the patient to an impella 5.5 (serial number (b)(6). upon explanation of the second impella cp, it was reported that the patient experienced a vascular injury at the impella cp access site after a perclose device failed. a covered stent was placed to resolve the bleeding. it was reported during impella 5.5 support, on (b)(6) 2025, there was an impella positioning issue which caused the patient to experience ectopy. the medical team was advised to reposition the impella, which resolved the ectopy. it was additionally noted that the patient¿s care was withdrawn, and the patient expired. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella cp, with serial number (b)(6). pump 3: impella 5.5, with serial number (b)(6). this mdr specifically addresses the impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | There is not enough evidence to exclude impella as an associated factor in the patient's demise. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 required intervention was inadvertently omitted from the initial report submitted and was added. g.3 the aware date on the initial report submitted was incorrect and should of reported 30-jul-2025. h.6 health effect - clinical code and health effect - impact codes were added as they were not on the previously submitted initial report. medical safety review: there was ectopy with no documentation of hemodynamic compromise due to the ectopy. the pump was repositioned without event. the patient expired 5 days later. although given these details, there is not enough to dissociate the impella 5.5 from the demise outcome the impella device was not returned, and the investigation has been completed. the device history review was completed and the pump passed all the post sterile inspection checks. the root cause of the vascular damage was most likely a use issue since the right femoral perclose failed and vascular damage was observed. the root cause of the arrythmia was unable to be determined due to insufficient clinical details.
6400 | 1220648-2025-46022 | 22926064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926064&pc=OZD | 08/01/2025 |  |  | 08/29/2025 |  |  | 08/01/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp lost the optical signal. the supporting automated impella controller (aic) was replaced to resolved the issue. later, the patient was declared do not resuscitate and expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. logs from the day of the event are consistent with the complaint and indicate a controller error alarm: controller error. root cause: the root cause for the reported controller error alarm and the loss of placement signal was due to the optical bench fw communication protocol anomaly. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6401 | 1220648-2025-45983 | 22926167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926167&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026689479 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 84 KG | Non Hispanic |  | Injury | Yes | No | Malposition of Device | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and distal limb ischemia. the issue was managed with reperfusion sheath and repositioning of the pump. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for analysis. the root cause of the reported ischemia was unable to be determined due to insufficient clinical details. the cause of the hemolysis was improper positioning as evidenced by issue resolution after repositioning. | D6b, date of explant, has been added.
6402 | 1220648-2025-46047 | 22926168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926168&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026680869 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 92 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost pulses in both legs. the patient was escalated to an impella 5.5 for continued support. | [5 empty] | D.4 udi number was revised as it was entered incorrectly on the initial report that was submitted. e.1 initial report phone number was added as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6403 | 1220648-2025-46009 | 22926170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926170&pc=OZD | 08/01/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026696916 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the placement signal of an implanted impella cp pump became erroneous and then dashed out during patient support. support continued uninterrupted despite the noted issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6404 | 1220648-2025-46046 | 22926184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926184&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025608730 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp encountered suction. the patient was being transferred from fairfield medical center to riverside, and upon arrival, continuous suction alarms were observed with equal max and min flows on impella and flattened motor current. there was no visualization of the impella in the left ventricle after numerous echo views. it was determined that the impella migrated out of the lv sometime during transport and after speaking with the family, the decision was made to explant the pump and treat the patient medically. | [5 empty] | No product was returned. the cause of the positioning issue was determined to be an unintentional use error as the pump was came out during transport of the patient. the root cause of the pump suction is most likely due to user error as the clinical states that the pump was out of position, and repositioning resolved the issue and data logs support this conclusion. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6405 | 1220648-2025-46045 | 22926191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926191&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026697664 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient implanted with and impella cp developed a hematoma. as a result, fem stop and coats were drawn, and the patient received transfusions of fresh frozen plasma, packed red blood cells, cryoprecipitate and protamine. the pump was removed shortly afterwards and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided.
6406 | 1220648-2025-46048 | 22926204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926204&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690215 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during removal of the impella cp a cut down to the right femoral artery was done to repair the vessel. the patient was escalated to an impella 5.5 pump and survived. | [5 empty] | The cause of the perforation was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6407 | 1220648-2025-46040 | 22926206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926206&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026657566 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female | 48 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after an impella cp was placed, the patient arrived at icu with a groin hematoma. the right groin had impella and venous cordis. manual compression and fem stop was given to stabilize the site. two units of packed red blood cells were transfused. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6408 | 1220648-2025-45986 | 22926216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926216&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025530621 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 155 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that attempts to implant an impella cp were unsuccessful due to the patient's anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D.6a, d.6b and e.1 initial report phone number were all added as they were inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned. due to anatomy of patient, the pump could not be implanted. the patient had tortuous vessel. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery. the device history review was completed and the device passed all post sterile inspection checks.
6409 | 1220648-2025-45958 | 22926353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926353&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  | 09/30/2025 |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024401630 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 63 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had oozing from the groin site. heparin was withheld and a purse string suture was placed to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | D.3 manufacturer name should not of contained numbers behind it on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.6 revised health effect - impact code as it was previously entered incorrectly on the initial report that was submitted. the investigation has been completed. the root cause of the access site bleeding issue was most likely use issue since heparin held, purse string suture placed and pressure held to achieve hemostasis. the device history review was completed and the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6410 | 1220648-2025-45978 | 22926518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926518&pc=OZD | 07/30/2025 |  |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026696254 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the groin site that was remedied by transfusing one unit of packed red blood cells and adjusting the dressings. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The customer's device was not returned for review. based on the clinical description, the root cause of the bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the thrombocytopenia cannot be determined as insufficient clinical details were provided.
6411 | 1220648-2025-45853 | 22926638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926638&pc=OZD | 07/31/2025 |  |  | 08/29/2025 |  |  | 07/31/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690928 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 128 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient endured bleeding from multiple sites during impella cp support. the patient was noted to be bleeding actively orally, through chest tubes, and from the impella site. pressure was applied to the site and blood products were given to manage the condition. support continued uninterrupted following the intervention. | [5 empty] | B.7 information was added as it was inadvertently omitted from the initial report that was submitted. d.3 manufacturer name should not have contained numbers behind it on the initial report that was submitted. manufacturer fax number was added as it was inadvertently omitted from the initial report that was submitted. e.1 initial reporter phone number and zip code extension were added as they were inadvertently omitted from the initial report that was submitted. the zip code field was updated as it was previously entered incorrectly in the initial reporter postal code field on the initial report that was submitted. e.4 selection was added as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturer contact fax number was added as it was inadvertently omitted from the initial report that was submitted. no product was returned for investigation. the device history review was completed and the pump passed all the post sterile inspection checks. access site bleeding : the root cause of the access site bleeding issue was most likely patient condition related since the patient is bleeding through the chest tube site, impella site, foley site, and orally. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6412 | 1220648-2025-45928 | 22926694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926694&pc=OZD | 07/30/2025 | 08/29/2025 |  | 08/29/2025 |  |  | 07/30/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026678514 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was attempted to be implanted with an impella 5.5 for mechanical circulatory support. during the procedure the pump was unable to pass beyond the innominate. after multiple failed attempts, the implant was aborted and a new impella cp pump was placed via right axillary access. there was no noted patient injury associated with the delivery issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6413 | 1220648-2025-46041 | 22926706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926706&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026635380 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 76 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Intracranial Hemorrhage | Ischemia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a cerebral bleed. the bleed was contained and no intervention was performed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. the cause of the infarction was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6414 | 1220648-2025-46044 | 22926708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926708&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026689033 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Death | Yes | No | Malposition of Device | [6 empty] | Tachycardia | Hematuria | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was placed in the setting of acute myocardial infarction and cardiogenic shock. in icu the patient developed red tinged urine and lactate dehydrogenase was high. in response, the impella was repositioned. additionally, the patient had an episode of ventricular tachycardia and amiodarone conversion was administered. it was further reported that the family decided to withdraw support and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned. the cause of the hematuria was not determined due to insufficient clinical details. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details.
6415 | 1220648-2025-46042 | 22926738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926738&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  | 10/31/2025 |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024422863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 197 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed thrombocytopenia. the pump was later explanted as care was escalated to the 5.5 pump. the patient survived. | [5 empty] | B.2 an incorrect selection was made on the initial report that was submitted. a new selection was added. d.3 manufacturer name should not of contained numbers behind it on the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the investigation has been completed. thrombocytopenia : the cause of the thrombocytopenia cannot be determined as insufficient clinical details were provided. the device history review was completed and the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6416 | 1220648-2025-46049 | 22926801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926801&pc=OZD | 08/02/2025 |  |  | 08/29/2025 |  |  | 08/02/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026694102 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 112 KG | Non Hispanic | White | Death | Yes | No | False Alarm | Restricted Flow rate | Malposition of Device |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella cp device was positioned too deep and generating alarms. despite repositioning the pump would return to a deep position. an echocardiogram revealed the patient had aortic stenosis causing the pump to be malpositioned. albumin was given for volume. suction alarms occurred necessitating a decrease in p level to p7. during explantation of the pump it was suspected that the patient had a retroperitoneal bleed and manual compression was held. ultimately, the patient coded and expired. | [5 empty] | The device was not returned for investigation. event logs were reviewed. the complaint was confirmed. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this file represents the pump the related manufacturer device report number 1220648-2025-46069 represents the introducer.
6417 | 1220648-2025-45954 | 22926804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926804&pc=OZD | 08/03/2025 |  |  | 08/29/2025 |  |  | 08/03/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026695557 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 97 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient's access site began to ooze and eventually saturated the dressing. the dressing was changed and another stitch was applied at the site to manage the condition. three days into support, the pump migrated into the aorta. attempts to recross the valve were unsuccessful and the pump was subsequently replaced in response to the issue. | [5 empty] | No product was returned for investigation. the cause of access site bleeding was most likely use issue related since hemostasis was achieved by adding additional forward tension suture and adjusting the angle. the cause of the positioning issue was unable to be determined as insufficient clinical details were provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6418 | 1220648-2025-46027 | 22926812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926812&pc=PYX | 07/24/2025 |  |  | 08/29/2025 |  |  | 07/24/2025 | 08/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026666701 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Tachycardia | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Impeller flex was weaned down to p4 while not on anticoagulation for a 30 minute period subtherapeutic partial thromboplastin time (ptt) noted the entire pump run in 30¿s to 40¿s. recommended act 160-180 which typically corresponds ptt 59-76. multiple flow fluctuations with suction events over the weekend including hematuria and hemolysis. pump was removed on day 10 of rp flex run. excellent run to help wean from venoarterial extracorporeal membrane oxygenation (va-ecmo) and provide right ventricle (rv) support. the pump was discarded and the patient was reported as stable. | The complainant reported that following impella rp flex implant, a patient went into ventricular tachycardia during a sternal dressing change. the patient was defibrillated and an amiodarone bolus was administered in response to the condition. nine days into support, the pump began experiencing intermittent suction issues and reduced flow rates. blood products were given to counteract the suction. hematuria was observed and the patient's labs hemolyzed by the following day. the developing condition prompted early removal of the impella rp flex pump. |  |  |  |  | The customer's device was not returned for analysis. the device data logs do confirm the reported issues. the cause of the pump flow, suction issues, and hemolysis were not determined as no product was returned and limited information was provided. the root cause of the reported issue of tachycardia was unable to be determined due to insufficient clinical details. | B5 updated with additional information. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6419 | 1220648-2025-46017 | 22926822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926822&pc=OZD | 08/03/2025 |  |  | 08/29/2025 |  |  | 08/03/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026635437 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Other | [27 empty] | Initial submission | Followup | [8 empty] | The complainant reported that a patient suffered a stroke during impella 5.5 support. support continued uninterrupted until planned weaning and explantation the following day. | [5 empty] | B.2 an incorrect selection was made on the initial report that was submitted. a new selection was added. d.3 manufacturer name should not of contained numbers behind it on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. stroke: the root cause of the stroke was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6420 | 1220648-2025-46050 | 22926858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926858&pc=OZD | 08/01/2025 |  |  | 08/29/2025 |  |  | 08/03/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026701228 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | False Alarm | Restricted Flow rate | Increase in Pressure | Patient-Device Incompatibility | Device Sensing Problem |  |  | Dyspnea | Myocardial Infarction | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp coded and required two rounds of cardiopulmonary resuscitation and defibrillation. the patient achieved return of spontaneous circulation and had a stent placed to the right coronary artery. the pump continued to have position in aorta alarms with repositioning, high purge pressure, and purge pressure blocked alarms. there were also ongoing suction alarms, ultimately leading to replacement of the pump. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6421 | 1220648-2025-45819 | 22926890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926890&pc=OZD | 07/22/2025 |  |  | 08/29/2025 |  |  | 08/04/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during impella 5.5 support, the patient became dizzy and alerted medical staff. the patient then went into cardiac arrest and required cardiopulmonary resuscitation to achieve a return of spontaneous circulation. it was noted that the automated impella controller did not alarm for suction alarms which occurred just prior to the patient¿s cardiac arrest. two days after this event, the patient expired. it was noted this was likely due to septic shock. there was no information available on the source of sepsis. this complaint involves one impella 5.5 pump (serial number (b)(6)) and one automated impella controller (serial number (b)(6)) this mdr specifically addresses automated impella controller, which is the subject of this report. a separate mdr will be submitted for the impella 5.5 associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6422 | 1220648-2025-45869 | 22926899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22926899&pc=OZD | 10/20/2023 |  |  | 08/29/2025 |  |  | 08/04/2025 | 08/29/2025 |  |  | 08/31/2025 |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024394054 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced hemolysis and hematuria. later, the patient was successfully weaned from the pump and survived. | [5 empty] | D.3 added manufacturer fax number as it was inadvertently omitted from the initial report that was submitted. e.4 added selection as it was inadvertently omitted from the initial report that was submitted. g.1 added manufacturer contact fax number as it was inadvertently omitted from the initial report that was submitted. the device was not returned and event logs were not returned. hematuria: the cause of the injury was not determined due to insufficient clinical details. hemolysis: reported due to increased value of free hemoglobin. the cause of the issue was not determined as no product or logs were returned and there was insufficient clinical information provided device history review: the pump passed all post sterile inspection checks | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6423 | 1220648-2025-45754 | 22927018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22927018&pc=OZD | 07/22/2025 |  |  | 08/29/2025 |  |  | 08/04/2025 | 08/29/2025 |  |  |  |  | 08/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657593 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female | 65 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiac Arrest | Sepsis | [11 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during impella 5.5 support, the patient became dizzy and alerted medical staff. the patient then went into cardiac arrest and required cardiopulmonary resuscitation to achieve a return of spontaneous circulation. it was noted that the automated impella controller did not alarm for suction alarms which occurred just prior to the patient¿s cardiac arrest. two days after this event, the patient expired. it was noted this was likely due to septic shock. there was no information available on the source of sepsis. this complaint involves one impella 5.5 pump (serial number (b)(6)) and one automated impella controller (serial number (b)(6)). this mdr specifically addresses impella 5.5, which is the subject of this report. a separate mdr will be submitted for the automated impella controller associated with this complaint. | [5 empty] | Based on the information available, it is unlikely that the sepsis was related to the use of the impella device; however, a device-related contribution cannot be definitively ruled out. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was not returned. the pump passed all post sterile inspection checks. the cause of suction issue was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the injury was unable to be determined due to insufficient clinical details.
6424 | 1220648-2025-45972 | 22910208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22910208&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported the complainant had a 5.5 pump support ongoing for the patient admitted in with cardiogenic shock/cardiomyopathy. the impella 5.5 was placed as a bridge to transplant/left ventricular assist device (lvad). approximately one after start of the impella support, the team noticed that the automated impella controller (aic) alarm for loss of placement signal. the team did not need to intervene, as the alarm was self-resolving. support continued uninterrupted. there was no report or harm to the patient as a result of this event. | [5 empty] | This console was tested after arriving at facilty. the root cause for the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. rtlog confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error there was no evidence in the logs to confirm the product experience code ¿ ¿optical sensor issue¿. the controller error was misreported as ¿optical sensor issue. | The device remains in use, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6425 | 1220648-2025-45927 | 22911002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22911002&pc=OZD | 07/29/2025 |  |  | 08/28/2025 |  |  | 07/29/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681869 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 78 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was escalated from an impella cp to an impella 5.5 for acute heart failure management while the patient awaited a heart transplant. during support a nurse reported the pump experienced high purge pressures with low purge flows, and the nurse presented the issue to patient care team. the physician decided to run two bags of tissue plasminogen activator per facility protocol, and the purge pressure and flows resolved. the patient support continued until the patient was successfully bridged to a heart transplant. | [5 empty] | The product was not returned for investigation. based on data log review and the clinical details provided, the cause of the high purge pressure was biomaterial. the device passed all post-sterile inspection checks. | Product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b, date of explant, has been updated.
6426 | 1220648-2025-46005 | 22911510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22911510&pc=OZD | 07/29/2025 |  |  | 08/28/2025 |  |  | 07/29/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025602035 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and two days after start of the support, multiple blood products were given due concern for retroperitoneal bleed was noted. following, the patient ended with a suspected compartment syndrome near the right thigh around the impella entry site. patient remained on extracorporeal membrane oxygenation (ecmo) when the impella cp device was explanted after a total of three days of mcs for concerns of retroperitoneal bleed and the aforementioned compartment syndrome. patient¿s neurological status and survival outcome at explant are unknown. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, retroperitoneal hemorrhage/ bleeding, the bleeding was reported due to the concern for retroperitoneal bleed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. pertaining to the limb ischemia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use (impella cp smartassist system) for the related event are as follows: section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6427 | 1220648-2025-46034 | 22911602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22911602&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025514305 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump experienced persistent placement signal low alarms during patient support. despite a dampened sensor, all other functions of the pump were noted to be working as intended. support continued uninterrupted. no patient harm reported due to the issue. | [5 empty] | The investigation was completed. no logs were returned. the returned product passed laser continuity and all of the sensor 5s parameters were within specification. the pump was put in the uson and it was able to hold pressure but was reading slightly higher than the set pressure. there were no damages or abnormalities found with the device. due to a lack of data logs returned and sufficient clinical information, the root cause of the optical sensor issue cannot be determined. the device history review was completed and passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6428 | 1220648-2025-45834 | 22911856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22911856&pc=OZD | 07/29/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674627 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that roughly nine days following impella cp implant, a low purge pressure alarm appeared unexpectedly. no leaks were visible in the assembly at the time of the alarm. both the automated impella controller and purge cassette were swapped to troubleshoot the alarm, but the issue persisted. when additional onsite measures failed to resolve the issue, the decision was made to monitor situation. the following day, both the purge flow rate and pressure stabilized without further action being required. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The pump was returned for investigation. no visual damage was observed on the returned product. the pump was then successfully primed and kept for 10 minutes, during which steady purge pressure and flow within specifications were observed. the data log shows the 82 hour gradual rise in purge flow followed by the purge pressure low alarm on (b)(6) 2025. as per the given clinical, purge cassette change did not resolve the low purge pressure alarm. and there was no indication of the product damage or the purge leak during reported issue of low purge pressure. later, purge parameters were seen returned to normal range without any troubleshoot. the cause of the low purge pressure was not determined without returned product investigation. no defect was found on the returned product. the cause of the purge pressure low issue was not determined due to insufficient clinical information provided. the pump passed all post sterile inspection checks. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6429 | 1220648-2025-46037 | 22912434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22912434&pc=OZD | 07/29/2025 |  |  | 08/28/2025 |  |  | 07/29/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1597534 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) failed to recognize purge fluid in the line. purge disc not detected and purge fluid not detected alarms occurred. a new aic was used to support the patient. there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6430 | 1220648-2025-45959 | 22914120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22914120&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026630620 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced oozing from the insertion site. heparin was withheld. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the pump passed all the post sterile inspection checks. | A4 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6431 | 1220648-2025-45875 | 22914339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22914339&pc=OZD | 08/13/2025 |  |  | 08/28/2025 |  |  | 08/13/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026679415 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella device experienced a pump shutdown. during therapy, the "pump shutdown current consumption rise" alarm occurred. it could not be reset. the impella device was explanted. | [5 empty] | D.3 manufacturer name should not have contained numbers behind it on the initial report that was submitted. e.1 added initial reporter first and last name as they were inadvertently omitted from the initial report that was submitted. e.3 revised as it was entered incorrectly on the initial report that was submitted. e.4 selection was added as it was inadvertently omitted from the initial report that was submitted. g.1 added manufacturer contact fax number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. pump stop: clinical details report a pump stop with increased motor current on 13/aug/2025. the patient was started on impella cp support on 25/jul/2025. the pump was explanted as the pump could not be restarted. data logs were reviewed, and pump metrics were generally stable through 12/aug/2025. on this day, purge pressure and flow began to show variability when the pump was increased to p-8. then, there was the first motor current (mc) spike to 1400 ma with a ms deviation. the pump did not stop, but the morning of 13/aug/2025, the first high mc pump stop alarm triggered with a mc spike and ms deviation. the pump was unable to be restarted. these patterns noted in data logs could be due to several potential causes. product was not returned for investigation. since the product was not returned and data logs were inconclusive, the cause of the pump stop could not be determined. the device history review was competed and the pump passed all post-sterile inspection checks. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6432 | 1220648-2025-45968 | 22916275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22916275&pc=OZD | 07/15/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026623551 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 77 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hemolysis. the pump was repositioning. impella support continues. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the reported bleeding was not determined due to insufficient clinical details and no product returned.
6433 | 1220648-2025-45980 | 22916288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22916288&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026690400 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant was placing an impella cp using a 14fr x13cm introducer sheath via the femoral. the introducer appeared to be damaged after the dilator was removed. the orange locking mechanism came off the sheath and blood loss occurred. the team removed the sheath and tightened down the perclose. hemostasis was achieved by the perclose and transfusion of packed red blood cells. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6434 | 1220648-2025-45987 | 22918934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22918934&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026682075 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  |  | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Arrhythmia | Stroke/CVA | Ischemia | Asystole | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced asystole, suction alarms, and mottled extremities. tomography showed a small stroke with no neuro response. later, the patient was successfully weaned from the impella and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product or data logs were returned for investigation. the root cause of the stroke/ischemia/arrhythmia was unable to be determined due to insufficient clinical details. the cause of the asystole cannot be determined as insufficient clinical details were provided. the cause of the pump suction cannot be determined due to no product or data logs available the pump passed all post sterile inspection checks.
6435 | 1220648-2025-45990 | 22918937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22918937&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced complete atrioventricular block requiring temporary right ventricular apical pacing. additionally, the patient's blood pressure continued to fall below 60 mmhg, so venaortica-ecmo (ecpella) circulation was established to provide circulatory support. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the phr could not be completed because the device was not returned.
6436 | 1220648-2025-45989 | 22918938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22918938&pc=OZD | 07/04/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026674515 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a thrombus in the left ventricle. the impella was explanted. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the thrombosis was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
6437 | 1220648-2025-45981 | 22918939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22918939&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026677126 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 73 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced limb ischemia. a fem fem bypass was done however, the patient coded and expired. | [5 empty] | The root cause of the ischemia was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6438 | 1220648-2025-45988 | 22918940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22918940&pc=OZD | 07/30/2025 |  |  | 08/28/2025 |  |  | 07/30/2025 | 08/28/2025 |  |  |  |  | 08/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026671969 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after implantation of an impella cp the pump was accidentally pulled out of the left ventricle to the descending aorta. after several unsuccessful attempts the impella could not advance into the lv. the physician exchanged the pump for another impella cp. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned. unable to place or position (positioning issue): the first one was pulled out of the left ventricle (lv) to the descending aorta accidentally and could not advance into the lv by multiples attempts. the cause of the positioning issue was determined to be an unintentional use error since the pump was pulled out of the lv to the descending aorta accidentally. the device history review was completed and the device passed all post sterile inspection checks.
6439 | 1220648-2025-45945 | 22901459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22901459&pc=OZD | 07/29/2025 |  |  | 08/27/2025 |  |  | 07/29/2025 | 08/27/2025 |  |  |  |  | 08/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1857010 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 102 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the 5.5 was placed as care escalation from a prior placed impella cp pump. after 5 days of support, the automated impella controller (aic) alarmed for 20 seconds for optical signal loss. the aic was replaced with back up aic without any issue. no patient harm was reported due to the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. | Investigation summary: data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the crc error. see the attached image rt_log.png. device analysis: the console was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the optical bench was evaluated for 5s parameters and the analog signals were checked for distortion, but no issues were found. additionally, the aic was tested with know good pump and purge cassette and no issues were found. root cause: the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" and optical signal issue was due to a firmware issue (optical bench fw anomaly).
6440 | 1220648-2025-45976 | 22902935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22902935&pc=OZD | 07/29/2025 |  |  | 08/27/2025 |  |  | 07/29/2025 | 08/27/2025 |  |  |  |  | 08/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607811 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to a heart transplant. it was reported that during support the patient experienced ventricular tachycardia (vt) and the impella was repositioned to resolve the complication. the following day on (b)(6) 2025, the patient experienced additional episodes of vt, however these episodes were attributed to the patient coughing and self-resolved. on (b)(6) 2025, there was a report of additional vt episodes with impella suction alarms. a stat echocardiogram was ordered and the impella position was evaluated. the impella appeared to be well positioned, and no adjustments were made. the patient is currently stable waiting for a suitable donor for transplant on remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The customer's device was not returned. the root cause of the positioning issue was unable to be determined as insufficient clinical details were provided. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details.
6441 | 1220648-2025-46012 | 22904359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22904359&pc=OZD | 07/29/2025 |  |  | 08/27/2025 |  |  | 07/29/2025 | 08/27/2025 |  |  |  |  | 08/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696469 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 110 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost distal pulses. the impella cp was explanted and replaced with an impella 5.5. distal pulses were restored. the complainant reported that a patient underwent single vessel coronary artery bypass graft surgery and right coronary artery feeding the left anterior descending artery. the plan was to rest and recover the heart as the patient tolerated surgery well. upon exiting the operating room the patient had cardiac arrest. cardiopulmonary resuscitation was performed and the patient was defibrillated multiple times. the decision was made to place an impella cp. due to significant calcific disease there was no flow found down the impella access leg. the patient was taken to the catheterization lab for placement of a pulmonary artery catheter and external femoral to femoral bypass due to the limb ischemia. on day two of impella cp support distal pulses were not detected and a decision was made to escalate the patient to an impella 5.5. the impella cp was successfully removed surgically and distal pulses returned. upon impella cp explant, the patient began decompensating. the patient continued to decompensate after the impella 5.5 was implanted and the patient was placed on support. the patient was then additionally cannulated for veno-arterial extracorporeal membrane oxygenation. it was further reported that the patient expired while on impella 5.5 support, 8 days after the impella cp was explanted. | [5 empty] | The relationship between the impella and cause of death is unknown. based on the information available the impella cp cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The customer's product was not returned for investigation. the device passed all post sterile inspection checks.
6442 | 1220648-2025-45944 | 22905631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22905631&pc=OZD | 07/27/2025 | 12/24/2025 |  | 08/27/2025 |  |  | 07/27/2025 | 08/27/2025 |  |  |  |  | 08/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583046 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 127 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that the patient implanted with an impella cp required repositioning the following day as it was deep at 5.5 cm after "aggressive" diuresing the patient the prior day. the impella cp device was pulled back 1-2 cm with bedside echocardiogram. notedly, there were no signs of hemolysis. the next day performance level was p-7 with flows at 3.4 l/min. the impella cp device was deep again at 6.5cm on transthoracic echocardiogram. the device was pulled back 2cm to 4.2cm. diastolic suction started, and performance level was reduced to p-6 with flows at 2.9 l/min. later this day, the patient was lying flat, and the pump transiently pulled into the aorta. the issue resolved itself when the patient sat up. the impella was repositioned again. overnight, urine was increasingly bloody; the patient developed "severe" hemolysis. consequently, the patient was weaned to p-3 with goal of impella removal. discussion for further mechanical circulatory support was made and noted longer required due to improving function on transthoracic echocardiogram. the device was explanted with a survival outcome. | [5 empty] | The customer's pump was not returned for investigation. the root cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position during patient movement. the data log shows a small motor current spike followed by a 3-minute gradual rise in purge pressure with a damped motor current waveform. the purge pressure trend was seen resolving with a gradual decreasing trend after the first instance of a high purge pressure alarm. motor current was seen gradually fluctuating later case run. the optical signal issue occurred 5 hours after the high purge pressure event. the cause of the high purge pressure issue was most likely biomaterial ingestion, based on data log review. the cause of the hemolysis issue was most likely biomaterial ingestion, based on data log review. data logs show fiber damage characteristics on all 5s optical signal parameters. the cause of the optical signal issue was most likely fiber damage, based on data log review. however, the exact location of the damage was not determined without the returned product investigation. | The pump was returned for investigation. the returned pump had a catheter kink near the 110 cm marking and failed the laser continuity test at the kinked catheter. the cannula was inspected and biomaterial was found on the impeller and purge gap. high purge pressure was able to be reproduced in the lab. the catheter was cut near the motor housing while purging, and purge fluid was able to come out from the catheter-side purge lumen, which proves that the blockage was in the motor-side purge line. no other physical abnormalities were reported on the returned pump. the root cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position during patient movement. the cause of the high purge pressure issue and hemolysis was biomaterial ingestion, based on data log review and returned product investigation. the cause of the optical signal issue was fiber damage, based on data log review and returned product investigation. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated.
6443 | 1220648-2025-46026 | 22909823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22909823&pc=OZD | 07/28/2025 |  |  | 08/27/2025 |  |  | 07/28/2025 | 08/27/2025 |  |  |  |  | 08/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660187 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 110 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | False Alarm | Difficult or Delayed Positioning | Restricted Flow rate | Inaccurate Flow Rate | Patient Device Interaction Problem |  |  | Fever | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Perforation of Vessels | Vascular Dissection | Embolism/Embolus | Thrombosis/Thrombus | Aortic dissection |  |  |  | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced complications requiring intervention. the patient developed acute onset chest pain and was syncopal at work. on emergency medical services arrival, the patient was diaphoretic, confused with low blood pressure. the patient arrived at the hospital and anterior/inferior st-segment elevation myocardial infarction was noted. the patient was emergently transferred to the catheterization lab, unable to lay flat, and was intubated. the impella cp was placed prior to the intervention. of note, the patient was cannulated for extracorporeal membrane oxygenation (ecmo) as well. the patient was sent to the unit on support. later that day, impella started alarming for impella in aorta, motor current flat, dropped to performance level p-2 and a transthoracic echocardiogram was enroute. the device was advanced and locked at 95cm and increased to support level p-4. the impella cp was noted to be in optimal positioning. following, the left ventricle was underfilled and therefore, the patient was given albumin 250 mls. had a rainbow of labs drawn, and lactate was down to 3.5. the patient was dyssynchronous with the ventilator. the purge solution was changed to a bicarbonate additive and anticoagulation was withheld until the labs resulted. overnight the patient had three episodes of back-to-back suction alarms and 1000 mls. of albumin in total. the impella cp pump required repositioning and was now measuring at 5cm depth towards apex of left ventricle and locked at 98cm. the patient¿s lactate had cleared to 2.1 last night and, however, but was up to 4.6 this morning. patient with transient pulsatility and treated with bivalirudin. now day three of ecmo and impella support, the patient was scheduled for escalation to impella 5.5 and possible ecmo decannulation. ¿gross¿ hemolysis was noted with lactate dehydrogenase greater than 2,000. overnight, the patient received one unit of packed red blood cells (prbcs) and albumin 250mls for parameters, suction alarm. pulsatility in all waveforms were with slightly higher pulse pressures. notedly, bivalirudin drip was off since the morning for the operating room. the patient was noted to be neurologically intact. now day three of support, and it was noted the patient had a change in neuro status yesterday that postponed the impella 5.5 placement. head computed tomography scan showed no cerebral vascular accident, and eeg was normal. the patient was able to wake up and follow commands this am with sedation weaned. the patient did become febrile yesterday, prompting a bronchoscopy, blood cultures, and intravenous antibiotics. bronchoscopy revealed a foreign body in lung suggesting aspiration. however, now after these three days of support on the impella cp, the arrived patient to the cardiovascular operating room intubated on peripheral venoarterial ecmo. the patient prepped and draped for left axillary impella 5.5 placement to avoid radial arterial line complications. at this time, transesophageal echocardiogram (tee) was performed showing notable organized thrombus formation in the distal root on the impella cp. the decision was made to continue with cp removal which was completed and the impella 5.5 implanted. the impella cp of note was removed via a left femoral cutdown and arteriotomy to remove impella cp with arterial thrombectomy. clot was evacuated and removed. distal flow was lost to left lower extremity. the arteriotomy was re-opened. under tee, possible dissection was noted to descending aorta. possible aortic intramural hematoma noted - not to be repaired at this time. fibrin casting surrounding the impella cp caused occlusion to distal lower extremity. an iliac stent was performed and a computed tomography planned for further evaluation. the patient¿s estimated blood loss was approximately 2.2 liters, receiving 5 units of prbcs, 2 units of platelets, 3 units of frozen fresh plasma, 500 mls. of albumin, and 29,000 units of heparin. the patient was eventually transferred to the unit on impella 5.5 support which continued for 18 days before the patient¿s demise. | [5 empty] | Additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related manufacturer report numbers: this is one of three medical device reports associated with the patient's implant experience. refer to the related manufacturer report numbers as noted in section h10. | The complaint coding was reviewed and updates to the health effect: clinical and device problem codes were made to better align to the voice of customer reported event. therefore, section h6 health effect: clinical and impact coding and medical device problem coding were corrected, repopulated accordingly. medical safety review: the patient was initially supported with an impella cp for three days and experienced complications and is a serious injury. the patient was escalated to an impella 5.5 and supported for approximately 18 days and experienced complications including but not limited to significant bleeding requiring transfusion 5 units. there is not enough evidence to exclude the impella 5.5 as associated factor to the outcome. however, it should be noted the patient presented in critical condition: diaphoretic, anterior/inferior st-segment elevation myocardial infarction with a 100% left main occlusion. the patient was unable to fully recover from the cardiac event. regarding the automated impella controller (aic), patient expired for reasons unrelated to the aic. there was no interruption in support as the aic does not appear to have been exchanged, or issue impacted the therapy. aic device remained in service after the controller failure and defective purge cassette encounter. the medical safety officer reviewed the event and noted the same in that impella 5.5 device cannot be dissociated from the demise outcome. and the impella cp and aic is a serious injury and product malfunction type of reportable event respectively. related medical device reports: this impella cp device report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device reports on the impella 5.5 which is associated with the demise outcome and automated impella controller. | The product with event logs were received and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. ischemia/thrombosis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. fever/aortic dissection: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. difficult to place/position: as per clinical review, impella in aorta alarms were observed. transthoracic echocardiogram (tte) confirmed pump had been appeared to be at 86.5 cm which was locked at 89 cm initially. it was advanced and locked again to 95 cm. later on (b)(6) 2025 it was noted the pump required repositioning at depth close to apex of left ventricle (lv) and locked at 98 cm. physician was aware of the impella cp depth and tte confirmed that patient has small lv cavity, several global hypokinesis. impella cp pump was visually inspected and no kinks, damages confirmed. no positioning issue with device defect was found. the cause of the positioning issue was due to patient condition since the patient was having small lv cavity. mechanical interaction with tissue: as per clinical, there was notable organized thrombus formation in the distal root on the impella cp. the fibrin cast in surrounding the impella cp caused occlusion to the distal lower extremity. arterial thrombectomy was performed, clot was evacuated and removed. the cause of vascular dissection was not determined as no defect with product and thrombectomy was performed to evacuate clot from femoral artery. mechanical interaction with blood: as per clinical, the patient experienced another suction alarm and received 1 unit of packed red blood cells (prbcs) and albumin 250 mls. lactate dehydrogenase (ldh) has bumped to 2000. gross hemolysis noted with ldh greater than 2,000. the patient was scheduled for axillary 5.5 placement (b)(6) 2025. it was noted patient had global hypokinesis. the return pump did not show any physical damages or abnormalities potentially causing hemolysis. hemolysis test was done and pump successfully passed test with mih values of 9.91. few occurrences of suction #14 & #73 alarms seen during case since start. ps was trending down/low during suction. low flows were seen on (b)(6) during suction and later p-4 on (b)(6). but ps became better post suction. no motor current spike, major placement signal trend. placement signal was slightly spiking/noisy but not trending low/down later. the cause of hemolysis/mechanical interaction with blood was likely patient condition related as no defect found with return product testing and patient having hypokinesis based on clinical information. pump suction: as per (b)(6) 2025 it was reported that overnight, the patient had 3 episodes of back-to-back suction alarms. the patient received 1000 mls. of albumin in total. on (b)(6) 2025 the patient experienced another suction alarm and received 1 unit of packed red blood cells (prbcs) and albumin 250 mls. data logs were review and 22 occurrences of suction #14 alarm was found on (b)(6) 2025 and 20 occurrences of suction #73 alarm were found on (b)(6) & (b)(6) 2025. few impella position in aorta alarms were sound on (b)(6) which resolved with repositioning and there were some occurrences which were found on (b)(6). several occurrences (880) of impella position unknown alarms #33 were also seen but were after (b)(6). patient was also on ecmo device support along with impella. ecmo decannulation was planned during 5.5 escalation. the return product did not show any biomaterial ingestion at outflow or wrapped at impeller. biomaterial was observed at the motor shaft of the cp pump. no broken blades or other issues found. low pump flow or suction did not reproduce when ran at p-9 in bucket of water. the cause of pump suction is likely patient related. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device reports on the impella 5.5 which is associated with the demise outcome and automated impella controller.
6444 | 1220648-2025-46025 | 22909834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22909834&pc=OZD | 07/31/2025 |  |  | 08/27/2025 |  |  | 07/31/2025 | 08/27/2025 |  |  |  |  | 08/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721787 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 110 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (mcs) and would experience complications. impella cp mcs was provided for a total of three days before being escalated to the impella 5.5 device. the next day after starting support on the impella 5.5 mcs, an air in the purge alarm was triggered and required a cassette change. three cassettes were used to resolve the issue. an automated impella controller error (aic) occurred as well and was resolved prompting the team to place a backup aic outside the patient¿s room. support continued with the aic uninterrupted for an additional 18 days. eventually, the patient was placed in palliative care and comfort care initiated. the patient would ultimately expire. | [5 empty] | The patient was initially supported with an impella cp for three days and experienced complications and is a serious injury. the patient was escalated to an impella 5.5 and supported for approximately 18 days and experienced complications including but not limited to significant bleeding requiring transfusion 5 units. there is not enough evidence to exclude the impella 5.5 as associated factor to the outcome. however, it should be noted the patient presented in critical condition: diaphoretic, anterior/inferior st-segment elevation myocardial infarction with a 100% left main occlusion. the patient was unable to fully recover from the cardiac event. regarding the automated impella controller (aic), patient expired for reasons unrelated to the aic. there was no interruption in support as the aic does not appear to have been exchanged, or issue impacted the therapy. aic device remained in service after the controller failure and defective purge cassette encounter. the medical safety officer reviewed the event and noted the same in that impella 5.5 device cannot be dissociated from the demise outcome. and the impella cp and aic is a serious injury and product malfunction type of reportable event respectively. related medical device reports: this aic device report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device reports on the impella cp and impella 5.5, which is associated with the demise outcome. | Based on the details at hand, there is not enough information to dissociate the impella 5.5 device from the demise outcome. this report on the automated impella controller is one of three medical device reports associated with the patient's implant experience. refer to the related manufacturer report numbers as noted in section h10 which includes the report number for the impella 5.5 device). device status: the automated impella controller device was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed. related manufacturer report numbers: this is one of three medical device reports associated with the patient's implant experience. refer to the related manufacturer report numbers as noted in section h10.
6445 | 1220648-2025-45925 | 22893301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22893301&pc=OZD | 07/24/2025 |  |  | 08/26/2025 |  |  | 07/24/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679379 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Failure to Advance | [6 empty] | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | Life Threatening | Other | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was prepped in the operating room (or) but was unable to advance past brachiocephalic trunk due to calcified stenosis. the patient continued to decompensate and an impella cp was prepped and urgently implanted instead of the impella 5.5. the impella 5.5 was saved in a sterile environment for possible escalation and the patient was sent to the cardiac catheterization laboratory (ccl) for a possible percutaneous coronary intervention (pci). in the ccl, the patient was found to have triple vessel disease, and the medical team was unable to perform the pci. the patient stabilized with the impella cp but remained in cardiogenic shock and the patient¿s lactate remained elevated. the medical team decided to escalate the patient to the impella 5.5 for increased support and the patient was taken back to the or for impella 5.5 insertion through the axillary artery. the impella 5.5 was successfully placed and support initiated. the patient was taken back to the icu for recovery and monitoring. the patient was noted to have been trending positively with lactate results down trending, however the patient acutely decompensated with no pulsatility. the patient continued to decompensate with profound hypotension and eventually went into ventricular fibrillation arrest and advanced cardiac life support was initiated as a result. the medical team was not certain of the cause of the acute decompensation, but suspected the cause was due to quickly declining right ventricular failure. | [5 empty] | B.2 added selections as they were inadvertently omitted from the initial report that was submitted. life threatening was selected incorrectly on the initial report that was submitted. the investigation was completed. the root cause of the delivery issue was determined to be patient condition as the patient had calcified stenosis preventing successful delivery of the impella. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the root cause of the hypotension cannot be determined as insufficient clinical details were provided. the device history review was completed and the device passed all post sterile inspection checks. | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6446 | 1220648-2025-45926 | 22893655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22893655&pc=OZD | 07/29/2025 |  |  | 08/26/2025 |  |  | 07/29/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607847 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiac Arrest | Pericardial Effusion | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during a percutaneous cardiac intervention (pci). it was reported that during impella cp support, the pump experienced suction events that were resolved by giving the patient intravenous fluids. after the pci, the patient rapidly decompensated, requiring increased vasopressors, the patient was noted to be hypoxic on arterial blood gas results, and there were recurrent suction alarms on the impella. the patient proceeded to have a pulseless electrical activity arrest. an echocardiogram was performed and revealed massive bilateral pleural effusions and pericardial effusion. the physician believed these to be injuries from a previous cardiopulmonary resuscitation and then was heparinized for the pci and impella implantation which possibly led to a second arrest. the family decided to withdraw care, and the patient expired. | [5 empty] | The pump was not returned for investigation. the data logs confirm the pump suction events. the root cause of the injury was unable to be determined due to insufficient clinical details. the cause of the pump suction issue was most likely related to patient condition based on data log review and provided clinical details. the pump passed all post sterile inspection checks. | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6447 | 1220648-2025-31175 | 22894322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22894322&pc=OZD | 07/24/2025 | 08/26/2025 |  | 08/26/2025 |  |  | 07/24/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1411749 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant started priming and inserting an impella cp for support in a patient admitted with an acute myocardial infarction/cardiogenic shock. a situation was noted to have occurred that was an "emergency" and at the time of starting the priming, the purge failed and the pump would not start. the team troubleshoot by replacing the automated impella controller (aic) but the second aic failed also so the team went back to the first aic and got the pump started for patient support. the pump supported for days and the patient was transferred from the community to the tertiary center. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6448 | 1220648-2025-45985 | 22897186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22897186&pc=OZD | 06/12/2025 |  |  | 08/26/2025 |  |  | 06/12/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675708 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Wrong Label | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported there was a data download issue encountered due to serial number mismatch involving the impella 5.5 device. there was no report or indication of adverse effects to the patient as a result of the event. further information is unavailable. | [5 empty] | Product-specific information was obtained and confirmed during investigation of the complaint (see below) and fields d.4 and h.4 were updated accordingly. the impella device was not return and therefore an analysis of the device was not possible. however, images of automated impella controller console screenshot were provided and showed that the pump serial number was reflecting as (b)(6) . however, the pump used in support was labelled as serial number (b)(6) . the consecutive serial number pumps were mismatched. the device history record review indicated that both pumps serial numbers: (b)(6) were reworked for wrong labelling in pouch label and pump handle during final inspection. this issue was presumed to be corrected during rework and released for use. later, it was investigated and found that the issue was not resolved and likely relabeled the wrong serial number label again on both consecutive pumps before sending out to the field. the pump serial number (b)(6) was scrapped, and the pump labeled serial number (b)(6) was used in the field. both pumps passed all functional testing during manufacturing. the only issue identified during post-sterile inspection was the label swap issue, and all other characteristics passed. applicable corrections have been identified. the issue was escalated to assess the pump that was used in the field. the cause of wrong label issue is label swap due to quality control deficiency. | Patient-and product-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6449 | 1220648-2025-45940 | 22897933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22897933&pc=OZD | 07/28/2025 |  |  | 08/26/2025 |  |  | 07/29/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569565 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Contamination | Difficult or Delayed Positioning | Fluid/Blood Leak | Deformation Due to Compressive Stress | Device Sensing Problem |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and it was reported there were issues with the pump position, alarms for impella in aorta, blood seen on the anti-contamination sleeve, and a catheter kink. the patient presented with to the hospital on (b)(6) 2025 in st elevation myocardial infarction with an occulted ramus which was perforated during a percutaneous coronary intervention. the surgeon was unable to resolve the perforation with balloon inflations, and cardiovascular surgery was consulted. an intra-aortic balloon pump (iabp) was placed, and the patient was taken for 2 coronary artery bypass grafts. the patient remained on the iabp for four days and then removed from support. after the removal of the iabp, the patient began showing signs of shock and went into cardiac arrest requiring one round of cardiovascular resuscitation, defibrillation, and was re-intubated. the patient was then emergently taken to the catheterization laboratory for an impella cp implant. it was reported that during the first day of impella cp support there was a persistent impella in aorta alarm. repositioning was performed and it was noted that the device was caught in the papillary muscle. the pump was freed from the muscle, but the alarms continued. on the second day of support, blood was found in the anti-contamination sleeve and a catheter kink was noted, with no noted intervention required. on the fourth day of support, care was withdrawn at the request of the family and the patient expired. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product or data logs were returned. contaminated device during use no product returned, the root cause of the decontamination during use cannot be determined. product damage (catheter kink) a catheter kink was noted on 7/29. due to a lack of clinical information and no product returned, the root cause of the product damage (catheter kink) cannot be determined. difficult to place or position based on the clinical details, the cause of the positioning issue was determined to be an unintentional use error as the pump came out of position after patient transfer. optical sensor issue there were ongoing impella position in aorta alarms after position was confirmed. due to lack of product and no data logs returned, the root cause of the optical sensor issue cannot be determined. pump 544978 passed all post sterile inspection checks. | Correction: section h6: the medical device problem code of a050401 was incorrectly reported in the initial report, which is now updated to correct code.
6450 | 1220648-2025-45946 | 22897938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22897938&pc=OZD | 07/28/2025 |  |  | 08/26/2025 |  |  | 07/28/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654166 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 112 KG |  | White | Injury | Yes | Yes | Increase in Pressure | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) during work up for left ventricular assist device placement. the impella 5.5 was successfully implanted, and the patient was sent to the unit on support. approximately five hours later, suction was encountered with point-of-care ultrasound showing the inlet into the mitral valve apparatus. the impella 5.5 was repositioned and the following day, position remained in good positioning according to the physician. however this day, the patient was taken back to the operating room to surgically address a hematoma. support continued; the patient ambulating waiting for the left ventricular assist device. day 8 of support it was noted purge system blocked alarm was persisting despite changing cassette and checking for kinks. switch to heparin purge was not an option as the patient had developed heparin-induced thrombocytopenia. heparin had been replaced with sodium bicarbonate. tissue plasminogen activator was added to the purge to address a purge system blockage. the patient¿s condition was noted unchanged otherwise and flowing ¿well¿ on the impella 5.5. six days later the impella 5.5 device was explanted with a surviving outcome. the patient was bridged to a left ventricular assist device. | [5 empty] | No product was returned for investigation. the cause of the high purge pressure was most likely biomaterial buildup as data logs showed a gradual rise in pp and no motor current spikes or trends. the cause of the positioning issues was not determined since insufficient clinical details were provided and product was not returned. the cause of the hematoma was not determined since insufficient clinical details were provided and no product was returned. the cause of the thrombocytopenia cannot be determined as insufficient clinical details were provided. the pump passed all post-sterile inspection checks. | Patient information for ethnicity (section a5) is unknown. device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6451 | 1220648-2025-45964 | 22898961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22898961&pc=OZD | 07/29/2025 |  |  | 08/26/2025 |  |  | 07/29/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681808 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported the patient was admitted for a high-risk percutaneous coronary intervention (pci) to the left main, left anterior descending and right coronary arteries and impella pre-pci was planned. the impella cp device was inserted, and the pump was started. a single access used for pci. after some time, the team noticed bleeding from peel-away sheath. the sheath did not appear cracked; however, oozing was noted from where the sheath connects to the hub/orange wings. the physician elected to exchange the short peel-away sheath for a long peel-away sheath and resumed the pci. the patient was eventually weaned from the impella device, and the device was successfully explanted with a survival outcome. | [5 empty] | Patient-specific information for section a4: weight is unavailable at the time of this mdr writing. device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two device reports associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the peel-away introducer. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the clinical details state that there was bleeding noted from the access site, specifically from the sheath connects to the orange hub. when the sheath was removed, the ooze persisted and it subsided when the pump was removed. leaking was reproduced against the introducer. there was no pump returned to confirm if there was any defects. due to the clinical details stating that the oozing resolved with pump removal and there was no product returned, the root cause cannot be established.
6452 | 1220648-2025-45967 | 22900210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22900210&pc=OZD | 07/29/2025 |  |  | 08/26/2025 |  |  | 07/29/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687550 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 52 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. approximately six days after starting impella mcs, the patient was cannulated for venovenous extracorporeal membrane oxygenation (vv ecmo). during vv ecmo cannulation a high purge pressure alarm on the impella was encountered. the team checked that all purge connections were secure. impella connect was reviewed and noted there was a sudden spike in motor current before alarm occurred. issue was resolved without any intervention. impella support continued. the following morning, the patient experienced atrial fibrillation. it was noted the patient was cardioverted the day prior. the next day, the patient was successfully weaned and impella device was explanted at bedside. the patient was then taken for trach placement and remained on vv ecmo. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for evaluation. only event logs were returned. arrythmia: clinical details noted that patient was cardioverted but was back in a-fib the following morning. no more clinical details were provided. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. high purge pressure: data logs were reviewed and a sudden motor current spike with speed deviation at time of spike in purge pressure to 1000mmhg. a "high purge pressure" alarm was triggered and purge pressure remained high for approximately 7 minutes before slowly resolving to normal purge pressure range over approximately 4 hours. at time of motor current spike, pump flows become briefly saturated at time of hpp alarm and is able to resolve over a few minutes. clinical details noted that high purge pressure alarm was spiked during vv ecmo cannulation. impella connect showed a motor current spike prior to alarm. issue was able to resolve without intervention. the cause of the high purge pressure was due to biomaterial ingestion based on returned data log analysis and clinical details. pump sn (b)(6) passed all post-sterile inspection checks.
6453 | 1220648-2025-45831 | 22900667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22900667&pc=PYX | 07/30/2025 |  |  | 08/26/2025 |  |  | 07/30/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659202 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported postoperative (coronary artery bypass graft with left arterial clip), the patient started to deteriorate in the intensive care unit requiring increased use of vasopressors. the patient was brought back to the operating room for exploratory surgery and impella placement (impella cp placed). the patient decompensated overnight requiring max vasopressors support and the decision was made to transfer patient for left ventricular assist device (lvad). the patient became unstable now requiring impella 5.5 and during an echocardiogram right ventricular dysfunction was noted. after escalation to and insertion of an impella 5.5 device, blood pressure ¿significantly¿ improved. however, few minutes later it alarmed suction and the blood pressure dropped. decision was made to place an impella rp flex which was completed. the patient was sent to the unit on bipella support. seven days into the support, the impella rp flex had high purge pressure. the purge cassette was replaced, but the issue persisted. tissue plasminogen activator was subsequently run through the purge to manage the issue. support continued for three more days, at which point care was withdrawn and the patient passed away. | [5 empty] | Patient information for sections a4, a5, and a6 are unknown. device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for evaluation. insufficient data logs during time of high purge pressure event were returned for analysis. the cause of the high purge pressure could not be definitively determined as no product and insufficient logs at time of high purge pressure event were returned and clinical information were provided.
6454 | 1220648-2025-45977 | 22901371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22901371&pc=OZD | 07/29/2025 |  |  | 08/26/2025 |  |  | 07/29/2025 | 08/26/2025 |  |  |  |  | 08/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701803 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced thrombocytopenia approximately two days after start of the support. it was noted the thrombocytopenia persisted; however, it was stable over interval. patient was placed on argatroban. support continued for another eight days. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the thrombocytopenia cannot be determined as insufficient clinical details were provided. the pump passed all the post sterile inspection checks.
6455 | 1220648-2025-45933 | 22882321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22882321&pc=OZD | 07/26/2025 |  |  | 08/25/2025 |  |  | 07/26/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1120449 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported during impella mechanical circulatory support, the automated impella controller sounded an alarm with an alarm message for purge pressure low. it was noted that the alarm was ongoing. actions taken to resolve the alarm issue is unknown as well as the status of the patient as a result of the event. however, there was no report or indication of any adverse effects to the patient. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
6456 | 1220648-2025-45787 | 22882731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22882731&pc=PYX | 07/29/2025 |  |  | 08/25/2025 |  |  | 07/29/2025 | 08/25/2025 |  |  | 02/28/2026 |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479484 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during delivery of an impella rp flex for patient support. despite successfully reaching the left ventricle with a .018 guidewire, the backloaded pump could not advance beyond the valve. the pump was removed and a .027 wire was placed for another attempt. the pump was then re-inserted and successfully advanced into the left ventricle. the pump was started at p-2, but immediately stopped. after multiple failed attempts to restart the pump, the device was removed and a new impella rp flex was placed. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The returned product did not reveal any damages on the pump. the root cause of the delivery issue was unable to be determined as insufficient clinical details were provided. due to a lack of clinical details provided, the root cause of the "pump did not start" cannot fully be determined as the issue was not able to be reproduced. the pump passed all post sterile inspection checks.
6457 | 1220648-2025-45830 | 22882867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22882867&pc=OZD | 07/29/2025 |  |  | 08/25/2025 |  |  | 07/29/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697730 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was unable to navigate the arch of the porcelain aorta during attempted implant. after multiple failed attempts, the impella placement was aborted and an intra-aortic balloon pump was placed for patient support. | [5 empty] | A.4 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The investigation has been completed. the pump, guidewire, and purge cassette returned. pump sn (b)(6), gw and purge cassette returned. unable to deliver or advance - during pump implanted procedure, the impella was unable to navigate the arch of the porcelain aorta due to noted vessel calcification. the returned pump showed no damage or other abnormalities. the root cause of the deliver issue was determined to be patient condition as the patient had calcified vessels preventing successful delivery of impella. the device history review was completed and the pump passed all post sterile inspection checks.
6458 | 1220648-2025-45829 | 22883954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22883954&pc=OZD | 07/28/2025 |  |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701752 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the patient experienced oozing at the access site and as a result of the bleeding, heparin was withheld. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the access site bleeding was most likely use related , heparin was stopped to achieve hemostasis as per the clinical description. the pump passed all the post sterile inspection checks.
6459 | 1220648-2025-31181 | 22884122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884122&pc=OZD | 07/28/2025 | 08/25/2025 |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692740 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 had a placement signal not reliable alarm after 21 days of supporting a patient. an echocardiogram confirmed that the impella 5.5 was in a good position. motor current and flows were unchanged. it was further reported that the impella 5.5 was exchanged for a new impella 5.5 due to clot in the cannula. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
6460 | 1220648-2025-31182 | 22884321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884321&pc=OZD | 06/16/2025 | 11/04/2025 |  | 08/25/2025 |  |  | 07/30/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | IMPELLA 5.5 |  |  | US |  | 1107442 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella 5.5 support in the intensive care unit and the patient lost pulsatility while coughing. it was noticed that on the automated impella controller, the placement signal disappeared. the placement signal did return on the impella connect shortly after it disappeared. no patient harm was reported. upon review of device data logs by abiomed investigators, it was found that the console may be the potential cause of the issue. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: review of data logs confirms the complaint of ¿placement signal disappeared¿, though no actual alarm was generated. at the time of the reported complaint, the imc logs show calculator placement #1045 errors. the rt logs show that at the same time, due to a crc error, all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values while the placement signal is represented as 3000. the cause of the aic issue was a 3rd party hardware/software issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated.
6461 | 1220648-2025-45852 | 22884380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884380&pc=OZD | 07/31/2025 |  |  | 08/25/2025 |  |  | 07/31/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676686 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced suction alarms upon start up. it was turned on with audio and immediately suction alarm occured. the position of the pump was confirmed via fluoroscopy and echocardiogram. the position was changed even thought it was confirmed correct. it did not help. on p-2, the pump was giving only 0.8 liters per minute and the suction alarm still occured. the pump was explanted and a new impella pump was implanted. the patient survived. | [5 empty] | The customer's device was returned and a visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the cause of the low flow was most likely biomaterial ingestion. | A.4, a.5, and a.6 are unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The pump was returned for investigation. a visual inspection found biomaterial wrapped around the impeller. no other issues were found on the rest of the pump. the cause of the low flow and suction events was biomaterial ingestion. the device passed all post sterile inspection checks.
6462 | 1220648-2025-31183 | 22884428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884428&pc=OZD | 07/21/2025 | 09/12/2025 |  | 08/25/2025 |  |  | 07/21/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662735 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 placed to support a patient with known cardiomyopathy, known chronic kidney disease on continuous renal replacement therapy and known congestive heart failure. the pump was placed but after close to a month of support, there were purge pressure and a purge system blocked alarm. the team made thoughts to infuse tissue plasminogen activator to the purge. the pump was not replaced and support continued. the patients outcome was stable. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the logs confirm high purge pressure alarm with no prior motor current spikes. the purge pressure rose for about 30 minutes before both purge system blocked and purge pressure-high alarms were triggered. there was also a simultaneous decrease in purge flow. high purge pressure was sustained for about an hour before troubleshooting efforts resolved issue. the root cause of the high purge pressure was most likely biomaterial as evidenced by the gradual increase in purge pressure in the logs simultaneous with purge flow decrease and issue resolution after p-level decrease b3 date of event was erroneously entered on the initial manufacturer device report. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6463 | 1220648-2025-45826 | 22884672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884672&pc=OZD | 07/28/2025 |  |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700855 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 114 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had been previously supported by a impella cp (serial number (b)(6)) and extracorporeal membrane oxygenation (ecmo), when the medical team decided to remove the first impella cp to allow for ecmo reconfiguration. the initial impella cp was removed from the left femoral artery and exchanged for ecmo arterial cannula. a new impella cp (serial number (b)(6)) was then implanted into the patient¿s left femoral artery. it was reported that after the impella was placed and support initiated, the patient required defibrillation to resolve persistent ventricular tachycardia. the patient remained on impella cp support for an additional 5 days and was then escalated to an impella 5.5 successfully for further long-term support. | [5 empty] | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6464 | 1220648-2025-31184 | 22884765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884765&pc=OZD | 07/08/2025 | 10/07/2025 |  | 08/25/2025 |  |  | 07/31/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133102 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 107 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 placed to support a patient with known cardiac disease and renal disease. the pump was supporting when there was an access site bleed treated by dressing changes. additionally, there was pain at the site. there was no treatment given or intervention was needed. the pump had also lost placement signal and it was "erratic" but the pump remained on for support. the investigating team at abiomed has deemed an investigation of the automated impella controller for the erratic placement signal was necessary. | [5 empty] | The investigation of optical signal issue has been completed. console logs were consistent with what was reported in the complaint. console was tested after arriving in field service. the reported placement signal issue was unable to be reproduced while testing the console with a known good purge cassette kit and pump. the optical bench 5s was checked but all parameters were within acceptable limits. distortion of the ccd array output was examined while analyzing the ccd from the optical bench. conclusion: the root cause of the placement signal issue was determined to be caused by a defective optical bench due to a distorted ccd array analog output. d9. device returned to manufacture date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6465 | 1220648-2025-45942 | 22884869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884869&pc=OZD | 07/28/2025 |  |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693676 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 138 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced positioning issues. an echocardiogram confirmed the pump was malpositioned. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D.4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned. positioning issue: no product was returned. echocardiogram was performed: depth 5.5 centimeters angled toward septum. no other alarms or any symptoms of malposition were observed. the root cause of the positioning issue was unable to be determined as insufficient clinical details were provided. the device history review was completed and the pump passed all post sterile inspection checks.
6466 | 1220648-2025-31185 | 22884897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22884897&pc=OZD | 06/19/2025 | 09/25/2025 |  | 08/25/2025 |  |  | 07/31/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1416992 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 59 KG | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal not reliable alarm appeared during impella 5.5 support. despite losing the placement signal and left ventricle readings, the motor current and flow rate remained stable. support was continued uninterrupted. during the investigation of the event, it was discovered that the automated impella controller experienced an intermittent loss of light, indicating a potential lamp failure. there was no resulting patient harm. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. the console and date logs were returned for evaluation. aic - display issue: no evidence was found in the logs indicating an ¿aic ¿ display issue,¿ and no display issues were observed during testing. most likely, the disappeared placement signal (ps) was misreported as ¿aic ¿ display issue. optical signal issue: the root cause of the intermittent placement signal not reliable (psnr) and the oscillating contrast was likely due to a defective lamp. d9 device returned to manufacturer date added g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31185. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6467 | 1220648-2025-45930 | 22885118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22885118&pc=OZD | 08/02/2025 |  |  | 08/25/2025 |  |  | 08/02/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638897 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was initially taken to the catheterization lab due to an absence of detectable pedal pulses on doppler examination. an angiogram was performed and it was noted that adequate flow was seen in the extremities. although ischemia was ruled out, a clot was found in the ascending to descending aorta. as the patient's left ventricular ejection fraction was normal, the decision was made to remove the pump. the patient remained supported via extracorporeal membrane oxygenation and was later implanted with an impella rp flex for right ventricle support. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
6468 | 1220648-2025-31186 | 22885293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22885293&pc=OZD | 06/20/2025 | 11/24/2025 |  | 08/25/2025 |  |  | 07/31/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1160669 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal on an impella cp pump became unreliable during patient support. the placement monitoring was disabled in response to the noted issue and support was continued uninterrupted. two days later, the patient was made palliative care by their family and passed away on support. the investigating team has requested the console to be investigated as well, beyond the complainant's request for pump investigation. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ic3137 ran cp pump (b)(6) from 6/17-6/22. on 6/20, the user received a placement signal not reliable alarm, which the user silenced, that did not resolve until the pump was removed on. after the pump was removed, the optical system reports light of 0.43v which indicates receiving only ambient light and none from the optical bench's lamp. ltlogs and rtlogs show that at the time of the placement signal not reliable alarm raw optical sensor dropped to -3276.8mmhg as placement signal went to 3000 mmhg and lamp went to 100% and contrast oscillated between +/-3276.8%. device analysis: the console was returned and elements of the failure mode reproduced and both contrast and snr were failing per pm though light was fine at this time. root cause: the placement signal not reliable and the oscillating contrast were most likely due to a failing lamp as seen by the bad lamp after the pump was removed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated. | The impella controller was returned for evaluation. upon completion of the investigation, a supplemental manufacturer device report will be filed.
6469 | 1220648-2025-45948 | 22885691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22885691&pc=PYX | 07/29/2025 |  |  | 08/25/2025 |  |  | 07/29/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590619 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 104 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex generated unexpected pulmonary artery sensor readings and decreased flow with spikes in motor current. the impella was removed and a clot was found at the impella outlet. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned. the cause of the low pump flow was most likely biomaterial ingestion base on the data log review. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
6470 | 1220648-2025-45941 | 22888807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22888807&pc=OZD | 07/26/2025 |  |  | 08/25/2025 |  |  | 07/27/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670848 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 105 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma and a thrombus. the pump was explanted and the artery was surgically repaired. the patient required blood product transfusions during surgery. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
6471 | 1220648-2025-45818 | 22888966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22888966&pc=OZD | 07/29/2025 |  |  | 08/25/2025 |  |  | 07/29/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693450 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the initial impella 5.5 failed to advance through the patient's artery after several attempts to pass resulting in buckling of the catheter. a second impella 5.5 was implanted to support the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The root cause of the delivery issue was determined to be patient condition as the patient had resistance in axillary artery preventing successful delivery of impella. the root cause of the catheter kink was determined most likely to be patient condition as the patient had resistance in axillary artery resulting in buckling/kinking of catheter near motor housing. the pump passed all the post sterile inspection checks.
6472 | 1220648-2025-45924 | 22889050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22889050&pc=OZD | 08/01/2025 |  |  | 08/25/2025 |  |  | 08/01/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685036 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 71 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Inflammation | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the right axillary access site began to swell. vascular surgery was consulted and it was discovered that the artery had been dissected. blood products were given and manual pressure was initially applied to help manage the condition. three days later, the patient was escalated to impella 5.5 support. upon removal, a stent was placed and surgical intervention was performed to repair the artery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The cause of the vascular dissection and inflammation cannot be determined as insufficient clinical details were provided. the device passed all post sterile inspection checks.
6473 | 1220648-2025-45947 | 22889108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22889108&pc=OZD | 07/29/2025 |  |  | 08/25/2025 |  |  | 07/29/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706539 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 121 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after successful implantation of an impella cp, the repo sheath was inserted and lots of bleeding and a large hematoma formed. the pump was explanted and the patient survived. | [5 empty] | The root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6474 | 1220648-2025-45906 | 22889189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22889189&pc=OZD | 07/26/2025 |  |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620342 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | No | Yes | False Alarm | Increase in Pressure | Device Sensing Problem |  |  |  |  | Multiple Organ Dysfunction Syndrome | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | It was reported that a patient implanted with an impella cp experienced placement signal unreliable alarms. proper pump position was confirmed and impella support continues. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. corrected information provided in h6. correction: the awareness date provided in the initial report was determined to be incorrect. the correct awareness date is 26-jul-2025. | D6b explant date provided. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represent the pump. refer to related reports in h.10 for the two reports representing the automated impella controllers. | The impella cp was returned for investigation. data logs were reviewed and did not confirm the complaint. no erroneous triggers of controller error or alarms were found. all optical metrics were stable and within range. the returned product had no visual abnormalities, and the pump passed all tests. the cause of the false alarms and optical signal issues could not be determined. there was a white/clear material under the impeller in the outflow cage of the returned pump. it is unclear if this was biomaterial or something else. there was also noted to a kink in the catheter. the gradual resolution of the purge pressure in data logs is indicative of biomaterial. however, the onset of the purge pressure rise/spike was fast but did follow a motor current spike, so it is unclear if it was an ingestion or buildup event. based on the patterns noted in data log review, the cause of the high purge pressure was determined to be biomaterial of an unknown nature. the device passed all post-sterile inspection checks. | The device was not returned for investigation. the data logs were reviewed and did not confirm the complaint condition described. therefore, the cause of the false alarms and optical signal issues could not be determined. based on the patterns noted in data log review, the cause of the high purge pressure was determined to be biomaterial of an unknown nature. the pump passed all post-sterile inspection checks.
6475 | 1220648-2025-45908 | 22889346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22889346&pc=OZD | 07/26/2025 | 10/29/2025 |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274026 |  | 0042-0000-UK |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced lots of issues with the automated impella controller (aic) and impella catheter. the controller issue was with a service loaner without connect issues. there was a controller failure. the controller error occurred when swapping the aic. controller error continued on the new controller. on both aics alarm would occur and resolve. the impella position in aorta, placement signal unreliable alarms would occur. position was noted to be fine on echocardiogram. impella support continues. | [5 empty] | Product complaint # (b)(4). b3 date of event has been corrected. g3 date received by manufacturer has been corrected. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the second automated impella controller. refer to related reports in h.10 for the report representing the pump and for the report representing the first automated impella controller. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6476 | 1220648-2025-45790 | 22889983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22889983&pc=OZD | 07/28/2025 |  |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702694 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that roughly four days into impella cp support, a patient endured persistent runs of ventricular tachycardia. multiple defibrillations were required to counteract the condition. support continued uninterrupted following the intervention. the following day, the bedside staff made the decision to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The device was not returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A.2, a.4, a.5, and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6477 | 1220648-2025-45910 | 22890781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22890781&pc=OZD | 07/25/2025 |  |  | 08/25/2025 |  |  | 07/25/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555853 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | False Alarm | Restricted Flow rate | [5 empty] | Atrial Fibrillation | Fever | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Tachycardia | Asystole | Heart Block | [6 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support to wean the patient from cardiopulmonary bypass. it was reported that the day after support was initiated, there were impella position unknown alarms likely due to the patient¿s low pulsatility, and the medical team was advised to correlate the impella position clinically. two days later, on (b)(6) 2025, it was reported that the patient developed paroxysmal atrial fibrillation with rapid ventricular response in the setting of ddi pacing. there were noted events of severe slow atrial fibrillation and pulsitility dropped in the setting of irregular rhythm. the patient was given a bolus of amiodarone to treat the arrhythmia. the next day, on (b)(6) 2025, the patient developed supraventricular tachycardia with complete heart block and periods of ventricular asystole. it was noted that the patient was experiencing a complex rhythm issue, sometimes observed in the setting of mitral valve regurgitation and aortic valve regurgitation. impella support was continued, and the patient was monitored. later that day, the patient developed ventricular tachycardia and was defibrillated twice. cardiopulmonary resuscitation (cpr) was initiated after the patient lost pulse. two rounds of cpr were performed and levophed was administered for low mean atrial pressure. suction alarms presented after pressures were restored, and the pump was lowered to p-7 from p-8, and it was noted the patient continued to experience intermittent ventricular tachycardia. amiodarone and lidocaine were administered, and a regular heart rhythm was achieved. on (b)(6) 2025, it was noted that there were no further arrhythmic events, however the patient¿s white blood count increased above normal range and the patient was febrile. on (b)(6) 2025, it was noted that the patient¿s hemaglobin dropped significantly with no apparent source of bleeding. heparin was discontinued and the patient was transfused with 2 units of packed red blood cells (prbc). blood cultures resulted and were positive for bacteria growth with gram-negative rods and gram-positive cocci bacteria identified. the patient was placed on antibiotics. the following day, on (b)(6) 2025, the source of bleeding was confirmed to be gastrointestinal as melena was observed in the stool. the patient received 5 more units of prbcs, heparin remained withheld. it was reported the patient had another episode of ventricular tachycardia that required cardioversion and amiodarone and lidocaine to resolve. on (b)(6) 2025, it was reported that there were false alarms for impella in aorta due to a loss of pulsatility and suctioning. on (b)(6) 2025, echocardiography was performed and the patient¿s left ventricle ejection fraction recovered to 25%, and a plan was made to explant. the patient was successfully weaned and explanted. upon explant, the impella site was cultured due to turbid drainage observed. the wound was irrigated and then closed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The product was not returned for investigation. the data log review confirmed the complaint condition. the cause of the pump suction is most likely patient condition, as clinical details as well as the low pulsatility in placement signal and motor current from data log review indicates poor or declining patient condition. the root cause of the atrial fibrillation, tachycardia, asystole, fever, and infection was unable to be determined due to insufficient clinical details. the cause of the retroperitoneal hemorrhage was not determined due to insufficient clinical details. the pump passed all post sterile inspection criteria.
6478 | 1220648-2025-45973 | 22890866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22890866&pc=OZD | 07/15/2025 |  |  | 08/25/2025 |  |  | 07/15/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687509 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella cp device to a patient in an acute myocardial infarction/cardiogenic shock was unsuccessful. the patient presented in cardiac arrest. return of spontaneous circulation was obtained and the patient was taken to the operating room, for extracorporeal membrane oxygenation (ecmo) placement and left ventricle unloading with the impella. however, multiple attempts to cross the aortic/bicuspid aortic valve with the impella cp were made unsuccessfully due to aortic stenosis. consequently, the impella cp implant was aborted. the outcome status of the patient was indicated as unknown. however, there was no report or indication of adverse effects to the patient. | [5 empty] | A4: patient weight is unknown. device status: the impella device was not received from the customer. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for the impella cp with smartassist system: section: potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The impella device was not returned, and therefore, an analysis of the device was not possible. a device history record review was completed and noted the device passed all post sterile inspection checks. the delivery issue root cause was determined to be patient condition as the patient had aortic valve resistance preventing successful delivery of impella. section h6 investigation coding has been updated accordingly based on the completed investigation.
6479 | 1220648-2025-45943 | 22891108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22891108&pc=OZD | 07/26/2025 |  |  | 08/25/2025 |  |  | 07/26/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701169 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 111 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was initially implanted with an impella cp device for mechanical circulatory support (mcs) and then escalated to an impella 5.5. for continuation in therapy. impella 5.5 was slowly increased to performance level p-5, impella cp was weaned and explanted. impella 5.5 measured under echocardiogram with inlet sitting free from structures, mid ventricular and measuring approximately 5cm from valve. the impella was secured, and the patient was transferred to the unit on support. approximately three days after starting the impella mcs, the patient experienced ventricular fibrillation (vf) and ventricular tachycardia (vt) which were treated with defibrillation. two days later, the impella position was found to be deep, and the device was repositioned. the patient continued to have vf and vt requiring 16 shocks and medication. the patient had ongoing multisystem organ failure and was ineligible for advanced therapies. day 18 of support, the patient expired. care was withdrawn. | [5 empty] | The customer's device was not returned. the device history review was done and the pump passed all post sterile inspection criteria. the root cause of the reported ventricular tachycardia and ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of the pump positioned deep was unable to be determined due to insufficient clinical details regarding how the pump went deep. | . device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6480 | 1220648-2025-45909 | 22891404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22891404&pc=OZD | 07/26/2025 |  |  | 08/25/2025 |  |  | 07/28/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1629362 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced lots of issues with the automated impella controller (aic) and impella catheter. the controller issue was with a service loaner without connect issues. there was a controller failure. the controller error occurred when swapping the aic. controller error continued on the new controller. on both aics alarm would occur and resolve. the impella position in aorta, placement signal unreliable alarms would occur. position was noted to be fine on echocardiogram. impella support continues. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the first automated impella controller. refer to related reports in h.10 for the report representing the pump and for the report representing the second automated impella controller. | B3 date of event has been corrected. g3 date received by manufacturer has been corrected. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6481 | 1220648-2025-45936 | 22891511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22891511&pc=OZD | 07/29/2025 |  |  | 08/25/2025 |  |  | 07/29/2025 | 08/25/2025 |  |  |  |  | 08/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676189 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 116 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support (and at this time intra-aortic balloon pump and venoarterial extracorporeal membrane oxygenation (ecmo) were removed). after approximately one month of impella 5.5 support, the patient was noted to have gastrointestinal bleeding. heparin was suspended and blood products were administered. it could not be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. due to continued decompensation and inability to wean off vent and veno-pulmonary ecmo, family and team ultimately decided to withdraw care due to patient becoming unresponsive. the patient expired. | [5 empty] | No product was returned for investigation. the root cause of bleeding was determined to be use issue because the patient was on aggressive pt/ot and heparin was stopped to manage the bleed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use: as noted in the impella ifu: section: potential adverse events (united states) - "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6482 | 1220648-2025-45860 | 22871308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22871308&pc=OZD | 08/02/2025 |  |  | 08/22/2025 |  |  | 08/02/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697743 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced some bleeding from the access site during impella cp support. the patient was given protamine to control the bleed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the complaint was confirmed. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6483 | 1220648-2025-45855 | 22871393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22871393&pc=PYX | 08/02/2025 |  |  | 08/22/2025 |  |  | 08/02/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026679972 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 127 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella rp flex underwent a procedure for an internal jugular line placement. after the procedure, the pump placement signal disappeared and a placement signal not reliable alarm appeared. no intervention was taken regarding the noted issue. it was additionally reported that the patient developed hemolysis during support. again, no intervention was taken in response to the condition. care was withdrawn later that day upon the family's request and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6484 | 1220648-2025-45919 | 22871842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22871842&pc=OZD | 07/24/2025 |  |  | 08/22/2025 |  |  | 07/24/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666911 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 83 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Thrombocytopenia | Epistaxis | [9 empty] | Required Intervention | Death |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was initially implanted with an impella cp for mechanical circulatory support (mcs) and experienced complications involving the peel-away introducer. the patient was escalated to an impella 5.5 after six days on the impella cp and was placed on extracorporeal membrane oxygenation support. during impella 5.5 support, the patient experienced suction events, hypotension, heparin-induced thrombocytopenia and bleeding requiring discontinuation of anticoagulation. the patient ultimately expired on support. three days after starting impella 5.5 mcs, a suction event alarm was triggered due to hypotension and preload. one unit of blood was administered. echocardiogram at bedside showed the impella 5.5 angled toward the septal wall measuring 4.25 centimeters away from the aortic valve area. the care team was made aware. three days later, there were occasional suction alarms. fluid was repleted and a second unit of packed red blood cells was transfused. the patient was in a hypovolemic state. the next day, the patient remained slightly seated, responded to verbal command and was making needs known. a suction event was again experienced related to preload. lascivious drip was held and central venous pressure was now 10. thrombocytopenia was noted persisting. heparin-induced thrombocytopenia panel was sent, and heparin was suspended with a noted partial thromboplastin time in the 50¿s with plans to start argatroban. however, of note, the purge composition was free of heparin with 5% dextrose in water and 25 milliequivalents per liter of sodium bicarbonate. now 15 days after starting on the impella 5.5 mcs, systemic anticoagulation has been put on hold again due to epistaxis and a noted gastrointestinal bleed confirmed as negative by colonoscopy. heparin was restarted. a couple of days later, suction alarms were noted again. lasix was decreased and suction alarms resolved. the patient was made do not resuscitate/do not intubate and had been weak with very limited mobility and had not gotten out of bed. the family discussed palliative care and the patient expired on support after 21 days of impella mcs. | Additional information was received. the team did not save the device or related equipment. also, the team determined there was no gastrointestinal bleed and the blood in stool was related to a nosebleed. the patient was restarted on systemic anticoagulation thereafter. this patient was very complex and not a candidate for advanced therapies per heart failure. |  |  |  |  | Medical safety officer reviewed the event and noted the impella 5.5 device is unlikely to be the cause of the demise outcome. the patient was in a hypovolemic state triggering the suction events. the patient was on a heparin free purge composition and heparin-induced thrombocytopenia may or may not have been induced by heparinizing the patient when the impella was used/inserted. epistaxis and other bleeding appear to be unrelated to the impella device. however, although unlikely related, there is not enough information to exclude the impella 5.5 as an associated factor in the overall event outcome. refer to related manufacturer device report number 1220648-2025-45918 for the report on the impella peel-away introducer. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical summary: suction alarms were triggered due to hypotension and preload and resolved by blood compensation. the patient also had thrombocytopenia, gastrointestinal bleeding, and epistaxis. the patient then expired. data review: data logs confirmed the failure mode ad clinical narrative. logs showed multiple intermittent suction alarms were triggered during the case corresponded with clinical description. there was no motor current or placement signal spike for the entire case. no other issues were found on the logs. the product was not returned for review. the cause of suction issue was most likely patient condition related since the clinical team confirmed the suction was due to hypotension and preload and resolved by blood transfusion. regarding the thrombocytopenia, gastrointestinal bleed and epistaxis and hypotension, the cause of the clinical events cannot be determined as insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6485 | 1220648-2025-45857 | 22872169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22872169&pc=PYX | 08/02/2025 |  |  | 08/22/2025 |  |  | 08/02/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479845 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 56 KG | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that roughly two days into impella rp flex support, a placement signal not reliable alarm appeared. despite the failure, both the motor current and flows remained unaffected. the alarm was subsequently disabled and support continued uninterrupted. the following day, the decision was made to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. the event logs were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The rp flex pump was not returned. data logs were reviewed and optical signal failure was confirmed. the cause of optical sensor issue was due to damaged sensor based on log analysis findings. but, the cause of damaged sensor is unknown since product was not returned/due to insufficient clinical information.
6486 | 1220648-2025-45782 | 22872268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22872268&pc=OZD | 07/28/2025 |  |  | 08/22/2025 |  |  | 07/28/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614320 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 123 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp implant, the patient experienced wide complex tachycardia. the patient was defibrillated multiple times to counteract the condition. it was noted that the patient had underlying atrial fibrillation with right bundle branch block leading up to the event. it was additionally reported that despite having absent bilateral lower extremity posterior tibial and dorsalis pedis pulses prior to impella implant, the condition continued to worsen with the extremities becoming very mottled. support continued despite the condition for two days, at which point the patient was moved to comfort care. the patient expired following pump removal. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the tachycardia and ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
6487 | 1220648-2025-45861 | 22872316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22872316&pc=OZD | 08/02/2025 |  |  | 08/22/2025 |  |  | 08/02/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704147 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 114 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that roughly two days into impella 5.5 support, a computed tomography scan revealed that the patient had experienced an embolic stroke. administration of low dose heparin was approved following the stroke. the patient was also on venoarterial extracorporeal membrane oxygenation and right ventricular assist device. later, the family withdrew care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The device was not returned for investigation. the complaint was confirmed. the device passed all post sterile inspection checks.
6488 | 1220648-2025-45929 | 22874697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22874697&pc=OZD | 07/31/2025 |  |  | 08/22/2025 |  |  | 07/31/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 83 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. following impella cp placement, there was minimal flow in the impella leg. a distal perfusion catheter was subsequently placed and flow to the extremity improved. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
6489 | 1220648-2025-45884 | 22876426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22876426&pc=DQX | 07/25/2025 | 08/22/2025 |  | 08/22/2025 |  |  | 07/25/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018"PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  | 9217120 |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues and guidewire product damage during implantation of an impella cp. once the impella was placed, the pigtail appeared rotated and the impella appeared kinked between the motor and the bend of cannula. the physician attempted to reposition the device but was unsuccessful. as a result, the impella was removed through the peel away sheath and the 0.018 guidewire was damaged during process. a 0.014 300 centimeter whisper wire was used in place as no 0.018 in guidewires existed at this facility. the impella was cleaned and the steps were repeated without issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | D.4 the serial number is unknown and therefore the serial number portion was left out of the primary udi number. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report represents the guidewire. related manufacturer device report number 1220648-2025-45893 represents the pump. | No product was returned for analysis. the cause of the guidewire damage was not determined as no product or images were returned for analysis. the guidewire batch passed all post sterile inspection checks.
6490 | 1220648-2025-45859 | 22876945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22876945&pc=OZD | 07/31/2025 |  |  | 08/22/2025 |  |  | 08/01/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1851056 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller failed to charge when plugged in. the power cord was swapped but the issue persisted. there was no patient involved. | [5 empty] | A.2 - a.6 were left blank as there was no patient involvement. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The automated impella controller (aic) was returned for further investigation. during testing, the reported issue was successfully reproduced. the power cord and fuses were inspected but no issues were found. the voltage was measured between the power supply and the power brick module was below the normal voltage. the dc power supply was replaced with a known good unit to resolve the issue. further inspection of the batteries and pbm revealed no issues. additionally, the console was tested with a known-good pump and purge cassette, and no problems were encountered. the reported issue that the aic wouldn't charge when plugged in was determined to be caused by a malfunction of the power supply module. the power supply module was replaced and a full functional check of the console was performed before returning the device to the customer.
6491 | 1220648-2025-45934 | 22877026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22877026&pc=OZD | 07/28/2025 |  |  | 08/22/2025 |  |  | 07/28/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689157 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 70 KG |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and the next day following start of the support, the patient experienced bleeding from the peel away introducer that had been left in place during the impella cp support. the peel away introducer was subsequently removed, and the repositioning sheath was advanced to achieve hemostasis. support continued uninterrupted. two days later suction events occurred, and the patient received a unit of packed red blood cells which helped resolved the suction. impella mcs continued for an additional three days before being successfully weaned and device was explanted with a survived outcome. | [5 empty] | A.5 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. refer to the related manufacturer report number as noted in section h10 for the report on the peel away introducer. | Additional information was received that provided further details regarding the event and noted the following: there was bleeding from the peel away sheath that was left in place. peel away sheath was removed and repositioning sheath was cleaned and inserted. impella secured per best practices. insertion site now begin. the impella cp device was not returned and therefore, an analysis of the device was not possible. however, event logs were received and analyzed. a device history review was completed and noted the device passed all post sterile inspection checks. the data logs showed there was suction seen throughout the case. the suction alarm that was triggered was accurately triggered as the motor current minimum was within the expected curve. however, the placement signal was trailing down during this time. all optical 5s parameters were stable and pump flows were within specification, although slightly on the lower end. the suction alarms resolved, which was in conjunction with the clinical details stating that a unit of packed red blood cells helped resolve the issue. based on the clinical details and data log analysis, the root cause of the suction is most likely patient related. section h6 investigation coding has been updated accordingly based on the completed investigation. refer to the related manufacturer report number as noted in section h.10 for the report on the impella peel-away introducer.
6492 | 1220648-2025-45932 | 22877516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22877516&pc=OZD | 07/29/2025 |  |  | 08/22/2025 |  |  | 07/29/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089162 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 84 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. during support, the automated impella controller had a placement signal low alarm, and then the customer repositioned the impella. however, at this time, there were an impella in aorta and impella in ventricle alarms, one right after another. additionally, impella connect had disconnected from this console in the middle of the night. the blue light however stayed on. the alarms sounds were described as an old school telephone ringing sound. however, once the alarms resolved, it was noted the impella cp ended up in the same original position. abiomed support team was contacted and advised to exchange out the console, which was noted to have been successfully completed. there was no harm to the patient as a result of the event. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. refer to the related report number as noted in section h.10 for the report on the impella cp. | D.4 an expiration date was entered on the initial report submitted and should of been left blank as this device does not have an expiration date. h.10 the related manufacturer report number referred to in the initial report submitted on august 28, 2025, was inadvertently omitted in section h.10 and was added.
6493 | 1220648-2025-45915 | 22877664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22877664&pc=OZD | 07/29/2025 |  |  | 08/22/2025 |  |  | 07/29/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695560 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 84 KG | Non Hispanic | White | Malfunction | No | Yes | False Alarm | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was an automated impella controller (aic) issue. the aic had a placement signal low alarm. therefore, the impella pump was repositioned. after repositioning, "impella in aorta" and "impella in ventricle" alarms occurred. additionally, impella connect disconnected from this console in the middle of the night but the blue light stayed on. the aic was successfully exchanged for a new aic. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. refer to the related report number as noted in section h.10 for the report on the automated impella controller. | Product complaint # (b)(4). the device was not returned; therefore, an evaluation/analysis of the device was not possible. event logs were received and analysis. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical details noted that there were placement signal low alarms. after repositioning, there were impella in aorta/ventricle alarms. once the alarms resolved while the pump was in its original position, they switched the controllers. data log review showed an unreliable placement signal at this time with inconsistent decreases in snr and contrast. the reason for the unreliable placement signal is unknown. motor current modulation was also below the threshold for triggering these alarms as expected but the reason for the low motor current modulation is unknown. once the console was exchanged, there were no more positioning alarms until pump explant. the cause of the false alarms was not determined since product was not returned and insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller.
6494 | 1220648-2025-45817 | 22877820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22877820&pc=OZD | 07/29/2025 |  |  | 08/22/2025 |  |  | 07/29/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691748 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 failed after several attempts to advance through the patient's axillary artery. an impella cp was implanted instead. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | No product was returned. the root cause of the delivery issue was determined to be patient condition as the patient had stenosis in the aorta preventing successful delivery of impella. the root cause of the product damage issue was most likely catheter kink due to the patient anatomy. the pump passed all the post sterile inspection checks.
6495 | 1220648-2025-45816 | 22878120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22878120&pc=OZD | 07/28/2025 |  |  | 08/22/2025 |  |  | 07/29/2025 | 08/22/2025 |  |  |  |  | 08/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1618084 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered an automated impella controller (aic) issue outside of clinical use. a visual inspection of the area around the outlet revealed that the cord was torn and that the cause of the tear was unknown. | [5 empty] | The automated impella controller (aic) was received from the customer and an investigation regarding the returned device has been completed. the data logs are not relevant to this complaint. however, no issues were observed while going through the logs. the aic was returned for investigation. during visual inspection, it was observed that the live (black) wire within the plug was disconnected. it appeared that the wire covering or outer sheath was improperly cut, as per the vendor wiring instructions. also, the clamps were not tight enough, and light clamping marks were visible on the outer sheath, which likely contributed to the wire becoming loose and detaching from the plug and causing a spark. the supplier quality team was made aware of the complaint and the investigation findings. the reported issue of the cord inside the plastic plug socket being torn was determined to be caused by loose wires within the alternating current cord plug. | A.2 - a.6 were left blank as no patient was involved. e.1 initial reporter title, first name and last name are unknown at the time of writing this report. the automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6496 | 1220648-2025-45905 | 22859679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22859679&pc=OZD | 04/02/2025 |  |  | 08/21/2025 |  |  | 04/02/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582864 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with a second impella 5.5 for continuation in therapy and experienced bleeding requiring discontinuation of anticoagulation, device sensing issue and hematoma. the patient was initially implanted with an impella cp device for mechanical circulatory support and then escalated to an impella 5.5 device. on the 12th day of support, the hospital staff member tripped over a cord which caused a separation from the proximal end of the red plug. a pump stop occurred and emergent replacement for continuation in therapy was needed. now on the secondly placed impella 5.5, the patient was receiving chemotherapy and choked and had a hypoxic episode. once back on the unit, the patient coded and had to be intubated. approximately 43 days after starting support on the secondly placed impella 5.5 device, the patient developed a gastrointestinal bleed that required discontinuation of anticoagulation. it could not be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the outcome of the bleed for the patient. seven days later, anticoagulation was restarted. the next day, placement signal was noted intermittently and caused confusion. an echocardiogram was performed. the pump was 3.4 centimeters in left ventricular and advanced by .5 centimeters. the patient continued on impella support and now 8 days later, the patient was noted to have fully recovered neurologically and was ready for a durable left ventricular assist device, planned for the following week. notedly, the patient was not a candidate for an orthotopic heart transplant due to malignancy. impella support was continued. however, subsequently, heparin was stopped again due to an ¿old¿ hematoma at the impella site. assessed by the vascular team, there was a hematoma at the site but now there was a mass. there were no signs of active bleeding and, therefore, heparin was restarted again. the second implanted impella 5.5 device was eventually explanted with a survived patient outcome. | [5 empty] | Additional information was received and noted a new, related complaint involving the automated impella controller was determined to be associated with the impella 5.5 assigned to this medwatch report. refer to the no product return investigation results below for details on its association with the impella 5.5 device sensing issue. no product return investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding adverse events: retroperitoneal hemorrhage: gastrointestinal bleed reported during pump support, the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. access site adverse event: clinical details noted that the 12 days after pump explant, a hematoma at old impella site was observed. the hematoma at access was found to be a mass. the cause of the access site adverse event could not be definitively determined as limited clinical details were provided. regarding device malfunction, optical sensor issue: clinical details noted that the placement signal would intermittently go out. echocardiogram was performed and pump measured at 3.7cm and advanced 0.5cm. data logs were reviewed and a sudden increase in placement signal to 3000 on day 52 of support. optical parameters are consistent with oscillating contrast of the automated impella controller console. upon review of data logs, it is apparent that the console is the potential cause of the issue. no defect was found with the pump based on returned data log analysis. h6: investigation: type, findings and conclusion codes and h6: component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5 pump 1 and automated impella controller related to/cause of the optical sensor issue. | Refer to related manufacturer device report number 1220648-2025-26713 for the first impella 5.5 that had the pump stop. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6497 | 1220648-2025-45920 | 22860182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22860182&pc=OZD | 07/23/2025 |  |  | 08/21/2025 |  |  | 07/23/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684322 |  | 000000000001000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 71 KG | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella 5.5 had a small leak at the filter. the leak was monitored but no change to the purge pressure was noted. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Clinical details state that about a week into support, a small leak was reported at the sidearm filter. purge pressure was unchanged and there were no other purge issues reported during the case. a data log review is not applicable, and no product was returned. the cause of the purge leak was most likely a leak from the sidearm but the exact location is unknown since the device was not returned. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | Review of the data log review did not show any low purge pressure events. the cause of the purge leak was most likely a leak from the sidearm but the exact location is unknown since product was not returned. the data log review did not show any high/low purge pressure events. there were multiple cassette changes throughout the case but no significant rises in pp were observed prior to any change. the cause of the high purge pressure was not determined since product was not returned, data logs did not show high purge pressure, and insufficient clinical details were provided. the pump passed all post-sterile inspection checks. | Additional information added, med dev problem code a141102.
6498 | 1220648-2025-45923 | 22861454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22861454&pc=OZD | 07/27/2025 |  |  | 08/21/2025 |  |  | 07/27/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697058 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma at the access site and hematuria. the impella was measuring a little deep so it was repositioned, and systemic heparin was withheld. the patient was successfully weaned from the pump and survived. | [5 empty] | The root cause of the access site bleeding issue was most likely use related since the patient was thrashing around on the table and developed a hematoma at the impella site. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6499 | 1220648-2025-45785 | 22861473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22861473&pc=OZD | 07/28/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702466 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Remove | Patient Device Interaction Problem | [5 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that resistance was met during impella cp removal. the device was successfully removed with tension noted. upon explant, biomaterial was observed on the pigtail and dissection of the artery was discovered. the patient was taken to the operating room for vascular repair of the dissection. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The customer's pump was returned along with a small container containing a clot. data logs are not applicable. a visual inspection found no issues on the pump. the returned clot was confirmed. no calcium crystal were found on the pigtail or inside the pump housing. the cause of removal issue was unable to be determined since no issue was found on the product and insufficient evidence from clinical review. the cause of the thrombosis was unable to be determined since no issue was found on the product and insufficient evidence from clinical review. the device passed all post sterile inspection checks. | A.4 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
6500 | 1220648-2025-45862 | 22861584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22861584&pc=OZD | 08/02/2025 |  |  | 08/21/2025 |  |  | 08/02/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692538 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a patient developed a hematoma over the inguinal ligament. the pump was subsequently removed in response to the condition. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation has been completed. hematoma: the cause of the hematoma is most likely use related since after transfer to the intensive care unit, a massive hematoma developed over the inguinal ligament in a cranial direction. the device history review was completed and the pump passed all the post sterile inspection checks. | Sections a.2 through a.6 were left blank as the information is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6501 | 1220648-2025-45794 | 22861764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22861764&pc=OZD | 07/26/2025 |  |  | 08/21/2025 |  |  | 07/26/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607160 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 70 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp encountered multiple complications. the patient's urine was noted to be red. the positioning of the pump was checked and the device was found to be facing towards/against the mitral apparatus. the urine cleared several hours following repositioning. two days later, the patient began to bleed from their access site. manual pressure was held and three units of blood were transfused to treat the bleed. the dressing was adjusted and angle matching was achieved at the site to resolve the condition. after another two days of support, the pump was once again found to be directed towards the mitral valve. the doctor opted to leave the pump in the suboptimal position while decisions were made on the plans for the patient. due to the patient's underlying poor condition, the patient's family made the decision to withdraw care. the patient passed away following explant. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The cause of access site bleeding was not determined due to insufficient clinical details. the cause of hemolysis was most likely patient condition based on clinical, log review. the cause of positioning issue was not determined due to inconclusive evidence from clinical. the pump passed all the post sterile inspection checks. | A.5 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6502 | 1220648-2025-45812 | 22861862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22861862&pc=OZD | 07/28/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684957 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported a patient implanted with an impella cp experienced a purge system blocked alarm. the purge cassette was changed however, it did not resolve the issue. the pump was replaced in the catheterization lab. | [5 empty] | Medical safety review was received. initial pump placement resulting in high pressure was replaced without incident and the patient continued on the impella cp support until escalation. the patient would subsequently expire on the impella 5.5 and cannot be dissociated from the demise outcome. refer to section h.10 for the related report numbers. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device history review was completed and the pump passed all sterile inspection checks. the pump was returned for investigation. no visual damage was noted on the returned product. the purge gap was found to be blocked with biomaterial. the pump was primed, and high purge pressure was reproduced. the pump was not able to be started in the bucket of water test. further investigation revealed blood in the purge line after the catheter was cut near the motor during priming. also, the purge fluid was expelled from catheter end when the cut was made, which justifies no blockage in the purge line inside the catheter. data log showed an 8-minute gradual rise in purge pressure followed by saturated pump flow and a high purge pressure alarm. motor current was also seen trending upward after the high purge pressure event. data log and returned product investigation suggest biomaterial buildup occurring during the case, which later allows blood to ingress through the purge gap. high or saturated pump flow: the cause of the high purge pressure issue was biomaterial buildup based on returned product investigation and data log review. high purge pressure: the cause of the high/saturated pump flow issue was biomaterial buildup based on returned product investigation and data log review. | A.4 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
6503 | 1220648-2025-31177 | 22861936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22861936&pc=OZD | 07/24/2025 | 11/21/2025 |  | 08/21/2025 |  |  | 07/24/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1365450 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 65 KG | Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a battery failure (red alarm) when in use in the operating room. zero percent charge after 24 hours of being plugged in was noted. the automated impella controller was replaced and support was successfully placed for the impella 5.5 surgical patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs from the reported day of the event are consistent with the complaint. the logs show battery failure (transport connection blown/missing #360) alarm triggered just after the aic was turned on. device analysis: console ic6414 was not returned to field service for further investigation. root cause: the root cause of the battery failure (transport connection blown/missing #360) alarm was not determined since aic was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact facility name corrected. g1 reporting contact address line 1 should have been left blank.
6504 | 1220648-2025-31179 | 22862264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22862264&pc=OZD | 07/24/2025 | 08/21/2025 |  | 08/21/2025 |  |  | 07/24/2025 | 08/21/2025 |  |  | 09/30/2025 |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412919 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was placed to support a percutaneous coronary intervention (pci) procedure. after the pci, the impella cp was weaned slowly for removal. the patient tolerated it well. the companion sheath and impella cp were removed without issue. perclose had been placed pre procedure so the peel away sheath was re-wired for perclose. once the peel away sheath was removed, the first perclose failed to deploy. a second perclose also failed. the third perclose was lodged into the vessel and the internal suture was not able to be removed. manual pressure was held as the patient was bleeding around the perclose. emergency transport was ordered to transfer the patient to another facility for vascular repair. a fem stop was applied for transfer. the vascular medical doctor removed the perclose suture from the vessel and repaired the artery surgically. the patient's outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
6505 | 1220648-2025-45877 | 22862281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22862281&pc=OZD | 08/03/2025 |  |  | 08/21/2025 |  |  | 08/03/2025 | 08/21/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1292324 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced a system self-check failure. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: the console logs indicated a communication issue with the power battery manager (pbm). due to this communication problem, the console rebooted automatically (as designed) to attempt another power cycle. after rebooting, the console displayed the "system self check failed" screen. this confirms the reported failure mode. device analysis: this console was tested and visual inspection confirmed the pbm contamination. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. device history review: the device passed all post sterile inspection checks
6506 | 1220648-2025-45814 | 22862335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22862335&pc=OZD | 07/28/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694452 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was positioned deep and oriented toward the mitral valve. multiple unsuccessful attempts were made to reposition the pump. eventually, a new impella cp was implanted to resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The customer's device was not returned for review. based on the clinical description, the cause of the positioning issue was most likely use related since the pump was pulled in to the aorta while repositioning of the pump.
6507 | 1220648-2025-45784 | 22862602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22862602&pc=OZD | 07/27/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622675 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 94 KG | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the placement signal readings of an implanted impella 5.5 pump became irregular during patient support. it was noted that the waveforms and alarms were showing that the pump was in the aorta, despite the pump positioning being verified as correct via echocardiogram. the alarm was subsequently disabled in response to the persistent issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The customer reported 'impella position in aorta' alarms which did not match the position confirmed by echo. neither the customer's device nor data logs were returned. the cause of the optical signal issue was not determined as no product or logs were returned.
6508 | 1220648-2025-45759 | 22862711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22862711&pc=OZD | 07/28/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686600 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon advancement of the repositioning sheath following impella cp implant, bleeding and a hematoma developed at the access site. the pump was subsequently removed and replaced with a new pump. no patient harm was reported. | [5 empty] | A.4, a.5, and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The pump was not returned for investigation. a data log review was not required for this case run. as per the clinical report, bleeding and hematoma occurred after the advancement of the repositioning sheath, and the issue was resolved after replacement of the pump using a dry seal sheath. the cause of the access site bleeding was not determined without returned product investigation and sufficient clinical information was not provided. the pump passed all post sterile inspection checks.
6509 | 1220648-2025-31176 | 22862861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22862861&pc=OZD | 07/24/2025 | 09/22/2025 |  | 08/21/2025 |  |  | 07/24/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT EART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694522 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 78 KG | Non Hispanic | Asian | Death | Yes | Yes | Defective Device | Deformation Due to Compressive Stress | [5 empty] | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Melena | [8 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during the insertion, it was hard to insert into the graft and there was damage to the 23 french sheath. a new sheath was used successfully. additionally, the patient had a hematoma at the access site. pressure dressing was applied and one unit of blood products were given to the patient. additionally, the patient had spike in temperature. the patient was put on broad spectrum antibiotics. it is unknown if the patient's infection was impella related. furthermore, the patient had black stool. anticoagulation was withheld due to the patient's stool. support continued. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The introducer was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of access site bleeding, infection/sepsis, vascular damage - peripheral vessel, introducer damage - mechanical, and product damage (sheath kink)has been completed. only 23fr x 6cm introducer was returned for investigation. data log revie was not required for this case investigation. access site bleeding - major: as per the clinical, act was 267 at the time of bleeding. the patient was on anticoagulants. the cause of the access site bleeding issue was most likely high patient activated clotting time values. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessel: as per the clinical, patient taken off anticoagulants due to black stool. the cause of the vascular damage issue was not determined without sufficient clinical information. the relationship to gi bleed and impella was unknown. introducer damage - mechanical and product damage (sheath kink): the returned 23fr x 6cm introducer was kinked along the scoreline. no other damage was noted on the returned product. as per the clinical report, there was difficulty in inserting the introducer into the 8mm terumo vascutek gel weave graft. the cause of the introducer damage issue was introducer sheath kink along the scoreline. the cause of the product damage (sheath kink) issue was most likely use-related since the doctor used the 8mm terumo vascutek gel weave graft. the abiomed recommended graft size was 10mm. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31176 b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31176. h1 type of reportable event was erroneously selected on the initial manufacturer device report number 1220648-2025-31176. h6 health effect - clinical code 1888 and health effect - impact code 4641 and 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31176.
6510 | 1220648-2025-31178 | 22863017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22863017&pc=OZD | 07/24/2025 | 08/21/2025 |  | 08/21/2025 |  |  | 07/24/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674414 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 83 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was placed to support a percutaneous coronary intervention (pci) procedure. after the pci, the patient was taken to the intensive care unit to recover. the patient experienced hemolysis and the laboratory was unable to process labs due to gross hemolysis. the impella was in optimal position. it was further reported that the patient expired after 72 hours of impella support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6511 | 1220648-2025-31180 | 22863398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22863398&pc=OZD | 07/24/2025 | 08/21/2025 |  | 08/21/2025 |  |  | 07/24/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680647 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was placed without issue though the short sheath. the pump was started on auto. approximately three minutes later, the medical doctor found the groin site bleeding profusely. the impella cp and sheath were explanted immediately. manual pressure was held. mass transfusion was ordered and delivered though a belmont. an angiogram revealed the side branch of the femoral artery was perforated. the patients outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
6512 | 1220648-2025-45781 | 22864107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22864107&pc=OZD | 07/27/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723012 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 61 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that some complications were encountered during impella cp support. following implant, the pump was found to be shallow and in the papillary muscle. the cardiologist at the time refused adjustments to the positioning and support continued. the following day, the support level was dropped due to suction and hemolysis. an attempt was then made to reposition the pump, but the pump was inadvertently pulled out of the left ventricle. a thrombus was then discovered in the left ventricle and the decision was made to remove the pump. the patient survived. | [5 empty] | No product was returned for evaluation. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the low pump flow could not be definitively determined as limited clinical details were provided at time of reported suction and no product was returned for evaluation. the cause of the hemolysis could not be definitively determined as limited clinical details were provided and no product was returned for evaluation. the root cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of lv during pump repositioning. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
6513 | 1220648-2025-31174 | 22864698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22864698&pc=OZD | 06/25/2025 | 11/24/2025 |  | 08/21/2025 |  |  | 08/11/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1478138 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient that underwent chronic total occlusion, percutaneous coronary intervention (pci), mid- om, presented for a high-risk percutaneous coronary intervention with impella cp assist. the impella was inserted using best practice. appropriate flows, placement signal not reliable, alarm noted at the start of case. motor current normal ranges. single access pci preformed utilizing companion sheath. patient was weaned hypotension was noted, right heart catheterization done, physician decided continuation of weaning, impella explanted using best practice. hemostasis achieved. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. analysis: cp pump 539746 was run on (b)(6) on 6/25. the duration of the time with the pump connected confirmed the complaint as the rt logs show the placement signal not reliable alarm as the placement signal jumped to 3000mmhg and optical sensor oscillated to -3276.8 mmhg and snr <1000 ((b)(6) rtlog 539746.png). the previous pump run on this console also had <1000 snr though the placement signal performed as expected. the console was not returned for further analysis. root cause: the cause of the optical signal issue could not be determined as the console was not returned for investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d4 unique identifier (udi) # updated. d10 concomitant medical products and therapy dates updated.
6514 | 1220648-2025-45825 | 22864817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22864817&pc=OZD | 07/28/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026699019 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 104 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted as the pump could not cross the atrioventricular with 5.5. an impella cp was successfully implanted instead. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The root cause of the delivery issue was determined to be patient condition as the patient had stenosis preventing successful delivery of impella. the pump passed all post sterile inspection checks.
6515 | 1220648-2025-45883 | 22865549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22865549&pc=OZD | 10/18/2024 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526825 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Anemia | Hemolysis | Respiratory Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information updated with new adverse event for acute respiratory failure with hypoxia. the relationship is listed as possibly related. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening anemia with a "possible" relationship to the impella device use. additionally, it was noted that the patient endured worsening hemolysis with a "probable" relationship to the impella device. no concern for active bleeding but this lead to impella removal and possible intra-aortic balloon pump insertion. the issue did not resolve. there was also a call regarding a placement signal low alarm. the team was advised to perform an echocardiogram and reposition the device. ultimately, care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Product was not returned for investigation, only event logs were received. data log confirms several placement signal low alarms with placement signal dropping from 80/40 to 40/0 for about an hour on 16-oct. motor current and pump flow showed some fluctuations and were found to slight decrease after the placement signal issue was resolved. snr dropped below 1500 and contrast was drop below 10, during the optical signal issue without affecting other 5s optical parameters. no placement signal low alarm issue was noted later during the case run. data logs also suggest gradual decrease in placement signal till second day of implant, without any noted major impact on motor current and pump flow. as per the clinical details, a low placement signal alarm was reported on 16-oct, with the pump position marked at 84 cm. this was a 3 cm difference compared to the reported position before and after the update in the patient report. additionally, at the time of the low placement signal, the cvp was reported to have decreased to 10 cm from the pump placement time, with reported rv dysfunction and lv offloading condition. the cause of the hemolysis issue was not determined since the product was not returned for investigation and the data log was inconclusive without sufficient clinical information. | A.4, a.5 and a.6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6516 | 1220648-2025-45931 | 22865670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22865670&pc=OZD | 07/29/2025 |  |  | 08/21/2025 |  |  | 07/29/2025 | 08/21/2025 |  |  | 08/31/2025 |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390881 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 78 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced pulmonary edema and ventricular tachycardia. the p level of the pump was increased and the patient was treated with a procainamide drip. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the pulmonary edema was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6517 | 1220648-2025-45789 | 22865889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22865889&pc=OZD | 07/29/2025 |  |  | 08/21/2025 |  |  | 07/29/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668829 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Female | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a thrombus was discovered in the left ventricle during impella cp support. the thrombus was noted to be floating and attached to the pigtail of the impella. there were no noted effects on the functionality of the pump and support continued uninterrupted. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The pump was returned for investigation. no issues were found with the returned product. the root cause of the injury was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
6518 | 1220648-2025-45851 | 22866607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22866607&pc=OZD | 07/28/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678487 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 complained of new acute pain and numbness in right upper extremity. an ultrasound scan showed a thrombus in the right brachial artery. the decision was made to remove the pump to prevent further complications. the patient survived. | [5 empty] | The root cause of the throbosis was unable to be determined due to insufficient clinical details. the cause of the reported pain cannot be determined as insufficient clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿.
6519 | 1220648-2025-45895 | 22866621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22866621&pc=OZD | 08/05/2025 |  |  | 08/21/2025 |  |  | 08/05/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025602083 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 77 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 endured runs of ventricular tachycardia. an ablation was planned to treat the condition, however it was later reported that the patient passed away due to a massive cardiac arrhythmia. the patient was on support at the time of their passing | [5 empty] | No product was returned for investigation. the root cause of the arrhythmia/tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | There is not enough evidence to exclude the impella as an associated factor in the patient's outcome. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6520 | 1220648-2025-45897 | 22867158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22867158&pc=OZD | 08/06/2025 |  |  | 08/21/2025 |  |  | 08/06/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692313 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient was previously cannulated for extracorporeal membrane oxygenation (ecmo) and was escalated to the impella 5.5 from an impella cp device. it was reported that an optical sensor error appeared during impella 5.5 support. it was noted that there was no placement signal at the time of the reported error. the device was explanted as a result of the failure. the patient continued on ecmo support, noted to be in critical condition and outcome of survived at time of explant of impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6521 | 1220648-2025-45813 | 22868575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22868575&pc=OZD | 07/28/2025 |  |  | 08/21/2025 |  |  | 07/28/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701800 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Rupture | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was escalated from an impella cp to an impella 5.5. after the impella 5.5 was placed and impella cp was explanted, it was identified that the patient had a ruptured papillary muscle with hyper dynamic mitral valve. the impella 5.5 was repositioned at the bedside. the impella was weaned successfully concurrently with increasing extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks.
6522 | 1220648-2025-45815 | 22868632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22868632&pc=OZD | 07/29/2025 |  |  | 08/21/2025 |  |  | 07/29/2025 | 08/21/2025 |  |  |  |  | 08/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657993 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma of the right femoral artery. the hematoma was treated with a femstop and manual compression. additionally, the patient's platelet count was down to 15 and there were concerns of heparin-induced thrombocytopenia. impella support continues. | [5 empty] | The device was not returned for investigation. the root cause of bleeding was determined to be use issue because the hematoma was formed during patient transportation. the cause of the thrombocytopenia was not determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6523 | 1220648-2025-45745 | 22848814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22848814&pc=OZD | 07/26/2025 |  |  | 08/20/2025 |  |  | 07/26/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704457 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 69 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced a hemorrhagic brain bleed. no noted intervention was required for the bleed. the patient's neurology status continued to decline and the patient passed away the following day. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The cusotmer's device was not returned for investigation. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. the device with passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6524 | 1220648-2025-45776 | 22848891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22848891&pc=OZD | 06/06/2025 | 08/20/2025 |  | 08/20/2025 |  |  | 06/06/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671954 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and was found to have finding of apical left ventricle thrombus. the patient was initially implanted with an impella cp device for hemodynamic stability and was escalated to the impella 5.5 device as a bridge to transplant. however approximately three days after start of the impella 5.5 support, an echocardiogram revealed a finding of apical left ventricle thrombus. it was noted physicians were aware of contraindication and potential risk; however, the decision was made to leave patient on impella mcs. the patient was noted to be stable following the event and remained on impella mcs or an additional 43 days before being bridged to transplant. | [5 empty] | The device was not returned, and therefore an evaluation/analysis of the device was not possible. device history record review was completed, and the device passed all sterile inspection checks. the investigation concluded in order to make a root cause determination on the thrombus, all essential clinical details are required. therefore, the root cause of the injury was unable to be determined due to insufficient clinical details. | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: potential adverse events- ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6525 | 1220648-2025-45747 | 22849188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22849188&pc=OZD | 07/26/2025 |  |  | 08/20/2025 |  |  | 07/26/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689278 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 128 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implantation of an impella cp, a hematoma developed at the left femoral artery site. manual pressure was applied. two liters of normal saline was used. three units of packed red blood cells and five units of albumin were transfused to manage the condition. shortly after, the patient was returned to the catheterization lab as their impella leg was mottled with no distal pulse. the pump was explanted and a fasciotomy was performed in response to the ischemia. | [5 empty] | The root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | A.5 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this file represents the pump. the related manufacturer device report number 1220648-2025-45746 represents the introducer.
6526 | 1220648-2025-45762 | 22849286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22849286&pc=OZD | 07/27/2025 |  |  | 08/20/2025 |  |  | 07/27/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686565 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a hematoma developed, and oozing was noted around the sheath. manual pressure was held in an attempt to manage the condition, but the hematoma continued to grow, and the decision was ultimately made to explant the pump which was completed in the intensive care unit. however, the patient was taken back to the catheterization lab for investigation of the femoral access where the physician went up and over to image the right femoral artery. a balloon was inflated in the right iliac and dry closure was completed to the access site with perclose. post this intervention, the access site was noted clean, dry, no bleeding or hematoma. | [5 empty] | E.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. the impella device was not returned, and the investigation has been completed. the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. the device history review was completed and the device passed all post sterile inspection checks. | Device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follow for the impella cp with smartassist system: section: general patient care considerations - ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer-¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events- ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6527 | 1220648-2025-45772 | 22851117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22851117&pc=OZD | 07/27/2025 |  |  | 08/20/2025 |  |  | 07/27/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677255 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 78 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella 5.5 implant, intermittent suction alarms and suboptimal pump flows were encountered. the physician's attempts to reposition the device were unsuccessful and the catheter subsequently kinked near the locking port. the pump could no longer be advanced and the decision was made to replace the device. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The pump was not returned for investigation. the cause of the pump suction issues was not determined as product was was returned, data logs were not recovered, and insufficient clinical details. the cause of the pump flow issues was not determined as product was was returned, data logs were not recovered, and insufficient clinical details. the cause of the cannula kink was not determined as product was was returned, data logs were not recovered, and insufficient clinical details. the cause of the positioning issues was not established as clinical does not mention details related to use or patient anatomy. the root cause of the positioning issue was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks.
6528 | 1220648-2025-45779 | 22851351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22851351&pc=OZD | 07/28/2025 |  |  | 08/20/2025 |  |  | 07/28/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679934 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 61 KG | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery. despite successfully wiring the left ventricle, it was documented that the outlet and motor were unable to make the acute turn in the aortic arch. after multiple failed attempts, the implant was aborted. | [5 empty] | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The root cause of the delivery issue was determined to be patient condition as the patient had tortuous anatomy preventing successful delivery of the impella. the pump passed all post sterile inspection checks.
6529 | 1220648-2025-45769 | 22851361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22851361&pc=OZD | 07/27/2025 |  |  | 08/20/2025 |  |  | 07/27/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670585 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced some oozing from their impella cp access site during and following implantation. the access site was previously used for an intra-aortic balloon pump and the bleeding was present prior to implantation of the impella cp. once implanted, the groin continued to ooze until the angle match positioning was adjusted. as support continued, the patient developed hematuria. the p level was reduced and the pump was found to be shallow. the pump was repositioning and the hematuria cleared. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The pump was not returned for investigation. the pump passed all post sterile inspection checks. the cause of access site bleeding was due to a use related issue based on clinical information. the cause of the mechanical interaction with blood was not determined since the product was not returned and inconclusive log and clinical review. the cause of positioning issues were not determined due to insufficient clinical details.
6530 | 1220648-2025-45914 | 22851476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22851476&pc=OZD | 07/28/2025 |  |  | 08/20/2025 |  |  | 07/28/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684967 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 103 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day exhibited bleeding at the femoral insertion site. the bleeding was treated with a suture. the patient's blood pressure was low so epinephrine and vaso were given. additionally, it was noted there were no additional suction events since the prior late night. the patient was noted to still be a candidate for a transcatheter aortic valve replacement after recovery from impella event. the plan was for the patient to rest, recover and team monitors the impella insertion site. however later this day in evening, the patient expired. care was ultimately withdrawn. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for analysis. the pump passed all post sterile inspection checks. the data logs confirm the reported suction alarms. the logs do not show any motor current spikes or positioning issues. the cause of the issue was not determined as no product was returned, no definitive log abnormalities were observed and limited clinical information was provided. the cause of the access site bleeding was likely use-related as evidenced by bleeding resolving after addition of sutures. the cause of the hypotension was not determined due to insufficient clinical details.
6531 | 1220648-2025-45917 | 22851531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22851531&pc=OZD | 07/25/2025 | 08/13/2025 |  | 08/20/2025 |  |  | 07/26/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637941 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 107 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 as a bridge to transplant. while on support, a hematoma was reported at the site of the impella insertion. additionally, the echocardiogram showed the impella 5.5 was pointing to the mitral valve; there is no information at hand that indicates actions outside of troubleshooting and compromise to patient's support. ongoing intermittent suction alarms were noted with no stated resolution. there was some bleeding noted at the surgical site. anticoagulants were paused. the impella dressing was changed frequently with oozing noted. hematoma was noted as stable. one unit of blood products was transfused. the patient remains on impella support on the date of this report. | [5 empty] | The investigation was completed. no product returned. the cause of the issue was not determined as limited clinical information was provided. no data logs were returned. the cause of the issue was not determined as no data logs or product was returned and limited clinical information was provided. a device history reviewed was completed and passed all post sterile inspection checks. | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6532 | 1220648-2025-45771 | 22851550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22851550&pc=OZD | 07/27/2025 |  |  | 08/20/2025 |  |  | 07/27/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703394 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient suffered from a stroke during impella 5.5 support. the stroke was discovered via computed tomography scan after the patient failed to wake from sedation. it was additionally noted that a blood clot was found in the left ventricle. anticoagulants were withheld as a precaution due to concerns of a cerebral hemorrhage and the decision was ultimately made to remove the pump. the patient remained supported via extracorporeal membrane oxygenation. | [5 empty] | The root cause of the reported stroke was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
6533 | 1220648-2025-45783 | 22851825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22851825&pc=OZD | 07/28/2025 |  |  | 08/20/2025 |  |  | 07/28/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508936 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 61 KG | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. despite successfully wiring the left ventricle, it was documented that the outlet and motor were unable to make the acute turn in the aortic arch. after multiple failed attempts to implant the pump, the impella 5.5 implant procedure was aborted. | [5 empty] | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as resistance was met at an acute turn at the aortic arch preventing successful delivery of impella. the pump passed all post-sterile inspection checks.
6534 | 1220648-2025-45788 | 22852242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22852242&pc=OZD | 07/28/2025 |  |  | 08/20/2025 |  |  | 07/28/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669092 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hematuria | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following foley placement in the coronary care unit, an impella cp patient's urine became dark red. it was noted that intermittent suction alarms were occurring leading up to the noted condition. despite repositioning the pump, the issue persisted and the patient's hemodynamic pressures continued to drop. the patient passed away on support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | E.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. the customer's device was not returned for investigation but event logs were available for review. a device history review was completed and the pump passed all post-sterile inspection checks. analysis summary: the clinical details reported that the patient had a foley placed and had dark red urine output. additionally, intermittent suction alarms correlated with reduction in levophed drops. the pump was confirmed to be in proper position via fluoroscopy, and no blood products were infused. early that morning, the medical doctor made attempts to reposition the pump, but suction did not resolve, and the patient expired on support. pump suction: data logs were reviewed and the intermittent suction was confirmed. there were no motor current (mc) spikes and no positioning alarms, supporting the clinical report that the pump was in adequate position. there were 111 triggers of suction, both intermittent and continuous, throughout support. it was also noted that there were several severe trends down in the placement signal (ps) during the case. this correlates with the clinical detail provided that the patient's pressure would drop when levophed drops were attempted to be weaned. during the first major trend down in ps, there was a clear drop in diastolic mc while at p-9, followed by the "impella flow reduced" algorithm activating, and suction alarms. this all indicates that the patient's condition had deteriorated. the version of suction that triggered the most during the case was the pulsatile algorithm (98 occurrences), which triggers when the mcmax value is below a threshold for the p-level. this indicates that in addition to low systemic pressures, the pump did not have sufficient volume to flow well even during systole. the product was not returned for investigation. the clinical details provided in conjunction with the trends seen on data logs indicate that the cause of the pump suction was patient condition related. hemolysis/hematuria: as detailed above, there were many triggers of suction alarms throughout the case, particularly continuous suction alarms, indicating that the pump did not have sufficient volume for typical mcmax values intermittently throughout the case. data logs did not have any positioning alarms, supporting that the pump was in proper position. there were also no mc spikes to indicate an ingestion event was causing suction. finally, the significant trends down in the ps support the clinical report that the patient was heavily reliant on the levophed drop to maintain systemic pressures. the product was not returned for investigation. the clinical details provided in conjunction with the trends seen on data logs indicate that the cause of the hemolysis was patient condition related. | A.4 and a.5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6535 | 1220648-2025-45902 | 22852964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22852964&pc=OZD | 07/29/2025 |  |  | 08/20/2025 |  |  | 07/29/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694909 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support, and approximately 18 days after start of the support, it was reported that an impella 5.5 showed incorrect blood pressure readings. no patient harm was reported. the patient continued on impella mcs for an additional 15 days before successful wean and device explant with a survived outcome. the patient received a heart transplant that was noted as "successful." | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The customer's device was returned in a damaged condition, making some device testing impossible. there was biomaterial in the outflow/inflow cages and cannula. data logs were reviewed and the placement signal showed a "linear" trend down from the start of the case. since the trends seen in data logs do not follow any known failure patterns and returned product was unable to be tested, the cause of the optical signal issue could not be determined.
6536 | 1220648-2025-45792 | 22853144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22853144&pc=OZD | 07/28/2025 |  |  | 08/20/2025 |  |  | 07/28/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665992 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced oozing from the groin site. systemic heparin was paused to manage the condition. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The device passed all post sterile inspection checks. the root cause of access site bleeding was most likely use issue related since the partial thromboplastin time was elevated and hemostasis was achieved after pausing the heparin. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6537 | 1220648-2025-45836 | 22853279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22853279&pc=OZD | 06/27/2025 |  |  | 08/20/2025 |  |  | 06/27/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681790 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a loqi database reported upon adverse event of impella site right femoral hematoma. the relationship is listed as "possibly" related. post cath, site oozed with hematoma. site was injected with epinephrine and lidocaine. hematoma was reduced and resolved upon discharge additionally, it was noted that the impella was weaned without issue and explanted. however, 1 perclose broken during tightening during procedure. hemostasis was achieved with single perclose, angioseal, manual compression. groin clean dry and intact no issues. | [5 empty] | The root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of removal issue is determined to be use issue because of excessive force used during tightening of perclose. the pump passed all the post sterile inspection checks. | Device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for the impella cp with smartassist system: section general patient care considerations - ¿assess access site for bleeding and hematoma.¿ section entering cardiac output use of the repositioning sheath and peel-away introducer - ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6538 | 1220648-2025-45888 | 22853709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22853709&pc=OZD | 07/01/2025 |  |  | 08/20/2025 |  |  | 07/01/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680865 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male | 62 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Low Blood Pressure/ Hypotension | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced complications requiring intervention. the device was implanted for approximately 55 days while waiting for a heart transplant surgery. four days prior to the impella 5.5 device implant, the patient was admitted with increased weakness, fatigue, nausea, vomiting, and weight loss over the last couple of months. on arrival to the emergency department, the patient was tachycardic and hypotensive and was started on an epinephrine 0.03 drip. an echocardiogram was completed, and plans were made for right heart catheterization and innominate impella 5.5 insertion which were successfully completed. of note, an impella rp device was available. however, the decision was made for an impella 5.5 insertion only based on the right heart catheterization. approximately 11 days after start of impella support, the patient was positive for heparin-induced thrombocytopenia. purge solution was switched to bicarbonate and systemic heparin was switched to bivalirudin. impella mcs continued. approximately 18 days later, the patient was taken to catheterization lab for a central venous catheter for therapeutic plasma exchange. notedly, during general anesthesia induction, there was right ventricular dysfunction. the impella 5.5 was also repositioned this day due to suctioning events. impella position was then verified, and no active alarms were present. the patient would continue mcs awaiting a heart transplant. on day 53 of impella 5.5 support, the patient experienced some orthostatic hypotension. the impella 5.5 support level was increased to p-8. prior (two days earlier), the patient had a slight fever and broad-spectrum antibiotics were given. blood cultures were negative, with no signs of infection. two days later, the patient underwent a successful heart transplant and impella 5.5 removal with a surviving patient outcome. the patient was extubated later this same day, and epinephrine was weaned off the following day. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding adverse events thrombocytopenia, pain, fever, and hypotension, the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. h6: investigation: type, findings and conclusion codes and h6: component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6539 | 1220648-2025-45893 | 22853888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22853888&pc=OZD | 07/25/2025 |  |  | 08/20/2025 |  |  | 07/25/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635466 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Failure to Advance | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues and guidewire product damage. after implantation of an impella cp, the pigtail appeared rotated and the impella cp appeared kinked between the motor and the bend of the cannula. the physician attempted to reposition the device but was unsuccessful. as a result, the impella cp was removed through the peel away sheath and the 0.018 guidewire was damaged during the process. a 0.014 300 centimeter whisper wire was used to successfully implant the pump. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this file represents the pump. the related manufacturer device report number 1220648-2025-45884 represents the guidewire. | The returned product confirmed cannula kink no patient anatomy complications were noted. the cause of the positioning issue was likely a use related issue. the cause of the cannula kink was likely use related as issue did not reproduce on second attempt at pump delivery and no patient issues were noted. the pump passed all post sterile inspection checks.
6540 | 1220648-2025-45899 | 22854120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22854120&pc=OZD | 07/07/2025 |  |  | 08/20/2025 |  |  | 07/07/2025 | 08/20/2025 |  |  |  |  | 08/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691872 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 126 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp had a brain bleed, high purge pressure and an optical signal issue. the patient had a computed tomography done which showed a brain bleed that worsened during impella support. no anti-coagulation was used during support. additionally, the pump showed issues with the optical signal and there was intermittent high purge pressure. no further intervention was noted and the patient remained stable. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The product was not returned. event logs were returned. due to a lack of sufficient clinical details provided and no product returned, the root cause of the optical signal issue cannot be determined. due to lack of product returned, the root cause of the high purge pressure cannot be determined. the root cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
6541 | 1220648-2025-45900 | 22854345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22854345&pc=OZD | 07/07/2025 |  |  | 08/20/2025 |  |  | 07/07/2025 | 08/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691749 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received: computed tomography angiography demonstrated a severe middle cerebral artery infarct with a low likelihood of recovery. the patient was in atrial fibrillation and required multiple vasopressors. | The complainant reported that a patient was implanted with an impella 5.5 and experienced a stroke. support continued until the patient was successfully weaned from the pump. the patient survived. |  |  |  |  | B5 additional information received and h6 codes were updated. | The investigation was completed. the root cause of the stroke was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6542 | 1220648-2025-45736 | 22837959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22837959&pc=OZD | 07/25/2025 |  |  | 08/19/2025 |  |  | 07/25/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1266277 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during preparation for support, an automated impella controller (aic) generated a blue screen error when the console failed system check. another aic was used to successfully complete the patient procedure. | [5 empty] | The customer's device was returned for evaluation. the device's console logs were reviewed and confirm the complaint. the reported system self-check failure was able to be reproduced. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. a visual inspection of the console's internal components confirmed that there was power battery manager contamination. the root cause of the ¿system self check failure¿ was determined to be power battery manager corrosion due to leakage of the electrolyte from the battery failure alarm super caps. the power battery manager was replaced, a full functional check on the console and optical system was performed, and the device was returned to the customer for use. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
6543 | 1220648-2025-45890 | 22838450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22838450&pc=OZD | 07/20/2025 |  |  | 08/19/2025 |  |  | 07/20/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700119 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 64 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient's hemoglobin levels dropped from 12 at baseline to 7.3 post-op during impella 5.5 support. the patient was transfused blood in response to the drop in hemoglobin levels. it was noted that the patient had a history of chronic anemia leading up to the event. it is unknown if impella support contributed to the condition. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the anemia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | Additional information provided in h6.
6544 | 22839190 | 22839190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22839190&pc=OZD | 08/06/2025 | 08/08/2025 | 08/08/2025 | 08/19/2025 | 08/08/2025 | 08/19/2025 |  | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | OZD | [5 empty] | ABIOMED ,INC |  |  | US | [5 empty] | Device was returned to manufacturer |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [6 empty] | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella purge cassette tubing was leaking purge fluid from the mounting site. the primary rn identified it and replaced the tubing. there was no harm to the patient. impella rep [redacted] collected the damaged tubing and took with her.
6545 | 1220648-2025-45828 | 22839577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22839577&pc=OZD | 07/26/2025 |  |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133094 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the placement signal disappeared during impella cp support. a placement signal unreliable alarm appeared as a result of the failure. the audio was subsequently disabled and the decision was made to continue with support. the staff was made aware of signs that would indicate the pump may need to be replaced. upon additional investigation, it was discovered that the automated impella controller may have contributed to the failure via a potential defective lamp. | [5 empty] | Product complaint # (b)(4). investigation summary ==> data analysis: review of the logs confirms the complaint. #1036 calculator placement signal error followed two minutes later by a ¿placement signal not reliable¿ alarm (see complaintalarm.jpg). placement monitoring is disabled by a user approximately 6 minutes later. the log also contains optical bench 5s parameters with values which suggest a failed bulb (see 5s_logs.jpg). the lt and rt logs show a sudden decrease in snr, gain, and light at the time of the complaint (see rt_log.jpg and lt_log.jpg). contrast values are oscillating between approximately +/- 3000 (ltlog_osccontrast.png). device analysis: when the console is powered on there is an immediate ¿controller error¿ alarm, and there is no light seen coming from the optical fiber at the center of the blue pump connector (see displayalarm.jpg and pumpconnectornolight.jpg). the optical bench 5s parameters are tested and show very low light and snr measurements, along with a gain of 0.92 (see bench5stesting.jpg). the console was opened and the optical bench was inspected for loose connections, damaged fibers, etc. with none found. the voltage powering the optical bench was measured at 4.936v. root cause: the placement signal not reliable alarm was the result of oscillating contrast due to a failed or failing bulb in the optical bench lamp assembly. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report represents the automated impella controller. the related manufacturer device report number 1220648-2025-45744 represents the impella pump. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6546 | 1220648-2025-45749 | 22840286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840286&pc=OZD | 07/26/2025 |  |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1763601 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered a "placement signal low" alarm on an automated impella controller (aic). the impella was confirmed to be in proper position and the patient's hemodynamics did not correlate with the alarm. as a result, the audio was disabled for the alarm. impella support continued. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D.6a. and d.6b. implant and explant dates should not have been reported on the first follow up report that was submitted. e.1. added zip code extension as it was inadvertently omitted from the initial report that was submitted. h.10. added related report number. the investigation has bee completed. the console and data logs were returned for investigation. the data logs confirm the reported issue. a visual inspection of the internal circuitry and cable connections of the console did not reveal any issues. the console was tested with a known good pump and purge cassette, and no anomalies were observed during testing. the root cause of the " placement signal low" alarms was not related to any issue with the automated impella controller (aic). no issues were found with the aic. | The automated impella controller was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6547 | 1220648-2025-45774 | 22840302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840302&pc=PYX | 07/27/2025 |  |  | 08/19/2025 |  |  | 07/27/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625109 |  | 1000323 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 124 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient inadvertently pulled an implanted impella rp flex pump out of their body in their sleep. manual pressure was held over the site. the decision was made to explant the device. no patient harm was reported due to the issue. | [5 empty] | D.4 added model as it was inadvertently omitted from the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. e.4 added selection as it was inadvertently omitted from the initial report that was submitted. h.6 codes b13, b18, g04105 and c19 were reported incorrectly on the initial report that was submitted. new codes were added. h.10 the additional manufacturer on the initial report that was submitted should of stated that the device was not returned versus discarded. the device history review was added as it was inadvertently omitted from the initial report that was submitted. the device history review was completed and the pump passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. according to the clinical details, the pump was accidentally pulled out during sleep. the root cause of the positioning issue was determined to be an unintentional use error, as the pump was accidentally pulled out of position while the patient was asleep.
6548 | 1220648-2025-45824 | 22840360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840360&pc=OZD | 07/26/2025 |  |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670702 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 102 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp, the patient developed ventricular tachycardia which quickly deteriorated to ventricular fibrillation. the patient was successfully defibrillated. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
6549 | 1220648-2025-31163 | 22840586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840586&pc=OZD | 07/21/2025 | 12/04/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1610582 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 87 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a console had an alarm but the nurse was unaware of the cause or significance as it resolved on its own before they could assess. through review of the impella connect, it was revealed that there was a first placement signal loss, followed by a yellow controller error alarm that resolved less than two minutes later. the next day, a placement signal disruption was witnessed which was corrected before any alarm was triggered. this second known incident resolved in about two minutes from initiation. the implanting physician and intensivist were made aware of these incidents and agreed to swap the controllers. the patient continued on support. | [5 empty] | Data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm triggered during the clinical procedure. additionally, error message "invalid sample due to ambient pressure out-of-range" and "algs suspended opm crc error" was also seen on (b)(6) 2025 at 08:45:29. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the crc error. device analysis: the console was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. additionally, upon inspecting the analog output of the ccd from the optical bench, intermittent distortions in the signal (envelope/fringe) were observed, while the measured "5s" parameters, were within the specified value outlined in the pm procedure (not related to the reported complaint). root cause: the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" was due to a firmware issue (optical bench fw anomaly). h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of two related reports. this report addresses the console. the pump is referenced under complaint (b)(4).
6550 | 1220648-2025-45740 | 22840603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840603&pc=OZD | 07/25/2025 |  |  | 08/19/2025 |  |  | 07/25/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655185 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female | 86 KG | Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon attempted delivery of an impella cp pump, both the pump and introducer kinked. the introducer was replaced, but the second introducer encountered the same problem. a gore sheath was subsequently placed and the pump was successfully placed for support. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation was completed. the pump reportedly kinked when attempting to pass through the first sheath kink. the cause of the cannula kink was not determined since the product was not returned. e.1 zip code extension was added as it was inadvertently omitted from the initial manufacturer report. e.4 unknown should of been selected on the initial report. therefore, it was added to this report. h.10 related report number for second introducer was added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report represents the impella pump. the related manufacturer device report number 1220648-2025-45741 represents introducer.
6551 | 1220648-2025-31164 | 22840701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840701&pc=OZD | 07/22/2025 | 09/19/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  | White | Death | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | Insufficient Information | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to replace a prior impella cp that was explanted with clot concerns. this pump was supporting a cardiogenic shock/cardiomyopathy patient when there was a purge leak and high purge pressure/low purge flow alarms. the pump remained on despite the purge issues for another two days when the patient had care withdrawn after just under nine days. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump and high purge pressure (hpp) has been completed. the impella device was not returned for evaluation. data log review showed a 45 minute decrease in purge pressure to 100 mmhg that was sustained for about two hours on (b)(6) 2025. after the two hours, purge pressure rose rapidly within a few minutes as described in the clinical details when the bypass was attempted. there was sustained hpp for 1.5 hours before a bag change (potentially when the tpa was added). pp then gradually decreased to 600 mmhg for the next 20-30 hours. care was withdrawn shortly after. there were no motor current spikes observed prior to the increases in pp. purge leak - pump: the cause of the purge leak was a leak from the sidearm but the exact location of the leak was unknown since product was not returned. high purge pressure: the cause of the high purge pressure was most likely biomaterial due to the rise in pp after two hours of 100 mmhg purge pressure but the type of biomaterial interaction is unknown.
6552 | 1220648-2025-45823 | 22840721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840721&pc=OZD | 07/26/2025 | 08/19/2025 |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678012 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 as a bridge to transplant. while on support, the patient received volume due to suction alarms, which resolved. an echocardiogram was obtained and there was an attempt to reposition the impella 5.5 to increase the flow. after changing the impella 5.5¿s position and verifying correct placement, flows did not increase above p4/p5 without suction alarms. support continued uninterrupted despite the reduced support level. | [5 empty] | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history review confirmed the device passed all the post sterile inspection checks. regarding, pump suction, volume was given along with the position verified but did not resolve. after reviewing the logs, logs showed noisy low placement signal, continuous suction and continuous position unknown alarms. the cause of the pump suction was most likely patient condition since the event logs confirmed low flow or suction along with placement signal trending down. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation.
6553 | 1220648-2025-31166 | 22840795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840795&pc=OZD | 07/21/2025 | 08/19/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694576 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and had cognitive decline on day four of impella support. a head computed tomography was performed and confirmed a subarachnoid hemorrhage. a decision was made to explant the impella 5.5 and discontinue anticoagulation. the impella 5.5 was explanted the next day and the patient's outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6554 | 1220648-2025-45833 | 22840816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840816&pc=OZD | 07/30/2025 |  |  | 08/19/2025 |  |  | 07/30/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640237 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 77 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced persistent suction issues coinciding with a decoupled placement signal. the pump position was verified via echocardiogram.there was adequate volume and right heart function. the alarm was silenced and support continued until the patient was successfully weaned from the pump. | [5 empty] | The device was not returned for investigation. data logs were reviewed and the cause of the pump suction couldn't be determined. there was no problem identified with the optical signal. the passed all post-sterile inspection checks. | The impella device was not received from the customer, but the event logs were received. therefore, the investigation of the device is not possible, but event logs evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
6555 | 1220648-2025-31167 | 22840817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840817&pc=OZD | 07/19/2025 | 08/19/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688071 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a guidewire got damaged as the physician curved the tip. a new guidewire was used and the impella cp was placed without any issues. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6556 | 1220648-2025-45744 | 22840937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840937&pc=OZD | 07/26/2025 |  |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674585 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the placement signal disappeared during impella cp support. a placement signal unreliable alarm appeared as a result of the failure. the audio was subsequently disabled and the decision was made to continue with support. the staff was made aware of signs that would indicate the pump may need to be replaced. upon further investigation, it was discovered that the console may have caused/contributed to the event via a potential defect lamp. | [5 empty] | No physical product was returned. data logs were reviewed and the placement signal not reliable alarm was confirmed. the placement signal not reliable alarm was the result of a failed or failing bulb in the optical bench lamp assembly. | The impella device was not received from the customer, but the event logs were received. therefore, the investigation of the device is not possible, but event logs evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this report represents the pump. the related manufacturer device report number 1220648-2025-45828 represents the automated impella controller. | No physical product was returned. data logs were reviewed and a placement signal not reliable alarm was confirmed. additionally, a defective lamp alarm was found after the pump was disconnected. this is a known failure related to automated impella controller. there was no defect found with the pump. the device passed all the post sterile inspection checks.
6557 | 1220648-2025-31165 | 22840940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22840940&pc=OZD | 07/22/2025 | 09/03/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150576 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complaint reported an automated impella controller (aic) was being used for training with the team and abiomed representative when there was a yellow controller alarm. the aic was troubleshooted by turning off and on. the alarm remained. the aic was therefore removed. no patient was involved. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation. console logs from the reported day of event are consistent with complaint and show controller error alarm (#1039, due to defective optical sensor lamp), preceded by optical bench detecting ¿bad_lamp¿ condition due to out-of-spec ¿5s¿ parameters. optical bench resets do not resolve the issue, neither did power-cycling of the console. console ic2933 was tested in danvers and issue was reproduced. it¿s been observed that values of light parameter of 0.65v to 0.7v (voltage at ccd detector in optical bench) were below required 2.6v, but still higher than ¿dark current¿ noise that usually produces up to 0.5v, suggesting that there was some remaining light at the ccd and the lamp was not defective. inspection of the optical pump connector fiber revealed contamination of fiber facet, which resulted in reduced internal reflection of light which upon illuminating the ccd produced insufficient voltage. after cleaning the optical fiber, issue was resolved, and ¿5s¿ parameters were within specification. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
6558 | 1220648-2025-31168 | 22841164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22841164&pc=PYX | 07/22/2025 | 09/11/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025523511 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp was unable to pass through the right ventricle due to the patient¿s anatomy. the impella rp was exchanged for an impella rp flex successfully. | [5 empty] | The investigation was completed and there was no product returned. the root cause of delivery issue is determined to be patient condition because the pump was unable to pas through rv due to difficult patient anatomy and later rp flex was inserted successfully. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
6559 | 1220648-2025-31169 | 22841195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22841195&pc=OZD | 07/22/2025 | 09/18/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701007 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 being placed to support a patient with planned high risk surgery. the pump failed to appropriately deliver to the left ventricle. it was removed and replaced due to the medical doctor having concerns of the pump malfunction and kinking. a new impella 5.5 was utilized. no harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and product damage (catheter kink)has been completed. pump was returned for investigation. upon inspection, it was noted that there is a kink on the cannula and the no other defects were noted on the returned product. delivery issue: the root cause of delivery issue is determined to be use issue because the second pump was implanted successfully without any issues. product damage(cannula kink): it was noted that the returned product have kink on the cannula and the root cause of cannula kink is determined to be use issue because the issue did not occur with the new pump and insertion was successful without any damage to the pump. d9 device returned to manufacturer date added.
6560 | 1220648-2025-45811 | 22841426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22841426&pc=OZD | 07/26/2025 |  |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629004 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 150 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding and thrombocytopenia. heparin was withheld and switched to argatroban. the patient was also febrile. impella support continued for an additional 10 days until successfully weaned and device explanted with a survived patient outcome. | [5 empty] | The root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. the cause of the thrombocytopenia and fever cannot be determined as insufficient clinical details were provided. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6561 | 1220648-2025-31170 | 22841436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22841436&pc=OZD | 07/22/2025 | 09/18/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656939 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 68 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and after the companion sheath was removed there was considerable bleeding around the peel-away sheath. the doctor pulled the peel away out and noted a big hole in the peel-away. no hematoma or other bleeding occurred after peel-a-way sheath was removed. the patient's outcome is stable. | [5 empty] | The investigation of access site bleeding and introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. access site bleed: product was not returned for investigation. based on the clinical details provided, a hole in the introducer sheath was a potential cause of the access site bleed. however, insufficient clinical details were provided and product wasn't returned to determine the nature of the hole and what caused it. for these reasons, the cause of the access site bleed could not be determined. introducer damage - mechanical: clinical details report a hole in the introducer sheath. product was not returned for investigation. a single access procedure was done, however, the clinical details provided do not make it clear what occurred to create the hole. since insufficient clinical details were provided and product wasn't returned for investigation, the cause of the introducer damage could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6562 | 1220648-2025-31171 | 22841683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22841683&pc=OZD | 07/22/2025 | 09/25/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696073 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 was being placed to support a patient with acute myocardial infarction/cardiogenic shock. the impella 5.5 could not be advanced beyond the aortic arch due to vascular anatomy. the aortic arch prevented placement. computed tomography scan showed a left subclavian artery diameter of only 6.5¿7 millimeters, making advancement unsafe. the decision was made to switch to an impella cp which was implanted percutaneously via left femoral artery successfully and without issue. the impella 5.5 was also removed without issue, with a non-irritating wound closure. the patient had been and continues to be successfully supported by the impella cp. the next plan is to implant a left ventricular assist device in the near future and continue to support the patient with the impella pump until then. | [5 empty] | The investigation of delivery issues has been completed. per returned product investigation, product damage (cannula kink) and product damage (catheter kink) were added as failure modes. delivery issues: clinical reported that the pump could not advance past the aortic arch due to vascular anatomy. the left subclavian artery was reported to have a diameter of 6.5-7 mm, which led to difficulty advancing the pump. the product was returned for investigation, and a kink was observed on the cannula near the motor housing, as well as a small kink on the catheter above the motor housing. the cause of the delivery issues was most likely patient condition, as the patient's artery was too small to safely advance the pump. product damage (cannula kink): clinical reported that the pump could not advance past the aortic arch due to vascular anatomy. the product was returned for investigation, and a kink was observed on the cannula near the motor housing. the cause of the cannula kink was most likely patient related, as the patient's artery was too small for the pump to advance through. product damage (catheter kink): clinical reported that the pump could not advance past the aortic arch due to vascular anatomy. the product was returned for investigation, and a kink was observed on the catheter above the motor housing. the cause of the catheter kink was most likely patient related, as the patient's artery was too small for the pump to advance through. | The impella device has been received, but the evaluation has not been completed. upon completion of the investigation, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6563 | 1220648-2025-45901 | 22841685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22841685&pc=OZD | 07/26/2025 | 08/19/2025 |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562123 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella sheath site was oozing and the site was sutured with a purse string suture. blood products given to the patient included two units of fresh frozen plasma, two units of red blood cells, two units of platelets, and one unit of cryoprecipitate. the patient continued to ooze without next steps noted. next, the patient was taken to the operating room for a non-impella related abdominal surgery for bleeding and the patient was given several units of blood. no further mention of bleeding was noted. the patient expired. | [5 empty] | Device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | Product complaint # (b)(4). no device received, investigation complete: the investigation into the reported event has been completed. no device was received for evaluation.
6564 | 1220648-2025-31172 | 22841798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22841798&pc=OZD | 07/22/2025 | 09/08/2025 |  | 08/19/2025 |  |  | 07/22/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528113 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during extracorporeal membrane oxygenation cannulation, an implanted impella cp pump was inadvertently pulled out of position and subsequently removed from the body. there was no patient injury. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most likely use related as the pump was inadvertently pulled out of position during extracorporeal membrane oxygenation (ecmo) cannulation per clinical description. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6565 | 1220648-2025-31173 | 22842475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22842475&pc=OZD | 07/21/2025 | 08/19/2025 |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611470 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 87 KG | Non Hispanic | Asian | Death | Yes | Yes | Fluid/Blood Leak | Detachment of Device or Device Component | [5 empty] | Ventricular Fibrillation | [12 empty] | Other | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had impella 5.5 support ongoing for a patient admitted in with acute myocardial infarction/cardiogenic shock. on day eight, there was a minor leak of dextrose solution with the purge filter and reservoir that was noted. the solution produced from the leak has been minimal and maybe estimated to be around one milliliter in 24 hours. the slow leak is from the protective housing and leaking onto the white cable. there have been no significant changes in the purge flows or pressures prior to or after noting the leak. on day nine, the contamination sleeve came apart from the locking hub. the physician repositioned the contamination sleeve over her and secured it with tegaderm. the pump remained on despite the malfunction as the patient was still in need of support. days later the patient suffered from a new ventricular tachycardia/ventricular fibrillation episode and the patient failed to recover and care was decided to be withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two related reports. this report addresses the pump. the console is referenced under manufacturing report number 1220648-2025-31163. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
6566 | 1220648-2025-45743 | 22843214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22843214&pc=OZD | 07/26/2025 |  |  | 08/19/2025 |  |  | 07/26/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705483 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 67 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was escalated from an impella cp to an impella 5.5 due to severe ischemic cardiomyopathy. during impella 5.5 support, the patient had a right middle cerebral artery infarct. care was ultimately withdrawn and the patient expired. | Additional information was received. the patient was noted to have a right middle cerebral artery infarct. after review with the team, it is believed to be from removal of the impella cp before escalation to the impella 5.5. |  |  |  |  | Medical safety review was completed and noted the following: the stroke with the demise outcome was reported on the impella 5.5. additional information was received and noted it is believed the stroke occurred after removal of the impella cp before escalation to the impella 5.5. as the impella 5.5 device was in use at time of expiration and timing of the stroke is indeterminate, the impella 5.5 association with stroke and demise outcome will remain unchanged. the associate director of post market clinical review noted without information on when the stroke occurred this is hard to determine definitely. however, it should be noted that the impella cp was placed through a 16fr cook sheath which is off label and could lead to increased risk of clot formation in the sheath. related medical device reports: this impella 5.5 report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B.5 additional information was received. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. e.4 selection was added as it was inadvertently omitted from the initial report that was submitted. the impella device was not returned, and the investigation has been completed. the root cause of the reported stroke was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks.
6567 | 1220648-2025-45708 | 22843759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22843759&pc=OZD | 07/28/2025 | 07/30/2025 |  | 08/19/2025 |  |  | 07/28/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648817 |  | 1000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered some oozing at the site, intermittent ventricular tachycardia (vt) and hemolysis during support. no treatment for the site oozing was mentioned. vt was experienced by the patient intermittently one night and hemolysis eventually cleared. | [5 empty] | Corrections: d1, d4, e4. the investigation of the reported failure mode has been completed. the device was received for investigation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. hemolysis: no physical damage was noted on the returned product. data log shows the placement signal and pump flow trending downward during the time of reported hemolysis. no positioning issues were noted. as per the clinical report, hemolysis was reported on (b)(6) and was stated to be resolving the next morning. clinical details suggest hemolysis was managed with volume monitoring. the cause of the hemolysis was most likely patient condition related to low blood volume. access site bleeding: no damage was noted on the returned product. the repositioning sheath passed leak test with a maximum pressure of 200 mmhg. the cause of the access site bleeding issue was not determined since no defect found on returned product and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6568 | 1220648-2025-45767 | 22845676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22845676&pc=OZD | 07/27/2025 |  |  | 08/19/2025 |  |  | 07/27/2025 | 08/19/2025 |  |  |  |  | 08/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697711 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after one day of impella cp support an attempt was made to wean the patient off bypass, however the pump could not be repositioned into an optimal position. due to the patient's hemodynamic needs, the decision was made to escalate to impella 5.5 support. no patient harm was reported. | [5 empty] | The product was not returned for investigation. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6569 | 1220648-2025-45739 | 22826125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22826125&pc=OZD | 07/25/2025 |  |  | 08/18/2025 |  |  | 07/25/2025 | 08/18/2025 |  |  | 10/31/2025 |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415351 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to the patient's anatomy. the guidewire could not pass due to calcification and tortuosity in the aorta. the patient was placed on extracorporeal membrane oxygenation instead and survived. | [5 empty] | The pump was not returned for investigation. the cause of the delivery issues was patient anatomy related that resulted in no adverse event for the patient related to the impella. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
6570 | 1220648-2025-45751 | 22826923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22826923&pc=OZD | 07/07/2025 |  |  | 08/18/2025 |  |  | 07/24/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686635 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Failure | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received regarding a respiratory failure with relationship listed as "possibly" related. the loqi stated: "intubated after impella procedure. after 4 days of support, family decided to withdrew care and impella pulled out of the left ventricle and turned off. patient taken to hospice unit to pass comfortably. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6571 | 1220648-2025-45730 | 22828871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22828871&pc=OZD | 07/26/2025 |  |  | 08/18/2025 |  |  | 07/26/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687475 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis. the pump was repositioned but that did not resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Data logs were reviewed and show that the pump ran without issue during support. no motor current spike or positioning issues occurred before the hemolysis was noted. there were suction alarms triggered during the case but they resolved quickly. the customer's device was not returned for review. the cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
6572 | 1220648-2025-45723 | 22829193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22829193&pc=OZD | 07/25/2025 |  |  | 08/18/2025 |  |  | 07/25/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695279 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 73 KG |  |  | Injury | Yes | Yes | Difficult or Delayed Positioning | Pumping Stopped | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock (pccs) was implanted with an impella cp device for mechanical circulatory support (mcs) and encountered positioning difficulty, hemolysis, limb ischemia and pump stop eventually resulting in device explant. the patient was admitted with congestive heart failure exacerbation; reported three weeks of progressive dyspnea on exertion. workup revealed severe mitral regurgitation and ejection fraction of 65-70%. one day prior to start of impella mcs, the patient underwent coronary artery bypass graft surgery x 3 and mitral valve replacement. pccs prompted the decision to place an impella rp with smartassist. an echocardiogram the following morning noted concern for poor left ventricular function, and the decision was made to place the impella cp for left-sided support which was successfully completed. post implant the same day, an attempt was made to reposition the impella cp under transesophageal echocardiogram (tee) and at one point, the pump dove in deep with the impella in the ventricle alarm followed by high motor currents. the impella stopped due to the high motor current. the physician pulled back and was able to restart the impella cp pump. however, prior to repositioning, the impella cp was running at performance level p-8 and after restart the performance level was unable to go above p-6 without suction alarm. the physician decided to maintain support at p-6 until the following day. the next day, the patient was noted to have hemolysis. lactate dehydrogenase was elevated with decrease in urine output which was red/tinged. the patient had been administered four units of packed red blood cells. the impella cp was repositioned under tee, and the physician was able to obtain an appropriate position. however, the performance level was unable to go above p-4. additionally, it was noted an echocardiogram was completed and left ventricle function appeared to have approved. however, the following day, there was concern for loss of pulses in left leg. therefore, the patient brought back to operating room for impella cp removal. ultimately, the device was explanted with a patient survived outcome and the patient continued on impella rp support with planned assessment of the patient¿s need for left side support. note: additional device required was for the extracorporeal membrane oxygenation. | [5 empty] | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Product complaint # (b)(4). investigation summary: pump and data stick returned. hemolysis: no product damage was observed on returned product. small biomaterial was present on tip of impeller blade. data logs were reviewed and suction alarms are present throughout support. a motor current spike occurred with drop in pump flows with a temporary pump stop occurring 30 seconds after motor current spike. pump was able to restart but low flows were present for remainder of case with suction alarms. clinical details noted that prior to pump insertion, patient was experiencing pink tinged urine. after pump insertion, clinical noted that urine was starting to clear up. the following day, patient had elevated ldh, decreased urine output and red tinge. the cause of the hemolysis was due to biomaterial based on returned product and data log review. ischemia: clinical details noted that patient had loss of pulses in left leg. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. difficulty to place or position: clinical details noted that pump was in poor position with inflow interacting with mv after patient was transferred to unit on first day of impella support. pump was attempted to be repositioned with tee but was unsuccessful. pump was pushed too deep during repositioning and a temporary pump stop occurred. pump was able to restart after pulled back but was unable to reposition properly. pump was attempted to be repositioned the following morning and was able to point pump more towards apex. the cause of the positioning issue could not be definitively determined as limited clinical details were provided on how pump initially came out of position. temporary pump stop: returned pump was visually inspection and pump damage were observed. small biomaterial was present on tip of impeller blade. motor cable electrical lines were intact and within spec. pump was run at p-9 for 5 minutes then down to p-6 for approximately 30 minutes without pump stop reproducing. data logs were reviewed and rt log at time of temporary pump stop shows a motor current spike with speed deviation and step down in pump flows. immediately after motor current spike, "impella in ventricle" alarm was triggered. 30 seconds after motor current spike, pump stop alarm #13 was triggered. pump restarted after approximately 2 seconds. multiple temporary pump stops were triggered after initial pump stop over approximately 2 minutes before pump was able to be restarted with low flows for remainder of case. clinical details noted that when attempting to reposition pump, pump was pushed deep with "impella in ventricle" alarms. pump was pulled back and able to be restarted but was not able to flow above p-6 without suction alarms. the cause of the temporary pump stop was most likely due to biomaterial ingestion likely as a result of improper positioning per clinical details and small biomaterial was present on returned pump. device history lot: device lot #: 1888618. device history batch: subcomponent lot #: n/a. device history review: pump sn (b)(6) passed all post-sterile inspection checks.
6573 | 1220648-2025-45878 | 22832116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22832116&pc=OZD | 08/01/2025 | 02/26/2026 |  | 08/18/2025 |  |  | 08/04/2025 | 08/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683204 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation | [7 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Aortic Valve Insufficiency/ Regurgitation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The impact study reported a patient supported by an impella 5.5 device experienced an access site infection. the principal investigator and abiomed safety stated the relationship to be definite to the impella procedure. the impact forms stated "antibiotic treatment: during impella device explantation swab was taken. microbiological analysis revealed infection with corynebacteria. antibiotics were administered." the patient survived the procedure. | [5 empty] | No product was returned for investigation. the root cause of the infection was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impact study revealed aortic regurgitation. h6 updated.
6574 | 1220648-2025-45795 | 22833117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22833117&pc=OZD | 07/25/2025 |  |  | 08/18/2025 |  |  | 07/25/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689276 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump experienced suction overninght while supporting the patient. two units of blood were given and the patient was subsequently escalated to impella 5.5 support. no patient harm was reported due to the issue. | [5 empty] | The pump was not returned for investigation. data logs were available. data logs were reviewed and the suction alarms were confirmed. since insufficient clinical details were provided regarding what was occurring at the time of the alarms and product wasn't returned for investigation, the cause of the suction alarms could not be determined. the device passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6575 | 1220648-2025-45791 | 22833199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22833199&pc=OZD | 07/25/2025 |  |  | 08/18/2025 |  |  | 07/25/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697708 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female | 48 KG | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp implant, the pump began to experience persistent suction issues. the staff was unable to resolve the noted suction and the decision was made to escalate to impella 5.5 support. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6576 | 1220648-2025-45810 | 22833348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22833348&pc=OZD | 07/24/2025 |  |  | 08/18/2025 |  |  | 07/24/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693656 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient encountered limb ischemia during impella cp support. the patient had a successful left femoral artery (lfa) external bypass conducted. it was confirmed during support that the bypass continued functioning well and that distal pulses remained present as a result. patient is stable post intervention. | [5 empty] | Additional information received indicated the physician decided to leave the peel-away sheath in (per the physician's decision) and opted to place an external bypass sheath in place to perfuse the leg distal to the peel-away sheath. the patient continued on support. condition improving; successful wean of the impella. patient was stable on no vasoactive medications. the impella was explanted in the catheterization lab with a survived patient outcome, returning to the unit in stable condition. patient information a4: weight of patient is unknown. device status: the impella device was not received from the customer, as it was reportedly discarded by the facility. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system - section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6577 | 1220648-2025-45880 | 22833355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22833355&pc=OZD | 08/04/2025 |  |  | 08/18/2025 |  |  | 08/04/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666709 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 58 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. the patient was sedated and was in a controlled coma. there was less brain activity and the patient passed away while on impella support. | The impact study reported a patient supported by an impella 5.5 device experienced an access site hematoma. the principal investigator and abiomed safety stated the relationship to be probable. the impact forms state ¿in the morning of (b)(6) 2025 the patient reported a swelling at the access site. surgical extirpation of a hematoma was performed¿. it was also reported from an impact study which stated the patient had bleeding at the access site. the impact forms state "new onset of bleeding at the impella implant site, surgical revision necessary". the patient continued on impella support and at time of this report remains on impella support with the same device. | Notifications were received from abiomed's impact and reported a patient undergoing high-risk coronary artery bypass graft surgery was implanted with an impella 5.5 for mechanical circulatory support. while on support, day 7, the patient experienced bleeding at the impella access site that required surgical revision. the following day a vascular access site had swelling, and a hematoma developed, which required surgical extirpation. both access site bleeding and hematoma events were determined probably related to the impella procedure (however, unrelated to the impella device). the patient continued on support. the purge system composition was bicarb. approximately 21 days later, the patient's condition became worse, and the patient was placed on veno-venous extracorporeal membrane oxygenation (vv ecmo). impella performance level was at full support on p-9. the impella was noted to be running well on p-9 without incidents and pulmonary was secured with vv ecmo. approximately two weeks later, the patient experienced a cerebral hemorrhage. the sedated patient was in a controlled coma, less brain activity. the cerebral hemorrhage event was determined possibly related to the impella device and the impella procedure. the patient ultimately expired two days later on support. |  |  |  | The impella device was not returned, and the investigation has been completed. the device history review was completed and the pump passed all the post sterile inspection checks. the cause of the injury was not determined since the product was not returned and insufficient clinical details were provided. sections b.2, b.5, d.6b, h.1 and h.6 were update due to new information received. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction report is being submitted as the adverse event of respiratory failure was erroneously reported and cerebral hemorrhage was unwittingly omitted as a reported. consequently, corrections were made to section b5 to reflect the correct reported event details and h6 health effect: clinical/impact and medical device problem codes.
6578 | 1220648-2025-45805 | 22833591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22833591&pc=OZD | 07/25/2025 |  |  | 08/18/2025 |  |  | 07/25/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698054 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 101 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on arrival to intensive care unit (icu), the patient had profuse bleeding from the groin due to poor dressing from team rushing to get to the icu. impella position was unclear and the patient was given seven units of packed red blood cells, seven units of fresh frozen plasma, and one unit of platelets. additionally, multiple units of blood were transfused throughout the day but the bleeding was unable to be stopped. the patient was too unstable to take for a computed tomography or to the operating room for exploration and resolution of the bleeding. the source of the bleed was unknown. however, the physician felt it was likely that there may have been a retroperitoneal bleed since the impella was placed during cardiopulmonary resuscitation. the family made the decision to withdraw care after extensive discussion with the medical team and the patient expired. | [5 empty] | The root cause of access site bleeding was most likely use issue since hemostasis was achieved by adding additional sutures. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
6579 | 1220648-2025-45758 | 22834351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22834351&pc=OZD | 07/26/2025 |  |  | 08/18/2025 |  |  | 07/26/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694640 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to advance beyond the innominate during attempted delivery. the pump was removed as a result of the resistance and the decision was made to re-attempt with a stiffer wire and new impella 5.5 device, which was successfully placed based on patient updates. no patient harm was reported. | [5 empty] | Patient information: a4 is unknown. device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added medical history/preexisting condition as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone and zip code extension as they were inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. the impella device was not returned, and the investigation has been completed. the root cause of the delivery issue was determined to be a use issue as the staff had used a non-abiomed product (boston scientific platinum plus guidewire) which most likely prevented successful delivery of impella. the device history review was completed and the device passed all post sterile inspection checks. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6580 | 1220648-2025-45760 | 22835018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22835018&pc=OZD | 07/26/2025 |  |  | 08/18/2025 |  |  | 07/26/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662627 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and the following day, the patent was noted hypotensive on p-7 support level. additionally, there was suspected hemolysis with blood in the urine. an echocardiogram was performed with impella cp measuring 3.7 cm to aortic valve annulus; the physician satisfied with its position. however, it was noted, as well, the patient was bleeding from the access site. c-shock team was called with plans to escalate to the impella 5.5, which was successfully completed, and ventricular assist device work up. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The pump was returned for analysis. returned product was investigated and there were no abnormalities or damages noted. the complaint could not be reproduced. the cause of the hemolysis was determined to be patient condition related. the cause of the bleed could not be determined. the cause of the hypotension cannot be determined as insufficient clinical details were provided. the pump passed all post-sterile inspection checks.
6581 | 1220648-2025-45756 | 22835167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22835167&pc=OZD | 07/24/2025 |  |  | 08/18/2025 |  |  | 07/24/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687919 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 89 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant a patient (unknown race) with history of acute myocardial infarction/cardiogenic shock requiring extracorporeal membrane oxygenation (ecmo) support was brought to the operating room (or) for escalation of venoarterial- venopulmonary ecmo with an impella cp (serial number (b)(6)) unloading to unloading with an impella 5.5 (serial number (b)(6)) for the purpose of weaning ecmo support with the intention to decannulate the following day. during the initial cutdown for the impella 5.5 axillary graft site, some bleeding complications were encountered. the graft was placed successfully, and multiple "quick clot" methods were used to stitch the site. following implant, protamine was used to reverse the effects of heparin and blood products were given to counteract the blood loss. support continued uninterrupted. the patient¿s left ventricular ejection fraction recovered to 45-50%, and the patient was taken to the or for the removal of the impella 5.5. the patient was stable during the wean and there were no issues reported at explant. the patient remained on venopulmonary extracorporeal membrane oxygenation after impella 5.5 was explanted. eight days later, the patient required mechanical circulatory support to assist in weaning from ecmo support. an impella cp (serial number (b)(6)) was prepped and primed, however imaging revealed the patient had severe thrombus and dissection and lack of perfusion below the right iliac ostia to the common femoral artery. as these are contraindications for the impella cp, the cp was aborted, and the patient was prepared to be taken to the or for an impella 5.5 device. prior to going to the operating room, it was noted that patient¿s hemoglobin had decreased and thromboelastography also showed need for fresh frozen plasma (ffp). one unit of ffp and 1 unit of red blood cells (rbc) were transfused prior to the trip to the or. during the initial cutdown and subsequent sewing of the graft, the patient became bradycardic, and blood pressure became unreadable. cardiopulmonary resuscitation was initiated and advanced cardiac life support protocol activated. the patient received 1 round of epinephrine and three ampules of sodium bicarbonate solution. the return of spontaneous circulation achieved in less than 4 minutes. the patient stabilized on increased presser support and impella implant procedure completed following abiomed best practices. the patient remained in critical state on max dose pressers and received 4 more units of rbcs in attempt to curb the patient¿s hypotensive, vasoplegic state. the patient was transferred to cardiovascular icu without incident. it was reported that on the second day of support with the second impella 5.5 device (serial number (b)(6)), the patient experienced ventricular fibrillation arrest while repositioning in bed. it was noted that the patient was having complications with ecmo support with suspected lysis from the ecmo device. the patient continued to deteriorate despite increasing support and began showing signs of multiple organ failure. the patient was unable to be weaned off of ecmo support, was no longer responding to max pressors, and was experiencing intermittent arrhythmias. the patient¿s family decided to withdraw care. care was withdrawn and the patient expired. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. related report numbers: this mdr specifically addresses pump 4, the impella 5.5 with serial number (b)(6), which is one of three pumps associated with complaints involving the reported event. refer to related report numbers for the report on the other two pumps. device status: the impella was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6582 | 1220648-2025-45757 | 22835259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22835259&pc=OZD | 07/26/2025 |  |  | 08/18/2025 |  |  | 07/26/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681779 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant bleeding was noted at the impella site. the perclose strings were tightened and an additional suture was placed at the access site in response to the bleed. also, the angle of entry was supported and manual pressure was held for roughly 40 mins to help manage the condition. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. the device was not returned for investigation. the root cause of bleeding is determined to be use related due to improper anticoagulation management because the activated clotting time values are above the therapeutic range than suggested as per instructions for use. hemostasis was achieved by manual pressure. the device history review was completed and the device passed all post sterile inspection checks.
6583 | 1220648-2025-45770 | 22835558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22835558&pc=OZD | 07/03/2025 |  |  | 08/18/2025 |  |  | 07/03/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681007 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 113 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown ethnicity) implanted with an impella 5.5 device for mechanical circulatory support experienced a left ventricle thrombus during impella support. the patient¿s anticoagulation was switched from heparin to bivalirudin, and support continued. the patient remained on impella 5.5 support until successfully bridged to a heart transplant. the patient was noted to have survived at impella explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6584 | 1220648-2025-45882 | 22835665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22835665&pc=OZD | 07/06/2025 | 10/30/2025 |  | 08/18/2025 |  |  | 07/06/2025 | 08/18/2025 |  |  |  |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1683279 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 89 KG |  | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered an automated impella controller (aic) issue. the aic touchscreen was compromised overnight. as a result, the aic was swapped. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6585 | 1220648-2025-45887 | 22835674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22835674&pc=PYX | 06/06/2025 | 08/18/2025 |  | 08/18/2025 |  |  | 06/06/2025 | 08/18/2025 |  |  | 06/30/2025 |  | 08/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024366105 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was escalated from an impella cp and implanted with an impella 5.5 and an impella rp for increased mechanical circulatory support. the patient remained in persistent shock, renal failure, and atrial flutter despite the additional support of impella cp and the decision was made to escalate the patient to bipella support with an impella 5.5 and impella rp. it was reported that the removal of the impella cp was complicated due to bleeding at the superficial femoral artery and it required extensive repair. it was noted that the heparin was reversed with protamine administration and the patient was resuscitated with blood products and fluids during acute bleeding and drop in mean arterial pressure and filling pressures. the purge solution for both impella rp and impella 5.5 were to hold heparin until the following morning. the day following the initiation of bipella support, it was reported that the patient experienced ongoing bleeding from the access site of the previous impella cp and heparin was continued to be withheld from purge solutions. it was also reported on the same day that the patient¿s laboratory tests for lactate dehydrogenase (ldh) and alanine transaminase were unable to be performed due to hemolyzed samples, it was however noted that the patient¿s lab draws had been hemolyzed for the entire duration of the patient¿s stay. additionally, it was noted that there was a significant drop in purge flow of the impella rp approximately 2 hours post implant of the device. recommendations to add tissue plasminogen activator (tpa) were given, however this was not accessible at the time of recommendation. it was also noted that there was a brief purge pressure alarm that occurred during the exchange of the impella cp and the 5.5 that self-resolved. it was advised to the medical team that the pump had potential for failure and recommendation for pump exchange if purge pressures were unable to be resolved. it was reported that during vascular repair of the impella cp access site, purge alarms were noted again the impella rp. to resolve the alarm, the pump flows were decreased to 0 and quickly increased back to p8, however purge flows were still lower than expected. the impella rp device was monitored for further red purge alarms, and when the alarms returned tpa was added to the purge solution. after tpa ran through the purge, the impella rp purge flows increased, and the purge pressure decreased to expected ranges. the following day, it was noted that the impella 5.5 had placement signal low alarms. echocardiography was performed at the beside and the impella 5.5 was pulled back. two days later, it was reported that the patient¿s lactate results were increased following the escalation of support to bipella, and the cause was likely due to right lower extremity ischemia from injury of the impella cp removal. after resolving the injury with surgical intervention, the lactate results were lower, near normal range, and there was a large improvement in cardiac index/cardiac output. it was noted that the previous day¿s positioning of the impella 5.5 had resolved the increased ectopy the patient had been experiencing and the suction alarms from the device. the next day, it was noted that the patient¿s lactate dehydrogenase lab draw was hemolyzed. the medical team began to wean the impella rp as the patient¿s right heart function had improved and a plan was made to explant the impella rp following a planned esophagogastroduodenoscopy. however, now four days later, the patient went into pulseless electrical activity arrest following the removal of swan. return of spontaneous circulation was obtained after 5 rounds of cardiopulmonary resuscitation (cpr) and the patient was re-intubated. following the cpr, the patient¿s hemoglobin and hematocrit began dropping and there was a swelling noted at the impella 5.5 access site. the next day, the patient was taken for computed tomography of the chest. a large hematoma in the right axillary was observed adjacent to the impella catheter with possible pseudoaneurysms at the catheter insertion site. it was also noted that the patient had a subarachnoid bleed. heparin was discontinued as a result of the noted access site hematoma and a concern for active bleed. the patient¿s family decided to withdraw care. impella 5.5 support was withdrawn, and the patient expired. | [5 empty] | Explant date corrected with new provided information. the investigation for the reported issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. device event logs were provided. high purge pressure: clinical details note that there were brief high purge pressure alarms during rp flex support at the time of 5.5 implant. the next day, there were more high purge pressure alarms and tissue plasminogen activator protocol was initiated. no more issues with purge were noted. data log review showed a gradual 3 hour rise in purge pressure (pp) from ~500 to >1000 mmhg at the beginning of the rp flex case. there were no motor current spikes/trends prior to the rise. the high pp was sustained for ~9 hours before gradually resolving over the next 20 hours. no product was returned. the cause of the high purge pressure was most likely biomaterial buildup since there was a gradual rise in pp over 3 hours and there were no motor current spikes prior to the rise. hemolysis: clinical details state that the patient's lactate dehydrogenase was hemolyzed on (b)(6) 2025. it is unknown if/when the hemolysis resolved. data log review showed suction alarms throughout the case, but most present between (b)(6) 2025 and (b)(6) 2025. motor current and placement signal are fairly stable throughout the entire case as well. no product was returned. the root cause of the hemolysis was not determined since no product was returned and insufficient clinical details were provided. arrythmia: clinical details state that the patient had significant ectopy and runs of nonsustained ventricular tachycardia (nsvt) after the impella was pulled back from 6.7 cm measurement shortly after insertion. the ectopy resolved when the pump was pushed in past 6 cm measurement again. it is unclear which pump (rp flex or 5.5) was being repositioned. patient had more runs of ectopy throughout the case. data logs are not applicable, and product was not returned. in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the ectopy/nsvt was unable to be determined due to insufficient clinical details. device history review. the pump sn (b)(6) passed all post-sterile inspection checks. | Note: actual date of explant is unknown. captured 30th as this date cannot be before the implant date. medical safety officer reivew: multiple pumps which included bipella support with impella rp and impella 5.5 supported the patient. impella cp and impella rp related events are serious injury type of reportable events. the demise outcome is associated with the impella 5.5 as there is not enough evidence to excluded the impella 5.5 device used as an associated factor to the patient's outcome; last device used and the patient had hemolysis, extensive repair and blood transfusion which is more likely to be related to the demise outcome. device status: the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. related reports: this complaint involves three pumps: this mdr specifically addresses impella rp with serial number (b)(6), which is the subject of this report. refer to the related report numbers for the regulatory reports for the other two devices with a complaint associated with this reported event.
6586 | 1220648-2025-45724 | 22814056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22814056&pc=OZD | 07/31/2025 |  |  | 08/15/2025 |  |  | 07/31/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692279 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during pump prep, there was a reported issue with the optical sensor not measuring the placement signal. it was reported that all steps to prepare the pump were followed correctly. the pump was replaced with a new impella cp to resolve the issue. there was no patient harm reported due to this issue. | [5 empty] | Investigation summary: pump sn (b)(6) returned and inspected. immediately upon start, issue with optical sensor, did not measure placement signal, steps were followed correctly, pump was replaced with a new pump. pump passed laser continuity testing. 5s parameters were nominal. no physical abnormalities were noted on the pump. the pump was uson tested at 0, 50, 100, 150, and 200mmhg of externally applied pressure. at each pressure gradient, the pump read as -10, 36, 87, 135, and 184 mmhg, respectfully. the actual minus expected readings were -10, -14, -13, -15, and -16 mmhg, respectfully. no offsets were applied. the root cause of the optical signal issue could not be determined as no abnormalities were found with the returned pump, data logs were not recovered, pump s/n (b)(6) passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
6587 | 1220648-2025-45850 | 22814165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22814165&pc=OZD | 08/01/2025 |  |  | 08/15/2025 |  |  | 08/01/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1380525 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered issues with the automated impella controller (aic). the aic was unable to finish the self-test. no further information was provided on the issue. no patient was involved . | [5 empty] | Product complaint#: (b)(4). investigation summary: data analysis: on the complaint date, on (b)(6) 2025, during initial boot-up, the console logs indicated a communication issue with the pbm. due to this communication problem, the aic console rebooted automatically (as designed) to attempt another power cycle. after rebooting, the console displayed the "system self check failed" screen". this confirms the reported failure mode. device analysis: this console was tested after arriving at fs. visual inspection confirmed the pbm contamination, which has impacted on a neighboring rs232 communication channel (pbm_(b)(6).png). root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. service & repair: before returning this console ((b)(6)) back to the customer, fs was instructed to perform the following, replace the pbm (0042-1125) due to the corrosion. perform sections 18, 19, and 20 of the pm (preventive maintenance) procedure (0042-7309). perform a full functional check (0042-7316). device history lot: n/a. device history batch: subcomponent name: pca, power/battery manager, rohs, tested. material number: 0042-1125, lot number: 2019329440, serial number: (b)(6). device history review: q1) service history review - are there an qns associated with the complaint device¿s most recent service report? no. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? no. q3) complaint history review - have there been any other complaints against this console? no. phr summary: there were no issues related to the complaint discovered when reviewing the product history of console (b)(6). | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6588 | 1220648-2025-45846 | 22814384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22814384&pc=OZD | 08/01/2025 | 02/24/2026 |  | 08/15/2025 |  |  | 08/01/2025 | 08/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693547 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information has been reported. the replacement 5.5 worked well, and the patient was stable. no more data is available due to hospital restrictions, and data collection is not allowed. | The eu complainant reported that a fluid leak was seen from the purge filter of an implanted impella 5.5 pump during patient support. despite the leak, the purge values remained within a normal range. the decision was made to replace the pump due to the noted leak. there was no patient injury. |  |  |  |  | B5 updated. | Purge leak - pump: pump sn (b)(6) and purge cassette were returned for investigation. no physical damage was observed on the returned product. the pump was then successfully primed without issues and tested for around 1 hour, during which no leak was observed. the side arm protection case was opened, and no damage was noted inside. to recreate the leak, the pump side arm exit was blocked with a clamp; however, no leak was observed on the returned product. all purge parameters were within specifications during testing. data logs were not provided or could not be retrieved from the remote server. no defect was found on the returned product. the cause of the purge leak-pump issue was not determined due to insufficient clinical information provided. the pump sn601923 passed all post sterile inspection checks. | Section a: patient information is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6589 | 1220648-2025-45775 | 22815634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22815634&pc=OZD | 07/24/2025 |  |  | 08/15/2025 |  |  | 07/24/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661723 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 58 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that roughly six days into impella 5.5 support, heparin was stopped due to the patient experiencing thrombocytopenia. no additional intervention was required and support continued uninterrupted. | [5 empty] | Product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use as noted: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6590 | 1220648-2025-45858 | 22815680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22815680&pc=OZD | 07/24/2025 |  |  | 08/15/2025 |  |  | 07/24/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696265 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male | 114 KG | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered high purge pressure and thrombus during impella cp support. a purge flow decrease was noted by the nurse and resolved by removing kink in purge tubing. additionally, thrombus was confirmed and observed in the left ventricle. the team was recommended to adjust the impella pump position. after a few days the echo demonstrated no definitive signs of left ventricle thrombus. patient is still on device support and stable. | Additional information was received. the patient expired. |  |  |  |  | Product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use. as noted in the impella ifu: section: potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B.2 additional information was received that the patient expired. the date of the death was added. required information was selected incorrectly on the initial report that was submitted as no intervention was done. serious injury should of been selected on the initial report that was submitted but new information has been received since then that makes this report a death. b.5 additional information was received. d.4 added model and explant date as they were not on the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. h.1 revised to reflect new information that was received that the patient died. h.6 code e0506 was reported incorrectly on the initial report that was submitted. code f02 was added to replace code f12 as the patient has died. the impella device was not received from the customer, but the event logs were received and the investigation has been completed. the device history review was completed and the pump passed all incoming inspection criteria. high purge pressure: no product was returned. data log analysis showed brief spikes in purge pressure but no high purge pressure alarm. the cause of the high purge pressure was a kinked purge line since removing the kink resolved the issue as reported in the clinical. thrombus: a left ventricle thrombus was seen on echocardiogram. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
6591 | 1220648-2025-45727 | 22816104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22816104&pc=OZD | 07/24/2025 |  |  | 08/15/2025 |  |  | 07/24/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691860 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 116 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was initially implanted with an impella cp device and escalated to an impella 5.5 as a bridge to decision (transplant or durable left ventricle assist device). approximately 21 days after start of the impella 5.5 support, the patient was positive for heparin-induced thrombocytopenia. the decision was made to bridge to a durable left ventricle assist device, which was successfully completed 11 days later. | [5 empty] | Product status: the impella device was not received from the customer as according to the site it was not saved. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation). | The product was not returned for investigation. a device history record review was completed and noted that the device passed all post sterile inspection checks. the cause of the thrombocytopenia cause was not determined due to insufficient clinical details. e.1 zip code extension was added as it was inadvertently omitted from the initial manufacturer report submitted. section h6 investigation coding has been updated accordingly based on the completed investigation. h.6 code b18 was reported incorrectly on the initial manufacturer report and the correct code was added to type of investigation section.
6592 | 1220648-2025-45808 | 22816818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22816818&pc=OZD | 07/23/2025 |  |  | 08/15/2025 |  |  | 07/23/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701079 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to not being able to pass calcified vessels. no patient harm was reported. | [5 empty] | The pump was not returned for investigation. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to calcification in the vessels. the pump passed all post sterile inspection checks. | Device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6593 | 1220648-2025-31162 | 22816993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22816993&pc=OZD | 07/18/2025 | 09/12/2025 |  | 08/15/2025 |  |  | 07/18/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693480 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 66 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an impella cp implanted from an axillary approach following an aborted impella 5.5 implant due to patient anatomy. a double barrel technique was used (femoral impella cp was in place and running at p6-7). axillary impella cp was ramped up to p6 as femoral impella cp was weaned down to p1 withdrawn into the descending aorta and placed into surgical mode. flows reached three liters per minute at p6 for approximately one minute until a suction alarm presented and flows dropped to one liter per minute. position was confirmed optimal and volume was administered. following the explant, clot like debris was visible in the proximal cannula and upon review, impella connect there was a spike in motor current. following loss of flow, the patient deteriorated requiring chemical support and cardiopulmonary resuscitation but ultimately expired on support. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of low pump flow and thrombus has been completed. the impella device was received for investigation. data log review showed motor current (mc) spike right after initial pump start mc spike that cause the pump to run at a lower mc. after that, flows at p-4 routinely below 1l/min, which is also below the expected range of 2.0-2.5l/min at p-4 for cp. suction alarm showed up right at pump start and did not resolve until pump was turned off. the pump was returned and inspected. after soaking the pump in water water and terga zyme, biomaterial was found wrapped around the impellor blade. low pump flow: the cause of the low pump flow was biomaterial due to the biomaterial wrapped around the impellor blade. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
6594 | 1220648-2025-45832 | 22817309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22817309&pc=OZD | 10/23/2024 |  |  | 08/15/2025 |  |  | 07/23/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539104 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Cardiac Arrest | Renal Impairment | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella device used. the patient was given a bolus lasix which helped the condition. additionally, patient endured anemia with a "possible" relationship to the impella device used. hemoglobin dropped down to 7.1 and one unit of blood was ordered. the urine then looked clear with no hematomas or ecchymosis. the cause of the anemia was unclear. patient also endured ongoing ventricular arrythmias with a "probable" relationship to the impella device used. it was reported that the patient coded on induction of anesthesia and cardiopulmonary resuscitation was performed for hypotension for 5-10 minutes. the patient stabilized with moderate to high doses of inotropes, dobutamine, epinephrine, vasopressin, and inhaled nitric oxide. in the left axillary an impella 5.5 was placed for mechanical circulatory support. the predominant underlying rhythm was sinus tachycardia with frequent ventricular ectopy being noted. there was an intraventricular conduction delay and the patient experienced episodes of ventricular tachycardia/ventricular fibrillation with loss of consciousness with coughing. the first episode occurred while the patient was out of bed, in a chair. the patient was sedated and cardioverted. the patient had additional episodes requiring shock after he had been lifted back to bed. the patient started on amiodarone and lidocaines. a bedside echocardiogram performed, and the impella was repositioned. the patient was electively intubated. the patient not recovered/not resolved. the patient ultimately expired. | A fourth notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that on (b)(6) 2024 regarding patient experiencing cardiac arrest with impella relationship listed as "definitely related". it was reported that after anesthesia induction, patient suffered a cardiac arrest. cardiopulmonary resuscitation was performed for about 5-10 minutes. during this time, femoral access was obtained for possible extracorporeal membrane oxygenation (ecmo) cannulation. patient regained rosc (return of spontaneous circulation) and stabilized on moderate-high dose inotropes. decision made to proceed with the planned axillary impella procedure. |  |  |  |  | Patient information: a4 patient weight unknown. device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use: as noted in the impella ifu section: potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information received from the clinical site and the following fields were updated. b5, h6. correction g2.
6595 | 1220648-2025-45773 | 22819787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22819787&pc=OZD | 07/24/2025 |  |  | 08/15/2025 |  |  | 07/24/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 65 YR |  | Female | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the access site experienced some oozing. a mattress suture was placed around the site and blood products were given to treat the condition. the following day, it was noted that the impella leg had limited pulsatility. no intervention was performed as plans were already made to wean and explant the pump the following day. | [5 empty] | Device status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use: as noted in the impella ifu section: potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Product was not returned for investigation. pre procedural adverse event: access site bleeding: the root cause of access site bleeding was determined to be use issue because the oozing was managed by placing mattress sutures around repo sheath and hemostasis was achieved. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
6596 | 1220648-2025-45786 | 22820122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22820122&pc=OZD | 06/30/2025 |  |  | 08/15/2025 |  |  | 07/24/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679375 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received regarding a patient who was implanted with impella 5.5 with worsening ventricular tachycardia (vt). the patient continued to have vt burden after the cardiac arrest. additionally, patient experienced left ventricle (lv) thrombus with relationship listed to be "possibly" related to the impella device. post impella, the tee (transesophageal echocardiography) showed lv thrombus and as a result, the patient was placed on heparin gtt. patient was successfully weaned of the device and is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the thrombus was unable to be determined due to insufficient clinical details. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
6597 | 1220648-2025-45778 | 22821687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22821687&pc=OZD | 07/24/2025 |  |  | 08/15/2025 |  |  | 07/24/2025 | 08/15/2025 |  |  |  |  | 08/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684313 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 81 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, the patient experienced a gastrointestinal bleed. as a result of the bleeding, heparin was discontinued. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella automated controller device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: h11 on the initial manufacture report narrative it erroneously called out controller instead of pump.
6598 | 1220648-2025-45801 | 22801610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22801610&pc=OZD | 07/20/2025 |  |  | 08/14/2025 |  |  | 07/20/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689391 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a thrombus. it is unclear how the thrombus was treated. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the thrombus was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
6599 | 1220648-2025-45822 | 22804626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22804626&pc=OZD | 07/16/2025 |  |  | 08/14/2025 |  |  | 07/16/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025601511 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient with an cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced some complications requiring intervention, the patient was initially implanted with an impella cp device for unloading for 11 days before being escalated to the impella 5.5 device. the patient was successfully weaned from extracorporeal membrane oxygenation (ecmo) and decannulated. the patient was on multiple pressers/inotropes, intubated, and sent to the unit to recover. on the unit it was noted the 12fr sheath would be removed when practical. the axillary pocket was left open and packed with two jackson-pratt drains. later this day, there was evacuation of a hematoma from the pocket and replacement of jackson-pratt drains. the patient continued on support. four days thereafter, the patient was noted to be having fevers due to cholecystitis and bronchial cultures were positive. the patient was started on antibiotics and tylenol. the patient was unable to be extubated and was trached. fevers continued; the patient was placed on a cooling blanket. now, approximately four days later, a fasciotomy closure took place and blood products were given. restart of heparin was planned for the following day. however, two days later a lower gastrointestinal bleed was noted. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. blood products were administered, the patient remained on performance level p-7 with plan to continue cardiac rehabilitation. | Additional information was received. the patient was transported to another facility. upon automated impella controller (aic) transfer, it was noted to have an impella defective alarm. the aic was attempted to be switched multiple times with no ability to restart. the impella 5.5 was explanted due to the event and the patient was medically managed. |  |  |  |  | A4. patient weight unknown. product status: the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, peri-procedural adverse events: fever, in order to make a root cause determination, essential clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. access site bleeding: the root cause of access site bleeding could not be determined since the product was not returned, and insufficient clinical details were provided. vascular injury: the root cause of vascular injury cannot be determined since the product was not returned, and insufficient clinical details were provided. regarding the pump stop: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. no logs are available. h6: investigation: type, findings and conclusion codes and h6: component code were updated accordingly based on the completed investigation. | Additional information was received from the field clinical rep and the following was noted: the cholecystitis and positive cultures were not related to impella. however, the cause was still unknown. an opportunity to obtain more of the patient's health did not present itself. section h6 investigation coding remains unchanged. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and noted the support concluded for the patient. consequently, the device explant date was shared. d.6b was updated accordingly. additional information was received. b.5 and h.6 codes were updated accordingly. d.3 manufacturer name should not have contained a number behind it on the initial manufacturer report. e.1 initial reporter address line 2 and zip code extension were added as they were inadvertently omitted from the initial manufacturer report.
6600 | 1220648-2025-45796 | 22804990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22804990&pc=PYX | 07/16/2025 |  |  | 08/14/2025 |  |  | 07/16/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025531520 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 75 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device followed by an impella rp flex device for bipella mechanical circulatory support and following bipella implant, the patient began to develop hemolysis. four units of blood were given to treat the condition. the following day, the patient was diagnosed with ischemia in both extremities. lactic remains elevated, contributing to limb ischemia. the impella cp pump was removed, a fasciotomy was performed, and the patient was escalated to impella 5.5 support, which was placed with some difficulty. support ramped up and position was stable. later that same day, the physician believed that the right leg was dead. additional intervention was not performed however. also, it was noted the belly kept getting bigger, a lot of blood seen around the liver on transesophageal echocardiogram. it was noted at that said time, there was no plan for any intervention. due to the patient's prognosis, care was withdrawn and the patient expired. | [5 empty] | Given the details of the event, there is not enough evidence to exclude the impella cp as an associated factor in the patient's demise outcome. hemolysis along with limb ischemia present that progressed and along with vascular surgical repair related to the impella cp device removal. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was inadvertently omitted from the initial report that was submitted. e. 1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. h6: investigation: component code and type of investigation codes were updated accordingly based on the completed investigation. h.11 the demise outcome is associated with the impella cp capture under complaint file (b)(4). the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. per the report, on the day of the impella implant, the patient had tinged urine with cause unknown and was given four units of packed red blood cells with no resolve. the cause of the hemolysis was not determined as the product was not returned, data logs were not recovered, and there is insufficient clinical information.
6601 | 1220648-2025-45845 | 22808866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22808866&pc=OZD | 07/16/2025 |  |  | 08/14/2025 |  |  | 07/16/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679212 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient presented to outside emergency department with chest pain and found to be in anterior st elevation myocardial infarction and diabetic ketoacidosis. the patient was transferred to the complainant hospital and was urgently brought to the cardiovascular laboratory. during left heart vascularization, the patient became unstable and required high dose pressors and intubation. the decision was made on (b)(6) 2025 to place an impella cp (serial number: (b)(6) prior to percutaneous intervention. on (b)(6) 2025, the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo). after 13 days of support, the impella cp abruptly stopped pumping and was unable to be restarted. the decision was made to replace the impella cp with an impella 5.5 (serial number: (b)(6), the patient remained also on va emco support. the patient was noted to be stable after impella 5.5 was implanted. the complainant reported that the patient experienced a large hemorrhagic stroke during impella 5.5. support. the patient¿s family decided to withdraw care. care was withdrawn and the patient expired. this complaint involves 2 pumps: pump 1: impella cp serial number: (b)(6). pump 2: impella 5.5, serial number: (b)(6). this mdr specifically addresses impella 5.5 with serial number: (b)(6) which is the subject of this report. a separate medwatch report will be submitted for the other device associated with this complaint. | [5 empty] | Given the details of the event, there is not enough evidence to exclude the impella cp as an associated factor in the patient's demise outcome. a pump stop related to the impella cp device removal cannot be excluded as an associated factor in the patient¿s outcome. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6602 | 1220648-2025-45820 | 22808875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22808875&pc=OZD | 07/16/2025 |  |  | 08/14/2025 |  |  | 07/16/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688268 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Thrombosis/Thrombus | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp via the right femoral artery and support initiated on performance level p-6. prior to the impella cp implant, the patient had recurrent ventricular tachycardia/fibrillation arrest and was shocked multiple times with return of spontaneous circulation. after implant, percutaneous coronary intervention to the right coronary artery was completed. subsequently, there were no distal pulses in the right leg, and the patient was taken immediately to the operating room for escalation to the impella 5.5 device which was successfully completed. the patient was sent to the unit on impella support. day 4 of support noted the patient was no longer experiencing ectopy, received 2l of fluid, and was hemodynamically stable with no suction alarms on p-4. the plan was to wean to p-2 and ambulate the patient to the chair. two days later it was noted the patient received two units of packed red blood cells, still having melena. a check for gastrointestinal bleed was confirmed. bivalirudin was suspended. it could not be determined if the use of the anticoagulant for the impella impacted or exacerbated the outcome of the bleed. however, the patient was hemodynamically stable on p-5, and revascularization was planned once hemoglobin stabilized. the next day bleeding slowed, and the patient was administered one unit of packed red blood cells this morning and taken to the catheterization lab for revascularization. the patient remained hemodynamically stable. three days later the patient was successfully weaned from the impella support and the impella 5.5 device was explanted without incident with a patient surviving outcome. | Initial report: the complainant reported that a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp via the right femoral artery and support initiated on performance level p-6. prior to the impella cp implant, the patient had recurrent ventricular tachycardia/fibrillation arrest and was shocked multiple times with return of spontaneous circulation. after implant, percutaneous coronary intervention to the right coronary artery was completed. subsequently, there were no distal pulses in the right leg, and the patient was taken immediately to the operating room for escalation to the impella 5.5 device which was successfully completed. the patient was sent to the unit on impella support. day 4 of support noted the patient was no longer experiencing ectopy, received 2l of fluid, and was hemodynamically stable with no suction alarms on p-4. the plan was to wean to p-2 and ambulate the patient to the chair. two days later it was noted the patient received two units of packed red blood cells, still having melena. a check for gastrointestinal bleed was confirmed. bivalirudin was suspended. it could not be determined if the use of the anticoagulant for the impella impacted or exacerbated the outcome of the bleed. however, the patient was hemodynamically stable on p-5, and revascularization was planned once hemoglobin stabilized. the next day bleeding slowed, and the patient was administered one unit of packed red blood cells this morning and taken to the catheterization lab for revascularization. the patient remained hemodynamically stable. three days later the patient was successfully weaned from the impella support, and the impella 5.5 device was explanted without incident with a patient surviving outcome. additional events: additional information was received via notifications from abiomed's long term outcome and quality indicator impella registry and noted the patient experienced the following additional events: deep vein thrombosis, bleeding with anticoagulant withheld, hematoma and infection requiring intervention. the patient would ultimately expire. right upper extremity deep vein thrombosis with a "possibly" relationship to the impella used. further information noted heparin drip was started and withheld due to gastrointestinal bleed and interventional radiology. the heparin was restarted the following day. pelvic hematoma with a "possibly" relationship to the impella device used. a repeat computed tomography showed ascites. a computer tomography was completed due to concern for ileus. gas forming with in hematoma with concern for superinfection. infectious disease was consulted, and the patient was treated with antibiotics such as meropenum. the site was drained, and the patient had a jp bulb for suction. infectious disease was consulted for monitoring of infection. however, the patient expired died prior to event resolving. pelvic hemoperitoneum with a "possibly" relationship to the impella device used. further information noted that trauma services was consulted for pelvic fluid collection collecting infected hemoperitoneum. the physician drained the pelvic collection in interventional radiology. a jp drain with bulb suction was placed. fluid had positive culture for heavy klebsiella/oxytoca/raoultella species. the patient was on meropenum, unaysn and cipro. the patient expired prior to event resolving. |  |  |  |  | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received noting additional patient events, and based on this additional information, the following sections were updated b5: event description. the following fields were repopulated a2: outcomes attributed, h1: type of reportable event, and h6: health effect - clinical and impact codes. (note: the actual date of death is unclear and therefore this field was left blank). investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the following adverse events: retroperitoneal hemorrhage: patient had upper gastrointestinal bleed and bival was kept on hold. it was unclear whether the bleeding was resolved based on the provided clinical information. the root cause of the injury was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. hematoma: the registry reported upon ae #8 for subject of a pelvic hematoma. the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. infection/ thrombosis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Product complaint # (b)(4). no device received: the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6603 | 1220648-2025-31161 | 22809314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22809314&pc=OZD | 07/17/2025 | 08/26/2025 |  | 08/14/2025 |  |  | 07/17/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589254 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient on impella 5.5 support exhibited signs and symptoms of a cerebrovascular accident (cva) and confirmed middle cerebral artery occlusion with head computed tomography (ct) scan. it was further reported that the patient was diagnosed with cva approximately 10 hours after receiving the impella 5.5. when the patient was taken for the ct scan, the neurologist determined that the cva was experienced 24 to 48 hours previous to the scan that was done at the 10-hour mark post procedure. due to these findings, the team determined it was not impella related. it was unknown the definitive cause of the stroke. potential anoxic brain injury secondary to hypoperfusion during resuscitative efforts was the common understanding. there is not enough information to exclude the impella device as an associated factor in the patient's cva. | [5 empty] | The medical safety officer reviewed the event after conservatively reporting to regulatory agency and determined the impella can be dissociated from the cerebrovascular accident (cva); there is no identifiable connection. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B1 adverse event erroneously entered on the previous manufacturer device report number should be left blank. h1 type of reportable event erroneously entered on the previous manufacturer device report number updated to n/a.
6604 | 1220648-2025-45802 | 22809424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22809424&pc=OZD | 07/21/2025 |  |  | 08/14/2025 |  |  | 07/21/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610480 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had bleeding form the access site. this was treated by transfusing one unit of packed red blood cells. the outcome status post the event was indicated as unknown. additionally noted, however, was the impella 5.5 device was subsequently explanted three days later per the physician's request so an mri could be completed to check neuro status. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks.
6605 | 1220648-2025-45807 | 22812006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22812006&pc=OZD | 07/21/2025 |  |  | 08/14/2025 |  |  | 07/21/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026715956 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 60 KG |  |  | Death | Yes | No | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. however, during removal of the impella cp inner dilator, the clear locking hub detached and could not be removed. the pump was implanted through the diaphragm and the clear ring was removed with surgical shears. the patient was sent to the unit on support. the next day it was noted the patient experienced hemolysis. the pump was repositioned to resolve the issue. later, the following day, care was withdrawn and the patient expired. | [5 empty] | Neither the customer's device nor data logs from the customer's device were available for investigation. the only data logs available from the case was an empty lt log from (b)(6) 2025. it was not reported how the pump had come out of position. data logs were not available for review. since insufficient clinical details were provided, data logs weren't available, and the product wasn't returned for investigation, the cause of the positioning issues could not be determined. based on the clinical details provided, it was not confirmed that repositioning the pump resolved the reported hemolysis. since insufficient clinical details were provided and data logs nor product were available for investigation, the cause of the hemolysis could not be determined. | Patient information: a5 and a6 are unknown. product status: the impella cp device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6606 | 1220648-2025-45803 | 22812196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22812196&pc=PYX | 07/22/2025 |  |  | 08/14/2025 |  |  | 07/22/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026675603 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced access site bleeding which was addressed with an additional hemostatic suture. the pump was explanted and the patient survived. | [5 empty] | The device passed all post sterile inspection checks. the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. | Patient information: a4 is unknown. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6607 | 1220648-2025-45804 | 22812817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22812817&pc=OZD | 07/22/2025 |  |  | 08/14/2025 |  |  | 07/22/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664679 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 101 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding and required a transfusion of red blood cells. the patient was transported for a peripheral angiogram and coded upon arrival in the lab, and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks.
6608 | 1220648-2025-45847 | 22812840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22812840&pc=OZD | 07/26/2025 | 08/12/2025 |  | 08/14/2025 |  |  | 07/26/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701691 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing coronary artery bypass surgery and mitral valve repair was implanted with an impella 5.5 device for mechanical circulatory support. while on support, some runs of ventricular tachycardia (vt) were managed in an unspecified way. lidocaine and amiodarone added for vt but there were still some premature ventricular contractions. tachyarrhythmia was managed. the clinical staff was unable to wean the patient off of vasopressors due to hypotension. the patient remained on impella support and was successfully weaned. | [5 empty] | B.7 added information that was inadvertently omitted from the initial report that was submitted. the investigation was completed and the device was not returned for investigation. the cause of the tachycardia/hypotension injury was not determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | The patient's ethnicity and race are unknown at this time. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6609 | 1220648-2025-45848 | 22812855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22812855&pc=OZD | 07/25/2025 | 08/13/2025 |  | 08/14/2025 |  |  | 07/26/2025 | 08/14/2025 |  |  |  |  | 08/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657783 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing right heart catheterization was implanted with an impella cp device for mechanical circulatory support. while on support, the patient suffered an ischemic stroke with minimal left sided weakness noted. support continued uninterrupted for two more days, at which point the patient's family made the decision to withdraw care. the patient expired following pump removal. | [5 empty] | Product status: the impella device was not received from the customer; there is uncertainty what happened with the device after the patient's demise. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are section: potential adverse events- ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions- ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation was completed however there was no device returned. the root cause of the injury was unable to be determined due to insufficient clinical details. a device history review was done and passed all the post sterile inspection checks.
6610 | 1220648-2025-31158 | 22793769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22793769&pc=OZD | 07/24/2025 | 12/11/2025 |  | 08/13/2025 |  |  | 07/25/2025 | 08/13/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1643465 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had an impella pump stop per the bedside nurse ¿without warning.¿ the bedside care team attempted to restart the impella on the original automated impella controller (aic) and received a ¿controller failure¿ alarm. the aic was swapped and attempted to restart the pump on a new aic. the pump would not restart per the bedside nurse. the patient expired while on support which is reported on manufacturer report number: 1220648-2025-31090. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - pump stop has been completed. console logs were consistent with what was reported in the complaint. console was tested after arriving in field service. the reported pump stop issue was unable to be reproduced. conclusion: 5.5 pump 586680 ran for 20 days on ic9544 before a pump stop occurred. multiple attempts to resolve the pump stops were made but were unsuccessful. this pump was ultimately transferred to ic9543 to continue support but the pump stops persisted with this console as well. there was no issue found with this console. the issue followed the pump based on data analysis. upon review of the data on the aic, it is apparent that the pump is the potential cause of the issue. see manufacturer report number (b)(4) for pump investigation results. d9 device returned to manufacturer date added. | Additional information was received in the form of medical safety review and noted the following: an adult patient presented in cardiogenic shock on (b)(6) and underwent multiple rounds of cpr. an intra-aortic balloon pump (iabp) was placed for initial support. on (b)(6), the patient was escalated to impella cp for higher hemodynamic support. due to worsening hemodynamics, va-ecmo was initiated for primary mechanical circulatory support, and impella 5.5 was implanted for left ventricular (lv) unloading. on (b)(6), the patient exhibited signs of infection (no documented results). on (b)(6), multiple deep vein thromboses (dvts) and a pseudoaneurysm at the axillary site were noted. on (b)(6), the impella 5.5 stopped unexpectedly; attempts to restart the device failed. the device was left in the aortic position. the patient expired on (b)(6) while supported on ecmo and the non-functioning impella 5.5. device involvement: impella cp functioned as intended; no device-related complaints. escalation was due to patient¿s clinical deterioration, not device malfunction. impella 5.5 was an unexpected pump stop on (b)(6) with failed restart attempts. device remained in situ until patient expired. considered a contributing factor to patient death, though underlying condition was severe. automated impella controllers (aic) conservatively reporting as malfunction. based on reported information, the impella cp performed as expected; the impella 5.5 exhibited pump stop during use. the patient¿s death primarily attributed to underlying clinical condition; device malfunction may have contributed. note: type of reportable event and investigation with coding remain unchanged. this is one of three device reports associated with the event. refer to the following report numbers for the related device reports: 1220648-2025-31158 (aic), 1220648-2025-31159 (aic) 24-46884-1 and 1220648-2025-31090 (pump).
6611 | 1220648-2025-31159 | 22794030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22794030&pc=OZD | 07/24/2025 | 12/11/2025 |  | 08/13/2025 |  |  | 07/25/2025 | 08/13/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1643465 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had an impella pump stop per the bedside nurse ¿without warning.¿ the bedside care team attempted to restart the impella on theoriginal automated impella controller (aic) and received a ¿controller failure¿ alarm. the aic was swapped and attempted to restart the pump on the new aic. the pump would not restart per the bedside nurse. therefore, the second controller is being reported on this report. the patient expired while on support which is reported on manufacture report number 1220648-2025-31090. | [5 empty] | Additional information was received in the form of medical safety review and noted the following: an adult patient presented in cardiogenic shock on 07/01 and underwent multiple rounds of cpr. an intra-aortic balloon pump (iabp) was placed for initial support. on 07/02, the patient was escalated to impella cp for higher hemodynamic support. due to worsening hemodynamics, va-ecmo was initiated for primary mechanical circulatory support, and impella 5.5 was implanted for left ventricular (lv) unloading. on 07/10, the patient exhibited signs of infection (no documented results). on 07/14, multiple deep vein thromboses (dvts) and a pseudoaneurysm at the axillary site were noted. on 07/24, the impella 5.5 stopped unexpectedly; attempts to restart the device failed. the device was left in the aortic position. the patient expired on 07/25 while supported on ecmo and the non-functioning impella 5.5. device involvement: impella cp functioned as intended; no device-related complaints. escalation was due to patient¿s clinical deterioration, not device malfunction. impella 5.5 was an unexpected pump stop on 07/24 with failed restart attempts. device remained in situ until patient expired. considered a contributing factor to patient death, though underlying condition was severe. automated impella controllers (aic) conservatively reporting as malfunction. based on reported information, the impella cp performed as expected; the impella 5.5 exhibited pump stop during use. the patient¿s death primarily attributed to underlying clinical condition; device malfunction may have contributed. this is one of three device reports associated with the event. refer to the following report numbers for the related device reports: 1220648-2025-31158 (aic), 1220648-2025-31159 (aic) 24-46884-1 and 1220648-2025-31090 (pump). | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported issue has been completed. data analysis: on the 20th day of support with pump sn: (b)(6), the console displayed ¿impella stopped (mechanical/electrical issue) ¿ alarm, event #115. the rt log confirms the pump stop. this confirms the reported failure mode. despite the user restarting the pump multiple times, it consistently stopped immediately. after transferring the pump to the second console, ic9543, the same issue persisted. device analysis: this console was tested after arriving at fs. unable to reproduce the reported failure mode - pump stop upon power cycling the console 20 times. each time, the console was able to run the test pump at p-level p9 without automatic stopping. there was a 20v supply at the impellatronic board connector j001.2, as intended. no hardware-related issues were found on the console. root cause: no issues were found with the console (ic9544). the root cause of the pump stop could not be determined due to the inability to reproduce the issue. h6 investigation type, findings and conclusions.
6612 | 1220648-2025-45799 | 22794504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22794504&pc=OZD | 07/14/2025 |  |  | 08/13/2025 |  |  | 07/14/2025 | 08/13/2025 |  |  |  |  | 08/13/2025 |  | PC-001938978 |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687456 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 102 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device and then escalated to an impella 5.5 device for continued mechanical circulatory support. two days after start of the impella 5.5 device support, the patient was running a fever and cultures were taken. the patient continued on support and was being worked up by advanced heart failure. two days later, the patient was being weaned from extracorporeal membrane oxygenation (ecmo) as tolerated and noted to be septic and coagulopathic. pressure dressing was applied to the site with no new drainage. however, there was bleeding from other sites and system anticoagulation was suspended. the next day bleeding from the site was noted and then the following day the bleeding was uncontrollable and leading the team towards impella 5.5 removal. noted update reflected "will resolve sepsis and reassess device configuration." patient was eventually decannulated and continued on impella mcs for an additional three days. the impella device was successfully weaned and explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, peri-procedural adverse event: access site bleeding, the root cause was not determined as the product was not returned and insufficient clinical details provided. for infection/sepsis, in order to make a root cause determination, essential clinical details would have to be provided. therefore, the root cause of the infection/sepsis was unable to be determined as well due to insufficient clinical details. h6: investigation: type, findings and conclusion codes were updated accordingly based on the completed investigation.
6613 | 1220648-2025-31160 | 22796969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22796969&pc=OZD | 06/10/2025 | 08/13/2025 |  | 08/13/2025 |  |  | 06/10/2025 | 08/13/2025 |  |  |  |  | 08/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613899 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported purge tubing kink and high purge pressure during support. it was reported that there was a purge system blocked alarm due to unsolvable kink in purge line. as a result, recommendation was given to change the purge system. the decision was made to get the new system from the cath lab and to call back support for step-by-step instructions on how to change purge system. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high risk pci section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ d.4 unique identifier (udi) # are unknown.
6614 | 1220648-2025-45821 | 22798419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22798419&pc=OZD | 07/23/2025 |  |  | 08/13/2025 |  |  | 07/23/2025 | 08/13/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629034 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, the patient experienced bleeding from the axillary access site of the impella. the patient was taken to the operating room for a left axillary exploration and washout to manage the condition. a suture hole was discovered in the artery causing the bleeding, and it was sutured to repair. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complications. | Additional update noted that the physician repositioned the pump to address the ectopy. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product was returned. major bleed: the cause of the major bleed was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. vascular injury: the root cause of vascular injury cannot be determined since the product was not returned, and insufficient clinical details were provided. arrythmia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review pump passed all post sterile inspection checks. | H6 medical device problem code a01 was inadvertently omitted. | Correction: section b5: the information in b5 was inadvertently left out in the initial report which has been added to this report. section h6 : the a ( problem code ) code and f (impact code ) code have been added to this report reflecting the update. section h3: to correctly indicate the device has not been evaluated by manufacturer as the section was erroneously marked "yes."
6615 | 1220648-2025-45761 | 22780077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22780077&pc=OZD | 07/26/2025 |  |  | 08/12/2025 |  |  | 07/26/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026694633 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after being diuresed overnight, suction alarms began to appear during impella 5.5 support. three units of blood were transfused response to the suction issues. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D.3 manufacturer name should not have contained a number behind it on the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. the impella device was not received from the customer, but the event logs were received. data logs showed the pump ran without issue during support. there were suction alarms triggered during the case but they were resolved. the cause of suction issue was most likely patient condition related since the suction was resolved by fluid compensation. the device history review was completed and the pump passed all post sterile inspection checks.
6616 | 1220648-2025-45765 | 22780127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22780127&pc=OZD | 07/26/2025 |  |  | 08/12/2025 |  |  | 07/26/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698106 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 86 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the sheath and access site and a hematoma. the sheath was repositioned and pressure was held however, the decision was made to remove the impella. the outcome status of the patient following the event is unknown. | [5 empty] | Product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system - section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The customer's device was not returned for investigation. the root cause of bleeding was most likely use issue related since the bleeding started to occur after the sheath was exposed and at the same time, it started to bleed more after the patient was moved to the icu the device passed all post sterile inspection checks.
6617 | 1220648-2025-45735 | 22781711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22781711&pc=OZD | 07/22/2025 |  |  | 08/12/2025 |  |  | 07/22/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553756 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 94 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during impella 5.5 support, the patient experienced ventricular tachycardia (vt) that required defibrillation. on (b)(6) 2025 it was noted that the patient experienced additional episodes of non-sustained. the following day, the family decided to withdraw care from the patient. care was withdrawn, and the patient expired. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 added selection as it was inadvertently omitted from the initial report that was submitted. d.4 model was added as it was inadvertently omitted from the initial report that was submitted. revised primary udi number as it was previously entered incorrectly on the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. h.6 added health effect - impact code as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. the pump was not returned for investigation. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all post sterile inspection checks.
6618 | 1220648-2025-45780 | 22782443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22782443&pc=OZD | 07/26/2025 | 08/11/2025 |  | 08/12/2025 |  |  | 07/26/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688321 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the impella was inserted and the peel away sheath was removed, the repositioning sheath was inserted via the right femoral artery access site. the doctor noticed bleeding from the site. the aggrastat drip was stopped and manual pressure was held at the site for twenty minutes with no improvement. the impella was removed as hemostasis could not be achieved at the sheath site. a 16 french sheath was placed to stop the bleeding and the patient was taken to the operating room for surgical repair. the complaint patient outcome was reported as stable. | [5 empty] | The patient's ethnicity is unknown at this time. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | E.1 revised initial reporter phone number as it was entered incorrectly on the initial report that was submitted. added zip code extension as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. the device was not returned for investigation. peri-procedural adverse event: access site bleeding: there was uncontrollable bleeding form the right femoral artery access site and hemostasis could not be achieved. the decision was made to remove the impella. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. a device history review was completed and the device passed all post sterile inspection checks.
6619 | 1220648-2025-31153 | 22783393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22783393&pc=OZD | 02/14/2023 | 10/27/2025 |  | 08/12/2025 |  |  | 02/14/2023 | 08/12/2025 |  |  | 09/30/2024 |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277317 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the placement signal was lost. the patient stayed hemodynamically stable and continued to receive effective support from the pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. the data logs were reviewed, and it was found that the trend of signal not reliable and light parameters was decreasing while lamp, gain and contrast were increasing. the root cause of the optical signal issue was most likely damaged optical fiber line due to kink in the fiber line based on the data log review. | All pertinent information available to abiomed has been submitted at this time. d6a: corrected the implant date. d6b: corrected the explant date. | G3 ¿date received by manufacturer¿ corrected.
6620 | 1220648-2025-31154 | 22783483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22783483&pc=OZD | 11/02/2023 | 08/12/2025 |  | 08/12/2025 |  |  | 11/02/2023 | 08/12/2025 |  |  | 08/31/2025 |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394407 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, the pump triggered repeated ¿placement signal low¿ alarms. the impella was initially repositioned in an attempt to resolve the issue; however, the alarm recurred. as correct positioning was confirmed via echocardiography, the decision was made to disable the alarm. support was maintained with the implanted pump, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was not returned for investigation. data logs were investigated and "placement signal low" was confirmed. as per clinical, the pump was repositioned after a dip in the placement signal was observed. once the pump was repositioned, the placement signal waveform is seen to be improving. the light, gain and lamp appear to be normal and within spec of the pump and contrast was a little variable along with decrease in signal not reliable during the failure which rose up soon after. per the provided clinical information and data log review, the root cause of the issue was not determined since product was not returned and data logs were inconclusive.
6621 | 1220648-2025-31155 | 22786217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22786217&pc=OZD | 07/18/2025 | 09/08/2025 |  | 08/12/2025 |  |  | 07/18/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633380 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with complains of shortness of breath, was cathed and patient declined surgery. taken to cardiac catheterization laboratory for percutaneous coronary intervention. during case dissection occurred distal to stent, able to place wire down lesion. patient quickly decompensated and required cardiopulmonary resuscitation (cpr). once cpr initiated, decision was made for impella cp. while patient was having cpr multiple checks for pulse revealed pulseless electrical activity. decision made to place impella, lfa while doing compressions. patient shocked multiple times with return of spontaneous circulation obtained. impella inserted with no issues and started on p-9. while removing the peel away sheath and insertion of repositioning sheath the catheter advanced into the ventricle. while in the ventricle, the impella kinked on itself. multiple attempts made to unkink the catheter. patient decompensated while impella was not working correctly. cpr started while impella was being removed. impella was pulled down through the descending aorta and the pigtail and inlet portion straightened out into the right common iliac artery. dr. went to remove the repositioning sheath and impella with a wire through the side port as vessel access, and the pigtail and inlet became dislodged from the cannula. the nitinol coil unraveled and came out with the impella, while the inlet and pigtail remained in the right common iliac artery. decision was made to leave it at this time. 14x25cm sheath was prepped and inserted over the wire to replace the impella access. new impella was placed and the patient later expired. the patient's outcome is reported on manufacture report number 1220648-2025-30890. | [5 empty] | The failure mode product damage (catheter kink) was changed to product damage (cannula kink) since it was reported that the device folded on itself within the left ventricle (lv) which can only occur on the cannula. the pump was returned cut and inspected which showed the unraveled nitinol coil of the cannula. product damage (cannula kink): catheter position is intended to be maintained during a sheath exchange. the cause of the product damage (cannula kink) was use related since the pump went deep in the lv during the sheath to repo unit exchange resulting in a the cannula kinking and folding on itself. cannula/pigtail detachment - great vessel/heart: the unravelling of the cannula indicates that there was no potential bond failure and that the detachment was a result of excessive force. the cause of the cannula detachment was use related since the pump went deep in the lv during the sheath to repo unit exchange resulting in a the cannula kinking and folding on itself. the detachment occurred due to the cannula folded and stuck on itself unable to be freed without excessive force. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6622 | 1220648-2025-31156 | 22786359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22786359&pc=OZD | 07/19/2025 | 09/15/2025 |  | 08/12/2025 |  |  | 07/19/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693526 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 had a "placement signal not reliable" alarm. aortic and left ventricle placement signal waveforms remained but were erratic. the following day the placement signal were working again. the patient was noted not to be improving therefore care was withdrawn and the patient expired. there is no information that reasonably suggest that the impella 5.5 pump cause or contributed to the demise outcome. | [5 empty] | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. the data logs reviewed and the cause of the optical signal issue could not be definitively determined as no product was returned for evaluation and limited clinical details were provided. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6623 | 1220648-2025-31157 | 22786636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22786636&pc=OZD | 07/23/2025 | 09/17/2025 |  | 08/12/2025 |  |  | 07/23/2025 | 08/12/2025 |  |  | 01/31/2026 |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536978 |  | 0550-0004 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed a impella 5.5 to support the patient with cardiogenic shock and known arrhythmia. the pump was supporting when after 3 days there was thrombus and purge pressure/purge flow issues. the team consulted with abiomed medical team and the purge had tissue plasminogen activator was added. purge pressure and purge flow in normal ranges within a few hours of the infusion being initiated. the pump remained on for support for just under 9 days. there was no success with ablation and therapy was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure and thrombus has been completed. the impella device was not returned for evaluation. high purge pressure: the data logs confirmed high purge pressure and purge system blocked alarms. the logs show a gradual rise in purge pressure and decrease in purge flow for about an hour before alarms were triggered. there was no motor current spike before high purge pressure the high purge pressure was sustained for about 8 hours before returning to normal values. the root cause of the high purge pressure was most likely biomaterial as evidenced by the gradual rise in purge pressure and decrease in purge flows observed in the logs, and the resolution after tpa addition. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
6624 | 1220648-2025-45777 | 22787064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22787064&pc=OZD | 07/20/2025 |  |  | 08/12/2025 |  |  | 07/20/2025 | 08/12/2025 |  |  | 02/28/2026 |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492515 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was noted that femoral insertion was not possible so the medical team opted for axillary insertion. it was reported that during impella cp support, the patient experienced a cold arm where the impella was inserted. as a result of the cold arm, the impella cp was removed. during the explant a thrombus in the brachial artery was observed. a thrombectomy was performed. patient outcome was survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the thrombus/ischemia was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
6625 | 1220648-2025-45768 | 22787718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22787718&pc=PYX | 07/26/2025 |  |  | 08/12/2025 |  |  | 07/26/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026630643 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device followed by an impella rp flex device for bipella mechanical circulatory support. one day after start of the bipella support, the patient lost pulses in the foot. upgrade to an impella 5.5 from the impella cp device was being considered. the following day hemolysis was noted. the physician decided to remove the rp flex device at the bedside. an echocardiogram completed and showed impella cp a approximately 5cm and was, therefore, repositioned to measure 3.4 cm. the next day thereafter, it was noted the patient had a gastrointestinal bleed; blood products were given. the bleed worsened the following day and heparin drip was suspended. this day it was also noted poor to absent pulses had persisted over the weekend. the impella side leg looked ischemic, and the patient required amputation in response to the ischemia. the patient's family decided to provide comfort care only until the patient expired. | [5 empty] | Based on the details of the events during support, there is not enough evidence to exclude the impella cp device as an associated factor to the patient's demise outcome. the impella rp flex was explanted, patient remained on impella cp support and experienced further complications one leading to amputation. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was not returned for investigation. data logs show no signs related to motor current spikes, suction, or positioning issues during the case run. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical information. the pump passed all post sterile inspection checks.
6626 | 1220648-2025-45738 | 22787753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22787753&pc=OZD | 07/25/2025 |  |  | 08/12/2025 |  |  | 07/25/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687483 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 98 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device followed by an impella rp flex device for bipella mechanical circulatory support. one day after start of the bipella support, the patient lost pulses in the foot. upgrade to an impella 5.5 from the impella cp device was being considered. the following day hemolysis was noted. the physician decided to remove the rp flex device at the bedside. an echocardiogram completed and showed impella cp a approximately 5cm and was, therefore, repositioned to measure 3.4 cm. the next day thereafter, it was noted the patient had a gastrointestinal bleed; blood products were given. the bleed worsened the following day and heparin drip was suspended. this day it was also noted poor to absent pulses had persisted over the weekend. the impella side leg looked ischemic, and the patient required amputation in response to the ischemia. the patient's family decided to provide comfort care only until the patient expired. | [5 empty] | No product was returned for evaluation. the cause of the retroperitoneal hemorrhage cannot be determined as insufficient clinical details were provided. the root cause of this ischemia was unable to be determined due to insufficient clinical details. the cause of the device in wrong position could not be definitively determined as insufficient clinical details were provided. the cause of the mechanical interaction with blood could not be definitively determined as data logs showed no definitive positioning or pump issues and clinical details did not note if repositioning resolved hemolysis. the pump passed all post-sterile inspection checks. | Based on the details of the events during support, there is not enough evidence to exclude the impella cp device as an associated factor to the patient's demise outcome. the impella rp flex was explanted, patient remained on impella cp support and experienced further complications one leading to amputation. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states)-¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions - including catheter position, pre-existing medical conditions, and small left ventricular volumes - may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6627 | 1220648-2025-45797 | 22790127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22790127&pc=OZD | 07/26/2025 | 08/12/2025 |  | 08/12/2025 |  |  | 07/26/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708778 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 128 KG |  | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, repositioning of the impella cp into a better position in ventricle was attempted as it was pointed into the mitral valve and under the papillary muscle. the doctor was unable to reposition the device under x-ray and ultrasound guidance. the doctor attempted to exchange the impella via the existing site. it was determined that the initial vessel selected was either the superficial femoral artery or the profunda femoris artery. the doctor removed that access site as it would not thread the new device. a large hematoma formed around the left femoral artery site, which required blood products. a new site was obtained in the left femoral artery and a new impella cp was placed. the patient¿s groin and hematoma did stabilize after manual pressure and blood products were given. the patient¿s outcome was stable following the procedure. | [5 empty] | The ethnicity of the patient is unknown at this time. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | H6 medical device problem code was inadvertently omitted on the initial manufacture report. the investigation was completed. no pump returned for investigation. clinical states impella cp was in the ventricle and was unable to be repositioned after patient was transported. the pump was explanted and swapped out. large hematoma formed at the access site and was managed with manual pressure and blood products. positioning: the cause of the positioning issue was determined to be an unintentional use error as the pump became malpositioned after the patient was transported per clinical. access site adverse event: the cause of the access site bleeding was not determined due to lack of clinical details and no product returned. a device history review was completed and passed all post sterile inspection checks.
6628 | 1220648-2025-45798 | 22790550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22790550&pc=OZD | 07/26/2025 | 10/10/2025 |  | 08/12/2025 |  |  | 07/26/2025 | 08/12/2025 |  |  |  |  | 08/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 86 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | Hematuria | [10 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's urine turned dark red and labs were hemolyzed. an echocardiogram was completed and the impella cp was measured deeply at 4.7 centimeters. a nurse practitioner repositioned the impella cp for a final measurement of 3.5 centimeters. the patient's urine was clear after repositioning. the left lower extremity was noted as cold and mottled. the clinical team was aware. care was withdrawn and the patient expired. | [5 empty] | Product status: the impella device was not received from the customer; there is uncertainty what happened with the device after the patient's demise. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported positioning issues, hemolysis, hematuria, and ischemia has been completed. the impella device was not returned for evaluation. data log review showed a small motor current spike on (b)(6) 2025 but no resulting changes in pump flow or motor current. the root cause of the positioning issues, hemolysis, hematuria, and ischemia could not be determined.
6629 | 1220648-2025-31131 | 22769482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769482&pc=OZD | 07/23/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 07/23/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687470 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 83 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/23/2025 |  |  |  |  | The user facility reported that following impella cp implant, the patient's labs hemolyzed and their urine turned dark brown. the p-level was reduced and the pump was repositioned via echocardiogram in response to the condition. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6630 | 1220648-2025-31132 | 22769539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769539&pc=OZD | 07/25/2025 | 09/22/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640529 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp placement, the repositioning sheath was placed partially in the peel away. the physician had to pull the repositioning sheath out of the peel away sheath, possibly with some force. when the repositioning sheath was placed in the groin after the peel away sheath was removed, there was oozing identified between the blue suture hub and the white piece with the abiomed logo on it. this oozing resolved after the physician tied it together with a suture. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. pump was returned back and issues were not upon investigation. product investigation has no impact in the failure investigation. the root cause of access site bleeding is most likely use issue because repositing sheath was placed into the peel away sheath before removing the peel away sheath which is not performed as per impella cp instructions for use and caused the bleed. bleeding was managed by applying suture. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6631 | 1220648-2025-31133 | 22769581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769581&pc=OZD | 07/25/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684326 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 84 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient developed limb ischemia. the patient lost pulses and developed a cold and modeled leg. the peel away sheath was left in place. the patient received a fasciotomy. later, the patient was never able to be weaned from max doses of epinephrine, vasopressin, levophed, and milrinone and continuous renal replacement therapy. the patient's liver and kidneys began to fail and severe acidosis set in leading to an eventual code where the patient expired on support. relationship between impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
6632 | 1220648-2025-31134 | 22769678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769678&pc=PYX | 07/25/2025 | 09/11/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026694434 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on central venoarterial extracorporeal membrane oxygenation (va ecmo) plus impella 5.5, had a redo mitral valve replacement. the patient went for decannulation of va ecmo and chest closure. the right ventricle looked decent on the transesophageal echocardiogram. the doctor placed an impella rp flex to bridge her through closing the chest as he has dealt with many right ventricles that decompensate after chest closure. va ecmo was removed. swan and introducer were removed. heparin was given and activated clotting time was 270. a new stick to right internal jugular was performed and an impella rp flex was placed under fluoroscopy and transesophageal echocardiogram guidance. the case went well. upon returning to the unit, the patient started having coagulopathy, urine turned tea colored, and labs were hemolyzing. two rounds of blood products were administered and protamine. the patient was taken back for a washout and looked for bleeding. no specific bleed was found, just generalized bleeding everywhere instead of the chest. the inlet of the impella rp flex was in the superior vena cava and outlet was three centimeters beyond the pulmonic valve. it was not thought to be a positioning problem from the impella rp flex. the impella rp flex never had suction either. three hours later, the patient was taken back for washout again. the patient continued to bleed from the chest tubes and stopped making urine. the impella rp flex site was not bleeding. four hours after the second washout, the surgeon removed the impella rp flex. there was no other changes that the medical doctor could think of that was leading to this generalized bleeding and the patient¿s labs showed gross hemolysis. the impella rp flex was removed. overnight, the bleeding stopped, and labs stopped hemolyzing. | [5 empty] | The investigation was completed. there was no product returned. clinical reports lab hemolyzing during support which resolved after explant. clinical confirmed good positioning via imaging with no patient anatomic abnormalities. blood products given with no resolution. the data logs show a motor current spike with speed deviation and shift in placement signal about 5 minutes after pump activation. there was also a drop in pump flow and purge flow with a rise in purge pressure. later on, during support, there was another ingestion pattern observed. the root cause of the hemolysis was most likely biomaterial as evidenced by the ingestion pattern observed soon after pump activation before reported hemolyzing. vascular damage - peripheral vessel: the root cause of the vascular damage - peripheral vessel was not determined as insufficient information related to failure was provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
6633 | 1220648-2025-31135 | 22769726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769726&pc=OZD | 07/25/2025 | 09/19/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686517 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 73 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was attempting to place an impella 5.5 for a patient with acute myocardial infarction/cardiogenic shock. the pump was unable to pass through anastomosis/graft and was replaced by an impella cp. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because pump was unable to pass through anastomosis graft due to stenosis in the vessels. 5.5 insertion was aborted and impella cp was used for implantation.
6634 | 1220648-2025-31136 | 22769782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769782&pc=OZD | 07/25/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685015 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 61 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was placed for higher-risk percutaneous coronary intervention (pci). the impella cp companion sheath was used for the pci and severe bleeding was noted after the companion sheath was placed. it appeared like a posterior score mark started opening when the companion sheath hubbed. the companion sheath was withdrawn a few centimeters and the issue was resolved. it looked as if multiple items in the peel away took up too much space and caused the score line to open. pci was completed and the impella cp was weaned and explanted successfully. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6635 | 1220648-2025-31137 | 22769814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769814&pc=OZD | 06/24/2025 | 11/10/2025 |  | 08/11/2025 |  |  | 08/07/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1681044 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with impella cp placement cardiogenic shock and ventricular septal defect vsd. however there was no optical sensor displaying on console. the physician was not comfortable with no waveforms, opted to remove and place new impella cp. however after a cross functional meeting with both the disposable and durable investigation teams the following actions have been identified. the reported optical signal issue(s) were found to be due in part to the console. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs from the reported day of event are consistent with complaint. cp pump 590658 is initially connected to ic10026 but there is a placement signal not reliable alarm and the user is requested to disconnect the cable and reconnect. the user follows the instructions, with a calculated g0 close to the factory and the pump is started, but immediately got a placement signal low alarm. the rtlogs show the first insertion had optical signals indicative of an air gap, and the second insertion despite measure optical signals as expected (and pulsatile), calculated placement signal was 0mmhg and did not change. device analysis: the 5s was measured and was as expected. additionally, the disconnect time of <1 second was attempted to be reproduced but was unsuccessful. similarly an air gap could be forced to occur by slightly disconnecting the optical connector but the placement signal flatlining at 0 could not be reproduced. investigation into a possible aic sw defect (gid-dft-3734730) concluded that the software operated as per design and correctly responded to all triggering criteria of pump disconnection and connection. root cause: the cause of the issue was not determined since issue could not be reproduced. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d4 unique identifier (udi) # updated. d10 concomitant medical products and therapy dates updated.
6636 | 1220648-2025-31139 | 22769982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22769982&pc=OZD | 07/24/2025 | 10/27/2025 |  | 08/11/2025 |  |  | 07/26/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691774 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | No Clinical Signs, Symptoms or Conditions | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered pump saturation and placement signal issues during impella 5.5 support. pump saturation was noted along with the patient decompensating. the staff then prepared to swap out the pump. while waiting the pump corrected itself and flows and pump metrics returned to normal. with the pump functioning normal the swap out was cancelled. overnight the placement signal began to give ¿in aorta alarms¿ despite being well placed. he flows were questionable. metrics became inconsistent with patient hemodynamics and the impella flows began to resemble pump saturation. as a result, the impella was exchanged to another impella 5.5. | The user facility reported an impella 5.5 was placed to support a patient with known heart failure/cardiogenic shock awaiting possible transplant. when admitted, there was a new facial droop and speech issues that prompted the team to have a computerized tomography scan done. there was no cerebrovascular accident observed. three weeks into the support, there was again cerebrovascular accident signs/symptoms that prompted imaging to reveal a hemorrhagic cerebrovascular accident. |  |  |  |  | The cause of the saturated flow was most likely due to biomaterial ingestion per returned product evaluation and data log analysis. | The initial manufacturer device report number 1220648-2025-31139 erroneously included the failure mode stroke against pump serial number (b)(6). the only failure modes that occurred while the patient was on support of impella cp serial number (b)(6) was optical signal issue and saturated flow. please see manufacturer device report number 30072025-0001904422 and manufacturer device report number. 17092025-0001934204 for serial number (b)(6) and the failure mode stroke with the outcome of death. the following corrections were made to manufacturer device report number 1220648-2025-31139. b1 product problem only b2 should have been left blank b3 date of event b5 patient narrative h1 type of reportable event; malfunction h6 health effect - clinical code, health effect - impact code, and medical device problem code corrected to reflect the failure modes reported. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6637 | 1220648-2025-31140 | 22770121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22770121&pc=OZD | 07/25/2025 | 09/22/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698768 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 53 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and the patient had heavily calcified iliac disease in femoral arteries. the pump was unable to pass calcified area and was removed. the patient's right iliac was then treated with shockwave and angioplastied. a new impella cp pump was then inserted successfully with no patient complications. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the product was returned for investigation and no kinks or abnormalities were observed. the root cause of the delivery issue was most likely determined to be patient condition related to the patient's calcified vessels.
6638 | 1220648-2025-31141 | 22770145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22770145&pc=OZD | 07/26/2025 | 09/22/2025 |  | 08/11/2025 |  |  | 07/26/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603576 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 92 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a patient admitted in with cardiac history, cardiomyopathy and hypoxia. the team observed signs of hemolysis. the urine color changed and labs hemolyzed. the team assessed pump positioning but the pump was appropriately placed within the left ventricle. the pump remained on with the hemolysis and family decision not to escalate. the care was withdrawn after under 76 hours. the patient expired. | [5 empty] | Data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. the investigation of hemolysis has been completed. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis." section: suction: "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure."
6639 | 1220648-2025-31142 | 22770514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22770514&pc=OZD | 06/14/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 07/28/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678954 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a placement signal not reliable alarm occurred. echocardigram confirmed proper positioning. per the investigation team, the data confirms a potential malfunction of the optical bench lamp. the automated impella controller was requested to be returned for evaluation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6640 | 1220648-2025-31143 | 22770633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22770633&pc=OZD | 07/28/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 07/28/2025 | 08/11/2025 |  |  | 07/31/2025 |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371858 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 65 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was placed for a coronary artery bypass graft/mitral valve replacement patient. the pump was placed successfully and heparin was held due to the post-operative nature until day two of support. the pump was explanted after just under four days of support and at that time there was a thrombus embolized down the arm and the team had to perform fogarty embolectomy. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
6641 | 1220648-2025-31144 | 22770820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22770820&pc=OZD | 07/30/2025 | 10/07/2025 |  | 08/11/2025 |  |  | 07/30/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1666954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had support ongoing when the automated impella controller (aic) had a frozen screen. the team was performing purge cassette replacement at the time. the aic was swapped/replaced per instructions for use recommendations. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - display issue has been completed. the impella device was returned for evaluation. the console logs are consistent with the reported issue. the root cause for the aic frozen on "please wait" screen with progress bar while purge cassette change was due to the defective purge unit driver (pud) pca. no issue with the aic were observed for the reported failure mode of aic - display issue.
6642 | 1220648-2025-31145 | 22770987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22770987&pc=OZD | 07/30/2025 | 11/25/2025 |  | 08/11/2025 |  |  | 07/30/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090363 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female | 112 KG | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) would not recognize the purge disk, multiple attempts were made to ensure the purge disk was inserted completely and correctly. a (second) new purge cassette was attempted as well with the same result, "purge disk not detected." the aic was replaced successfully. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. imc logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) several times. device analysis: the console was examined after arriving in fs. a broken purge disc flag was identified. conclusion: the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. es2023-0090 documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. d10 concomitant medical products and therapy dates updated.
6643 | 1220648-2025-31146 | 22771070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22771070&pc=OZD | 07/25/2025 | 09/19/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697054 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a patient on impella 5.5 support. while on support placement signal not reliable alarm occurred. the audio was disable. it was further reported that on day 18 of impella support there was a pump stop however the pump was able to restart. it was noted that the impella was not secured correctly on the patient and the catheter was kinked. the impella was exchanged and there was no patient harm. | [5 empty] | The investigation of optical signal issue, temporary pump stop, and product damage (catheter kink) has been completed. optical signal issue: clinical details report that the placement signal not reliable (psnr) alarm triggered on (b)(6) 2025. neither the pump nor console were replaced due to the issue. no further details were provided regarding the issue for the remainder of support. data logs were reviewed and the alarm triggered on (b)(6) 2025 as reported. at this time, snr dropped to zero, contrast dropped but increased in amplitude, lamp and gain increased, and light dropped. the alarm did not resolve for the remainder of the case. returned product was confirmed to have a kink in the catheter just distal to the red pump plug/kink protector. when the pump was connected to the console, all 5s parameters were still out of specification. the pump failed the laser continuity test with light leaking from the location of the catheter kink. there were no other light leaks noted on the pump, indicating that was the sole break. based on the signatures noted in data logs and the kink in the catheter/optical fiber line on returned product, the cause of the optical signal issue was determined to be a damaged fiber line. temporary pump stop: clinical details report that the pump stopped on (b)(6) 2025. the pump was able to be restarted, however, the decision was made to explant the pump and replace it with a new one. it was reported that the field believed a kink in the catheter was the cause of the pump stop. the pump was not secured properly to the patient. data logs were reviewed and the pump stop alarm was confirmed. there was a sudden trigger of alarm #119 with a motor current spike to 1250 ma. there were no abnormalities leading up to the event. the pump was able to be restarted, however, there were 6 more triggers of the same pump stop alarm before the pump was turned off. purge pressure remained stable throughout this time. returned product was investigated, and without manipulating the catheter at all, all motor cable lines were testing within specification. however, when testing the resistance in motor cable 3, kinking the catheter at the location where there was evidence of a previous kink would create an open circuit. the pump was then connected to a console, and the pump was able to start and run, however, when the kink in the catheter was exacerbated, the pump stop reproduced. this aligns with the intermittent nature of the pump stop alarms on field data logs. there were no other abnormalities on returned product: there were no signs of biomaterial, and the impeller was able to rotate without issue. based on the signatures seen on data logs and the intermittent open motor cable line on returned product, the cause of the temporary pump stop was determined to be an open electrical line. product damage (catheter kink): clinical details report that the catheter/pump was not properly secured to the patient, resulting in the catheter kink. returned product confirmed there was a kink in the catheter just distal to the red pump plug/kink protector. based on the clinical details provided that the catheter kinked due to it not being properly secured to the patient, the cause of the catheter kink was determined to be a use issue. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6644 | 1220648-2025-31147 | 22771371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22771371&pc=OZD | 07/25/2025 | 09/22/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678003 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support the patient admitted in with a cardiac history and st-segment elevation myocardial infarction. it was reported that while on support there was episodes of suction that seems to be unrelated to change in position or volume. abrupt separation in waveforms occurred with flows falling from 3.5 to 0.5. suction alarm only stopped at p-1. an echocardiogram performed and no obvious signs of obstruction or malposition noted. the patient then taken to cath lab for attempt at repositioning. at this time impella was advanced and pulled back to several positions and still unable to advance past p-1 without suction. the impella cp was exchanged for impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. data review: data logs confirmed the failure mode and clinical narrative. logs showed pump was started and run without issue for 69 hours, low flow occurred with motor current (mc) spike that triggered auto suction response & suction alarms. this event remained to the rest of the case. logs also showed pump was in aorta corresponded with the time per clinical description when pump was advanced and pulled back to several positions during reposition of the pump. pump was returned for analysis. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. conclusion: the root cause of the low flow was biomaterial ingestion. d9 device returned to manufacturer date added
6645 | 1220648-2025-31148 | 22771415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22771415&pc=OZD | 07/25/2025 | 09/19/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657658 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The us complainant had a bipella support placed of the impella 5.5 and rp flex. the patient had the 5.5 pump placed via direct access. while on support hematuria noted by bedside nurse, echocardiogram showed 5.5 to be shallow and outlet looked to be close to the graft. 5.5 was advanced 2cm and measuring 4.2cm now and urine is clearing up. the patient remains on support. |  |  |  |  | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. data review: data logs confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in aorta. pump then back to the good position and pump ran without issues to the rest of the case. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position since hemolysis was resolved after reposition of the pump. b1 updated to include death and date of death. b5 updated to patient outcome of death. h1 type of report updated to death based on patient outcome. h6 health effect - impact code 1802 added based on patient outcome. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6646 | 1220648-2025-45764 | 22771826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22771826&pc=OZD | 07/28/2025 | 10/23/2025 |  | 08/11/2025 |  |  | 07/28/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET, EU |  |  | OZD |  |  | GUIDE WIRE 0.018" |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025588739 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Device Damaged by Another Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered issues during pump placement. the guide wire could not be removed from cannula, and the guidewire was kinked. as a result the pump and the wire were removed and replaced. a new wire and a new pump were used. no patient harm was reported due to the issue. | [5 empty] | Section d9 and h6 has been updated as device was returned for investigation. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). additional information received that the guidewire with the issue reported in this event was non- abiomed guidewire. it belonged to a different manufacturer, so this report is deemed a not reportable. mdr is not required as the complaint was found against a non-abiomed product.
6647 | 1220648-2025-31149 | 22772640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22772640&pc=OZD | 07/25/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700945 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had some bleeding from the access site. pressure was held and the bleeding was resolved. the next day, there was a large hematoma at the access site. vascular surgery removed the impella in the or. the right femoral cutdown was done and the vessel was looped into the femoral artery for stabilization. after impella was removed, the femoral artery was repaired, hematoma evacuated, and groin closed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
6648 | 1220648-2025-31150 | 22772711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22772711&pc=OZD | 07/25/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 07/25/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616563 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp to support the patient admitted in with acute myocardial infarction and cardiogenic shock. the pump was placed but here were signs of hemolysis observed. the hemolysis was seen as urine color and urine production changes/reduction. the impella was slightly deep and was pulled back 1 cm. the following day the urine output was still dark therefore needing increased pressor and inotrope support, started on vaso and dobutamine. packed red blood cells were given to the patient as well. later weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6649 | 1220648-2025-45766 | 22772729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22772729&pc=OZD | 07/14/2025 |  |  | 08/11/2025 |  |  | 07/14/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681684 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered positioning issues, gastrointestinal bleeding (gi bleed), and hypotension during impella 5.5 support. impella 5.5 was initially prepped, implanted into optimal position and started. however, following chest closure and suction alarm there was concern for impella position and depth. transesophageal echocardiogram (tee), evaluated showed impella against the mitral valve and deep leading the physician to attempt optimization. due to poor tee imaging, left ventricle access was lost with impella in the aorta. impella placed in surgical mode while the physician attempted to recross without success. it was at this point when the chest had to be reopened with again difficult manipulation requiring explant. the physician then was able to cross the valve, exchange for .018 wire and successfully advance the impella into position. impella was restarted with optimal flows and good position. on (b)(6) 2025, initial plans were made to wean and explant the patient from the impella 5.5, however overnight the patient had significant episodes of hypotension and loss of pulsatility requiring increased impella support. epinephrine and dobutamine medications were restarted as a result. it was noted that the patient experienced a gi bleed during this time, and heparin was discontinued. the patient was transfused with two units of packed red blood cells (prbcs). the bleeding subsided the following day and heparin was restarted. on (b)(6) 2025, it was reported that the patient had a loss of consciousness after sitting in a chair and moved back to bed. after this episode, the patient was re-intubated. the patient was hypotensive and impella support was increased, however suction alarms would occur if the impella was increased above p5. the patient experienced shortness of breath and left pleural effusion. the patient was transfused with 1 unit of prbcs. the following day, the patient was transfused with an additional 2 units of prbcs. on (b)(6) 2025, heparin was discontinued due concerns of ongoing bleeding and a plan was made for impella 5.5 explantation. the patient was successfully weaned and explanted. the patient remained on vasoactive medication for support as recovery continued. | [5 empty] | The impella device was not received, and an investigation was completed. the root cause of the positioning issue was determined to be an unintentional use error the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6650 | 1220648-2025-45753 | 22773346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22773346&pc=OZD | 07/15/2025 |  |  | 08/11/2025 |  |  | 07/15/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the patient experienced hemolysis. the pump was repositioned twice to resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The customer's device was not returned for investigation. a review of the device's data log confirms the allegation and the observed motor current spike suggests biomaterial ingestion during the case run. the cause of the hemolysis was most likely biomaterial ingestion based on data log review. the root cause of the positioning issue could not be determined as an evaluation of the returned product was not performed and sufficient clinical information was not provided. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6651 | 1220648-2025-31151 | 22777342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22777342&pc=OZD | 07/02/2025 | 11/26/2025 |  | 08/11/2025 |  |  | 08/07/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1416992 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant was placing the impella 5.5 pump to support the patient admitted in with planned high risk surgery/coronary artery bypass grafting. the pump was found to have malfunctioning optical signal sensor before patient use. due to concerns with sensor/pump after insertion the pump was replaced. the new 2nd pump was placed. the investigating engineer of the impella 5.5 determined the issue likely stemmed from the automated impella controller and is requesting it to be returned for investigation. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: on the complaint date, july 2, 2025, the initial pump sn: (b)(6) was connected for 15 minutes, ideally without running. during this period, the rt log shows that the snr was low, below 2000. this confirms the reported failure mode - optical signal issue. although the initial pump was replaced with sn: 59312, the lt log indicates that the snr remained below 2000 throughout the case. device analysis: the console was placed on the bench and a test pump was ran at p-level 5 for several minutes. in this time, two ¿placement signal low¿ alarms were generated. the 5s parameters were evaluated using a fixed length test cavity and all values were within the ranges specified in preventative maintenance procedure 0042-7309, with the exception of contrast which was marginally low at 38%. the console was opened and a general inspection for damage, loose cables/connectors, etc. was performed with no issues noted. the analog output of the optical bench ccd array was evaluated with an oscilloscope and distortion was noted in multiple locations. root cause: the root cause of the placement signal issue was an optical bench with a distorted ccd waveform. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9: updated. h3: updated. d6a / d6b should have been left blank in the initial report. d10: concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6652 | 1220648-2025-31152 | 22777612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22777612&pc=OZD | 06/10/2025 | 08/11/2025 |  | 08/11/2025 |  |  | 08/08/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672997 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported by the abiomed investigation team that there was mechanical damage to introducer of the impella cp. during the case the patient remained in stable condition. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation. clinical details state that during a single-access procedure using the 7fr companion sheath, there was a continuous drip of blood from the sheath hub. the patient and hgb were stable. the sheath was returned and evaluated. leak testing was performed with the dilator inserted, and continuous leaking was reproduced from the hub below the 285-mmhg specification. no visual abnormalities were found with the valve. the cause of the introducer damage was most likely a damaged valve since the leak was reproduced below the 285-mmhg specification.
6653 | 1220648-2025-45752 | 22777671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22777671&pc=PYX | 07/14/2025 |  |  | 08/11/2025 |  |  | 07/14/2025 | 08/11/2025 |  |  |  |  | 08/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625510 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female | 84 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient (unknown ethnicity) with postcardiotomy cardiogenic shock was implanted with an impella 5.5 device followed by the impella rp flex for bipella mechanical circulatory support. reporting for the impella rp flex, the patient experienced ventricular fibrillation requiring intervention. the patient was admitted after coming to the emergency department with complaints of left-sided chest pain and shortness of breath for several days. the patient underwent a left heart catheterization and was diagnosed with multi-vessel disease. surgical consult was initiated. the decision was made to implant an impella 5.5 which was successfully completed. an axillary cut down made and a 10 x 30-millimeter hemashield platinum was anastomosed. the left ventricle was accessed, and a guidewire exchange occurred. the impella 5.5 was advanced across the aortic valve. the wire was pulled, and support was initiated at performance level p-2. support slowly ramped up and then had to go back onto bypass. during the procedure, when trying to separate from pump, the patient became pulseless, defibrillated multiple times and crashed back onto the pump in which the previously placed graft had to be repaired. the patient received ten units of packed red blood cells, two units of platelets, six units of fresh frozen plasma and two units of cryoprecipitate. an impella rp flex was placed and then flows titrated up. echocardiogram measured mid flow 4.3 centimeters from the aortic valve annulus. the graft was cut down and sutured into place. the impella was secured and pocket packed and chosen to be left open. the patient had to be defibrillated two more times before leaving the operating room. the patient was planned for a transfer to a different facility. the following day, the patient began to have seizure like activity and spiked a temperature. the patient was having multiple suction events and devices performance level was decreased. the patient was getting albumin and discussion was made for the need for volume verse increasing vasopressors and inotropes. the patient, at this time, was tolerating impella rp flex at p-7 with 2.9 liters per minute (l/min) of flow and impella 5.5 is at p-6 with 3.5 l/min of flow. the patient continued on support with plans remaining for transfer to an outside facility. the patient remained intubated and sedated tolerating p-9 on the impella rp flex with 3.1 l/min of flow and impella 5.5 at p-9 with 5.1 l/min of flow. on day three of support, overnight and in the morning, the patient had to be shocked multiple times due to going into ventricular fibrillation (vf). bedside, echocardiogram was performed. ejection fraction was approximately ten percent, and the right ventricle was not really moving. the position was verified at 5.15 centimeters. a noted plan was to do a bedside washout and an electroencephalogram to determine neurology status. the following day, the family was called and discussed withdrawal of care within the next day or so as the physician noted suspicion of a massive neurological injury. the patient was too unstable for scans. the physician was unsure if it was a bleed versus ischemic. core temperature was 40 celsius, left ventricular ejection fraction was five percent and right ventricular ejection fraction was ¿worse.¿ the following day care was withdrawn and the patient expired. | [5 empty] | Medical safety reviewed the event and determined there is not enough evidence to exclude the impella 5.5 as an associated factor in the patient's demise outcome (given this device-related adverse events of vascular/damage bleeding, temperature, ventricular fibrillation and seizure). refer to section h10. related report number. events (e.g., ventricular fibrillation) occurring with the impella rp flex are to be reported as serious injury. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H.2 correction was inadvertently omitted from the previously submitted follow up report. h.11 a.4 and e.1 zip code extension were inadvertently omitted from the initial report submitted and were added to the previous follow up report that was submitted. this information was inadvertently omitted from h.11 on the previously submitted follow-up report. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, peri-procedural adverse event: ventricular fibrillation, in order to make a root cause determination, essential clinical details would have to be provided. therefore, the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. h.6 component code and type of investigation codes were updated accordingly based on the completed investigation.
6654 | 1220648-2025-45733 | 22746503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22746503&pc=OZD | 07/14/2025 |  |  | 08/08/2025 |  |  | 07/14/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608348 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that the patient implanted with an impella cp encountered suction alarms and low pump flow. the patient was successfully weaned from the pump and survived. | [5 empty] | The device passed all post sterile inspection checks. neither the device nor device logs were returned for investigation. the root cause of the pump flow was unable to be determined as insufficient clinical details were provided. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A4 is unknown. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella cp and the purge cassette were returned for investigation. no logs were returned. the root cause of the low pump flow was most likely use issue related due to cannula kink found in the pump. the device passed all post sterile inspection checks.
6655 | 1220648-2025-45734 | 22746576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22746576&pc=OZD | 07/17/2025 |  |  | 08/08/2025 |  |  | 07/17/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669116 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 105 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Necrosis | Thrombosis/Thrombus | [9 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple complications in conjunction with impella cp support. following implant, some initial oozing was seen at the site. the dressing was changed to manage the condition and support continued. during planned escalation the following day, some heavy bleeding occurred for which factor seven had to be given. roughly three days after explant, the patient had no pedal pulse in their right foot. the patient was taken to the operating room and a thrombectomy was performed to remove a clot in the right groin at the impella cp site. despite the successful surgery, the right pedal pulse remained absent. discussions were initiated regarding the need to potentially amputate the right foot due to necrotic toes / partial necrotic foot. no decision was made regarding additional intervention. the patient survived. | [5 empty] | The root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the limb ischemia and thrombus was unable to be determined due to insufficient clinical details. the cause of the necrosis cannot be determined as insufficient clinical details were provided the device passed all post sterile inspection checks. | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6656 | 1220648-2025-45707 | 22746604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22746604&pc=OZD | 07/22/2025 |  |  | 08/08/2025 |  |  | 07/22/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694535 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after support was initiated, it was reported that the patient experienced absent pedal and posterior tibial pulses to the distal extremity of the impella leg, but the medical staff was able to obtain popliteal pulses with doppler. it was noted that the patient continued to experience self-terminating ventricular tachycardia with minimal pulsatility and the patient¿s neuro prognosis appeared to be poor with seizure activity noted, and overall prognosis guarded. the patient¿s neuro status remained unchanged in the following days and noted that the patient did not respond to pain stimulus. the patient¿s family decided to withdraw care, care was withdrawn and the patient expired.
6657 | 1220648-2025-31105 | 22747633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22747633&pc=OZD | 07/11/2025 | 09/05/2025 |  | 08/08/2025 |  |  | 07/11/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692795 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Convulsion/Seizure | Unspecified Vascular Problem | [8 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support and experienced a number of complications and would subsequently expire. the patient was admitted after coming to the emergency department with complaints of left-sided chest pain and shortness of breath for several days. the patient underwent a left heart catheterization and was diagnosed with multi-vessel disease. surgical consult was initiated. the decision was made to implant an impella 5.5 which was successfully completed. an axillary cut down made and a 10 x 30 millimeter hemashield platinum was anastomosed. the left ventricle was accessed, and a guidewire exchange occurred. the impella 5.5 was advanced across the aortic valve. the wire was pulled and support was initiated at performance level p-2. support slowly ramped up and then had to go back onto bypass. during the procedure, when trying to separate from pump, the patient became pulseless, defibrillated multiple times and crashed back onto the pump in which the previously placed graft had to be repaired. the patient received ten units of packed red blood cells, two units of platelets, six units of fresh frozen plasma and two units of cryoprecipitate. an impella rp flex was placed and then flows titrated up. echocardiogram measured mid flow 4.3 centimeters from the aortic valve annulus. the graft was cut down and sutured into place. the impella was secured and pocket packed and chosen to be left open. the patient had to be defibrillated two more times before leaving the operating room. the patient was planned for a transfer to a different facility. the following day, the patient began to have seizure like activity and spiked a temperature. the patient was having multiple suction events and devices performance level was decreased. the patient was getting albumin and discussion was made for the need for volume verse increasing vasopressors and inotropes. the patient, at this time, was tolerating impella rp flex at p-7 with 2.9 liters per minute (l/min) of flow and impella 5.5 is at p-6 with 3.5 l/min of flow. the patient continued on support with plans remaining for transfer to an outside facility. the patient remained intubated and sedated tolerating p-9 on the impella rp flex with 3.1 l/min of flow and impella 5.5 at p-9 with 5.1 l/min of flow. on day three of support, overnight and in the morning, the patient had to be shocked multiple times due to going into ventricular fibrillation (vf). bedside, echocardiogram was performed. ejection fraction was approximately ten percent and the right ventricle was not really moving. the position was verified at 5.15 centimeters. a noted plan was to do a bedside washout and an electroencephalogram to determine neurology status. the following day, the family was called and discussed withdrawal of care within the next day or so as the physician noted suspicion of a massive neurological injury. the patient was too unstable for scans. the physician was unsure if it was a bleed versus ischemic. core temperature was 40 celsius, left ventricular ejection fraction was five percent and right ventricular ejection fraction was ¿worse.¿ the following day care was withdrawn and the patient expired. additional information was received. the reason for the ventricular fibrillation was not associated with the pump but was due to poor cardiac output/perfusion along with electrolyte imbalances. | [5 empty] | The investigation of vf/vt/af/at, low pump flow, infection/sepsis, stroke/cva, and vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. low pump flow: the data log shows several suction alarms and low pump flow at the start of the case on (b)(6) 2025, along with a trending placement signal. the placement signal continued to show some trends the following day, during which low pump flow and additional suction alarms were reported. no motor current spikes or positioning issues were reported during the case. the patient underwent a blood transfusion during the pump implantation and received multiple blood products. the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessel: the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. | The medical safety officer reviewed the event and determined there is not enough evidence to exclude the impella 5.5 as an associated factor in the patient's demise outcome (given this device-related adverse events of vascular/damage bleeding, temperature, ventricular fibrillation and seizure). events (temperature and ventricular fibrillation) occurring with the impella rp flex are to be reported as serious injury. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6658 | 1220648-2025-31106 | 22747794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22747794&pc=OZD | 07/07/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672886 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 97 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Thrombocytopenia | Aortic Valve Insufficiency/ Regurgitation | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella 5.5 support and experienced thrombocytopenia. on day 14 of impella support, the patient started decompensating and had atrial fibrillation with rapid ventricular response. additionally, that patient had severe aortic insufficiency (ai) with a possible leaflet tear noted on transesophageal echocardiogram. the physician felt it may have started as a small perforation as a result of wire insertion during impella implant and had grown larger over time. however, we will not know for sure until aortic valve replacement (avr) is complete and the valve can be inspected. the impella insertion and left ventricle access was smooth without repeated attempts to cross the atrioventricular with stiff or multiple wires. it was further reported that the patient was not deemed a candidate for aortic valve replacement avr and care was withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (unites states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6659 | 1220648-2025-31107 | 22747842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22747842&pc=OZD | 07/23/2025 | 11/25/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1616687 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that placement signal and left ventricle waveform disappeared from the screen three times. each time prompting a "controller failure" alarm. support continued with no changes to flow. a decision was made to perform automated impella controller (aic) to aic transfer. the alarm has not been triggered since the aic was exchanged. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. data analysis: imc logs confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. multiple imc-ahp rfid errors are noted on the complaint date, but are not present when tested here with a known good purge cassette. device analysis: console (b)(6) was sent back to fs for further investigation. the console was opened and all cables and connectors to the purge unit were checked with no loose connections found. optical bench 5s parameters were taken and all values were within specification. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed.
6660 | 1220648-2025-31108 | 22747902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22747902&pc=OZD | 07/24/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611386 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 78 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp placed to support an acute myocardial infarction/cardiogenic shock patient. after 24 hours, the team made the decision to escalate to an impella 5.5. the impella cp access site was not achieving hemostasis after manual hold and so vascular surgery was consulted. the team made an open surgical closure and observed the side wall puncture. repair allowed for hemostasis and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6661 | 1220648-2025-31109 | 22748032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748032&pc=OZD | 07/24/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616404 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 70 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with a history of peripheral artery and vascular disease had an impella cp placed in the right femoral artery. in the catheterization lab and cardiovascular intensive care unit, there were already monophasic doppler sounds to both feet. after removal of the impella cp at bedside, the physician placed a femstop over the site and a bulb was inflated but slowly deflated over 15 minutes. doppler was checked at the patient¿s feet during that time with no doppler sounds. once the femstop was removed, there continued not to have any doppler sounds to the patients right foot. vascular surgery was consulted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6662 | 1220648-2025-31110 | 22748223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748223&pc=OZD | 07/24/2025 | 09/19/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704364 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 46 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was difficulty in placing the impella 5.5 via the right axillary artery. after two plus hours of attempts, even with multiple wires, placement was unsuccessful. the 0.018 wire kinked. the impella 5.5 got stuck where the innominate meets the aorta with a very acute bend. all troubleshooting techniques were attempted with no success. the left axillary was not an option. the device ultimately buckled with a significant kink on the catheter right before the motor housing. the case was aborted, and the patient was sent to the intensive care unit with intra-aortic balloon pump support to wait for transplant. | [5 empty] | The investigation of delivery issues, product damage (catheter kink), and product damage (guidewire kink) has been completed. according to the clinical report, resistance was encountered at the aortic arch, and the patient was reported to have a low body surface area. during the delivery procedure, the catheter and guidewire were found kinked. the cause of the delivery issue was most likely patient condition related to small anatomy. the cause of the product damage (catheter kink) issue was most likely patient condition related to small anatomy. the cause of the product damage (guidewire kink) issue was not determined since product was not returned and sufficient clinical information/image was not provided to confirm the damage. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6663 | 1220648-2025-31111 | 22748302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748302&pc=OZD | 07/24/2025 | 09/19/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696969 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was prepped and inserted via the left subclavian but would not advance into the aorta over the wire due to major tortuous anatomy. the impella 5.5 was removed and kinks were noted in the catheter and cannula. additionally, a dissection of the left subclavian artery was noted after the impella 5.5 was removed. the artery was successfully repaired surgically. a new impella was placed via the right subclavian artery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues, product damage (cannula kink), product damage (catheter kink), and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. delivery issues: impella would not advance into the aorta over the wire due to major tortuous anatomy. impella was removed and kinks were noted in the catheter and cannula of the device. based on the clinical details provided, the root cause of the delivery issues is most likely due to the patients condition. vascular damage - peripheral vessel: when the impella was removed, a dissection of the left subclavian artery was noted. due to limited clinical details surrounding the removal of the device, the root cause of the vascular damage - peripheral vessel cannot be determined. product damage (cannula kink) and product damage (catheter kink): the impella would not advance over the wire due to tortuous anatomy, and when the impella was removed, kinks were noted on the impella cannula and catheter. based on the clinical details, the root cause of the product damage (cannula kink) is most likely due to the patients condition. the root cause of the product damage (catheter kink) is most likely due to the patients condition.
6664 | 1220648-2025-31112 | 22748411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748411&pc=OZD | 07/25/2025 | 09/22/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519519 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the had a cp pump placed for a critical patient with prior coronary artery bypass graft (cabg), myocardial infarction (mi), st-segment elevation myocardial infarction (stemi), and hypotension with ventricular fibrillation and defibrillation. the pump was placed and the pump was persistently in poor position and interfering with the mitral valve. the team also observed signs of hemolysis. the pump repositioning and other troubleshooting measures failed. the labs were hemolyzed. when imaging for the multiple organ failure the team also observed thrombus and embolic event to the superior mesenteric artery. the care was withdrawn after 2 days of support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of positioning issues, hemolysis, and positioning issues has been completed. the impella device was not returned for evaluation. data review: data logs showed pump ran without issue during support. hemolysis: based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. positioning issue: the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
6665 | 1220648-2025-31113 | 22748462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748462&pc=OZD | 07/25/2025 | 09/26/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506297 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp stopped during support. the impella was exchanged. the patient is stable. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. data logs were not available for investigation. returned product was evaluated and the pump passed electrical testing of the motor cables. the outflow cage was visually inspected, and the purge gap was enlarged. the pump was connected to a console and run in a bucket of water. pump stop alarms did not reproduce; however, pump flows were initially saturated. both of these things can be expected with an enlarged purge gap. there were no other relevant abnormalities noted on returned product. per design trace matrix, impella cp design indication is 8 days of use. based on the fact that the pump was used for more than 8 days, and product was returned with an enlarged purge gap, the cause of the pump stop was determined to most likely be due to bearing wear.
6666 | 1220648-2025-45726 | 22748543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748543&pc=OZD | 07/11/2025 |  |  | 08/08/2025 |  |  | 07/11/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635363 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Restricted Flow rate | [6 empty] | Bacterial Infection | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs). the patient was escalated to an impella 5.5 while awaiting orthotopic heart transplant which was successfully completed. the patient was sent the unit, rest and recover, on impella mcs. however, later that evening intermittent suction at performance level p-6 was encountered. bedside echo to be completed. intermittent suction continued overnight, resolved at p-4. remains intubated and continued impella mcs. five units of blood were given overnight. the following morning the patient planned for colonic embolization in interventional radiology for gastrointestinal bleed. at this time patient was status 7 for transplant. two days later, the patient remained sedated and intubated with inhaled veletri and open abdomen with plans for potential abdominal closure the following day. impella 5.5 device at p-4 with unsuccessful attempts to increase the performance level due to suction each time. there were no plans to change the support until after abdominal closure. the next afternoon the abdominal closure was completed and an attempt to increase performance level to p-5 was unsuccessful as there was, again, immediate suction events. point of care ultrasound/formal echocardiogram was ordered. overnight oxygenation issues were encountered and improved with diuresis and vent management. the patient was still status 7 for transplant. however, there were now noted concerns for bacteremia. the patient was placed on antibiotics. the current plan changed to potentially wean pressors as tolerated and potentially start weaning support of impella as patient is potentially changing to bridge to recovery as opposed to heart transplant with poor prognosis. an emergent chole at bedside the next morning due to sludge noted on ct scan. treatment for possible sepsis related to gastrointestinal bleed was noted. the patient was delisted from heart transplant list and noted to be critically ill with a recent episode of profound hypotension requiring cardiopulmonary resuscitation. care goals were discussed with the family, the next day care was withdrawn, and the patient expired. | [5 empty] | Medical safety reviewed the event and determined that there is not enough evidence to exclude the impella 5.5 device as an associated factor in the patient¿s demise outcome. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The impella pump was not returned for evaluation and analysis. however, event logs were received and analyzed. device history record review was completed, and the pump passed all sterile inspection checks. pump suction: event logs show suction alarms whenever pump was above p-4 aligning with clinical details. there were no motor current spikes outside performance level changes and no abnormalities related to pump positioning. no patient anatomy or volume issues were noted. the cause of the pump suction was not determined as no product was returned for analysis, no log abnormalities were observed and limited clinical information was provided. bacterial infection/sepsis and retroperitoneal hemorrhage: in order to make a root cause determination, all of the respective clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details and/or return of the actual device used.
6667 | 1220648-2025-31114 | 22748581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748581&pc=OZD | 07/14/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627534 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female | 74 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured surgical wound breakdown with a "unknown (undetermined)" relationship to the impella device used: ¿pt presented for transplant outpatient visit on (b)(6) 2025 and to check her old impella site for an incision check for signs of infection. impella placed (b)(6) 2025 followed by oht on (b)(6) 2025. discharged and doing well at home. pt noticed a tract had formed at her prior impella site on her upper right chest w/some brown discharge around the site. per physician, pt admitted for possible infection and surgical wound breakdown at impella site. cultures of wound drainage obtained, results pending. iv daptomycin & cefepime started. labs cultures collected. at admission, open tract noted w/no active discharge, no erythema, swelling or pain. pt denies fevers, chills, chest pains, sob, orthopnea, or extremity swelling. id transitioned iv abx to po abx. no growth on cultures. pt discharged to brent house (b)(6) 2025. abx to end course (b)(6) 2025.¿. the patient recovered with sequelae. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6668 | 1220648-2025-31115 | 22748650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748650&pc=OZD | 07/25/2025 | 08/26/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084303 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had a purge disc issue upon set up. the aic was not recognizing the purge disc and would not go to the next step. this process was repeated multiple times with same issue. the aic was replaced. | [5 empty] | The investigation has been completed. the automated impella controller was returned for investigation. the imc logs from the complaint date revealed that the console issued purge disc not detected alarm. the console was examined after arriving in field service. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. the root cause of the reported issue was a broken purge disc flag. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6669 | 1220648-2025-31117 | 22748761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748761&pc=OZD | 07/25/2025 | 09/10/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697112 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had a 5.5 pump placed as care escalation from the cpthat was placed for the cardiogenic shock/cardiomyopathy patient. the 5.5 was seen to have a swelling/hematoma at the site. the team observed a hemoglobin drop from 9.3 to 5.8 and so 2 units of red blood cells and 2 units of platelets were given. the 5.5 was explanted after 161 hours of support. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since non-abiomed guidewire was used during insertion which is not recommended per instructions for use. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6670 | 1220648-2025-31118 | 22748810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748810&pc=OZD | 07/25/2025 | 11/25/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1544031 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 71 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) had an issue when the console was connected to the impella. an optical signal error with a controller error occurred. it is possible that the impella plug was not pressed all the way in. the plug was removed and reinserted and no issues after. support was continued. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation has been completed. analysis: imclogs from rlfs show that cp pump (b)(6) was initially connected to (b)(6) at 1:35 on (b)(6) 2025. while going through case start, the user was asked to disconnect the pump and reconnect due to an optical sensor system error. the pump was reconnected and the pump was run until 3:18 when the pump was moved over to (b)(6). the rtlogs show for the first connection, the raw optical sensor is -3276.8 as snr is near 0 and placement signal is 3000 mmhg indicative of an air gap. after the pump was reconnected, the rt logs were as expected. the complaint was confirmed in the logs. root cause: the cause of the optical signal failure was most likely an air gap between the aic and the pump as this was resolved by more securely connecting the pump. d10 concomitant medical products and therapy dates updated.
6671 | 1220648-2025-31119 | 22748835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748835&pc=OZD | 07/25/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681981 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that impella cp was placed in the middle of the percutaneous coronary intervention (pci) procedure as the patient started to decompensate. after the percutaneous coronary intervention (pci) procedure, medical doctor noticed a lack of bilateral femoral pulses on the patient. medical doctor decided to wean and explant the impella in cardiac catheterization laboratory. patient tolerated wean and pump was successfully explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
6672 | 1220648-2025-31116 | 22748842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22748842&pc=OZD | 07/24/2025 | 09/29/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655660 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an impella cp placed for support. the pump was placed but flows were low so after just under six hours, the pump was explanted and replaced. the patient survived and the second pump continues to support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. per returned product investigation, product damage (cannula kink) was added as a failure mode. low pump flow: data logs were not returned for this pump. the product was returned for investigation and there was no evidence of biomaterial ingestion, but a small kink was observed on the cannula near the motor housing. a flow test was performed on the pump to evaluate its flow levels at p-9 for 10 minutes in a bucket of water. the low flow did not reproduce, as the pump maintained a consistent flow of 3.7-3.8 l/min, which is within its specification. kinks in the cannula do not typically reproduce low flows during testing because there are no applied forces on the cannula like there would be in the patients anatomy. the cause of the low pump flow is most likely the cannula kink, as the kink would lead to low flows when external forces are acting on it inside the patients anatomy. product damage (cannula kink): clinical reported that the pump was showing reduced flows that resolved when it was replaced with a new pump. the product was returned for investigation and a small kink was observed on the cannula near the motor housing. the cause of the cannula kink was not determined due to insufficient clinical details.
6673 | 1220648-2025-31120 | 22749334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22749334&pc=OZD | 07/23/2025 | 09/15/2025 |  | 08/08/2025 |  |  | 07/23/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702600 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that at the beginning of the impella procedure, the sheath was opened prior to the peripheral angiogram being performed. angiogram showed severe peripheral vascular disease and the impella placement was aborted. no patient harm was reported. | [5 empty] | The investigation of deivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6674 | 1220648-2025-31121 | 22749362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22749362&pc=OZD | 07/25/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687255 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  |  | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient had a 5.5 pump placed for the patient who was deteriorating after the intra-aortic balloon pump was placed, patient was cardioverted and was unstable with blood pressures hypotensive at 50s/30s. the team escalated to the 5.5 and extracorporeal membrane oxygenation. at the insertion the delivery was complicated and the kinking caused the pump to "jump" and cause subclavian artery tearing. the team gave 8 units of blood, vascular surgery intervened, and then family made decision to withdraw care after just over 4 hours of support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
6675 | 1220648-2025-31122 | 22749440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22749440&pc=OZD | 07/24/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585959 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 50 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 and impella rp flex placed for bipella support of the surgical patient. the team also had extracorporeal membrane oxygenation and intubation was placed. the patient had a chest dressing change done at the sternum and went into ventricular tachycardia (vt). the patient had to be defibrillated for the vt. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6676 | 1220648-2025-31123 | 22749582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22749582&pc=OZD | 07/25/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660189 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 80 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella 5.5 support had placement signal not reliable alarm. the patient had been on impella 5.5 support for over a month and was listed for a heart transplant. no harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6677 | 1220648-2025-31124 | 22749898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22749898&pc=OZD | 07/25/2025 | 09/22/2025 |  | 08/08/2025 |  |  | 07/25/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697928 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the team noticed the left main was dissected by the guide catheter. support was continued. additionally, the patient developed a hematoma with oozing at the access site. no blood products were given to the patient. manual pressure was held. support was continued. additionally, the inlet was found about 2.54 centimeters into the left ventricle. the team decided to pull the impella back into the descending aorta and scheduled the patient for removal in the catheterization lab. the pump was eventually removed. the patient survived. | [5 empty] | The investigation of positioning issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. vascular damage peripheral vessel: the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. positioning issue: the root cause of positioning issue was most likely use issue since the console began showing ¿impella in aorta¿ position alarm while changing impella dressing and echo confirmed inlet about 2.54 cm into left ventricle (lv). instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31124. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6678 | 1220648-2025-45742 | 22750443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22750443&pc=OZD | 07/11/2025 |  |  | 08/08/2025 |  |  | 07/11/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025584574 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that the impella 5.5 implanted for cardiogenic shock. during 5.5 support it was reported that the patient had significant blood loss from jackson-pratt drain. received 2 packed red blood cells, 2 fresh frozen plasma and 2 cryoprecipitate and heparin was stopped. bleeding had subsided with stabilization of hemoglobin and hematocrit. additionally, it was noted that the pump was weaned to p-2 with small collapsed left ventricle visualized on echocardiogram. active upper gastrointestinal bleed was present. blood was transfusing. the patient remained on support with a duration of 599.90 hours. ultimately care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks.
6679 | 1220648-2025-45732 | 22750940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22750940&pc=OZD | 07/12/2025 |  |  | 08/08/2025 |  |  | 07/12/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026664360 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 87 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced gastrointestinal bleeding. the patient was transfused one unit of red blood cells and heparin was withheld. the patient was eventually successfully bridged to a durable left ventricular assist device and the impella 5.5 was explanted. | [5 empty] | The device passed all post sterile inspection checks. the root cause of vascular damage and gi bleed cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6680 | 1220648-2025-31125 | 22751188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22751188&pc=OZD | 07/24/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686440 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 75 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella 5.5 to support a high-risk percutaneous coronary intervention patient, admitted in with atrial fibrillation and non-st-segment elevation myocardial infarction. the patient was escalated from the intra-aortic balloon pump. the pump was placed. however, there were drops in labs and platelets were dropping. heparin was held and bivalirudin was begun as hematology was consulted. the patient additionally developed a hematoma and platelets were given. after eight days, the pump was weaned and explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6681 | 1220648-2025-31126 | 22751333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22751333&pc=OZD | 07/23/2025 | 09/29/2025 |  | 08/08/2025 |  |  | 07/23/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667368 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male | 72 KG |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump support ongoing for the surgical patient who had a bentall procedure. the pump was supporting when purge pressure rose and the team troubleshot with the use of lysis. additionally, there were signs of hemolysis noted on day five of support. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of high purge pressure and hemolysis has been completed. the pump was returned for investigation. no visual damage was noted on the returned product. no biomaterial was observed on the impeller. the pump was further tested in a bucket of water with priming, and no issues were noted. the pump was operated according to specifications. data logs show that on (b)(6), there was a gradual rise in purge pressure over approximately 25 minutes, followed by a spike in motor current. this purge issue trend resolved after about 7 hours with a gradual decreasing trend. on (b)(6), a similar pattern was observed, with a gradual rise in purge system pressure accompanied by a purge system block alarm. no motor current spike was observed during this high-pressure event, and it was resolved in approximately 9 hours with a gradual decrease trend. another motor current spike was reported at the end of the case, indicating possible biomaterial ingestion during the procedure. according to clinical notes, lysis administration helped stabilize the purge parameters and was also suspected to have addressed hemolysis caused by a clot in the pump. high purge pressure: the cause of the high purge pressure issue was most likely biomaterial buildup, based on data log review. hemolysis: the cause of the hemolysis issue was most likely biomaterial ingestion, based on data log review. thrombus failure mode was added as investigated failure mode based on data log review. the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9. device returned to manufacture date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6682 | 1220648-2025-31127 | 22752426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22752426&pc=OZD | 07/24/2025 | 09/22/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677105 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp to support the patient admitted in with a st-segment elevation myocardial infarction. the pump was placed but, an access site hematoma developed. the team treated this with red blood cells. as the percutaneous coronary intervention took place there were positional issues noted with the cp. the pump was interacting with the papillary. to troubleshoot the pump was pulled from the left ventricle and then redelivered for the support. the pump supported for 4 days and was successfully explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. positioning issues: the root cause of positioning issues was most likely use issue since the pump was originally placed in the papillary by the physician and the second implant attempt was done in good position. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6683 | 1220648-2025-31128 | 22754604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22754604&pc=OZD | 07/22/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/22/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1493604 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that upon becoming unplugged from the wall, the automated impella controller went to 50% power then back to 100%. the console was removed and transferred to a loaner. the console will be sent in for inspection. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6684 | 1220648-2025-31129 | 22754614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22754614&pc=OZD | 07/24/2025 | 09/15/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693820 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding from all access sites, including the pa line site. the team believed that the patient was in disseminated intravascular coagulation. one unit of packed red blood cells and fresh frozen plasma were administered. the patient was successfully transplanted. | [5 empty] | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding issue was most likely patient condition related since the patient on impella cp support developed bleeding from all access sites. vascular damage : the root cause of the vascular damage issue was most likely patient condition related since the patient on impella cp support developed bleeding from all access sites, including the pa line site. the team believed that the patient was in disseminated intravascular coagulation (dic). | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6685 | 1220648-2025-31130 | 22754666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22754666&pc=OZD | 07/24/2025 | 08/08/2025 |  | 08/08/2025 |  |  | 07/24/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700718 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Low Blood Pressure/ Hypotension | Ischemia | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support had doppler popliteal pulse in the right leg but both dorsalis pedis and posterior tibial pulses were absence. it was further reported that the patient experienced atrial fibrillation with rapid ventricular response. also, the impella was increased from p3 to p4 for hypotension. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
6686 | 1220648-2025-45750 | 22755897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22755897&pc=OZD | 07/11/2025 |  |  | 08/08/2025 |  |  | 07/11/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026658591 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during the initial cutdown for the impella 5.5 axillary graft site, some bleeding complications were encountered. the graft was placed successfully, and multiple "quick clot" methods were used to stitch the site. following implant, protamine was used to reverse the effects of heparin and blood products were given to counteract the blood loss. support continued uninterrupted. the patient¿s left ventricular ejection fraction recovered to 45-50%, and the patient was taken to the operating room for the removal of the impella 5.5. the patient was stable during the wean and there were no issues reported at explant. the patient remained on venopulmonary extracorporeal membrane oxygenation after impella 5.5 was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6687 | 1220648-2025-45711 | 22766219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22766219&pc=OZD | 07/13/2025 |  |  | 08/08/2025 |  |  | 07/13/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701174 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 85 KG | Hispanic |  | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, persisting positioning and low pump flow issues were encountered. it was reported that the team had attempted repositioning of impella but had failed due to lack of volume, enlarged left ventricle muscle, and diminished cavity. approximately three days later, the patient transported for transfer and during automated impella controller (aic)- to-aic transfer, it was noted when switching to another controller, a message that catheter was connected and resume settings was not received. the team indicated consoles remained on during transfer and the attempt to switch back to the original as well as start a new console for a switch was without success as the impella catheter was not recognized. it was stated the patient had been just cannulated for va ecmo and on p6. the team planned to pull the catheter back. however, it appears an aic worked with the impella catheter as the patient continued on impella mechanical circulatory support (mcs). thereafter, a few suction alarms were encountered, which were resolved after administering the patient one unit of albumin and packed red blood cells (prbc's). support continued. it was noted there were no alarms or issues noted, and support was decreased to p-3 from p-5 with plans to stop heparin the following morning. a couple of days later, there was bleeding in the abdomen. an exploratory laparoscopy was started at bedside but stopped due to the bleeding. the next day the patient underwent exploratory bowel surgery due to a high concern of acute bowel ischemia. the abdomen was left open; however, it was closed the next day. ongoing discussion regarding care goals were made with plans of weaning the impella pump as the patient was not a candidate for a durable left ventricle assist device. the family decided to transition the patient to comfort care, and the patient would expire after a total of 21 days on impella mcs. | [5 empty] | The patient was on heparin for the use of the impella and it cannot be determined if the use of heparin impacted/exacerbated the outcome of the bleed for the patient. medical safety reviewed the event and determined there is not enough evidence to exclude the impella pump as an associated factor in the patient's demise outcome given its malfunction and related adverse event. the automated impella controllers involved is a product malfunction type of reportable event. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the hemodynamic instability could not be determined without sufficient clinical information. the pump was not returned for investigation. data logs showed several suction alarms on (b)(6) 2025. the pump flow dropped below 2 l/min while running on a steady p-level but returned to normal. according to the clinical report, administration of blood products resolved the suction issue. patient had an enlarged left ventricle muscle, and diminished cavity. the cause of the suction issue was most likely patient condition-related, specifically low blood volume, based on the given clinical information. as per the clinical report, the impella was unable to be repositioned due to low volume, an enlarged left ventricular muscle, and a diminished ventricular cavity. the cause of the positioning issue was most likely patient condition-related to anatomy.
6688 | 1220648-2025-45719 | 22766228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22766228&pc=OZD | 07/16/2025 |  |  | 08/08/2025 |  |  | 07/16/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302904 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 85 KG | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, persisting positioning and low pump flow issues were encountered. it was reported that the team had attempted repositioning of impella but had failed due to lack of volume, enlarged left ventricle muscle, and diminished cavity. approximately three days later, the patient transported for transfer and during automated impella controller (aic)- to-aic transfer, it was noted when switching to another controller, a message that catheter was connected and resume settings was not received. the team indicated consoles remained on during transfer and the attempt to switch back to the original as well as start a new console for a switch was without success as the impella catheter was not recognized. it was stated the patient had been just cannulated for va ecmo and on p6. the team planned to pull the catheter back. however, it appears an aic worked with the impella catheter as the patient continued on impella mechanical circulatory support (mcs). thereafter, a few suction alarms were encountered, which were resolved after administering the patient one unit of albumin and packed red blood cells (prbc's). support continued. it was noted there were no alarms or issues noted and support was decreased to p-3 from p-5 with plans to stop heparin the following morning. a couple of days later, there was bleeding in the abdomen. an exploratory laparoscopy was started at bedside but stopped due to the bleeding. the next day the patient underwent exploratory bowel surgery due to a high concern of acute bowel ischemia. the abdomen was left open; however, it was closed the next day. ongoing discussion regarding care goals were made with plans of weaning the impella pump as the patient was not a candidate for a durable left ventricle assist device. the family decided to transition the patient to comfort care, and the patient would expire after a total of 21 days on impella mcs. | [5 empty] | Medical safety reviewed the event and determined there is not enough evidence to exclude the impella pump as an associated factor in the patient's demise outcome given its malfunction and related adverse event. the automated impella controllers involved is a product malfunction type of reportable event. product status: the automated impella controller device was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6689 | 1220648-2025-45731 | 22766230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22766230&pc=OZD | 07/16/2025 |  |  | 08/08/2025 |  |  | 07/16/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398680 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 85 KG | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, persisting positioning and low pump flow issues were encountered. it was reported that the team had attempted repositioning of impella but had failed due to lack of volume, enlarged left ventricle muscle, and diminished cavity. approximately three days later, the patient transported for transfer and during automated impella controller (aic)- to-aic transfer, it was noted when switching to another controller, a message that catheter was connected and resume settings was not received. the team indicated consoles remained on during transfer and the attempt to switch back to the original as well as start a new console for a switch was without success as the impella catheter was not recognized. it was stated the patient had been just cannulated for va ecmo and on p6. the team planned to pull the catheter back. however, it appears an aic worked with the impella catheter as the patient continued on impella mechanical circulatory support (mcs). thereafter, a few suction alarms were encountered, which were resolved after administering the patient one unit of albumin and packed red blood cells (prbc's). support continued. it was noted there were no alarms or issues noted and support was decreased to p-3 from p-5 with plans to stop heparin the following morning. a couple of days later, there was bleeding in the abdomen. an exploratory laparoscopy was started at bedside but stopped due to the bleeding. the next day the patient underwent exploratory bowel surgery due to a high concern of acute bowel ischemia. the abdomen was left open; however, it was closed the next day. ongoing discussion regarding care goals were made with plans of weaning the impella pump as the patient was not a candidate for a durable left ventricle assist device. the family decided to transition the patient to comfort care, and the patient would expire after a total of 21 days on impella mcs. | [5 empty] | Medical safety reviewed the event and determined there is not enough evidence to exclude the impella pump as an associated factor in the patient's demise outcome given its malfunction and related adverse event. the automated impella controllers involved is a product malfunction type of reportable event. product status: the automated impella controller device was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6690 | 1220648-2025-45728 | 22766232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22766232&pc=OZD | 07/16/2025 |  |  | 08/08/2025 |  |  | 07/16/2025 | 08/08/2025 |  |  |  |  | 08/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 85 KG | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, persisting positioning and low pump flow issues were encountered. it was reported that the team had attempted repositioning of impella but had failed due to lack of volume, enlarged left ventricle muscle, and diminished cavity. approximately three days later, the patient transported for transfer and during automated impella controller (aic)- to-aic transfer, it was noted when switching to another controller, a message that catheter was connected and resume settings was not received. the team indicated consoles remained on during transfer and the attempt to switch back to the original as well as start a new console for a switch was without success as the impella catheter was not recognized. it was stated the patient had been just cannulated for va ecmo and on p6. the team planned to pull the catheter back. however, it appears an aic worked with the impella catheter as the patient continued on impella mechanical circulatory support (mcs). thereafter, a few suction alarms were encountered, which were resolved after administering the patient one unit of albumin and packed red blood cells (prbc's). support continued. it was noted there were no alarms or issues noted, and support was decreased to p-3 from p-5 with plans to stop heparin the following morning. a couple of days later, there was bleeding in the abdomen. an exploratory laparoscopy was started at bedside but stopped due to the bleeding. the next day the patient underwent exploratory bowel surgery due to a high concern of acute bowel ischemia. the abdomen was left open; however, it was closed the next day. ongoing discussion regarding care goals were made with plans of weaning the impella pump as the patient was not a candidate for a durable left ventricle assist device. the family decided to transition the patient to comfort care and the patient would expire after a total of 21 days on impella mcs. | [5 empty] | Medical safety reviewed the event and determined there is not enough evidence to exclude the impella pump as an associated factor in the patient's demise outcome given its malfunction and related adverse event. the automated impella controllers involved is a product malfunction type of reportable event. product status: the automated impella controller device was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6691 | 1220648-2025-31087 | 22736237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22736237&pc=OZD | 07/23/2025 | 09/18/2025 |  | 08/07/2025 |  |  | 07/23/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700750 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was admitted in with a planned high risk surgery of the root enlargement/aortic valve replacement/septal myectomy. the impella 5.5 failed to deliver across the valve as expected. while troubleshooting the delivery, the patient developed ventricular fibrillation and was shocked. the wire was replaced and the pump delivery was attempted again. the second delivery attempt was successful and the impella 5.5 supported the patient for just under eight hours until care was withdrawn. the patient expired. | [5 empty] | The investigation of delivery issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to be patient condition because the surgeon able to get near aortic root, but unable to cross av due to tortuous anatomy of the patient. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31087. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6692 | 1220648-2025-31088 | 22736292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22736292&pc=OZD | 07/23/2025 | 08/07/2025 |  | 08/07/2025 |  |  | 07/23/2025 | 08/07/2025 |  |  | 08/31/2025 |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392999 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 77 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support had multiple episodes of ventricular tachycardia (vt) overnight with significant drop in mean arterial pressures during episodes. the patient received an amiodarone bolus and remained on an infusion. it was further reported that the patient expired on support. additonal information was received. upon review of the episodes of vt, they appeared to be spontaneous and unrelated to impella position or impella related events and could be attributable to the patient¿s underlying pathology and continued hypoperfusion/ischemia. echocardiogram was reviewed with bedside nurse and advised that the impella position be re-assessed to ensure that was not the cause. it appeared in good position on echocardigram per providers. echocardigram images were not the best due to patient¿s body habitus. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6693 | 1220648-2025-31089 | 22736391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22736391&pc=OZD | 07/24/2025 | 08/07/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683876 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female | 82 KG |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Insufficient Information | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella cp pump placed to support a patient admitted in with cardiogenic shock/acute myocardial infarction. the pump was placed but the pump came out of position and with the mispositioning, the pump was removed and the patient was medically managed. the patient did expire after 41 hours. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
6694 | 1220648-2025-31090 | 22736526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22736526&pc=OZD | 07/24/2025 | 12/11/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655828 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 79 KG | Non Hispanic |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient had an impella 5.5 pump placed for escalated chest pain and extracorporeal membrane oxygenation. the impella 5.5 was supporting and allowed for extracorporeal membrane oxygenation to be removed and plan for the left ventricular assist device in the future. the impella 5.5 was supporting for weeks when there were complications of deep vein thrombosis and infection as well as neurological changes. the pump remained on despite these issues until day 22. the pump stopped and was not able to be restarted. the stopped pump also had multiple alarms. the family knew the patient was not a candidate to escalate to other measures and so the patient expired. | [5 empty] | Additional information was received in the form of medical safety review and noted the following: an adult patient presented in cardiogenic shock on (b)(6) and underwent multiple rounds of cpr. an intra-aortic balloon pump (iabp) was placed for initial support. on (b)(6), the patient was escalated to impella cp for higher hemodynamic support. due to worsening hemodynamics, va-ecmo was initiated for primary mechanical circulatory support, and impella 5.5 was implanted for left ventricular (lv) unloading. on (b)(6), the patient exhibited signs of infection (no documented results). on (b)(6), multiple deep vein thromboses (dvts) and a pseudoaneurysm at the axillary site were noted. on (b)(6), the impella 5.5 stopped unexpectedly; attempts to restart the device failed. the device was left in the aortic position. the patient expired on (b)(6) while supported on ecmo and the non-functioning impella 5.5. device involvement: impella cp functioned as intended; no device-related complaints. escalation was due to patient¿s clinical deterioration, not device malfunction. impella 5.5 was an unexpected pump stop on (b)(6) with failed restart attempts. device remained in situ until patient expired. considered a contributing factor to patient death, though underlying condition was severe. automated impella controllers (aic) conservatively reporting as malfunction. based on reported information, the impella cp performed as expected; the impella 5.5 exhibited pump stop during use. the patient¿s death primarily attributed to underlying clinical condition; device malfunction may have contributed. note: type of reportable event and investigation with coding remain unchanged. this is one of three device reports associated with the event. refer to the following report numbers for the related device reports: 1220648-2025-31158 (aic), 1220648-2025-31159 (aic) and 1220648-2025-31090 (pump). | The investigation of pump stop has been completed. the pump was returned cut at catheter. upon visual inspection, a large biomaterial was present on and around impeller blade. data logs were reviewed and real time (rt) log at time of issue showed a sudden pump stop with alarm 115 triggered. motor current spike greater than 30000 ma 2 minutes later when pump was attempted to be restarted. prior to pump stop, pump flow and placement signal were variable prior to pump stop. clinical details noted that a sudden pump stop occurred onto support. pump was unable to be restarted and backup console was attempted to restart pump without success. the cause of the pump stop could not be definitively determined as pump was returned cut and was unable to be tested fully during investigation and limited clinical details were provided. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6695 | 1220648-2025-31091 | 22736646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22736646&pc=OZD | 07/24/2025 | 09/23/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697872 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female | 61 KG |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump met resistance during attempted delivery and kinked. it was noted that the patient had a small left ventricle and the kinking occurred in the aorta shortly after the bifurcation. as a result of the resistance caused by the patient's vasculature, the decision was made to abort the implant. | [5 empty] | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of the delivery issue was most likely patient condition related, before implanting the impella it was also difficult to bring in the pig tail in the left vernicle (lv), because of cardiac and vascular anatomy. product damage (cannula kink) : the root cause of the product damage issue was most likely patient condition related, kinking in aorta shortly after bifurcation and a slight aneurysm in aorta ascends. difficult vascular and cardiac anatomy. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6696 | 1220648-2025-31092 | 22736818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22736818&pc=OZD | 07/24/2025 | 08/07/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  | 11/30/2025 |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436187 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp for the support of a patient admitted in with an acute myocardial infarction complicated by cardiogenic shock. the situation was noted to be an "emergency" and at the time of start the priming/purge failed and pump would not start. the team troubleshot by aic replacement but the second aic failed and so the team went back to the first aic and got pump started for the support. the pump supported for days and the patient was transferred from the community to the tertiary center. patient survived the pump explant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6697 | 1220648-2025-31093 | 22736861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22736861&pc=OZD | 07/23/2025 | 09/15/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed access site bleeding. a second figure of 8 suture was placed and bleeding was resolved. patient survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related since a second figure of 8 suture was placed and bleeding was resolved. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6698 | 1220648-2025-31094 | 22737226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22737226&pc=OZD | 07/24/2025 | 09/26/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703464 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported cp pump placed to support the patient admitted in with acute myocardial infarction complicated by cardiogenic shock, diabetes, and renal insufficiency as well as other cardiac and systemic comorbidities. the pump was placed and then escalated to this 5.5 pump. the patient experienced a hematoma at the site and signs of hemolysis. the team intervened with red blood cells infused. after just under 9 days the pump was explanted and patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding and renal failure has been completed. product was returned for review. visual inspection found no issues on the pump. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no problem was found on the pump. the cause of the renal failure was not determined due to insufficient clinical details. d9 device returned to manufacture date added.
6699 | 1220648-2025-31095 | 22737377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22737377&pc=OZD | 07/24/2025 | 08/07/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701063 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient with acute myocardial infarction complicated by cardiogenic shock was escalated from impella cp to impella 5.5. in intensive care unit the patient had no pulse on the right foot. the medical doctor requested an arterial doppler for right leg and foot. the next day, there was still no pulse in the patient¿s right foot. vascular was notified. unknown intervention. it was further reported that the patient developed ischemic bowel and the family decided to withdraw care. the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
6700 | 1220648-2025-31096 | 22737541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22737541&pc=OZD | 07/24/2025 | 09/18/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693501 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had a cp pump being prepped for delivery and had a user caused failure. the wire exited incorrectly and was therefore replaced. the team placed a new cp pump for the high-risk percutaneous coronary intervention patient. upon inspection the wire had damaged the lumen. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. impella cp with smartassist system: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the pump was returned for investigation, but the 0.018 guidewire was not. in its returned state, the lumen was coming out of the inlet cage, but the lumen is supposed to stay inside the inlet and come out of the pig tail. a small kink was observed on the red lumen but there were no punctures found. a sample 0.018 guidewire was threaded through the red lumen with slight resistance due to the kink. based on the clinical details, it is possible that the lumen was pulled out of the pigtail when the 0.018 guidewire was placed in the lumen because of resistance from the kink. once the red lumen is pulled out of the pigtail, it is not able to be threaded back through, which is why the lumen may have come out of the inlet cage instead. the cause of the delivery issues was most likely red lumen issue, as a kink was found in product investigation, which led to resistance when inserting the 0.018 guidewire.
6701 | 1220648-2025-31097 | 22737852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22737852&pc=OZD | 07/24/2025 | 09/16/2025 |  | 08/07/2025 |  |  | 07/24/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694086 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient experienced access site bleeding a couple of hours after being in the intensive care unit. manual pressure was held, and heparin was stopped. additionally, the patient developed red urine. the impella cp was turned down to p7 from p8. under echocardiogram it looked like the impella was too deep, so it was repositioned. it was further reported that care was withdrawn and the patient expired. | [5 empty] | The investigation of access site bleeding, hemolysis, and positioning issues has been completed. the impella device was not returned for evaluation. data review: data logs confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position is unknown. no other issues were found on the logs. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely use related since accessing site was secured without angle matching. the root cause of hemolysis was most likely due to pump was not in the proper position since impella repositioning unit was not locked in place. the root cause of positioning issue was most likely use related since impella repositioning unit was not locked in place. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6702 | 1220648-2025-31098 | 22737909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22737909&pc=OZD | 06/01/2024 | 08/07/2025 |  | 08/07/2025 |  |  | 06/01/2024 | 08/07/2025 |  |  | 02/28/2026 |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | White | Death | Yes | No | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced multiple ¿placement signal low¿ alarms. an echocardiogram showed the impella was positioned deep so the pump was repositioned. the patient experienced ventricular tachycardia requiring cardoversion and increased vasoactive requirements. care was ultimately withdrawn and the patient expired. | [5 empty] | A4 and a5, ethnicity, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
6703 | 1220648-2025-45729 | 22738143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22738143&pc=OZD | 01/26/2024 |  |  | 08/07/2025 |  |  | 07/22/2025 | 08/07/2025 |  |  | 09/30/2025 |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401636 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received regarding a patient that endured death with a "unknown (undetermined)" relationship to the impella device used. it was found that during support, the patient developed sustained ventricular tachycardia (vt) and was cardioverted multiple times with defibrillation shock, however, the patient kept converting back to vt. lidocaine drip was started for arrhythmia. despite more stable rhythm, patient remained hypotensive. eventually patient care was withdrawn as was not a candidate for any other advance therapies. patient passed away. per the registry, the patient was discharged home post procedure and expired at home. | Additional information was received that the patient chemically coded and they made the patient comfort care and palliative care which would mean the patient passed in the hospital. |  |  |  |  | Additional information is being requested for clarification regarding the location and date of death. should any clarifying or new relevant information be received, a supplemental mdr will be submitted accordingly. produt status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use as noted in the impella: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device history review was completed and the device passed all post sterile inspection checks. the root cause of hypotension and ventricular tachycardia/ventricular fibrillation was unable to be determined due to insufficient clinical details. b.5 additional information was received regarding the patient's demise and was added as it was inadvertently omitted from the initial report that was submitted. b.7 additional information was added that was inadvertently omitted from the initial report that was submitted. e.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted.
6704 | 1220648-2025-31099 | 22738756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22738756&pc=OZD | 07/23/2025 | 09/15/2025 |  | 08/07/2025 |  |  | 07/23/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a pump being placed for the surgical cabg patient. the 5.5 was not placed via direct insertion and instead a femoral impella cp pump was to be placed. the placement of the cp failed despite all measures due to an "unspecified anomaly" of the aortic valve. the anatomy precluded placement of the cp therefore was aborted and an intra-aortic balloon pump was placed instead. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿.
6705 | 1220648-2025-31100 | 22739264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22739264&pc=OZD | 07/23/2025 | 09/11/2025 |  | 08/07/2025 |  |  | 07/23/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703433 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 was unable to turn at innominate junction given tortuosity. after several unsuccessful attempts a decision was made to abort impella placement due to the patient¿s anatomy. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since attempting direct insertion approach decision made to abort given difficult anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
6706 | 1220648-2025-31101 | 22739310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22739310&pc=OZD | 07/23/2025 | 09/15/2025 |  | 08/07/2025 |  |  | 07/23/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026699470 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed in the community for the support of a critical patient with a lv rupture. the patient was transferred to the tertiary center for surgical intervention and placement of a impella 5.5 pump. it was reported that the pump attempted to cross valve wirelessly with success. however, unable to torque pump into a deep mid ventricular position that would allow the outflow of the 5.5 to sit in the ascending aorta and not the graft. the surgeon attempted to reposition the impella but damaged one of the aortic valve leaflets in the process. impella 5.5 was removed and surgeon decided to go back on bypass and arrest for aortic valve replacement. goal was to move the aortic cannulas to the groin follow aortic valve replacement to make room for impella graft, however, after valve replacement, surgeons attempted to dry out patient and come off bypass multiple times without support and with no success. patient was emergently recannulated for bypass and transferred over to venoarterial extracorporeal membrane oxygenation to recover over night in intensive care unit. | [5 empty] | The investigation of positioning issues and heart valve damage has been completed. the impella device was not returned for evaluation. positioning issue: the root cause of the positioning issue was most likely use-related as evidenced by issue occurrence during repositioning with no outlined patient or device related abnormalities. heart valve damage: no patient anatomy issues or evidence of excessive force or issue related to catheter anchor was noted. the root cause of the heart valve damage was unable to be determined because though issue occurred during repositioning, limited clinical information or contributing factors to issue were provided to confirm cause. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
6707 | 1220648-2025-31102 | 22739777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22739777&pc=OZD | 07/23/2025 | 08/07/2025 |  | 08/07/2025 |  |  | 07/23/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682837 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Failure to Advance | [5 empty] | Hematoma | Asystole | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had a second impella cp pump being placed to support the non-st-segment elevation myocardial infarction patient. impella cp delivery was attempted but unsuccessful. impella went in but the patient continued to be in asystole and impella flows unable to reach >1 liters per minute. the patient also had hematoma requiring blood products and manual pressure. the family made decision to withdraw care when patient went asystolic. the patient expired in the procedural area. the death was captured under the first impella ((b)(6) serial number). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the first impella cp and the death is reported on manufacture report number 1220648-2025-31069. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
6708 | 1220648-2025-31103 | 22740202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22740202&pc=OZD | 07/23/2025 | 09/18/2025 |  | 08/07/2025 |  |  | 07/23/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665407 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being placed for the ventricular tachycardia (vt) ablation patient. it was reported that during set up, it was noticed that the travel perfusionist was jamming and struggling and forcing the white power cord plug in and out of the automated impella controller (aic) after he was told to wait on plugging that in until set up and case start prompts were followed. after plugging it in, there was an optical sensor yellow alert saying to remove plug and check if light was at eye level, which it was, and prompts were followed and sensor was cleared. the rest of set up and insertion went unproblematic. the aic and same pump were used. no alerts or alarms were had as pump was turned on or ramped up, everything appeared normal. it was when the patient was having the stat locks secured and heading up to the icu when the placement signal unreliable alarm was prompted and the aorta placement signal and left ventricle waveforms and numbers went away. they never returned. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. data log review showed ps railing to 3000mmhg after 1 hour on support. during placement signal not reliable (psnr) alarms, snr dropped to zero, light decreased to 0.5, contrast barcoded, and gain decreased. the pump was returned and inspected. 5s parameters were abnormal. laser continuity testing showing damage to the fiber line within the red pump plug. the cause of the optical signal issue was a damaged optical fiber line found during returned product inspection.
6709 | 1220648-2025-31104 | 22741030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22741030&pc=OZD | 06/28/2025 | 09/19/2025 |  | 08/07/2025 |  |  | 07/08/2025 | 08/07/2025 |  |  |  |  | 08/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Remove | Device Damaged by Another Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a repositioning sheath was inserted into the 14 fr peel away sheath prior to removing the .018 guidewire. this made removal of the wire difficult as it was pinched in the sheath. the staff was eventually able to remove the wire, but the tip of the wire was damaged in the process. the staff attributed the damage to user error. there was no resulting injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. | The investigation of removal issues and product damage (guidewire damage - great vessel) has been completed. the impella device was not returned for evaluation. removal issues: clinical reported repositioning sheath was inserted into the 14 fr peel away sheath prior to removing the .018 guidewire. this made removal of the wire difficult as it was pinched in the sheath. the staff was eventually able to remove the wire, but the tip of the wire was damaged in the process. excessive force reported. the cause of the issue was most likely use-related as evidenced by the improper technique of not removing the 0.018 guidewire prior to inserting repositioning sheath. product damage (guidewire damage - great vessel): clinical reported guidewire unraveling with excessive force used. the cause of the issue was not determined as no product/images were returned for analysis.
6710 | 1220648-2025-31040 | 22725106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725106&pc=OZD | 07/23/2025 | 09/25/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683209 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that two days after impella 5.5 implant, an impella in aorta alarm appeared. the pump was verified to be in the correct position at the time of the noted issue. support was able to continue with the implanted pump without interruption. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. the data logs show 'impella position in aorta' alarms with no log abnormalities indicating alarm was correct such as non-pulsatile ao-ps or motor current. all optical 5s parameters were in normal range. no placement signal not reliable (psnr) triggered. the cause of the issue was most likely patient condition related as clinical details noted that the pump position was verified via imaging and data logs do not show any hard failures of the optical sensor.
6711 | 1220648-2025-31041 | 22725245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725245&pc=OZD | 07/23/2025 | 08/06/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674592 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 70 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient admitted in with postcardiotomy cardiogenic shock had an impella cp and impella rp flex placed for bipella support. the patient was escalated from an intra-aortic balloon pump. the impella cp was supporting when there was limb ischemia noted on day two. the leg was revascularized in the cardiac catheterization laboratory. the pump remained on rather than escalating and explanting due to the patients critical nature and discussions ongoing with the family. the team withdrew care and patient expired after 61 hours. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6712 | 1220648-2025-31042 | 22725347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725347&pc=OZD | 07/11/2025 | 08/06/2025 |  | 08/06/2025 |  |  | 07/11/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026628645 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Use of Device Problem | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the implant of an impella 5.5 device to a patient with cardiomyopathy for mechanical circulatory support was unsuccessful. difficulty getting the motor housing to pass anastomosis was encountered. straightening graft and a buddy wire attempt were unsuccessful. there was concern that motor could have been damaged while attempting to cross anastomosis. therefore, the decision was made to explant and replace the impella 5.5 device which was successfully completed. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6713 | 1220648-2025-31043 | 22725419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725419&pc=OZD | 07/20/2025 | 08/06/2025 |  | 08/06/2025 |  |  | 07/20/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636036 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Device Sensing Problem | [6 empty] | Arrhythmia | Septic Shock | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal became erratic during impella 5.5 support. placement monitoring was disabled as a result of persistent impella in aorta alarms despite confirmed proper positioning. both the flow rate and motor current remained within the expected limits at the time of the alarm. the following day, the patient endured runs of ventricular tachycardia. support continued uninterrupted despite the noted condition. the patient was showing signs of septic shock. it is unknown if the sepsis was impella related. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: table 3.2 impella catheter components: "the infusion filter prevents bacterial contamination and air from entering the purge lumen." section: warnings & cautions: warnings: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
6714 | 1220648-2025-45720 | 22725443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725443&pc=OZD | 07/24/2025 |  |  | 08/06/2025 |  |  | 07/24/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701761 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock (pccs) was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced low hemoglobin and runs of ventricular tachycardia requiring intervention. the patient was admitted shortness of breath and elevated cardiac markers and diagnosed with non-st segment elevation myocardial infarction. the patient underwent coronary artery bypass surgery (x3) with a preoperative ejection fraction of 30%. grafts open with successful revascularization of all three vessels. however, the physician struggled to take the patient off bypass and realized that myocardium suffered reperfusion injury and stunning. the decision made to place impella 5.5 direct approach. graft was tunneled to right supra-clavicular area. the impella was successfully placed, and the patient was transferred to the unit on mcs. however overnight the patient experienced runs of ventricular tachycardia; amiodarone was started. later in the morning hemoglobin and hematocrit were low with concern of a gastrointestinal bleed. two units of packed red blood cells were given and heparin postponed. impella support continued four additional days until successfully weaned and device explanted with a patient survived outcome. | [5 empty] | The impella device was not received from the customer as the customer reported it has been discarded. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the event, clinical details noted that patient had runs of ventricular tachycardia overnight on first night of support and following morning; low hemoglobin hematocrit was noted with concerns for gastrointestinal bleed. in order to make a root cause determination, essential clinical details would have to be provided. the cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the cause of the gastrointestinal bleed could not be definitively determined as insufficient clinical details were provided. h6: investigation: type, findings and conclusion codes were updated accordingly based on the completed investigation.
6715 | 1220648-2025-31044 | 22725453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725453&pc=OZD | 07/23/2025 | 09/18/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698813 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient had a cp pump placement occurring to support a high-risk percutaneous coronary intervention. the patient had microvascular disease coronary artery disease but, despite all efforts the pump failed to deliver. the introducer kinked and pump could not be placed and so the cp was aborted and the groin site was closed. the vessel's native anatomy was noted to be tortuous. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues, introducer damage ¿ mechanical, product damage (sheath kink), and difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. delivery issues: the pump was unable to be advanced due to a sheath kink from tortuous vessels. delivery was aborted. based on the clinical details provided, the root cause of the delivery issues is most likely due to the patients condition as their tortuous and calcified vessels caused the sheath to kink and pump unable to be placed. introducer damage - mechanical: the peel away sheath kinked once the dilator was removed due to vessel tortuosity and calcification. based on the clinical details provided, the root cause of the introducer damage - mechanical is most likely due to the sheath kink as their tortuous and calcified vessels caused the sheath to kink and pump unable to be placed. product damage (sheath kink): based on the clinical details provided, the root cause of the product damage (sheath kink) is most likely due to the patients condition as their tortuous and calcified vessels caused the sheath to kink and pump unable to be placed. difficulty inserting/removing pump through the introducer: based on the clinical details provided, the root cause of the difficulty inserting/removing pump through the introducer is most likely due to the sheath kink as their tortuous and calcified vessels caused the sheath to kink and pump unable to be placed. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31044.
6716 | 1220648-2025-31045 | 22725537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725537&pc=OZD | 07/23/2025 | 09/18/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706519 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that following successful impella cp placement, flows immediately dropped and suction alarms appeared. troubleshooting was attempted but the issue remained. a fluoroscopy was performed and showed a kink in the cannula near the aortic valve. repositioning was attempted, but the issues remained and the decision was made to remove the pump. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of low pump flow and product damage (cannula kink) has been completed. product damage (cannula kink): clinical reported a cannula kink observed in fluoroscopy and difficulty navigating the patients anatomy and tortuous vessels. the product was returned for investigation and a small kink was observed. the cause of the cannula kink was most likely patient condition, as it was stated in the clinical details that there was difficulty positioning the pump through the patients anatomy and tortuous vessels. low pump flow: clinical reported an immediate suction alarm and drop in flows after the pump was switched on, as well as difficulty navigating the patients anatomy and tortuous vessels. the pump was returned for investigation and a small kink was observed near the outlet cage. data log review shows a slight downward trend in placement signal, but no motor current spikes outside of p-level changes. flow level is low from the start of the case, staying below 2 l/min. the cause of the low pump flow is most likely the cannula kink, as a kink was observed in physical product investigation, and clinical reports the kink occurred while positioning through tortuous anatomy.
6717 | 1220648-2025-45721 | 22725597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725597&pc=OZD | 07/25/2025 |  |  | 08/06/2025 |  |  | 07/25/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1683914 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the automated impella controller (aic) set off a controller failure alarm when turned on during prep for support of a patient undergoing high-risk percutaneous coronary intervention. consequently, another aic was used for the case, and there was no harm to patient as a result of this event. | [5 empty] | The automated impella controller data logs and console were returned for analysis and evaluation. data analysis demonstrated there was a controller failure alarm. the console was powered off. three days later the console was rebooted with the controller failure alarm persisting. on 25th july 2025, though the enterprise performance management (epm) hardware revision was correct as a known good epm hardware revision, the console still received the controller failure (epm1 sw-software-corruption) - alarm upon booting. the console was powered off and rebooted again on (b)(6) 2025. however, the event persisted. there were no battery related issues or alarms were seen in the event logs. it was noted; controller failure alarm could have occurred due to enterprise performance management software corruption or the incompatible hardware revision. device analysis demonstrated the console was returned to field service for further investigation. the reported failure mode was able to be reproduced. upon booting, the console received the controller failure alarm. a firmware check conducted via telnet confirmed that the issue was caused by a corrupted firmware on the impellatronic printed circuit assembly (pca) board. this was determined by comparing the extracted firmware with a known good version. the impellatronic pca was then replaced with a known good unit. following the replacement, the controller failure alarm was no longer triggered. additionally, the console was operated using a known good pump and purge cassette, as well as batteries were inspected, and no issues were observed during testing. in conclusion, the root cause for the controller failure (epm 1 sw corruption #1029) was due to a impellatronic pca firmware corruption. however, the cause of the firmware corruption was not determined. h6: investigation: type, findings and conclusion codes and additionally component coding were updated accordingly based on the completed device/event log analysis and investigation. | Accordingly. furthermore, additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. | The automated impella controller device was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6718 | 1220648-2025-31047 | 22725601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725601&pc=OZD | 07/20/2025 | 09/08/2025 |  | 08/06/2025 |  |  | 07/20/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697053 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 137 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 placed in a st-segment elevation myocardial infarction patient with ventricular tachycardia/ventricular fibrillation and prior placed extracorporeal membrane oxygenation. the pump was placed but after three days, the pump came out of position and was alarming for being in the aorta. the pump was removed and replaced. the patient remained stable and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most likely use related per clinical that the patient underwent cardiopulmonary resuscitation intervention leading to the pump becoming mispositioned.
6719 | 1220648-2025-31046 | 22725604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725604&pc=OZD | 07/23/2025 | 09/18/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693829 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had a cp pump placed to support the planned high-risk percutaneous coronary intervention. the patient was to have coronary microvascular disease treated when patient had declined the coronary artery bypass graft. the pump was placed and when the case completed the groin suffered a bleed. the bleed was treated by pressure being held, explant, vascular consult, perclose, and additionally an angioseal. patient survived. | [5 empty] | The investigation of access site bleeding has been completed. access site bleeding - major: clinical reports a large hematoma was observed around the sheath at the left groin. manual pressure was applied and p-level was dropped to p-2 after finishing intervention. the product was returned for investigation and no damage or abnormalities were observed. the cause of the access site bleeding was not determined due to lack of clinical details and no abnormalities observed in the product investigation. d9 device returned to manufacturer date added. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6720 | 1220648-2025-31048 | 22725635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725635&pc=OZD | 07/23/2025 | 09/17/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639774 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed hemolysis due to the impella being trapped under the papillary muscle. the physician did not want to reposition as the patient was scheduled for an left ventricular assist device the next day. left ventricular assist device was re-scheduled and the impella was repositioned two days later. patient survived. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle. no other issues were found on the logs. hemolysis: product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump being trapped in improper position. positioning issue: the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
6721 | 1220648-2025-31049 | 22725765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22725765&pc=OZD | 07/20/2025 | 08/06/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656679 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient arrived at the hospital after multiple codes. the patient was brought to cardiac catheterization laboratory and impella was placed. percutaneous coronary intervention while code completed. the groin was bleeding from the impella access site. the impella was removed, and a covered stent was placed. the patient was sent to the unit where the patient expired soon after arrival. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6722 | 1220648-2025-31050 | 22726013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22726013&pc=OZD | 01/10/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 01/10/2023 | 08/06/2025 |  |  | 08/31/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268484 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, pump lost optical signal during support. the signal loss did not cause any harm or injury, and the device remained in use without adverse effect to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the field log was reviewed and placement signal low, suction, and positioning issues were observed during the run. no other abnormalities were observed. the returned product was visually inspected and no abnormalities were observed. the cause of the optical signal issue was most likely patient condition related since the returned product operated to specification.
6723 | 1220648-2025-31051 | 22726154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22726154&pc=OZD | 07/23/2025 | 09/11/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had a cp pump being placed for the patient admitted in with plan for a high-risk percutaneous coronary intervention. the pump failed to deliver despite all measures and attempts to place via the introducer. the native anatomy precluded placement. no harm and plan to return and attempt an axillary site versus the femoral. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related.
6724 | 1220648-2025-31052 | 22726279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22726279&pc=OZD | 02/01/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 02/01/2023 | 08/06/2025 |  |  | 10/31/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285483 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump triggered a placement signal failure during use. however, the pump remained in use for support. no harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the product was not returned. the logs were reviewed and optical metrics were steady. the trending placement signal is indicative of a patient condition interaction with the sensor. the cause of the issue was most likely patient condition related.
6725 | 1220648-2025-31053 | 22726756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22726756&pc=OZD | 07/21/2025 | 09/09/2025 |  | 08/06/2025 |  |  | 07/21/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669862 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 131 KG | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed to support a patient admitted in from the community to the tertiary center. care was escalated from an intra-aortic balloon pump to an impella cp but the pump came out of position and could not be redelivered and so it was explanted and replaced. the patient expired on the second pump. however, we cannot disassociate the patients outcome from the first pump. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most likely use related as the physician pulled the pump across the valve when inserting the repositioning sheath.
6726 | 1220648-2025-31054 | 22726961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22726961&pc=OZD | 07/21/2025 | 09/09/2025 |  | 08/06/2025 |  |  | 07/21/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1484040 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) had a controller failure alarm. the aic was exchanged. there was no harm to the patient. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. console logs from the reported day of event are consistent with complaint and show controller failure alarm (#1029, due to epm software corruption), immediately preceded by indication of bad crc value during periodic epm software integrity verification. few minutes earlier, an error message was logged epm: error_float_conversion. couple minutes later, pump 600491 had been unplugged and transferred to backup console. console was tested, but the issue was not reproduced, and there was no indication of epm software corruption. software was extracted from the epm microprocessor to be compared with programming code, and it was identical. digital communication between epm and epm was monitored, but no anomalies were observed, and each epm sw verification showed identical values of epm software crc stored in device memory and calculated. analysis of epm software source code suggested that the value of hardware revision in code space might have been corrupted, which was communicated by intentionally mismatched crc value. the perceived epm software corruption was most likely caused by an undetermined epm software or hardware anomaly.
6727 | 1220648-2025-31055 | 22726982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22726982&pc=OZD | 07/09/2025 | 09/05/2025 |  | 08/06/2025 |  |  | 07/09/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574004 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Use of Device Problem | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | Hemodynamic instability | [12 empty] | Other | Death |  |  |  |  | IMPELLA 5.5 (PUMP 2) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella 5.5 device to a patient in postcardiotomy cardiogenic shock was unsuccessful. the patient underwent planned coronary artery bypass graft surgery (x2 vessels) and was weaned initially off bypass. however, while closing chest the patient continued to decompensate and was placed back on and off pump four times before decision made to place an intra-aortic balloon pump. the physician again attempted to wean off bypass and again the patient continued to decompensate with st changes noted on ekg. the decision was then made to place an impella 5.5 device via direct aorta approach. the 10mm graft was sewn and tunneled right supraclavicular introducer/graft locks in place. a 0.035 guidewire and al1 were used to cross aortic valve under fluoroscopy and then the guidewire was exchanged for an 0.018 guidewire. the impella wire backloaded onto impella 5.5 device and attempted to place into left ventricle; however, the device was unable to pass the anastomosis. a decision made to exchange wire. the impella was, therefore, removed and while attempting to exchange wires the device became unsterile. another impella 5.5 device was used to support the patient. however, after approximately 14 hours, the patient expired. no further details were noted. | [5 empty] | There is not enough evidence to exclude the initially attempted impella 5.5 device as an associated factor to the outcome given the delay in therapy for the decompensating patient. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since impella wire backloaded onto impella and attempted to place into left ventricle but was unable to pass anastomosis.
6728 | 1220648-2025-31056 | 22727231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22727231&pc=OZD | 07/23/2025 | 09/18/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692744 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during placement of an impella 5.5 in the patient for cardiogenic shock bailout post transcatheter aortic valve replacement, a left ventricular perforation occurred. due to prolonged left main (coronary artery) ischemia and left ventricle (lv) friability, a decision was made not to repair and to cease lifesaving efforts. therefore, the patient expired. additional information was received. the physician feels that the lv perforation was from extensive cardiac massage performed at various times after the chest was opened. | [5 empty] | The investigation of ventricle perforation has been completed. the clinical details state that during placement of the 5.5, an left ventricle (lv) perforation occurred as there was also prolonged lm ischemia and lv friability, so decision was made to cease life saving efforts. the returned product did not have any damaged or abnormalities. based on the clinical description, the root cause of the ventricle perforation was most likely due to the patients condition. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
6729 | 1220648-2025-31057 | 22727288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22727288&pc=OZD | 09/08/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 09/08/2023 | 08/06/2025 |  |  | 02/28/2025 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024333459 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the impella cp pump triggered "placement signal not showing" alarms. the pump remained successfully implanted in the patient. no known harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned and after looking at the logs, it is possible that the plug or pump connector could have caused a fiber air gap but also incomplete and intermittent electrical connection. without the pump returned, the root cause of the optical signal issue cannot be determined as an air gap or a defect cannot be ruled out.
6730 | 1220648-2025-31058 | 22727342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22727342&pc=OZD | 12/13/2024 | 08/06/2025 |  | 08/06/2025 |  |  | 12/13/2024 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554005 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the impella cp displayed an a-line pressure and position waveform that appeared abnormally out of sync, and a 'placement signal unstable' alarm was triggered. proper positioning was confirmed, and support was maintained with the implanted pump. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. impella cp was returned for evaluation. upon visual inspection, no damage was observed. optical parameters were all found to be within normal range. no alarms were reproduced with returned pump. data logs were reviewed. upon review of the data logs and returned product, it is apparent that the console is the potential cause of the issue (please refer to sister record "24-39808-2" for automated impella controller (aic) . no problem was found with the pump during the case.
6731 | 1220648-2025-31059 | 22727458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22727458&pc=PYX | 07/23/2025 | 08/06/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625103 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 98 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was placed on impella rp flex support. while on support, it was reported that after exchange of the peel away to a reposition sheath, a hematoma was noted. manual pressure was held and three units of packed red blood cells were given. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
6732 | 1220648-2025-31060 | 22727541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22727541&pc=OZD | 07/24/2025 | 12/04/2025 |  | 08/06/2025 |  |  | 07/24/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1649323 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller had a system self check failure. no patient involvement was reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check error¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation.
6733 | 1220648-2025-31061 | 22727792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22727792&pc=OZD | 07/24/2025 | 09/29/2025 |  | 08/06/2025 |  |  | 07/24/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571231 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced a pump stop. the pump could not be restarted despite trying a new supporting automated impella controller. the pump was explanted and no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation of pump stop has been completed. the impella cp was returned for evaluation. no physical abnormalities were found on the returned product. large purge gap was found which is due to bearing wear. data logs were reviewed, and a high motor current pump stop with alarm #13 occurred, on day 16. multiple attempts were made to restart the pump without success. the root cause of the pump stop was due to bearing wear as pump was run past indicated design life. d9 device returned to manufacture date added g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31061.
6734 | 1220648-2025-31062 | 22728442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22728442&pc=OZD | 01/27/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 01/27/2023 | 08/06/2025 |  |  | 09/30/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281693 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp triggered an ¿impella in aorta¿ alarm despite proper device placement. the placement signal (ps) was malfunctioning; however, the pump remained on for continued support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. impella cp was not returned for evaluation. the data logs were reviewed and the parameters were stable, showing no functional issues with the optical sensor. a reduction in motor current pulsatility was observed before and during the time of the impella in aorta alarm. this is most likely due to poor patient condition. the root cause of the optical signal issue is most likely patient condition.
6735 | 1220648-2025-31063 | 22728495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22728495&pc=OZD | 07/21/2024 | 08/06/2025 |  | 08/06/2025 |  |  | 07/21/2024 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519217 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump for support. during support, the pump alarmed with a 'placement signal not reliable' alert. the audio alarm was disabled. no harm to the patient was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the product was returned and evaluated. the returned pump catheter was cut and therefore could not be tested. there was biomaterial found around the impeller and near the optical sensor. no other abnormalities were found. data review matched the clinical description of the placement signal in the beginning of the case that eventually resolves. the root cause of the optical signal issue on the pump was not determined since optical parameters could not be tested on the returned product.
6736 | 1220648-2025-31064 | 22728567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22728567&pc=OZD | 09/04/2024 | 08/06/2025 |  | 08/06/2025 |  |  | 09/04/2024 | 08/06/2025 |  |  | 02/28/2026 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella cp pump for support. approximately 15 minutes after the start of the case, the impella representative observed that the aortic placement waveform and left ventricular placement signal had significantly increased. the case proceeded without incident, and patient support was not interrupted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned for investigation. no data logs provided. no pump returned for analysis. the failure mode was changed to 'optical sensor issue' as cp pump does not have a dp sensor which the failure mode "placement signal issue" is for. the root cause of the optical signal issue was not determined as no product or logs were returned for analysis.
6737 | 1220648-2025-31065 | 22729009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729009&pc=OZD | 02/10/2025 | 08/06/2025 |  | 08/06/2025 |  |  | 02/10/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025606537 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella 5.5 pump for support. during support, after ecmo was explanted, the impella 5.5 experienced a failure of the optical sensor. the device¿s position was verified, and no harm to the patient was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned for investigation. data logs were not returned as well. without the logs and the product, the failure cannot be investigated. therefore, the root cause of the optical signal issue was not determined since product was not returned and insufficient clinical information was provided.
6738 | 1220648-2025-31066 | 22729218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729218&pc=OZD | 04/11/2024 | 08/06/2025 |  | 08/06/2025 |  |  | 04/11/2024 | 08/06/2025 |  |  | 12/31/2025 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449398 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during implantation of the impella cp, a 'placement signal not reliable' alarm occurred. a backup console was used to resolve the issue. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. pump was not returned for investigation. data log shows the placement signal not reliable alarm along and all parameter was down which indicate hard failure of the sensor. as per the clinical details, placement signal was lost immediately after the shockwave device used. the cause of the optical signal issue was damaged sensor due to the shockwave device interaction, based on given clinical details and data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. refer to the dhr checklist for indication of: "the following qns associated with this device, which are unrelated to this failure mode. 1. (b)(4) associated with the delo discovered on lumen on all 10 units of motor housings and single pin subassemblies. all 10 unites have been de-kitted and scrapped out from work order. 2. (b)(4) associated with the deviation related to 100% inspection in pigtail-inflow cage icls and final visual icls for debris in the inflow cage." this review revealed that this product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.
6739 | 1220648-2025-31067 | 22729312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729312&pc=OZD | 04/29/2024 | 10/02/2025 |  | 08/06/2025 |  |  | 04/29/2024 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025478117 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the automated impella controller displayed a 'placement signal is not reliable' alarm. however, all device metrics remained within normal ranges, and no abnormalities were observed. the pump continued to provide circulatory support and was not removed as a result of the alarm. no patient harm was reported. | Additional information care was eventually withdrawn, and the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. pump was not returned for investigation. clinical reported aic displayed 'placement signal is not reliable' alarms. data logs confirm 'placement signal not reliable' alarms. the root cause of the optical signal issue was not determined because no product was returned and no failure pattern was identified. | Corrections that were made from the initial manufacturer device report number 1220648-2025-31067. b1 adverse event updated. b2 death and death date updated . b5 updated to include death description. h6 updated to include death coding.
6740 | 1220648-2025-31068 | 22729497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729497&pc=OZD | 05/22/2024 | 08/06/2025 |  | 08/06/2025 |  |  | 05/22/2024 | 08/06/2025 |  |  | 07/31/2025 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443272 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. it was noted that the optical sensor stopped functioning when the shockwave was delivered. motor current and flow remained within normal limits, and no other issues were noted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned for investigation. data logs did not report any issues related to placement signal. no trend was observed in the optical signal parameters. according to clinical reports, the optical sensor went out during the shockwave. no issues were noted with the motor current or pump flow. the cause of the optical signal issue was not determined because the product was not returned and the data logs did not confirm any issue related to the optical signal.
6741 | 1220648-2025-31069 | 22729655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729655&pc=OZD | 07/23/2025 | 09/19/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629194 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 81 KG | Hispanic | White | Death | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hematoma | Ventricular Fibrillation | Asystole | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp being placed to support a non-st-segment elevation myocardial infarction patient with multivessel coronary artery disease and acute myocardial infarction/cardiogenic shock. the primary registered nurse reported a large hematoma. manual pressure was held but the hematoma returned. a large unresolvable hematoma in the right groin was reported. additionally, the impella position was unable to be optimized. the patient was intubated for airway protection and was placed back on the catheterization lab table. a repo unit was seen to have not be able to reach the vessel. the doctor attempted to advance but resistance was being met on advancement of the impella. inlet at the level of aorta valve in the ventricle. visible kink from repo at the level of the groin. the medical doctor attempted to advance again but was unsuccessful. the doctor rewired the reaccess sheath to attempt reinsertion. however, the impella cp was removed and explanted due to the complication. an additional impella was used. a new introducer sheath was placed. the patient coded and the medical doctor tried to access the left groin. advanced cardiac life support measures were continued for several minutes. return of spontaneous circulation (rosc) achieved after several rounds of cardiopulmonary resuscitation. massive blood transfusion was called. the patient continued to go into ventricular fibrillation. drips were started and maxed out. the patient continued to be in asystole and impella cp flows were unable to reach greater than one liter per minute. the patient expired. additional information from the representative was provided. "they mainly attribute the death to the internal bleed from the repo sheath not reaching the vessel. the md attempted to advance the sheath back but was unsuccessful due to stenosis, so the only other option was to reinsert it again. he tried getting contralateral access but unfortunately unsuccessful due to stenosis also in the left cfa. the patient didn't tolerate the remove and reinsertion of the impella. the blood loss, severe disease, and pause of support all contributed to the death of the patient." | [5 empty] | The investigation of access site bleeding, product damage (catheter kink), delivery issues, and vf/vt/af/at has been completed. per physical product investigation, product damage (cannula kink) was added as a failure mode. access site bleeding - major: the product was returned for investigation and a small puncture mark was observed near the end of the repositioning sheath. the cause of the access site bleeding was not determined due to insufficient clinical details. product damage (catheter kink): the product was returned for investigation, and a small puncture was observed on the sheath of the repositioning unit, but there were not any kinks observed along the catheter. the cause of the puncture was not determined due to lack of clinical details. delivery issues: the product was returned for investigation, and a cannula kink was observed. the cause of the delivery issues was most likely patient condition, as clinical reported disease in the aorta that caused difficulty advancing the repositioning unit to the vessel. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. product damage (cannula kink): the product was returned for investigation and kinking was observed close to the outlet cage on the cannula. the cause of the cannula kink was most likely patient condition, as clinical reported resistance met when trying to advance the impella, as well as disease found in the aorta. b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31069. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer - ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6742 | 1220648-2025-31070 | 22729912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729912&pc=OZD | 10/16/2024 | 08/06/2025 |  | 08/06/2025 |  |  | 10/16/2024 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544040 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella 5.5 pump for circulatory support. during support, the device generated ongoing 'placement signal not reliable' alarms. pump position was confirmed to be correct. the audio alarms were disabled. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of placement signal issues, per fda recommendation. the investigation of placement signal issue has been completed. the product was not returned for investigation. data logs were reviewed, and the alarms were confirmed. upon review of data logs and previous pump performance on the console, it is apparent that the console is the potential cause of the issue. please refer to complaint#: (b)(4) for an investigation into the console.
6743 | 1220648-2025-31071 | 22729951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729951&pc=OZD | 06/03/2010 | 08/06/2025 |  | 08/06/2025 |  |  | 09/13/2011 | 08/06/2025 |  |  | 01/19/2009 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 0030765 |  | 004671-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that two hours into support an impella 2.5 pump experienced a sudden spike in motor current followed by a pump stop. the pump could not be restarted and was removed. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
6744 | 1220648-2025-31072 | 22729964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729964&pc=OZD | 11/24/2024 | 10/02/2025 |  | 08/06/2025 |  |  | 11/24/2024 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512870 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the use of an impella 5.5 pump for circulatory support. during support, the pump triggered a placement signal (ps) alarm indicating that the placement signal was not reliable. the nurse inspected the lines and confirmed there were no kinks in the cable. additionally, the patient was not moving at the time the alarm occurred. an echocardiogram was ordered to verify device position, and the audio for the 'ps not reliable' alarm was disabled. there was no patient harm. | Additional information care was eventually withdrawn, and the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issue, per fda recommendation. the investigation of optical signal issue has been completed. neither data logs nor product were available for investigation. clinical details do not suggest any potential causes for the alarm and do not mention any methods of troubleshooting that may have resolved the issue. due to limited information, the root cause of the optical signal issue could not be determined. | Corrections that were made from the initial manufacturer device report number 1220648-2025-31072. b1 adverse event updated. b2 death and death date updated. b5 updated to include death description. h6 updated to include death coding.
6745 | 1220648-2025-31073 | 22729988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22729988&pc=OZD | 07/23/2025 | 09/18/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693839 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp placement was aborted due to the patient¿s aneurysm and tortuosity in the aorta. there was no reported patient harm. the patients outcome is stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due aortic aneurysm and tortuosity in the aorta. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
6746 | 1220648-2025-45722 | 22730294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22730294&pc=OZD | 07/26/2025 |  |  | 08/06/2025 |  |  | 07/26/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704412 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male | 86 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device as elective placement for mechanical circulatory support and then subsequently experienced bleeding and expired. the patient was a transfer in for advance care and found to be in cardiogenic shock with severe aortic insufficiency, mitral regurgitation, tricuspid regurgitation, and pulmonary stenosis. the patient underwent aortic valve replacement (avr) and pulmonary valve replacement (pvr). during this surgery, surgeons also found a left main stent was in the aorta and was subsequently removed. after this stent removal, the left main was found to be patent, and coronary artery bypass graft surgery was not required. post avr and pvr, the impella 5.5 device was successfully implanted via the right axillary artery. the impella support was initiated on performance level p-2, the patient slowly weaned from cardiopulmonary bypass (cpb), and the impella was then dialed up to p-8. flows were approximately 4.7lpm with inlet placed approximately 5.2 cm from the aortic valve annulus and midventricular. shortly after coming off cpb, impella suctioning occurred. right ventricular failure was noted on echocardiogram, and there was a concern for failed pulmonary valve replacement. a decision made to reconfigure to venous-arterial extracorporeal membrane oxygenation (va ecmo). the patient continued to have excessive bleeding and required multiple blood products. the patient was eventually transferred to the unit on impella and ecmo support. however, approximately one day later, the patient was noted to have expired. care was withdrawn. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, peri-procedural adverse event, bleeding: the root cause was most likely patient condition since the patient was coagulopathic. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Additional comments by the physician stated the bleeding is related to extracorporeal membrane oxygenation and the patient¿s predisposed coagulopathy issues with underlying congenital condition. however, there is not enough evidence to exclude the impella device as a contributing factor in the patient¿s outcome. product status: the impella device was not received from the customer; there is uncertainty what happened with the device after the patient's demise. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6747 | 1220648-2025-31074 | 22730332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22730332&pc=OZD | 06/14/2025 | 08/26/2025 |  | 08/06/2025 |  |  | 06/14/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661764 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Material Split, Cut or Torn | [5 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support the patient awaiting heart transplant. while on support the nurse called and reported that the back end of sterile sleeve separated from 2nd luer lock. advised to clean with betadine and dress/secure sterile dressing. later it was reported that impella in aorta alarm however it self resolved. additionally it was reported that the patient blood cultures were positive from unknown origin and had been started on vancomycin. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of infection/sepsis, product damage (sleeve damage), and optical signal issue has been completed. the impella device was not returned for evaluation. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. product damage (sterile sleeve damage): no product was returned for analysis. the root cause of the product damage (sterile sleeve damage) was not determined as no product/images were returned optical signal issue: clinical reported impella position in aorta alarm which self resolved. the data logs do not show any 'impella position in aorta' alarms but show 'impella position unknown' which occurred outside complaint reporting clinical timeline (alarms on 6/6, issue reported 7/21). the 5s parameters were stable and within expected ranges. the logs show a steady increase in placement signal during support and no motor current spikes outside p-level changes. the root cause of the optical signal issue was most likely patient condition related as evidenced by clinical account of self resolution and no related abnormalities observed in data logs d6a implant date was erroneously entered on the initial manufacturer device report number 1220648-2025-31074. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31074. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock . section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6748 | 1220648-2025-31075 | 22730347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22730347&pc=PYX | 12/20/2024 | 10/02/2025 |  | 08/06/2025 |  |  | 12/20/2024 | 08/06/2025 |  |  | 02/28/2026 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517941 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump triggered a 'placement signal not reliable' alert. although the pump was advanced, the alert persisted. despite this, the pump continued to provide circulatory support but was unable to calculate flow rates. no patient harm was reported. | Additional information care was eventually withdrawn, and the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of placement signal issues, per fda recommendation. the investigation of placement signal issue has been completed. no product was returned for evaluation. data logs were reviewed. the root cause of the placement signal issue was most likely due to sensor drift. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. please refer to dhr checklist for indication of: "there were 5 other product malfunction complaints (b)(4) in subcomponent lot#: 2024320708 related to this failure mode with a total lot quantity of (b)(4)." this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. | Corrections that were made from the initial manufacturer device report number 1220648-2025-31072. b1 adverse event updated. b2 death and death date updated . b5 updated to include death description. h6 updated to include death coding.
6749 | 1220648-2025-31077 | 22730588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22730588&pc=OZD | 07/12/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 07/12/2023 | 08/06/2025 |  |  | 03/31/2025 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346463 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during use, the pump experienced a loss of placement signal; however, it remained in use to provide support. no harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the returned product was investigated and damage to the sensor face was observed. the damage was consistent with shockwave interaction. the parameters were tested prior to sensor removal and were found to be outside of normal range. the log shows a change is parameters consistent with damage to the sensor occurring when there was a loss of placement signal. the cause of the optical signal issue was a damaged sensor due to shockwave interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.
6750 | 1220648-2025-31076 | 22730600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22730600&pc=OZD | 07/23/2025 | 09/15/2025 |  | 08/06/2025 |  |  | 07/23/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701669 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 was advanced to aortic valve (av) but was unable to cross. the wire was lost wire and the left ventricle was accessed. the right axillary was attempted four times but was unable to cross the av. the catheter behind the motor was going down descending. the procedure was aborted and was switched to left axillary with same steps as above. the anatomy was still tortuous but better. attempted inspiratory was held to help advance but was unable to. the wire stayed in the left ventricle but was unable to advance the catheter through the patients anatomy. the case was aborted. the patients outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of the size mismatch from native to donor. the catheter was unable to advance catheter through anatomy due to tortuous and the insertion was aborted.
6751 | 1220648-2025-31078 | 22730723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22730723&pc=OZD | 01/26/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 01/26/2023 | 08/06/2025 |  |  | 06/30/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234311 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the pump triggered a false 'impella in aorta' alarm. proper pump positioning was confirmed via echocardiography. no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the product and logs were not returned. the cause of the false position in aorta alarm was patient condition related.
6752 | 1220648-2025-31079 | 22731248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22731248&pc=OZD | 02/07/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 02/07/2023 | 08/06/2025 |  |  | 12/31/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292237 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the pump triggered a 'placement signal not reliable' alarm both during setup and support. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the field log was reviewed and upon connection, placement signal not reliable triggered and persisted. pump provided sufficient support. the returned pump was connected to an automated impella controller (aic) and the placement signal not reliable alarm reproduced. the sensor was removed and damage to the sensor was observed. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.
6753 | 1220648-2025-31080 | 22731707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22731707&pc=OZD | 03/06/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 03/06/2023 | 08/06/2025 |  |  | 10/31/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023272399 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the pump triggered a 'placement signal lost' alarm after more than 70 days of runtime. pump position was confirmed via echocardiography. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned. data logs showed the "placement signal not reliable" alarm triggered; the waveform remains out of specification for the remainder of support. the optical sensor parameters at the time of the failure were characteristic of optical fiber damage due to torque. therefore, it was determined that the cause of the optical signal issue was most likely a damaged optical fiber line.
6754 | 1220648-2025-31081 | 22731835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22731835&pc=OZD | 05/17/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 05/17/2023 | 08/06/2025 |  |  | 07/31/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252341 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump triggered multiple 'impella in aorta' alarms despite confirmed proper positioning. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump and logs were not returned. the cause of the false position in aorta alarm was most likely patient condition related.
6755 | 1220648-2025-31082 | 22731865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22731865&pc=OZD | 07/06/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 07/06/2023 | 08/06/2025 |  |  | 12/31/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024287865 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump triggered placement signal issues. the pump was not removed. support was continued, and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned. as per clinical, continued impella in aorta alarm was observed even after pump position was confirmed and the alarm was disabled. data logs were reviewed and aorta alarm were observed. the cause of the optical signal issue was most likely patient condition related.
6756 | 1220648-2025-45725 | 22731869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22731869&pc=OZD | 04/25/2025 |  |  | 08/06/2025 |  |  | 07/22/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654486 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | Thrombosis/Thrombus | [10 empty] | Hospitalization | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Notifications were received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that experienced ventricular tachycardia, access site hematoma and thrombus all requiring intervention. post implant, same day in the evening, the patient experienced sustained ventricular tachycardia in which the patient was direct cardioversion back to normal sinus rhythm with a "definite" relationship to the impella device used. the impella device was repositioned at bedside under echo guidance, and thereafter this patient remained electrically quiet. the patient was listed as status 2 for orthotopic heart transplant. two days later, intermittent runs of ventricular tachycardia were experienced to which maps remained within normal limits; converted once patient vagels. the impella 5.5 device remained at p-8; milrinone and nipride on-going. days thereafter, increased ventricular tachycardia burden was noted. lidocaine was added. ventricular arrhythmia appropriately suppressed on iv lidocaine infusion. the patient recovered with no sequelae. now day 7 of impella mechanical circulatory support, the patient experienced a right axillary hematoma with a "definite" relationship to the impella device used: the hematoma appeared bigger and cardiothoracic surgery evaluated; noted to be stable. heparin drip was suspended for 4 hours. the patient recovered with no sequelae. on day 21 post start of the impella mcs, the patient experienced a nonocclusive superficial venous thrombosis in the left cephalic vein with a "possible" relationship to the impella device used. eliquis was started over the weekend. notes reflected, ¿no indication for ac. stop heparin.¿ the patient recovered with no sequelae. the patient was eventually bridged to transplant and the impella 5.5 device was successfully explanted with a patient survived outcome. | [5 empty] | Note: explant date and thrombosis event date are being confirmed to determine if said event occur while on or post impella support. product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use as noted in the impella 5.5: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6757 | 1220648-2025-31083 | 22731887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22731887&pc=OZD | 02/03/2023 | 10/27/2025 |  | 08/06/2025 |  |  | 02/03/2023 | 08/06/2025 |  |  | 10/31/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023274738 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. after 40 days of support, the pump triggered placement signal issues. the pump was not removed. support was continued, and no harm occurred to the patient. | [5 empty] | G3 ¿date received by manufacturer¿ corrected. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned. the log shows that at the time of the placement signal not reliable alarm when placement signal is lost. the cause of the optical signal issue was most likely a damaged optical fiber line caused likely by a kink in the catheter. | All pertinent information available to abiomed has been submitted at this time.
6758 | 1220648-2025-45713 | 22731938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22731938&pc=OZD | 07/22/2025 |  |  | 08/06/2025 |  |  | 07/22/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653711 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 64 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient had gastrointestinal bleeding (gi bleed) and low pump flow issues during impella 5.5 support. heparin was held due to the bleeding. discussion was held on product replacement, however, the patient is still currently on support with the same device. additionally, the patient experience impella suction and the p level was dropped to p5. two units of blood and volume were given to the patient. patient is stable and on support so far. | It was reported that several days of support the patient expired. |  |  |  |  | A.4 added weight as it was inadvertently omitted from the initial report that was submitted. d.6b explant date was added. e.1 zip code extension was added as it was inadvertently omitted on the initial report that was submitted. h.6 code f01 was added as it was inadvertently omitted from the initial report that was submitted. the customer's device was not returned for investigation. event logs were returned and reviewed. the logs show intermittent instances of suction alarms correlating with the patient updates. the suction alarms are accurately triggered as there was slight lower motor current values during this time, correlating with changes in pressure. there were no motor current spikes or drastic changes in flow. based on the clinical information and data logs, the root cause of the suction is most likely due to a patient related issue. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received and the following fields were updated: b2, b5, h1, h6.
6759 | 1220648-2025-31084 | 22733150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22733150&pc=OZD | 03/23/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 03/23/2023 | 08/06/2025 |  |  | 11/30/2024 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278056 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. after 89 days the pump experienced placement signal drift. support was continued, and no harm occurred to the patient. | [5 empty] | D6: explant date unknown. this report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the field log was reviewed and the optical metrics were steady. the root cause of the optical signal failure was most likely an air gap from between the automated impella controller (aic) and the patient cable plug or within the middle of the red handle ferrule.
6760 | 1220648-2025-45715 | 22733637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22733637&pc=OZD | 07/22/2025 |  |  | 08/06/2025 |  |  | 07/22/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680853 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had a gastrointestinal bleed (gi bleed) during impella 5.5 support. the patient had recurrent gi bleeding requiring blood transfusion. in the setting of mixed shock picture and significant gi bleeding, the patient was sedated, extubated, vasoactive drips were stopped and the impella was tapered off and stopped. as such, the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. vascular damage : the root cause of the vascular damage gi bleed issue was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6761 | 1220648-2025-31085 | 22733864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22733864&pc=OZD | 05/04/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 05/04/2023 | 08/06/2025 |  |  | 03/31/2025 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324307 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. a 'position in aorta' alarm occurred intermittently but appeared to be false. echocardiography confirmed good left ventricular positioning of the impella, and motor current demonstrated good pulsatility. no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump and logs were not returned. the cause of the false position in aorta alarm was patient condition related.
6762 | 1220648-2025-31086 | 22734336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22734336&pc=OZD | 06/23/2023 | 08/06/2025 |  | 08/06/2025 |  |  | 06/23/2023 | 08/06/2025 |  |  | 04/30/2025 |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333079 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced a persistent 'placement signal low' alarm. the impella was pulled back 1 cm, but the alarms persisted. support was continued and no harm occurred to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned for evaluation. data logs were returned and impella position unknown, placement signal low alarms were observed. the cause of the optical signal issue was most likely patient condition related with placement signal low alarms.
6763 | 1220648-2025-45706 | 22734342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22734342&pc=OZD | 06/30/2025 | 07/30/2025 |  | 08/06/2025 |  |  | 06/30/2025 | 08/06/2025 |  |  |  |  | 08/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689019 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient who had been weaned from an initial impella 5.5 and explanted for coronary artery bypass surgery was attempted to be implanted with a new impella 5.5 device for mechanical circulatory support as the patient was removed from bypass support. it was reported that the physician was advised to place a new graft for the insertion of the new impella 5.5, however the physician declined to place a new graft and attempted to insert the impella 5.5 into the used axillary graft. the physician was unable to advance the device through anastomosis and the implant was aborted. there was no patient harm reported. | [5 empty] | Product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action.
6764 | 1220648-2025-30987 | 22713101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713101&pc=OZD | 07/21/2025 | 09/29/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688986 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during prepping, the impella cp purge system failed to prime. the automated impella controller (aic) was exchanged but the issue was not resolved. a new impella cp and a new aic were used. no harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of priming issue has been completed. the pump and purge cassette were returned for investigation. no visual abnormalities were reported on the returned products. the pump and purge cassette were primed without issues, with purge fluid seen exiting from motor end. the pump was tested for an additional 10 minutes without any observed issues. the data logs requested from the remote server were not sufficient to provide detailed case start information. the cause of the priming issue was not determined since no issue found on the returned product and sufficient clinical information was not provided including data log.
6765 | 1220648-2025-30988 | 22713187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713187&pc=OZD | 07/21/2025 | 10/07/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143273 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during prepping, the impella cp purge system failed to prime and there was an automated impella controller (aic) failure. the aic was exchanged but the issue was not resolved. a new impella cp and a new aic were used. no harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of priming issue has been completed. the impella device was returned for evaluation. since no priming issues were found in the associated case logs for either of the console and the console functioned as intended, the reported failure mode is most likely a use issue.
6766 | 1220648-2025-30989 | 22713245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713245&pc=OZD | 07/21/2025 | 09/17/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688182 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the anticontamination sleeve ripped. no further information was provided and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage) has been completed. the clinical details state that there was a sleeve tear. the returned product showed that the sleeve was detached from the side closest to the blue butterfly. there was tegaderm on it to seal the sleeve back to the ring. due to lack of clinical details regarding how the sleeve became torn, the root cause of the product damage (sleeve damage) cannot be determined. d9 device returned to manufacturer date added.
6767 | 1220648-2025-30990 | 22713338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713338&pc=OZD | 07/21/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1690147 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a console was being used for training when there was a system self check failure observed. a "change console immediately" alarm was triggered when the console was turned on. there was no patient involvement as this occurred during an in-service. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. console logs were consistent with what was reported in the complaint. powermod_1 communication issues were observed in the logs. console was tested and the reported system self-check failure was able to be reproduced while testing on an engineering lab bench. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection of the consoles internal components confirmed that there was power battery manager (pbm) contamination which has impacted a neighboring rs232 communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
6768 | 1220648-2025-30991 | 22713453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713453&pc=OZD | 07/21/2025 | 09/17/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695497 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 63 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp support had hemolysis. the medical doctor was concerned about a drop in platelets, elevated lactate dehydrogenase and low haptoglobin. one unit of platelets was given. the medical doctor wanted to wean and remove the impella cp. the pump and explanted and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation of hemolysis and thrombocytopenia has been completed. the impella device was not returned for evaluation. hemolysis: no product was returned for analysis. the data logs do not show any motor current spikes outside p-level changes or any trends or abnormalities with a relation to hemolysis. data logs show no positioning issue. there were a few suction alarms overnight leading into day rising ldh levels were reported. no information on any clinical event before rising ldh was reported. good pump position was not confirmed via imaging. the root cause of the hemolysis was unable to be determined as no product was returned for analysis, no related data log abnormalities were seen and limited clinical information was provided thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details.
6769 | 1220648-2025-45714 | 22713516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713516&pc=OZD | 07/10/2025 |  |  | 08/05/2025 |  |  | 07/10/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655829 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Tamponade | Thrombosis/Thrombus | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device as elective placement for mechanical circulatory support (mcs) and developed thrombus requiring intervention. a little over a month prior, the patient was transferred from an outside facility in acutely decompensated heart failure. the patient was in cardiogenic shock supported by inotropes and an impella 5.5. a transesophageal echocardiogram prior to a planned heartmate 3 left ventricular assist device (lvad) revealed dilated ventricles, ejection fraction of 30-35% with mildly reduced right ventricular function, aortic and tricuspid regurgitation. the lvad implantation was, therefore, aborted. impella 5.5 was removed and milrinone was discontinued. later while on the step-down unit the patient was on dobutamine, difficult to get off inotropes, which was thought to be from valve disease. therefore, a decision made to perform aortic and tricuspid valves replacement with elective placement of the impella 5.5. later that evening, a protek duo right ventricular assist device was emergently placed as the patient had been suctioning with low left ventricular filling pressure and only able to flow at performance level p-5 at 3.0l. after the rvad placement, the impella 5.5 support level was able to be increased to p-7 with maximum flows at 3.0l. hemoglobin was low, and the patient was administered two units of packed red blood cells. central venous pressure was now at 12; patient began making more urine. patient critically ill continued on impella 5.5 and rvad therapies. the next morning the patient had a drop in rvad flows. potential thrombus was noted at the intra jugular site. transesophageal electrocardiogram showed external compression of the right atrium or clot. the impella 5.5 was noted to be in optimal position. later that evening the patient underwent an emergency washout due to progressive tamponade picture and decreased rvad flows. significant amounts of clot were removed. the patient returned to the unit noted to be doing better. the patient continued to improve over the next couple of days with plans to increase the impella 5.5 flow and explant of the rvad, which was removed after a total of 6 days on this support. the impella 5.5 device support continued for an additional eight days before being successfully weaned and device explanted with a patient survived outcome. | [5 empty] | Product status: the impella device was not received from the customer. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The product was not available for return by the customer as it was not saved. event logs were received and analyzed however. device history record review was completed and revealed the device passed all post sterile inspection checks. cardiac tamponade and thrombus: in order to make a root cause determination, all essential clinical details as outlined would have to be provided. the root cause of the cardiac tamponade and thrombus was unable to be determined due to insufficient clinical details. low or blocked pump flow/suction: according to the clinical report, the protek rv assist device was used to improve flow, and blood products were administered during the episode of low pump flow on (b)(6) 2025. the pump was implanted on (b)(6) 2025 at 11 pm. event data log confirms that pump flow was at the lower specification limit during the first day of implant, with placement signal and pump flow trends while running on a steady p-level. the following day, the pump was turned to p-6 with improved flow and suction alarms. no positioning issues or motor current spikes were reported during the case run. the cause of the low pump flow and suction was most likely patient condition-related, based on data log review and provided clinical details.
6770 | 1220648-2025-30992 | 22713518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713518&pc=OZD | 07/25/2025 | 09/25/2025 |  | 08/05/2025 |  |  | 07/25/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646686 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had a cp pump placed to support the elderly patient admitted into the tertiary center after being transferred from the community. the patient was in acute failure, intubated, and cp place to allow for valve surgery. the pump was placed but the pulse pressure rose with pulmonary fibrosis dropping, and troubleshooting discussions were correct in that there was ingested clot observed when the pump was explanted and replaced. | [5 empty] | The investigation of high purge pressure and thrombus has been completed. high purge pressure: data logs were not recovered. the pump was returned and inspected and biomaterial was found in the purge gap. the cause of the high purge pressure was biomaterial due to biomaterial found in the purge gap. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
6771 | 1220648-2025-30993 | 22713605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713605&pc=OZD | 07/21/2025 | 09/16/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662649 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after one day of impella cp support, the patient began to develop hemolysis evidenced by hematuria. the patient's urine was red in color. the positioning was checked and it was noted that the pump was deep and angled towards the mitral valve. the pump was successfully repositioned without issue. however, roughly thirty minutes later, a spike in motor current occurred followed by suction alarms and low flows. the pump was removed in response to the failure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation reported for hemolysis, low pump flow, and temporary pump stop has been completed. logs showed after pump started & run for 15 hours without issue, mc spiked followed low flow & temporary pump stop that triggered impella stopped-mc high, auto suction response & suction alarms. the low flow then remained to the rest of the case. visual inspection found biomaterial inside the inflow cage. no other issues were found on the rest of the pump. pump was conducted hemolysis test. no low flow occurred during testing & the pump passed hemolysis test. the root cause of hemolysis, low pump flow, and temporary pump stop was most likely biomaterial ingestion since biomaterial was found on the pump, data log meet the hemolysis, low pump flow, and temporary pump stop pattern & no issue was found on the pump.
6772 | 1220648-2025-30994 | 22713695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713695&pc=OZD | 07/16/2025 | 08/14/2025 |  | 08/05/2025 |  |  | 07/17/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673827 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support had ectopy episodes on day two of impella support. on day four of impella support, the physician attempted to advance 1cm and the patient had some ventricular tachycardia, so the cp was pulled back to original positioning. additionally, the patient had pocket bleeding at the axillary cutdown site. the patient required wound exploration and washout. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. after further review the bleeding reported in the initial report was for the impella 5.5 axillary bleeding and not related to the impella cp. the impella 5.5 bleeding was captured under complaint manufacturer device report 1220648-2025-45821. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-30994. b5, h6, h10 (instructions for use) erroneously included the access site bleeding/1888 on the initial manufacturer device report number 1220648-2025-30994. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6773 | 1220648-2025-30995 | 22713732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713732&pc=OZD | 07/20/2025 | 09/11/2025 |  | 08/05/2025 |  |  | 07/20/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682506 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support had bleeding from the groin site. in intensive care unit, the forward tension sutures were redone. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related since forward tension sutures redone to achieve hemostasis. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6774 | 1220648-2025-30996 | 22713783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713783&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679770 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced an optical signal malfunction. placement signal not reliable alarm ongoing with ¿placement monitoring disabled/audio off¿ was noted since the operating room per the nurse. impella 5.5 support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6775 | 1220648-2025-30997 | 22713870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713870&pc=OZD | 05/22/2025 | 11/25/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1727631 |  | 0042-0040-UK |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was being prepped for an impella cp implant and during the setup, the device was not recognized. it kept asking for grey-to grey connection like with the old cp. it was a cp smart assist. it eventually then seem to rectify itself so the staff continued with set up. the investigating engineer found the issue could be due to the automated impella controller (aic). | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. data analysis: review of the imc data logs confirms the complaint of ¿device was not recognised on set up for ages, it kept asking for grey-to grey connection like with the old cp.¿ there are entries suggesting that the impellatronic pcba was reset twice before the pump was eventually detected. these resets are characteristic of a software solution that was put in place to mitigate a pump detection issue that is secondary to a quartz crystal on the impellatronic pcba either failing to start, or to properly sustain oscillation. device analysis: the console was not returned for analysis despite being requested. root cause: the cause of the issue was not determined since product was not returned. d10 concomitant medical products and therapy dates updated.
6776 | 1220648-2025-30998 | 22713891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713891&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586182 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced multiple episodes of ventricular fibrillation requiring defibrillation. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6777 | 1220648-2025-30999 | 22713939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22713939&pc=OZD | 02/03/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Abiomed japan received a publication from italian authors and is entitled "mechanical external compression devices to manage bleeding during mechanical circulatory support with impella cp" which includes a case of impella cp in which "after the revascularization, for patient, a clinically relevant, access-site bleeding occurred" the patient is a 65 yo male. the 65 year-old patient had bleeding at the access site. the first attempt to stop the bleeding was through the pull-up stitching technique with gauzes, but it was unsuccessful. then, we used local tranexamic acid at the site of the repositioning sheath and partially locked the knots of the two proglides. nevertheless, external manual compression was the only way to limit the bleeding. a safeguard device was applied. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a.1- a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
6778 | 1220648-2025-31000 | 22714235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22714235&pc=OZD | 07/21/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619822 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that patient had a 5.5 pump placed to support surgical patient. the pump was supporting while bridging to a left ventricular assist device. the pump was explanted after 16 days with high purge pressure. the malfunction caused no harm and patient survived the weaning and explant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6779 | 1220648-2025-31001 | 22714520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22714520&pc=OZD | 07/21/2025 | 09/25/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 202666690 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient had a 5.5 pump support placed for the patient admitted in for surgery. the pump was supporting when after a month of support there was a need to add lytics to the purge fluid to troubleshoot the elevated purge pressure. the pump was not explanted and remains on despite the malfunction. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. pump was not returned for investigation. data log shows a 3-hour gradual rise in purge pressure up to 1100 mmhg, with a gradual drop in purge flow. only the purge flow decrease alarm was triggered during the purge pressure rise trend. the purge trend returned to normal after 20 hours of pump run with a gradual decreasing trend. the doctor mentioned that the lysis protocol successfully resolved the purge issue. the cause of the high purge pressure issue was most likely biomaterial buildup during the case run, based on data log review and clinical troubleshoot method. d6b explant date added.
6780 | 1220648-2025-31002 | 22714854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22714854&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667892 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump was placed but there was pump movement at the blue hub of the repo sheath. the medical doctor was unaware of how the pump malposition occurred. there was additionally access site bleeding noted by oozing which was treated with three units of packed red blood cells. it was reported that the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6781 | 1220648-2025-31003 | 22715187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22715187&pc=OZD | 07/19/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/20/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536175 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an impella cp was being prepped and the purge cassette was not properly seated in the automated impella controller and was damaged while closing the door prior to insertion of the device. the doctor requested that a new pump kit be opened. the damaged purge cassette was replaced and the doctor recognized that there was not an issue with the original pump. the replacement pump was used successfully. the original pump was opened but never connected or implanted. there was no patient harm and the device was discarded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
6782 | 1220648-2025-31004 | 22715633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22715633&pc=OZD | 07/14/2025 | 08/19/2025 |  | 08/05/2025 |  |  | 07/23/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669813 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Hemolysis | Liver Damage/Dysfunction | Renal Failure | Sepsis | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received an impact study regarding a patient that endured atrial fibrillation with a "possible" relationship to the impella device used: ¿patient has multiple episodes of new onset post operative af treated with amiodarone/heparin and temporary pacing wires¿. additionally, the complaints team received an impact study regarding a patient that endured staph epidermis with meca gene with a "possible" relationship to the impella procedure used: "this patient has been very sick w/ low cardiac index and inotropic support w impella 5.5 post coronary artery bypass graft and mitral valve replacement." | The complaints team received details from an impact study regarding a patient who endured multiple adverse event. it was reported that the patient endured acute kidney injury with a "possible" relationship to the impella device and the impella procedure used: "fluid overload; patient on diuretics." the patient endured hyperbilirubinemia with a "possible" relationship to the impella device and the impella procedure used: "hyperbilirubinemia due to hemolysis "ruqus with sludge, and small stone but no signs of cholecystitis or biliary obstruction". the patient endured ventricular ectopy with a "possible" relationship to the impella device and the impella procedure used: "frequent ventricular ectopy requiring medications (amio, lido + pressors). impella position was revised multiple times to minimize the intraventricular stimuli." the patient endured transaminitis with a "possible" relationship to the impella device and the impella procedure used: "pt has elevated liver enzymes resulting in dx transaminitis". |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock . section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Corrections to the initial mfr report: b5 updated adverse events. h6 clinical codes 2041, 1886, 1954, 1721 and impact codes 4644 added.
6783 | 1220648-2025-31005 | 22715680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22715680&pc=OZD | 07/21/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671970 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female | 112 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had episodes of ventricular tachycardia and amio was started. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6784 | 1220648-2025-31006 | 22715787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22715787&pc=OZD | 07/22/2025 | 09/11/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612966 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 131 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted. the pump could not pass due to heavy calcification in the patient's ascending aorta creating stenosis. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since attempted impella 5.5 implant and due to the patient's heavy calcification in the ascending aorta creating stenosis, the pump could not pass. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6785 | 1220648-2025-31007 | 22715878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22715878&pc=OZD | 07/21/2025 | 09/12/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685633 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery. it was noted that resistance was met as a result of the small vessel size of the patient's iliac artery. the implant was aborted and an intra-aortic balloon pump was placed for ongoing support. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because pump was unable to be delivered due to small iliac artery. d6b explant date was erroneously entered on the initial manufacturer device report number 1220648-2025-31007.
6786 | 1220648-2025-31008 | 22715934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22715934&pc=OZD | 12/24/2024 | 08/05/2025 |  | 08/05/2025 |  |  | 12/24/2024 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526834 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, the pump generated low placement signal alarms. however, proper positioning was confirmed, the alarms were disabled, and the pump continued to provide circulatory support. no harm to the patient was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned for investigation. data logs were reviewed and the alarm did appear. since product was not returned, the root cause of the optical signal issue could not be determined.
6787 | 1220648-2025-31009 | 22716044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716044&pc=OZD | 08/19/2023 | 08/05/2025 |  | 08/05/2025 |  |  | 08/19/2023 | 08/05/2025 |  |  | 05/31/2025 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357210 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump in use to support a patient undergoing high-risk percutaneous coronary intervention. during support, a placement signal unknown alarm was present. the alarm was silenced by disabling the placement signal audio. the pump remained in use to provide continued support. no harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no pump was returned for analysis. data logs indicate console was the issue (please see sister record "23-17789-2" for automated impella controller (aic)). the root cause of the optical signal issue could not be determined, as no problem was identified.
6788 | 1220648-2025-31010 | 22716196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716196&pc=OZD | 01/30/2024 | 08/05/2025 |  | 08/05/2025 |  |  | 01/30/2024 | 08/05/2025 |  |  | 11/30/2024 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023271971 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a 5.0 pump to support a patient. during support, the pump experienced unreliable placement signals. pump position was confirmed via echocardiography, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the complaint pump and data logs were not returned for investigation. the cause of the placement signal issue was not determined since product and data logs were not returned.
6789 | 1220648-2025-31011 | 22716433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716433&pc=OZD | 03/24/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 03/24/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594249 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted. during support, the pump intermittently triggered 'impella in aorta' alarms. echocardiographic imaging confirmed correct pump positioning, and support continued uninterrupted. there was no patient harm, and the pump continued to provide circulatory support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs were reviewed and the alarms were confirmed. product was not returned for investigation. since limited clinical details were provided that would suggest a source of optical signal noise, data logs did not reveal a definitive cause, and product was not returned for evaluation, the root cause of the optical signal issue could not be determined.
6790 | 1220648-2025-31012 | 22716487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716487&pc=OZD | 02/21/2023 | 08/05/2025 |  | 08/05/2025 |  |  | 02/21/2023 | 08/05/2025 |  |  | 10/31/2024 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290514 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the implanted impella cp pump lost its placement signal following a 'placement signal not reliable' alarm during patient support. device positioning was confirmed to be correct via echocardiography, and support was continued with the implanted pump. no harm to the patient was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned for investigation. the data log shows the placement signal. the cause of the optical signal issue was most likely a damaged optical fiber line.
6791 | 1220648-2025-31013 | 22716491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716491&pc=OZD | 07/22/2025 | 09/18/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616493 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Device-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp, the physician initially wanted to use a single access technique. however, they accidentally punctured the proximal part of the sheath with the needle and had to removed the peel away sheath and place the repositioning sheath with good hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of single access issue and introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. single access issue: physician accidentally punctured the proximal part of the sheath with the needle and had to removed the peel away sheath and place the repositioning sheath with good hemostasis. the cause of the issue was use-related as evidenced by accidental puncture of sheath introducer damage - mechanical: physician accidentally punctured the proximal part of the sheath with the needle and had to removed the peel away sheath and place the repositioning sheath with good hemostasis. the cause of the issue was use-related as evidenced by accidental puncture of sheath.
6792 | 1220648-2025-31014 | 22716620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716620&pc=OZD | 07/22/2025 | 09/15/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664689 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 59 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed hemolysis. the pump was repositioned to resolve the issue, and later the impella cp was escalated to an impella 5.5. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed and there were no abnormalities that could be definitively connected to the complication. there were no suction alarms, no positioning alarms, no motor current spikes, and pump flows were in specification for p-9 throughout support. during the first couple hours of support, pump flows were on the lower end of specification, and there was some variability in the placement signal (ps) followed by a period of little ps pulsatility. however, all three of these things resolved by 18:00 on 21/jul, suggesting that it was not related to the hemolysis reported 14 hours later. product was not returned for investigation. since several troubleshooting methods were reported in clinical details, there were no abnormalities in data logs, and product wasn't returned for investigation, the cause of the hemolysis could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: hemolysis: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
6793 | 1220648-2025-31015 | 22716628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716628&pc=PYX | 10/25/2024 | 10/02/2025 |  | 08/05/2025 |  |  | 10/25/2024 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025507452 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an rp pump was placed to support a patient. approximately six hours after implantation, the placement signal of the impella rp flex began to drift. shortly thereafter, both the pulmonary artery placement signal and flow calculations disappeared from the display. pump positioning was verified via echocardiography, and support from the implanted pump remained unaffected. no patient harm was reported. | Additional information care was eventually withdrawn, and the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the returned product was inspected and there were no physical abnormalities. the pump passed electrical testing. the data logs confirm placement signal not reliable alarms and show pulmonary artery placement signal drifting downwards before going out of range. the root cause of the placement signal issue was most likely sensor drift. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. please refer to dhr investigation checklist for indicated " 4 pumps in batch #1811597 with sn: (b)(6) were scrapped for motor issues, purge flow issues, snr, and sensor fail issues respectively. " this review revealed that this product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. | Corrections that were made from the initial manufacturer device report number 1220648-2025-31015. b1 adverse event updated. b2 death and death date updated. b5 updated to include death description. h6 updated to include death coding.
6794 | 1220648-2025-31016 | 22716693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716693&pc=OZD | 10/12/2023 | 10/27/2025 |  | 08/05/2025 |  |  | 10/12/2023 | 08/05/2025 |  |  | 05/31/2025 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024340914 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the use of a cp pump to support a patient. during support, the pump was experiencing abnormal placement signals. despite the alarm, the patient remained hemodynamically stable and adequately supported. it was believed that the alarm may have been inaccurate due to a faulty or non-functioning placement signal. after the procedure was completed, the impella pump was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the impella 5.5 pump was returned. the pump was returned in a cut open state at the red impella plug connector. the optical bench 5s parameters or uson test could not be verified since pump was returned as a cut open device. data logs were reviewed, and the placement signal drift trend. the cause of the optical placement signal issue was most likely due to sensor wear with placement signal drift observed per data log review and field investigation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. only thing of note from the dhr checklist was "all 57 samples of batch # 1715047 were accepted per subcomponent inspection checks." | G3 ¿date received by manufacturer¿ corrected. | All pertinent information available to abiomed has been submitted at this time. d6a: corrected implant date. d6b: corrected explant date.
6795 | 1220648-2025-31017 | 22716709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716709&pc=OZD | 07/22/2025 | 09/18/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635443 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 115 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Tachycardia | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 exhibited symptoms of stroke. a computed tomography showed hypodensity in the post cerebellum. additionally, the patient had ectopy, ventricular tachycardia, anuria, and short activated clotting time requiring transfusion of two units of red blood cells. the patient was successfully weaned from the impella 5.5 and survived. | [5 empty] | The investigation of stroke/cva and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6796 | 1220648-2025-31018 | 22716727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716727&pc=OZD | 07/22/2025 | 11/21/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232826 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) had a controller failure with yellow flashing. the alarm would appear and self resolve. the console was replaced to resolve the issue. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: the case related to the reported complaint was initiated on july 15, 2025, with the pump sn: (b)(6). during support on the complaint date, july 22, 2025, the console displayed ¿controller error (pmd voltage out-of-spec) ¿ alarm,¿ event #503, due to the pmd vss being out of specifications. this issue was automatically resolved but reoccurred immediately throughout the rest of the case (imc_logs_(b)(6).pdf), confirming the reported failure mode. the pump was then moved to the second console, (b)(6), and support continued for 8 days and 15 hours, during which there was no ¿controller error¿ alarm. device analysis: this console was tested after arriving at fs and reproduced the reported failure mode. while running the console with the test pump at p-level p9, after 5 hours, controller error - event #503 occurred due to the pmd vss being out of specifications, and it was on and off until the console was shut down. the vss signal on the impellatronic board corresponds to the 3.3vanalog_dsp2. upon measuring the test point tp106.2, it was observed that the 3.3v was fluctuating, whereas in the known-good console, it remained stable. root cause: the root cause of the controller error was a malfunction of the impellatronic board. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
6797 | 1220648-2025-31019 | 22716926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716926&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680069 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male | 123 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon arrival in the intensive care unit, an echocardiogram revealed what appeared to be small air bubbles in the left ventricle during impella cp support. there were no noted air bubbles in the right ventricle or the purge line. the purge line was de-aired per protocol and a later echocardiogram revealed no further evidence of air bubbles in the left ventricle. it was additionally reported that the patient developed hematuria, for which the pump was repositioned, volume was added, and the p-level was turned down. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6798 | 1220648-2025-31020 | 22716940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22716940&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  | 01/31/2026 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528678 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male | 54 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that after implantation of an impella cp the patient experienced ventricular fibrillation. defibrillation was delivered to restore the patient to normal sinus rhythm. later, the patient experienced non-sustained ventricular tachycardia with no hemodynamic compromise. iv amiodarone was given. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
6799 | 1220648-2025-31021 | 22717037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22717037&pc=OZD | 04/16/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 04/16/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612408 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the placement signal was completely lost. however, motor current remained stable. alter the procedure was completed, the impella pump was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. due to lack of clinical details provided and no product or data logs returned, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. this review revealed that this product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.
6800 | 1220648-2025-31022 | 22717173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22717173&pc=OZD | 07/22/2025 | 09/23/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610514 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER SERIAL NUMBER (B)(6) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had a "disconnected" alarm on the supporting automated impella controller. the pump was replaced to resolve the error. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The cause of the optical signal issue was not determined since product was not returned and insufficient clinical details were provided.
6801 | 1220648-2025-31023 | 22717389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22717389&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656247 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed an access site bleed. multiple blood products were transfused. the blue suture hub was resutured in the operating room and dressing angle was matched. the dressing was completely soaked when checking into the intensive care unit. the dressing was removed and pulsatile bleeding was noted. pressure was held. a femstop was applied. coagulations and labs were sent. the purge was changed from five percent dextrose in water to bicarbonate at bedside. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
6802 | 1220648-2025-31024 | 22717613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22717613&pc=OZD | 07/21/2025 | 09/29/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668860 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support and there was no impella flow. the impella was exchanged and there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. data was not returned for review. product was returned in damaged condition (catheter was cut off). visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. low flow: the root cause of low flow was biomaterial ingestion. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacture date added.
6803 | 1220648-2025-31025 | 22717653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22717653&pc=OZD | 07/21/2025 | 09/15/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688409 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 84 KG |  |  | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | High Blood Pressure/ Hypertension | Pleural Effusion | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had ectopy after the impella placement. the pump was pulled back and was torqued. the ectopy continued but at lower frequency. additionally, the patient had pleural effusions. the patient had heparin drip prior to the effusions. blood products were given. moreover, the patient had hypotension coupled with one suction event. support continued. | [5 empty] | The investigation of vf/vt/af/at, positioning issues, vascular damage - peripheral vessel, and low pump flow has been completed. the impella device was not returned for evaluation. vt/vf/at/af: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. positioning issues: clinical reported pump was repositioned from septal wall so inlet was by midventricular wall and was pulled back 1cm and torqued. it is unknown when the pump became mispositioned and what led to it. the root cause of the positioning issue was not determined as limited clinical information was provided. vascular damage - peripheral vessel: the root cause of the vascular damage was not determined as no product was returned and limited clinical information was provided. low pump flow: the data logs do not show any motor current spikes. the logs confirm suction alarms and 'impella position unknown' alarms. placement signal was also noisy throughout support. the root cause of the low pump flow was most likely patient condition related as evidenced by clinical details mentioning issue occurrence due to coughing and ambulation and resolution after volume and stopped ambulation. d6b explant date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6804 | 1220648-2025-31026 | 22717660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22717660&pc=OZD | 07/23/2025 | 09/25/2025 |  | 08/05/2025 |  |  | 07/23/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693622 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 pump placed to support a patient with cardiogenic shock and cardiomyopathy. the pump had a purge leak on day 8 but the pump remains on for support due to the decisions still going for either a left ventricular assist device or transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". ¿clean the connector cable with 70% isopropyl alcohol.¿. | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. due to a lack of data logs returned and no product, the root cause of the purge leak - pump cannot be determined. d6b explant date added.
6805 | 1220648-2025-31027 | 22718751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22718751&pc=OZD | 07/21/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684968 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support an acute myocardial infarction/cardiogenic shock patient. during support, the nurse noted hematuria and the pump was repositioned. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6806 | 1220648-2025-31028 | 22718779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22718779&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667476 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 83 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that there had been pump damage to an implanted impella cp. the sleeve broke at the distal and tegaderm was used to keep it secure. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6807 | 1220648-2025-45717 | 22718874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22718874&pc=OZD | 06/11/2025 |  |  | 08/05/2025 |  |  | 07/24/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655909 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | Native Hawaiian Or Other Pacific Islander | Injury | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient (unknown ethnicity) experienced low pump flow and required device replacement. the patient was on a non-monitored floor preparing for discharge and had a syncopal episode. the patient hit her head (developed a subarachnoid hemorrhage) and arrested. multiple rounds of cardiopulmonary resuscitation took place. complete heart block was found, and a transvenous pacer and swan-ganz catheter were placed. the impella cp device was then inserted for mechanical circulatory support and low pump flow was encountered. suction alarm continued to sound regardless of the device position. consequently, the impella cp device was explanted and replaced with a new impella cp device. coronaries arteries were assessed and noted to be clear. the patient was transferred to the intensive care unit for further monitoring and care. the patient was reported to be in stable condition following the event. | [5 empty] | Medical safety reviewed the event and noted the following regarding the patient's outcome: the event describes low pump flow which resulted in replacement of the device. this malfunction is attributable to the first impella cp. the second impella cp was placed and the patient developed limb ischemia determined by the registry to be probably related to the second impella cp, which was removed due to concerns for arterial obstruction. and post-removal the patient had improved perfusion. later this day the patient expired. the device was removed and the limb ischemia event improved. the patient subsequently expired not on impella support. in this case, the patient's underlying condition contributed most to an outcome of death. however, the second impella cp cannot be dissociated from the outcome of the patient. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the 2nd impella cp, which is associated with the demise outcome. | The impella cp device was not returned and therefore an evaluation of the device was not possible. device history review was completed, and the pump passed all post sterile inspection checks. the root cause of the low flow was unable to be determined as limited clinical details were provided, and no product was returned for analysis. | Subsequent information noted the patient experienced limb ischemia on the replacement impella cp device and eventually expired; care was withdrawn. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6808 | 1220648-2025-45716 | 22718900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22718900&pc=OZD | 06/12/2025 |  |  | 08/05/2025 |  |  | 07/24/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659953 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A notification via abiomed's long-term outcome and quality indicator (loqi) impella registry regarding a patient (unknown ethnicity) experienced limb ischemia on the impella replacement device leading to explant of the device. the limb ischemia had a probable relationship to the replacement impella cp device. further information in addition to the study notification noted the patient subsequently expired. the patient was on a non-monitored floor preparing for discharge and had a syncopal episode. the patient hit her head (developed a subarachnoid hemorrhage) and arrested. multiple rounds of cardiopulmonary resuscitation took place. complete heart block was found, and a transvenous pacer and swan-ganz catheter were placed. the first impella cp device was inserted, and low pump flow was encountered. consequently, this initially implanted impella cp device was explanted and the replacement impella cp device was implanted. coronaries arteries were assessed and noted to be clear. the patient was transferred to the intensive care unit for further monitoring and care. patient updates noted there was troubleshooting for thrombosed swan catheter. subarachnoid hemorrhage was being closely monitor pending initiation of heparin. also, this day in the morning, the nurse noticed the patient's right lower extremity was cool, and the right foot was without pedal/post tibial pulses. the physician was notified. the right lower extremity became duskier and colder in the afternoon, and the physician was then called to the bedside. consequently, the replacement impella was removed due to concerns for arterial obstruction. post-removal the patient had improved perfusion to right lower extremity with 1+ doppler pedal pulse. per the registry, the patient recovered with no sequelae. patient updates further noted the patient was still 100% on fio2. lactate was clear, pressor/inotrope requirements were still high. arterial blood gas, chest x-ray and continuous renal replacement therapy were pending. however, later this same day care was withdrawn, and the patient expired. | [5 empty] | Although, the patient experienced serious adverse event with swan catheter and care was withdrawn, there is not enough evidence to exclude the replacement impella cp device from the patient's outcome given the impella-related limb ischemia event. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation). | Medical safety reviewed the event and noted the following regarding the patient's outcome: the event describes low pump flow which resulted in replacement of the device. this malfunction is attributable to the first impella cp. the second impella cp was placed and the patient developed limb ischemia determined by the registry to be probably related to the second impella cp, which was removed due to concerns for arterial obstruction. and post-removal the patient had improved perfusion. later this day the patient expired. the device was removed and the limb ischemia event improved. the patient subsequently expired not on impella support. in this case, the patient's underlying condition contributed most to an outcome of death. however, the second impella cp cannot be dissociated from the outcome of the patient. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the 1st impella cp. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the limb ischemia in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
6809 | 1220648-2025-31029 | 22718913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22718913&pc=OZD | 07/22/2025 | 09/18/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667331 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 64 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Pericardial Effusion | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a large pericardial effusion. the fluid was drained at the bedside to resolve the issue. the pump was malpositioned with the pigtail wrapped in the papillary. repositioning of the pump was unsuccessful. the patient became agitated and pulled the pump across the atrioventricular, breaking the sterile sleeve. the physician pushed the impella back into position and repaired the sleeve with tegaderm. bleeding was noted at the sheath site. entry angle matching was performed and three units of packed red blood cells were transfused. the patient's foot became ischemic and improved without intervention. upon explantation of the pump the original perclose sutures failed and had to be replaced, but the replacement sutures closed off the superficial femoral artery. the artery was surgically repaired. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma." section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of access site bleeding, positioning issues, product damage (sleeve damage), limb ischemia, and vascular damage - great vessel has been completed. the impella device was not returned for evaluation. access site bleeding: due to lack of clinical details, the root cause of the access site bleeding cannot be determined. positioning issue: due to lack of clinical information provided on how the pump became mispositioned, the root cause of the first positioning issue cannot be determined. positioning issue: about an hour after the first positioning issue, the patient became agitated and wrapped his foot around the catheter and pulled it across the av. based on the clinical details provided, the root cause of the positioning issue is most likely due to patient movement as the patient pulled the pump out. product damage (sleeve damage): when the physician attempted to push the pump back in, it was noted that the sleeve became torn during when the pump was pulled out. based on the clinical details provided, the root cause of the product damage ( sleeve damage) is most likely due to use as it became ripped when the pump was pulled out of position. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. vascular damage - great vessel: due to lack of clinical details, the root cause of the vascular damage - great vessel cannot be determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-31029.
6810 | 1220648-2025-31030 | 22718929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22718929&pc=OZD | 07/21/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702836 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support had right lower extremity mottled but present doppler pulses in the dorsalis pedis. the next day the right foot was swollen, dusky and purplish but pulses via doppler. it is unknown what caused the mottled right lower extremity. the team believed it was from multiple pressors, which they were unable to wean. the patient remained on support for an additional two days. weaning was successful. the device was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6811 | 1220648-2025-31032 | 22719154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22719154&pc=PYX | 03/04/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 03/04/2025 | 08/05/2025 |  |  | 01/31/2026 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025477198 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an rp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, 'placement signal unreliable' alarms occurred. the placement signal never resolved to display positive pressure values. pump position was confirmed via echocardiography, and support continued. no known patient harm or injury occurred. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned for evaluation. after review of the returned data logs, it is indicated that the cause of the issue is a result of the sensor. clinical details noted that a alarm was noted in the field with pressures reading extremely negative and clinical staff believed the sensor was unreliable. the root cause of the placement signal issue was most likely due to sensor drift determined from returned data log analysis. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. 7 parts from lot #2024320712 (174 total) failed visual inspection due to scratches and were scrapped. there were 6 other product malfunction complaints, (b)(4), in subcomponent lot #2024320712 (174 total) related to this failure mode.
6812 | 1220648-2025-31031 | 22719155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22719155&pc=OZD | 07/22/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485773 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 116 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia and nose bleeds. the patient was defibrillated to obtain return of spontaneous circulation and the bleeds were cauterized. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6813 | 1220648-2025-31033 | 22719185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22719185&pc=OZD | 07/21/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653867 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Septic Shock | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support lost pedal pulse and a distal perfusion catheter was placed. the patient also was in septic shock. the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6814 | 1220648-2025-31034 | 22719454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22719454&pc=OZD | 07/21/2025 | 09/12/2025 |  | 08/05/2025 |  |  | 07/21/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668963 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella 5.5 implant and the pump would not advance passed innominate artery. the implant was aborted and an impella cp was placed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because 5.5 was unable to be pass through the vasculature due to small vessel size.
6815 | 1220648-2025-31035 | 22719692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22719692&pc=OZD | 07/22/2025 | 09/11/2025 |  | 08/05/2025 |  |  | 07/22/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686411 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed a hematoma at the impella access site. pressure was applied, heparin was withheld, and the dressing was re-dressed with higher angle matching to resolved the issue. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related since the dressing was re-dressed with a higher angle matching with resolution to the hematoma.
6816 | 1220648-2025-31036 | 22719845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22719845&pc=OZD | 07/23/2025 | 09/24/2025 |  | 08/05/2025 |  |  | 07/23/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696228 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 purge side arm was leaking. the purge flow and purge pressure were in range. the impella 5.5 was weaned and explanted that same day. the patients outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. no data logs were provided for analysis. the returned pump was inspected and there was no damage observed or other physical abnormalities. the pump was purged for over 48 hours and no leak was observed and purge values were within expected range. when changing the purge tubing to the sidearm, it is a wet to wet connection. this can lead to fluid into the sidearm. the root cause of the purge leak - pump is most likely due to use as the issue was not able to be reproduced in the lab. d9 device returned to manufacturer date added.
6817 | 1220648-2025-31037 | 22720130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22720130&pc=PYX | 04/06/2025 | 08/05/2025 |  | 08/05/2025 |  |  | 04/06/2025 | 08/05/2025 |  |  |  |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025599099 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an rp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump generated multiple 'placement signal unreliable' alarms. however, there were no spikes in motor current, and pump functionality remained unaffected. the physician was aware of the issue and was confident in continuing support without relying on the placement signal. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the returned pump was inspected and there were no physical abnormalities. no abnormalities observed on sensor face and pump passed electrical testing. the data logs from the field show stable placement signal and pump flow readings for about 30 hours of support after which placement signal starts drifting down and pump flow starts drifting up before placement signal goes out of range and pump flow calculation is lost. the root cause of the placement signal issue was determined to be sensor drift as evidenced by the identified log pattern and the sensor drift reproduction on the returned product. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. 5 pumps from lot#: 1878292 were scrapped (sn: (b)(6) for sensor fail, calculation, osp, and purge flow issues. all pumps in the batch were reworked for motor cable length issues. also noted: qn: 210034644 revealed 57 pumps in this subcomponent lot had a labelling issue which was mentioned to not affect batch because label was not required, and 25-40707-1 was another product malfunction complaint in this subcomponent lot.
6818 | 1220648-2025-31038 | 22720616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22720616&pc=OZD | 03/02/2024 | 08/05/2025 |  | 08/05/2025 |  |  | 03/02/2024 | 08/05/2025 |  |  | 11/30/2025 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428667 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted. during support, the pump lost placement signal, but remained on for support. no harm to the patient was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned for investigation. the data logs were returned. there was a rapid drop in placement signal to negative followed. this is characteristic of a durability failure pattern for the optical sensor. per the logs investigation, the root cause of the optical signal issue was sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.
6819 | 1220648-2025-31039 | 22720809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22720809&pc=OZD | 03/15/2024 | 08/05/2025 |  | 08/05/2025 |  |  | 03/15/2024 | 08/05/2025 |  |  | 07/31/2025 |  | 08/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442778 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during use, placement signal alarms were triggered, including a 'placement signal not reliable' alarm. despite this, the account continued to monitor the patient¿s hemodynamics. there was no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no pump was returned. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint #(b)(4) for an investigation into the console.
6820 | 1220648-2025-30959 | 22699525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22699525&pc=OZD | 07/22/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587366 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 54 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a non-st-segment elevation myocardial infarction patient was in for a coronary artery bypass graft procedure when there was a left ventricle (lv) perforation. the lv perforated during impella 5.5 repositioning. the perforation was repaired and the patient came off the bypass pump and was placed on venoarterial extracorporeal membrane oxygenation. the patient outcome was reported as stable. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
6821 | 1220648-2025-45712 | 22699543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22699543&pc=OZD | 07/21/2025 |  |  | 08/04/2025 |  |  | 07/21/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026697028 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a gastrointestinal bleed during impella 5.5 support. heparin was put on hold in response to the noted condition. plans were made to continue to monitor the bleed as impella support continued uninterrupted. | [5 empty] | The root cause of gastrointestinal bleed cannot be determined since the product was not returned and insufficient clinical details were provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6822 | 1220648-2025-30960 | 22699929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22699929&pc=OZD | 07/22/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632702 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 118 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following initiation of impella cp support, flows began to ramp up in autom mode and plateaued at 1.5 - 1.8 liters per minute. the pump was switched from auto to p-level with suction above p-5. the volume and positioning were verified, but the issue persisted. the impella cp was repositioned by recommendation. as a result of the ongoing suction, the decision was made to replace the pump. there was no patient injury reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6823 | 1220648-2025-30962 | 22700606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22700606&pc=OZD | 07/22/2025 | 10/30/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663929 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 63 KG | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Arrhythmia | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia that was medically managed. after transferring the patient from one automate impella controller to another, the pump failed to restart. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | Pump not detected: the returned pump was evaluated and the impella defective alarm was reproduced when the pump was plugged into a console. the red handle was opened and blood was observed in the red handle. a hole was located in the catheter at the 30 cm mark along with scratches which could have been from improper suturing technique. the cause of the pump not detected was blood ingress into the red handle due to a hole in the catheter resulting from improper use. vt/vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. product damage (hole in catheter): failure mode product damage (hole in catheter) was added since a hole was observed in the catheter on the returned product relating to the pump not detected issue. the cause of the hole in the catheter was use related. d9 device returned to manufacturer date added.
6824 | 1220648-2025-30963 | 22700915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22700915&pc=OZD | 07/22/2025 | 09/18/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 11849905 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 63 KG | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death |  |  |  |  | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after transferring a patient who had an impella cp implanted from one supporting automated impella controller to a second, the impella cp would not restart. two additional automated impella controllers were tried without success. the patient expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - pump stop has been completed. the impella device was returned for evaluation. based on the data and device analysis the root cause of the aic-pump stop, i.e., the impella defective alarm, was not related to any issue with the aic, and no problem was found with the aic. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30963. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-30963. h8 usage of device was erroneously selected on the initial manufacturer device report number 1220648-2025-30963.
6825 | 1220648-2025-30964 | 22701283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22701283&pc=OZD | 07/22/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663911 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 56 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was prepped and inserted without complication for a primary percutaneous coronary intervention procedure. after obtaining access to the coronaries, the patient began to get restless and required urgent anesthesia/intubation. during this time the patient reported leg pain. the medical doctor was unable to doppler pedal pulses at this time. the medical doctor attempted removing the peel-away and placed a repo-sheath, but still no flow distally. the provider elected to explant the impella cp. the patients outcome was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6826 | 1220648-2025-30965 | 22701475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22701475&pc=OZD | 07/21/2025 | 09/05/2025 |  | 08/04/2025 |  |  | 07/21/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580409 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 149 KG |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 lost placement signal. an impella rp flex was also implanted but after 47 hours patient care was withdrawn and patient expired. | The medical safety office (mso) has reviewed the death and determined the ectopy though not documented as clinically significant is reportable. |  |  |  |  | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. the root cause of the optical signal issue was most likely patient condition related as evidenced by alarm triggering during intra-aortic balloon pump (iabp) inflations with self resolution and imaging confirming good position with no data log abnormalities observed. the medical safety office (mso) has reviewed the death and determined the ectopy though not documented as clinically significant is reportable. the investigation of ectopy is pending should additional information become available a supplemental report will be submitted the medical safety office (mso) has reviewed the death and determined the ectopy though not documented as clinically significant is reportable. b2 date of death was erroneously entered on the initial manufacturer device report number 1220648-2025-30965. b5 mso statement h6 health effect - .clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-30965.
6827 | 1220648-2025-30966 | 22701638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22701638&pc=OZD | 02/18/2023 | 08/04/2025 |  | 08/04/2025 |  |  | 02/18/2023 | 08/04/2025 |  |  | 11/30/2024 |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023284266 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced no placement signal waveform upon starting the impella, even though the motor current and pump flow were within normal limits. as a result, the decision was made to explant and replace the device. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs confirmed placement signal not reliable alarm. the parameters were characteristic of a broken sensor. the placement signal not reliable alarm reproduced when the returned pump was connected to a lab console. the pump passed optical continuity testing. upon further inspection, the sensor face was found to be damaged. therefore, it was determined that the cause of the optical signal issues was a damaged sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.
6828 | 1220648-2025-30967 | 22701650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22701650&pc=OZD | 07/21/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 07/21/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686547 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the companion sheath was caught on the inside rim of the 14fr peel-away. the sheath was damaged and was unable to pass through it. a new companion sheath was opened. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6829 | 1220648-2025-30968 | 22701961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22701961&pc=OZD | 07/22/2025 | 10/17/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617100 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure which was treated by adding lytics to the purge. no patient harm was reported. | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure which was treated by adding lytics to the purge. no patient harm was reported. |  |  |  |  | The investigation of high purge pressure has been completed. the pump was returned and was investigated. it was seen that there was large amounts of biomaterial around the impeller cage and inside the cannula wrapped around the impeller blades. the pump was purged and run; no high purge pressure alarms were reproduced. a catheter window was cut near the motor and no blood in purge line was seen. no purge line damage or kinks were seen. no data logs were returned for analysis. high purge pressure: based on the clinical review, it was seen that the purge pressure high alarm was observed during the case, due to which the impella pump was replaced. the root cause for the high purge pressure was biomaterial around the impeller. it was unable to be determined whether it was biomaterial ingestion or biomaterial buildup due to no returned data logs. d9 device returned to manufacturer date added | A4 and a5, ethnicity, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6830 | 1220648-2025-30969 | 22702005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22702005&pc=OZD | 02/23/2023 | 08/04/2025 |  | 08/04/2025 |  |  | 02/23/2023 | 08/04/2025 |  |  | 09/30/2024 |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282945 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump triggered "placement signal low" and false "impella in aorta" alarms during patient support. the pump remained in use for continued support. no harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the product was not returned. the log was reviewed and parameters were stable throughout the run. pump position in aorta and placement signal low alarms were confirmed. the cause of the false position in aorta alarm was most likely biomaterial. optical sensor was functional.
6831 | 1220648-2025-30970 | 22702011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22702011&pc=OZD | 07/22/2025 | 09/18/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670899 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced positioning issues and the patient's urine was red. the pump was repositioned to resolve the issue. no patient harm was reported. | [5 empty] | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. the data log showed several impella position alarms in the aorta/ventricle after 30 minutes of use, with some pump flow saturation lasting for about one hour. the placement signal waveforms were damped during the positioning issue, which was later resolved. according to clinical notes, the tuohy needle was not locked. the patient immediately relaxed after repositioning. the pump was also repositioned the following day, 22-july, and hemolysis began to resolve on the next day. the cause of the hemolysis issue was most likely improper positioning of the device based on data log review and provided clinical details. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30970. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-30970. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
6832 | 1220648-2025-30971 | 22702178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22702178&pc=OZD | 03/03/2023 | 08/04/2025 |  | 08/04/2025 |  |  | 03/03/2023 | 08/04/2025 |  |  | 08/31/2024 |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262796 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention (hrpci). the pump triggered an "unreliable placement signal" alarm, which did not correlate with invasive monitoring. the pump remained in use for continued support. no harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was returned. the root cause of the optical signal issue is most likely due to patient condition. when conducting the product evaluation in the lab, the optical signal issue did not reproduce and the placement signal was in normal range when the pump was run. the root cause of the placement signal issue is most likely due to patient condition.
6833 | 1220648-2025-30972 | 22702230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22702230&pc=OZD | 07/22/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668880 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male | 80 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. due to hemolysis the pump was explanted after 68 hours of support. the hemolysis was noted to be caused by the impella or extra corperal membrane oxygenation (ecmo); however, the pump was explanted and ecmo remained. the patient outcome was reported as stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6834 | 1220648-2025-30973 | 22703303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22703303&pc=OZD | 04/07/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 04/07/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630036 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella 5.5 pump intermittently experienced a 'placement signal unreliable' alarm during patient support. device positioning was verified via echocardiography, and support continued uninterrupted. there was no patient harm reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no pump was returned for analysis. data logs indicate a sudden loss of placement signal accompanied by alarms. the root cause of the optical signal issue could not be determined, as no problem was identified.
6835 | 1220648-2025-30974 | 22703522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22703522&pc=OZD | 07/22/2025 | 09/22/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577939 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed in a st-segment elevation myocardial infarction patient. the patients condition prompted transfer to a tertiary center via medical flight. during the automated impella controller (aic) to aic transfer, the pump failed to be detected. the pump malfunction/stop prompted the pump to be explanted after eight hours and the patient was medically managed while decisions were made for an impella 5.5 or an impella rp support placement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of pump did not start has been completed. impella cp returned for evaluation. no abnormalities observed upon visual inspection. pump electrical was all within spec. pump was connected to console and was able to run at p-6 without issue for approximately 2 hours. no alarms were triggered. no data logs were returned for evaluation. clinical details noted that during console to console transfer after patient transferred hospitals, when pump was connected to new console, an "impella defective" alarms was triggered. troubleshooting was attempted on 3 consoles without success. decision was made to explant pump. the cause of the pump did not start could not be definitively determined as no logs were returned for evaluation and issue did not reproduce with returned pump.
6836 | 1220648-2025-30975 | 22703531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22703531&pc=OZD | 04/15/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 04/15/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604512 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, a persistent 'placement signal not reliable' alarm occurred. despite comprehensive troubleshooting and successful maintenance of impella hemodynamic support, the absence of a reliable placement signal continued. given that the issue arose during a critical stage of the procedure, the physician elected to continue and complete the procedure. the pump remained operational until weaning and explant. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs were not recovered. pump was returned and inspected. the cause of the optical signal issue was a damaged sensor due to the reported placement signal failure. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.
6837 | 1220648-2025-30976 | 22704084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22704084&pc=PYX | 04/10/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 04/10/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025588064 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during patient support, placement signal issues occurred. the pump position was verified, and support continued without interruption. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned for investigation. the data log was reviewed and suggested that there was an optical sensor issue. all optical sensor parameter trends were indicating a sensor break trend. the root cause of the optical signal issue was most likely damaged sensor, based on data log investigation.
6838 | 1220648-2025-30977 | 22704377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22704377&pc=OZD | 04/21/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 04/21/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655484 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was using a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the placement signal was intermittent; however, the pump remained operational throughout the procedure and was explanted as planned after just under two hours. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the data logs and pump were returned. the pump was returned and there were no physical abnormalities. the parameters were within specification. based on the data logs and returned product, the root cause of the optical signal issue is most likely due to an offset error, however the conditions present to cause that offset error are unknown.
6839 | 1220648-2025-30978 | 22704475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22704475&pc=OZD | 07/22/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673840 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 127 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that upon explant of an impella cp the patient had bilaterally ischemic limbs. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
6840 | 1220648-2025-30979 | 22704795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22704795&pc=OZD | 07/21/2025 | 09/12/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666038 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male | 68 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 in left axillary artery. the patient was on brillinta at time of the surgical implant. the patient also had a gastrointestinal bleed over the past couple of days. a jp drain was placed intra procedure at the impella 5.5 insertion site. post procedure. the orogastric tube had significant blood output and the jp drain had filled up multiple times. estimated blood loss was around two liters. mtp protocol was issued overnight and eight units of plasma, six units of packed red blood cells, four units of platelet, and two units of cryoprecipitate were given to the patient to stabilize. later, the pump was explanted and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. vascular damage - peripheral vessel: vascular damage failure mode was added for continuous gi bleed. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
6841 | 1220648-2025-30980 | 22706041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22706041&pc=OZD | 04/21/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 04/21/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634367 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during patient support, the impella cp experienced an active placement alarm. the audio and placement monitoring are off. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned. data log review showed placement signal railing. the cause of the optical signal issue was not determined since product was not returned, insufficient clinical details details, and an inconclusive data log review.
6842 | 1220648-2025-30981 | 22706224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22706224&pc=OZD | 10/15/2023 | 08/04/2025 |  | 08/04/2025 |  |  | 07/16/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | Other | [5 empty] | 76 YR |  | Male | 84 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | [9 empty] | It was reported via a medwatch report (mw5172995) that a patient implanted with an impella pump decompensated. the medical staff and the patient's family decided to withdraw care and the patient expired. | [5 empty] | A1 is unknown. product details are unknown in section d4. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6843 | 1220648-2025-30982 | 22706796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22706796&pc=OZD | 05/03/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 05/03/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592904 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted. during patient support, the placement signal was lost, and no waveforms were observed. the white cable was checked and confirmed to be securely connected. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned. data logs were not recovered. the cause of the optical signal issue was not determined since product was not returned, data logs were not recovered, and insufficient clinical details.
6844 | 1220648-2025-30983 | 22707375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22707375&pc=OZD | 07/22/2025 | 09/10/2025 |  | 08/04/2025 |  |  | 07/22/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702545 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 76 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced aortic valve insufficiency and hemolysis. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation of hemolysis and heart valve damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
6845 | 1220648-2025-45718 | 22707454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22707454&pc=OZD | 06/12/2025 |  |  | 08/04/2025 |  |  | 07/21/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670597 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ischemia | Sepsis | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that on the first day of impella cp support, the patient experienced a mottled and cool leg where the impella was implanted and was noted to be improving. it was also reported on the first day of support that the patient had signs of sepsis and three days later (b)(6) 2025, the patient remained septic from an unknown source. while there was no allegation of impella related infection, impella cannot be ruled out as a possible source. a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received regarding hemolysis with relationship to pump as "possibly" related. it was noted that the patient¿s plasma free hemoglobin (pfhgb) decreased each day after the start of impella cp support. there was no baseline pfhgb drawn. on (b)(6)2025, the result of pfhgb 46.8 mg/dl and it was 8.8 mg/dl on (b)(6) 2025. it was reported that during support, the patient became hypotensive because of blood loss and associated decreased mixed venous oxygen saturation. the patient¿s norepinephrine was increased as a result, and the impella setting was increased from p5 to p6. the patient received a unit of packed red blood cells (prbc). a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received regarding the adverse event of left ventricle (lv) thrombus with relationship to be "possibly" related to the pump. post impella, patient showed sluggish filling of the left ventricle apex showing thrombus formation adjacent to anterior apex wall, for which aspirin was loaded. additionally, the patient was switched to heparin from bivalirudin for the clot. a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received regarding the adverse event of right femoral artery impella site bleeding with a relationship listed as "possibly related". it was reported that the impella site oozed overnight, and the patient received a unit of prbc. the impella was weaned to p4 and the patient received another unit of prbc the next day, (b)(6) 2025. the patient was successfully weaned, the impella cp explanted, and noted as outcome at explant as survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | G3. date received by manufacturer corrected to 6/12/2025.
6846 | 1220648-2025-45703 | 22708214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22708214&pc=OZD | 07/27/2025 | 07/30/2025 |  | 08/04/2025 |  |  | 07/27/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026698330 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in heart failure and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported during use of the impella 5.5, the impella was observed to have high purge pressure and low purge flow following a red alarm on the automated impella controller for high purge pressure. recommendations were to verify there were no kinks in the purge line, and then to add tissue plasminogen activator (tpa)to the purge solution. no kinks were observed and tpa was added to the purge solution, however this did not resolve the issues. it was then recommended to replace the pump, however the medical staff decided to keep the pump in place until the patient was able to be bridged to a durable left ventricle assist device (lvad) on (b)(6) 2025. there was no patient harm reported due to the high purge pressure malfunction. the impella 5.5 was successfully explanted and the patient was successfully placed on the lvad. | [5 empty] | D.9 updated as the pump was returned for investigation. h.6 updated investigation codes as the pump was returned for investigation. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | E.1. zip code was filled in on the initial report submitted when the code should have been in the initial reporter postal code field. the code has been added to the initial reporter postal code field. g.1. manufacturer site zip code was filled in on the initial report submitted when the code should of been in the manufacturer site postal code field. the code has been added to the manufacturer site postal code field. investigation summary: the impella 5.5 device was returned. (cut: no serial number was to be confirmed). data logs were not available. high purge pressure: clinical details reported a high purge pressure alarm on (b)(6) 2025. there were no kinks noted in the purge line. tissue plasminogen activator (tpa) was then added the purge, without changing the purge cassette. after 11 hours, the purge pressure had still not resolved. the pump was explanted on (b)(6) 2025. data logs were not available for review. the product was returned with a cut in the catheter, and only the motor/cannula side was returned. there were no kinks noted in the length of the catheter returned. after soaking the pump, there were signs of biomaterial in the outflow cage/purge gap and around the impeller. a window was cut in the catheter just proximal to the motor housing, and there were no signs of blood ingress. since the pump was returned cut, it could not be run to attempt to reproduce high purge pressure. though the biomaterial on returned product could have been the cause, clinical details report tpa did not resolve the complication and data logs nor complete product were returned to rule out a kinked purge line. since data logs were not available for investigation and product was returned cut, the cause of the high purge pressure could not be determined. device history review: the pump passed all post-sterile inspection checks. h.6. coding was updated to reflect investigation summary and findings. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6847 | 1220648-2025-30984 | 22708291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22708291&pc=OZD | 05/03/2025 | 08/04/2025 |  | 08/04/2025 |  |  | 05/03/2025 | 08/04/2025 |  |  |  |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was using a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. during shockwave therapy of the artery, the impella sensor was damaged, resulting in the loss of metrics. no other functional issues were noted. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned. data logs were reviewed and the alarm was confirmed. per the clinical details provided and data log review, the root cause of the optical signal was determined to most likely be a damaged sensor due to shockwave interaction.
6848 | 1220648-2025-30985 | 22709409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22709409&pc=OZD | 03/22/2025 | 10/27/2025 |  | 08/04/2025 |  |  | 03/22/2025 | 08/04/2025 |  |  | 09/30/2024 |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024304543 |  | 004413 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant was using a 2.5 pump to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the placement signal became non-functional. no patient harm was reported. | [5 empty] | G3 ¿date received by manufacturer¿ corrected. | All pertinent information available to abiomed has been submitted at this time. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the returned pump was visually inspected and no abnormalities were observed. the cause of the issue was most likely patient condition related since motor current was reported to be unaffected and the returned pump operated to specification.
6849 | 1220648-2025-30986 | 22709448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22709448&pc=OZD | 08/18/2023 | 08/04/2025 |  | 08/04/2025 |  |  | 08/18/2023 | 08/04/2025 |  |  | 09/30/2023 |  | 08/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022133726 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported using a 5.0 pump to support a patient. during support, the placement signal failed, and flow estimation became active. the device remained in use and was not removed. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product or data logs returned for investigation. the cause of the placement signal issue was not determined since product and data logs were not returned.
6850 | 1220648-2025-30911 | 22685176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685176&pc=OZD | 07/14/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/22/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2026649612 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that a patient with pulmonary congestion presented for percutaneous coronary intervention with impella cp support. during support, the patient developed a subarachnoid hemorrhage. the decision was made to withdraw the procedure due to poor neurological prognosis caused by the impella-induced subarachnoid hemorrhage. the patient passed away two days after removal. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | B.2 the exact date of death is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6851 | 1220648-2025-30912 | 22685261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685261&pc=PYX | 07/03/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/03/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625511 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp then followed by an impella rp flex device for bipella mechanical circulatory support (mcs). the patient experienced abdominal bleeding, and it is unknown if the use of heparin for impella support impacted the outcome of the abdominal bleed. the patient presented with chest pain and was emergently taken to the catheterization lab. the patient subsequently coded and was shocked eight times before proceeding with the impella cp inserted via the right femoral artery during a code blue procedure. flows were initiated and eventually the patient stabilized with return of spontaneous circulation. the culprit was a 99% occluded proximal right coronary artery which was treated with angioplasty and drug-eluting stenting. the disease to the left anterior descending and diagonal arteries were planned to be treated on a different day. a right heart catheterization was performed, and the physician proceeded with impella rp flex placement via the right femoral vein into the left pulmonary artery without issue. right-sided support was started with flows of 2.5-3.0 liters per minute achieved. the patient was "extremely" fluid responsive and no marked improvement with the impella rp flex placement. echocardiogram in the procedure area showed left ventricular function of approximately 55% and right ventricular function slightly diminished, however, not dilated. the patient was sent to unit on bipella (impella cp and impella rp flex) support. approximately one day into bipella mcs, systemic heparin drip was ordered with goal of activated clotting time above 200 seconds. the patient had decreased pulses in both extremities with more noticeable symptoms on the impella side. a femoral-femoral bypass was performed to manage the ischemia. the patient continued to have worsening hemodynamics. in addition, abdominal bleeding was observed for which multiple blood products were given. a bedside abdominal tap revealed blood in the abdomen that was firm. the patient continued to deteriorate from the underlying condition and the decision was made to withdraw care. | [5 empty] | The impella rp flex device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the impella rp flex is one of two devices associated with the event. there is not enough evidence to exclude the impella cp as an associated factor in the patient's demise outcome. refer to manufacturing report number 1220648-2025-30792 for the report on the impella cp.
6852 | 1220648-2025-30913 | 22685313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685313&pc=OZD | 07/20/2025 | 09/12/2025 |  | 08/01/2025 |  |  | 07/20/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696438 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 65 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support placed for a st-segment elevation myocardial infarction patient who had sepsis and chest pain. the pump was placed but was explanted after just two hours due to an access site hematoma. the field representative later updated the abiomed team that the patient did expire shortly after the case. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
6853 | 1220648-2025-30914 | 22685356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685356&pc=PYX | 07/20/2025 | 09/03/2025 |  | 08/01/2025 |  |  | 07/20/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025602825 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 90 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and during support, the pump was unable to repositioned and the decision was made to explant the pump. | [5 empty] | The investigation has been completed. the pump was not returned for investigation. the cause of the positioning issue was most likely patient movement per clinical that patient began to wake up, cough, and gag and the pump was then unable to be repositioned. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6854 | 1220648-2025-30915 | 22685388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685388&pc=OZD | 07/20/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/20/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696252 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed a bleed from their access site during impella 5.5 support. the patient received nine units of blood and a pressure dressing was applied to the site. when the bleed persisted, the patient was taken to the operating room for a washout and it was discovered that the patient was experiencing a muscle vessel bleed. additional blood products were given and the vessel was repaired. support continued uninterrupted. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6855 | 1220648-2025-45710 | 22685449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685449&pc=OZD | 07/01/2025 |  |  | 08/01/2025 |  |  | 07/02/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688070 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Break | Positioning Failure | Pumping Stopped | Failure to Read Input Signal | Material Twisted/Bent |  |  | Arrhythmia | Fever | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and would experience a fever requiring antibiotic, arrhythmic events and pump stop resulting in device replacement. the patient was readmitted for decompensated heart failure and was formally declined for a transplant for various reasons which included social settings. the patient required escalation to an impella in which an impella 5.5 was successfully implanted. approximately one day after the impella support started, the patient was noted being worked up for transplant and was febrile. the following days the patient was still febrile with 38.6 at the highest temperature. culture ran and labs drawn (results were unnoted). however, the patient was started on 1 dose of antibiotics, and a peripherally inserted central catheter (picc) line was removed. the next day it was noted that the patient was no longer febrile. the picc line was placed again and impella mcs continued. the patient was listed for transplant. now two days later, the patient experienced some episodes of ventricular tachycardia; patient noted as stable, however, continued impella mcs. the patient was able to ambulate three times per day with having some ¿occasional¿ ectopy. as the patient continued to wait for the heart transplant, the patient experienced a suction, ectopy and position unknown alarm event on day 16 of support. point-of-care ultrasound was completed and noted the impella was 7.0cm from aortic valve area. consequently, the impella 5.5 device was repositioned, pulled back to 5.6 cm and all related issues were resolved. however, the patient would have more suction; went to support level p-4 to break and would quickly suction again upon return to p-5 support level. albumin was given and 30 minutes later support level was increased to p-7 with no new suction alarms. the next day, the patient had a coughing spell resulting in a placement signal low alarm. patient was, however, stable and getting good flows. the physician made the decision to reposition pump and was damaged during this repositioning. additionally, it was noted the impella 5.5 device pump suddenly stopped during the repositioning and was unable to be restarted. notes indicated the physician ¿over manipulated¿ cannula and had torque built up in catheter. the decision was made to unplug impella 5.5 device cord from automated impella controller to remove the loops of the plug. however, when the impella 5.5 device cord was reconnected and started back up, an ¿impella defective¿ alarm triggered. an attempt to restart again was made with support level up to p-5; however, the ¿impella defective¿ alarm triggered again. the device was therefore removed and replaced in the operating room. the patient was noted to be in critical condition following the event. | [5 empty] | Additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instruction for use: for the related event are as follows: section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | Post coding alignment with product engineers, complaint coding for health effect - impact and medical device problem were updated to better reflect the reported event, consequently, a redo of section h6 for all health effect, medical device problem and component coding is being completed as part of this follow-up 2 supplemental reporting.
6856 | 1220648-2025-30916 | 22685462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685462&pc=OZD | 07/20/2025 | 09/29/2025 |  | 08/01/2025 |  |  | 07/20/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663423 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 58 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, there were placement signal not reliable alarms. the placement signal returned and placement monitoring was re-enabled. an echocardiogram showed the impella cp was aimed toward the mitral measuring a little shallow at 2.5 centimeters and inflow was not in contact with the mitral valve apparatus. the pump was left as is and the staff monitored. later, weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of optical signal issue and positioning issues has been completed. optical signal issue: data log review showed placement signal (ps) issues occurred upon pump start with placement signal not reliable (psnr) triggered shortly after. snr was low at 250 right before pump start and went to 0 on pump start. when psnr resolved 45 minutes later, snr increases to a mean of 560, contrast decreases, and gain decreases. lamp and light remain stable. the pump was returned and inspected. psnr unable to be reproduced. a kink was noted near the red pump plug on the catheter end but had no impact to 5s parameters, even when manually kinked. biomaterial was found wrapped around the impellor blades. the cause of the optical signal issue was not determined due to inconclusive data log and product review with insufficient clinical details provided. positioning issues: the pump was returned and inspected. the pump showed no physical abnormalities that would impact positioning. the cause of the positioning issue was not determined due to insufficient clinical details. d9 device returned to manufacture date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6857 | 1220648-2025-30917 | 22685527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685527&pc=OZD | 07/19/2025 | 09/12/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642035 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that delivery of an impella cp was aborted due to patient anatomy. the pump could not be advanced into the arteriotomy. the patient's outcome was stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6858 | 1220648-2025-30918 | 22685545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685545&pc=OZD | 07/19/2025 | 08/29/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1648111 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an aic failed to recognize an installed purge disc during an aic to aic transfer. after multiple failed attempts, the patient was switched back to the original aic. | [5 empty] | The investigation has been completed. imc logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) several times. the purge disc detection issue was not able to be reproduced during testing. after inspecting the console there were signs of residue found around the purge sensor assembly. the root cause of the purge disc detection issues could not be conclusively determined due to the inability to be reproduced. the console operated as intended while testing on an engineering lab bench. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6859 | 1220648-2025-30919 | 22685656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685656&pc=OZD | 07/19/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677876 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient had a cp patient escalated to a 5.5 pump for the st-segment elevation myocardial infarction patient. the 5.5 was supporting when there was optical signal loss. the 5.5 supported for 25 days when explanted due to heart transplant. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6860 | 1220648-2025-30920 | 22685658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685658&pc=OZD | 07/19/2025 | 09/11/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026647398 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after one day of impella 5.5 support, a purge system block appeared. the purge cassette was replaced but the issue remained. tissue plasminogen activator was subsequently added to the purge and the issue resolved within two days. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of high purge pressure (hpp) has been completed. the impella device was not returned for evaluation. data log review showed purge system blocked alarms and no motor current (mc) spike outside of p-level changes. purge pressure build up over 6 hours prior to first purge system blocked alarm. purge flows decrease over the same period. logs show hpp resolving after 2 days. the cause of the high purge pressure was most likely biomaterial buildup due to purge pressure increases and purge flow decreases that take place over 6 hours before purge blocked alarms trigger. issue resolves 2 days after beginning tissue plasminogen activator (tpa). | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6861 | 1220648-2025-30921 | 22685696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685696&pc=OZD | 07/19/2025 | 09/16/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | Defective Device | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella cp implant and upon peeling away the peel away sheath, the pump was pulled out of the left ventricle. the procedure was done over again and patient was quickly back on support. furthermore, the patient had a small hematoma at the access site. staff believed it was getting worse. integriln drip was paused. additionally, the patient's pulse in their right leg diminished. the patient also had hematuria, position was checked and p level was lowered. after lowering the p level, the hematuria cleared. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of positioning issues, difficulty inserting/removing introducer, access site bleeding, limb ischemia, and hemolysis has been completed. positioning issues: data logs were reviewed and there were no position in aorta alarm triggers. however, there was a significant drop in motor current (mc) modulation for a brief moment just before the pump p-level was dropped and the pump was removed from the body, likely confirming the reported clinical details. based on the clinical details provided, the pump was pulled back inadvertently by the physician when trying to peel away the introducer sheath. therefore, the cause of the positioning issue was most likely a use issue. difficulty removing introducer: clinical details report that the physician required excessive force to peel away the introducer sheath. there were no reports of damage to the sheath. no further details were provided and product was not returned for investigation. since limited clinical details were provided and product wasn't returned for investigation, the cause of the difficulty removing the introducer could not be determined. access site bleed: since insufficient clinical details were provided and product wasn't returned for investigation, the cause of the access site bleed could not be determined. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. hemolysis: since limited clinical details were provided regarding the troubleshooting methods used to resolve hemolysis, data logs were inconclusive, and product wasn't returned for investigation, the cause of the hemolysis could not be determined.
6862 | 1220648-2025-30922 | 22685759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685759&pc=OZD | 07/19/2025 | 12/03/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681627 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 81 KG | Non Hispanic | White | Death | No | Yes | Device Sensing Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | No Clinical Signs, Symptoms or Conditions | [8 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, care was withdrawn and the patient expired. | It was reported that a patient implanted with an impella 5.5 experienced a placement signal not reliable alarm. placement signal waveforms and numbers were present.the alarm was disabled and support continues. no patient harm was reported. |  |  |  |  | Pump was not returned for investigation. data log shows low snr (below 2500) and placement signal values were lost during the first few days of case start. psnr alarm was seen during the ps issue. there was no issue recorded on remaining 5s optical signal parameters during the ps issue. later, placement signal and snr returned to normal without affecting any other 5s optical parameters. the cause of the optical signal issue was not determined since product was not returned for investigation. also, it is apparent that the console could be potential cause of the issue. mso review statement: medical safety reviewed the event and noted the following: the patient was brought to the operating room for placement of an impella 5.5 device as a bridge to coronary bypass grafting (cbg). upon evaluation, the axillary artery was determined to be too small to accommodate the impella 5.5. an impella cp was implanted without issue. post-implant, hemolysis was noted, which is considered a reportable adverse event. following cabg, mitral valve replacement (mvr), and tricuspid valve replacement (tvr), the decision was made to place an impella 5.5 directly. after 24 hours of support, the impella 5.5 generated a placement signal not reliable alarm. placement signal values and waveforms were present, and the alarm did not impact hemodynamic support. the alarm was disabled. during support, the patient developed volume overload and acute kidney injury (aki), requiring initiation of continuous renal replacement therapy (crrt). on day 9 of impella 5.5 support, the patient began to decompensate and experienced a gastrointestinal (gi) bleed. the patient was suspected to be in septic shock. after continued support, it was determined that the patient was no longer a candidate for permanent lvad therapy. the care plan shifted to continued impella 5.5 support with the goal of native heart recovery. on (b)(6) 2025, the impella 5.5 was repositioned, resulting in improved flows. despite ongoing support, the patient expired after 41 days of impella support. the impella cp hemolysis is a reportable adverse event. impella 5.5 placement signal not reliable alarm are considered device malfunction reportable events. the renal insufficiency, infection and gi bleed are reportable patient adverse events for the impella 5.5. the patient¿s death will be conservatively reported on the impella 5.5, although it is unlikely that the device contributed to the cause of death. the patient¿s deteriorating condition and non-candidacy for permanent lvad were likely contributing factors. note: the automated impella controller will remain a reportable product malfunction accordingly. d6b updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6863 | 1220648-2025-30923 | 22685846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685846&pc=OZD | 07/19/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696927 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was being prepped and the sheath kit was opened while gaining access. the patient developed a hematoma at the access site. the implant was aborted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6864 | 1220648-2025-30924 | 22685913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22685913&pc=OZD | 07/19/2025 | 09/15/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693437 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced purge system open and purge pressure low alarms occurred. the purge cassette and dextrose concentration were changed but the alarms did not resolve. additionally, a drop in purge pressure occurred. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the impella 5.5 was returned for evaluation. product damage (hole in catheter): upon visual inspection of the returned pump, blood was present in the red handle and dried blood present around red handle to catheter kink protector. a hole in the catheter was present at the 30cm mark. an additional failure mode of "product damage (hole in catheter) was added to the investigation per returned product findings. clinical details noted that a blood-tinged leak was coming from the catheter side of the red impella plug soon after pump was inserted. the cause of the product damage was most likely a use related issue based on returned product analysis. purge leak - pump: returned pump was purged with a syringe and it was found that purge was leaking from the purge tubing at the 30cm marking. a hole in the catheter was also present at the 30cm marking. data logs were reviewed and a sudden drop in purge pressure with increase in purge flow was observed approximately 10 minutes after pump was initiated. no motor current spikes were observed at time of drop in purge pressure. clinical details noted that 10 minutes into support, a sudden drop in purge pressures was observed. physician was removing the impella cp pump at the time of drop in purge pressure. prior to this, purge pressures were within a normal range. additionally, blood-tinged leak was observed coming from the catheter side on the red impella plug. the cause of the purge leak - pump was most likely due to damaged purge tubing based on returned product and data log analysis. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
6865 | 1220648-2025-30925 | 22686048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686048&pc=OZD | 07/19/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  | 11/30/2025 |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429017 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an impella cp for support due to acute myocardial infarction. it was noted that the impella cp serial number (b)(6) was opened and prepped. during set up and prior to insertion, the purge cassette was not properly seated and cracked when attempting to close the purge door. once a new purge cassette was replaced, the physician realized it was a purge issue and not a pump issue. pump serial number (b)(6) was connected and implanted for duration of the case without issue. pump serial number (b)(6) was never connected and discarded. during support it was reported that on bedside echocardiogram the impella looked to be pointed at the mitral valve. urine was light pink. the pharmacy was getting purge solution with sodium bicarbonate added and pedal pulses were operable. the site was clean, dry and intact. additionally, the patient was taken to the intensive care unit and while being checked in began to code. the patient was losing volume rapidly out of the chest tubes. over three liters were lost during the code and the patient remained in pulseless electrical activity for the duration of the code. rhythm was not obtained back and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6866 | 1220648-2025-30926 | 22686099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686099&pc=OZD | 07/19/2025 | 09/11/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677192 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support had bleeding around the repo sheath. pressure was held and heparin was stopped. the patient also had actively bleeding from the right femoral. impella side is the left femoral. blood products were administered. hemoglobin dropped from 12.3 to 9.3 and is back up to 12.4. the patient is stable. | [5 empty] | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding : the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time (act) values, heparin was stopped to achieve hemostasis as per the clinical description. vascular damage : the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6867 | 1220648-2025-30927 | 22686149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686149&pc=OZD | 07/19/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  | 12/31/2025 |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446034 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 84 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an impella cp support due to methamphetamine toxicity and was found to be in cardiogenic shock. the pump was placed and distal perfusion was placed due to limb ischemia. the support was ongoing when after five and a half hours, the patient had cardiac arrest and cardiopulmonary resuscitation was initiated. the family made the decision to withdraw care and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6868 | 1220648-2025-30928 | 22686271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686271&pc=OZD | 07/20/2025 | 09/12/2025 |  | 08/01/2025 |  |  | 07/20/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655080 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after impella cp was placed there was oozing at the impella access site and distal perfusion was compromised. distal reperfusion sheath was placed in the lsfa and right fa sheath was placed for external fem/fem bypass. figure eight suture placed on impella access site and pressure dressing applied for the bleeding. it was further reported that the family decided to withdraw care and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was determined to be use issue because the bleeding was managed by placing figure 8 suture at place and applying manual pressure. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30928.
6869 | 1220648-2025-30929 | 22686481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686481&pc=OZD | 07/19/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693519 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient had thrombocytopenia. anticoagulation was withheld. support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6870 | 1220648-2025-30930 | 22686682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686682&pc=OZD | 07/18/2025 | 10/27/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660928 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported that impella had placement signal not reliable alarm. the audio was disable. | The user facility reported that the patient on impella 5.5 support had some numbness and loss of dexterity in the left arm. the patient had a cerebrovascular accident. the treatment and patients outcome are unknown. |  |  |  |  | B1 revised to reflect both adverse event and product problem based on the additional information received from the clinical site. b5 revised to reflect additional information received from the clinical site. d6b explant date added. h6 medical device problem code 2917 added to reflect the additional failure mode reported. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva and optical signal issue has been completed. the impella device was not returned for evaluation. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. optical signal issue: the data logs confirm placement signla not reliable (psnr) alarms on 04-aug-25 about 38 days after support was initiated. the logs show a drift down of the placement signal for about an hour until placement signal goes out of range and psnr alarms trigger. the pattern identified has been characterized as a wear of the optical sensor. the 5s parameters before sensor wear are stable after which they are affected. the design life of the impella 5.5 s2 according to design traceability matrix is 60 days and this failure occurred within the design life. the cause of the optical signal issue was due to a sensor silicone wear of the optical sensor.
6871 | 1220648-2025-30931 | 22686941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686941&pc=OZD | 07/19/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594107 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was being placed to support a patient admitted in with a st-segment elevation myocardial infarction. the pump was placed and the patient went into ventricular tachycardia (vt). the pump was repositioned. the patient was shocked and the case was proceeded. the vt was caused by the left ventricle positioning. the support was for 69 hours and then explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6872 | 1220648-2025-30932 | 22686999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22686999&pc=OZD | 07/19/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/19/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696183 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 87 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella 5.5 support had a run ventricular tachycardia which required one shock. the patient was started on lidocaine drip. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6873 | 1220648-2025-30933 | 22687686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22687686&pc=PYX | 07/04/2025 | 08/29/2025 |  | 08/01/2025 |  |  | 07/04/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026675497 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was referred to surgery for coronary artery bypass graft. however, the patient was showing signs of being septic with diagnosis of urosepsis and bacteremia. the surgeon recommended antibiotics and for the patient to follow up in 6 weeks. the patient was discharged to a rehab facility. at the rehab facility, the patient started to experience chest tightness and shortness of breath. the patient was brought back to hospital and admitted to the intensive care unit. the patient's status continued to decline. the patient required intubation and developed bi-ventricular failure. the patient was noted to be maxed out on multiple pressors. care team spoke with family on plan of care. the family elected to take to cardiac catherization and place bipella support. during support it was noted that the mattress suture broke. the physician attempted to place multiple new mattress sutures, but was unsuccessful. large amounts of bleeding was noted at the sheath site. manual pressure was held. vascular surgeon was able to suture the site. hemostasis was achieved. | The medical safety officer reviewed and determined given the massive bleeding in the groin requiring intervention there is not enough evidence to dissociate the impella rp flex device from the patient outcome. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: isection: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, since vascular surgeon called and able to suture site to achieve hemostasis. the medical safety officer (mso) reviewed and determined given the massive bleeding in the groin requiring intervention there is not enough evidence to dissociate the impella rp flex device from the patient outcome. the following fields were corrected on the initial manufacturer device report number 1220648-2025-30933 based on mso review. b2 death and date of death. b5 patient outcome. h1 type of report. h6 health effect - impact code 1802.
6874 | 1220648-2025-45704 | 22688448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22688448&pc=OZD | 07/22/2025 | 07/30/2025 |  | 08/01/2025 |  |  | 07/22/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094547 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that when the automated impella controller (aic) was powered on for use during an emergent case, and the medical team was alerted to an error message "system self-check failed." a back-up aic was used instead, and no patient harm was reported. | [5 empty] | Section a: patient information is unknown. product status: the automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
6875 | 1220648-2025-30934 | 22688869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22688869&pc=OZD | 02/22/2023 | 08/01/2025 |  | 08/01/2025 |  |  | 02/22/2023 | 08/01/2025 |  |  | 11/30/2024 |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2024289693 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported abnormal placement signal values during support. switching the pump to a new automated impella controller (aic) console did not resolve the issue. support continued with no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the returned pump was investigated and the "placement signal not reliable" alarm occurred when the pump was dry. when the sensor was wet there was a placement signal. the pump was investigated and the 5s were out of range. the pump was unable to hold the correct pressure in the uson. the sensor was removed and found to be damaged. the imc logs showed an incorrect offset as the pump had been programmed under a wet condition and due to the damage to the sensor, operated differently under a dry condition when returned. the incorrect offset applied by the console caused the change in the optical signal parameters which then caused the divergence in the field between the pump position waveform and the a-line. the cause of the optical signal issue was a damaged sensor. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the thime the product was manufactured and released to the customer. with the available information, there is no indication that there was is a product malfunction or deficiency.
6876 | 1220648-2025-30935 | 22689138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22689138&pc=OZD | 05/28/2023 | 08/01/2025 |  | 08/01/2025 |  |  | 05/28/2023 | 08/01/2025 |  |  | 12/31/2024 |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2024304059 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a loss in placement signal during impella support. there was no patient harm resulting from the noted issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. returned pump plugged into the automated impella controller (aic) and placement signal was -27/-27mmhg. biomaterial observed on sensor. sensor cleaned with sporicidin and placement signal returned to -1/-1 mmhg. the cause of the issue was biomaterial. refer to (b)(4) for additional biomaterial pattern details. with the available information, there is no indication that there was/is a product malfunction or deficiency.
6877 | 1220648-2025-30936 | 22689391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22689391&pc=OZD | 02/13/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 02/13/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135599 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump placed to support a patient with acute myocardial infarction and cardiogenic shock. the pump lost its optical signal; however, this did not prompt a replacement. the pump remains in use without causing harm or injury. the incident has led the team to also investigate the console. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. console logs from the reported day of the event are consistent with the complaint. the logs show multiple instances of the following errors. the console was sent back to field service for further investigation. the reported complaint could not be reproduced. however, upon inspecting the analog output of the ccd from the optical bench, intermittent distortions in the signal (envelope/fringe) were observed. while the signal was distorted, the measured "5s" parameters were not within specifications (low snr). the root cause of the issue was identified as intermittent distortions in the optical bench signal, likely caused by a malfunction in the optical bench or its components. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. please refer to investigation checklist for "qn: (b)(4): dv deviation. not related to the complaint" and prior aic related complaints. corrective actions: before returning the console to the customer, the following actions are instructed to perform: replace the optical bench. perform a full functional check on the console and optical system.
6878 | 1220648-2025-30809 | 22689407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22689407&pc=OZD | 07/14/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 07/14/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567052 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp to promote left ventricular functional recovery. during support he patient developed an infection. it was noted that left ventricle and right ventricle functions recovered, however the patient was still having an infection. it was noted that as the infection resolves they'll definitely wean and remove impella. the impella was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6879 | 1220648-2025-30937 | 22689427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22689427&pc=OZD | 01/09/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 01/09/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550391 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump in use for 35 hours to support a patient undergoing high-risk percutaneous coronary intervention (hrpci). the pump lost its optical signal, but it remained operational until weaning and explant. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs were provided and console was returned. returned product was investigated and there were no abnormalities noted. the pump passed laser continuity, and when connected to the console and run in a bucket for 2 hours, all optical signal metrics were stable the entire period. clinical details provide that the patient's map had dropped during the first few hours of support but had recovered upon explant, which supports the ps trend noted in data logs. based on all the above mentioned details, the root cause of the optical signal issue was determined to be patient condition. with the available information, there is no indication that there was a product malfunction or deficiency.
6880 | 1220648-2025-30938 | 22689479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22689479&pc=OZD | 01/17/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 01/17/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596931 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of the implanted impella 5.5 pump consistently registered lower than the arterial line and blood pressure cuff readings throughout patient support. there was no noted impact on hemodynamic support, and the patient remained stable and supported throughout the planned surgery. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the root cause of the optical signal issue was unable to be determined as no product or data logs were returned / available for this investigation.
6881 | 1220648-2025-30939 | 22689653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22689653&pc=OZD | 01/22/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 01/22/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547322 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during the initial priming of the impella cp, the placement signal sensor alarmed. the case proceeded, and the pump was initially run in auto mode despite no placement signal appearing on the console. a second console was obtained, and the impella cp was transitioned to it, at which point the placement signal functioned properly. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was returned and evaluated. all optical bench parameters passed and no placement signal alarms were triggered. no problem was found with the pump. with the available information, there is no indication that there was a product malfunction or deficiency.
6882 | 1220648-2025-30940 | 22689861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22689861&pc=OZD | 01/29/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 01/29/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025606277 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, 'placement signal unreliable' alarms were triggered. support continued without interruption. no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned. the data logs show a loss in placement signal corresponding. these changes are indicative of a fiber break. the loss in placement signal occurred while transferring the patient to bed. the root cause of the optical signal issue was most likely a damaged optical fiber line as data logs follow the pattern and the loss in placement signal occurred during patient transfer.
6883 | 1220648-2025-30941 | 22690132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22690132&pc=PYX | 03/24/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/24/2025 | 08/01/2025 |  |  | 08/31/2025 |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025574582 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during use, placement signal issues occurred. pump position was verified, and support continued. no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs were reviewed and the alarms were confirmed. no product was returned for analysis. the root cause of the optical signal issue was likely as a result of an interaction with another device, in this case an interaction between the wires and the optical sensor leading to placement signal not reliable alarms as evidenced by data log analysis and clinical information with the available information, there is no indication that there was a product malfunction or deficiency.
6884 | 1220648-2025-30942 | 22690156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22690156&pc=OZD | 03/28/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/28/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561737 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was placed. during support, placement signal issues occurred. despite this, support continued, and no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump and data logs not returned. the cause of the optical signal issue was not determined due to no product or data logs available for investigation.
6885 | 1220648-2025-30943 | 22690193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22690193&pc=OZD | 04/03/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 04/03/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584663 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella 5.5 pump experienced a loss in placement signal during patient support. no interruption in support occurred as a result of the loss in signal. despite this, support continued, and no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the returned device was inspected and there were no abnormalities. the 5s parameters were within specification. the data logs confirmed placement signal not reliable alarms. a pattern for an oscillating contrast failure was observed. upon review of the data logs and returned pump, it is apparent that the console is the potential cause of the issue. please refer to complaint#: (b)(4) for an investigation into the console.
6886 | 1220648-2025-30944 | 22690339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22690339&pc=OZD | 03/11/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/11/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620149 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal issues were observed, including false 'impella in aorta' alarms. motor current remained pulsatile, and echocardiography confirmed correct impella positioning. the clinical team elected to disable placement monitoring. no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. pump was not returned. data log review showed pulsatile parameters during the impella position in aorta alarms. the false impella position alarms were confirmed to be triggered relative to pulsatile signals. the cause of the optical signal issue was most likely patient condition related due to false position alarms. with the available information, there is no indication that there was a product malfunction or deficiency.
6887 | 1220648-2025-30945 | 22690476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22690476&pc=OZD | 02/24/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 02/24/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592870 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella 5.5 pump experienced persistent optical placement signal issues during patient support. despite the unreliable signal, support continued without interruption. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the impella 5.5 pump was not returned. data logs indicated alarm after placement signal drifting upwards with relative to snr gradually drifting down but within specified range. no other major abnormalities found from log review. further investigation on the subcomponent build revealed that there was no flaw with coating on the sensor. the root cause of the optical signal issue was not determined since product was not returned and no conclusive evidence from data logs on the failure pattern. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. please refer to investigation checklist for identified markers indicating "the rework associated to this device was delo rework which is unrelated to this failure mode." and "25-42011 had similar optical failure complaint related to subcomponent lot # 1803573."
6888 | 1220648-2025-30946 | 22690586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22690586&pc=OZD | 03/17/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/17/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578618 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was using a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. prior to starting the implanted impella cp pump, a 'placement signal not reliable' alarm appeared on the console. the positioning and motor current were verified, and support proceeded without interruption. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. clinical details report that a alarm triggered prior to starting the pump. no further details were provided. the provided data log has recorded only a "case start" notification. the pump was not returned for investigation and only a portion of imc logs were available for review. since limited clinical details were provided and insufficient product and data logs were returned for investigation, the root cause of the optical signal issue couldn't be determined.
6889 | 1220648-2025-30947 | 22690938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22690938&pc=OZD | 03/18/2025 | 10/27/2025 |  | 08/01/2025 |  |  | 03/18/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585989 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The customer reported that the device exhibited optical signal issues. | The complainant reported that an impella 5.5 pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. motor current remained pulsatile, and flow rates were appropriate for the set p-level. support continued uninterrupted, and no patient harm was reported. |  |  |  |  | G3 ¿date received by manufacturer¿ corrected. | B5: added additional information d6b: corrected the explant date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. pump was not returned. data log reviewed showed placement signal alarms at the beginning of the case. the cause of the optical signal issue was not determined since product was not returned, real time logs were not returned for the date the complaint occurred, and insufficient clinical details.
6890 | 1220648-2025-30948 | 22691050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22691050&pc=OZD | 03/23/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/23/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609618 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, placement signal suddenly changed and was abnormal. support continued without interruption. no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. pump was not returned. data logs confirmed low alarms but did not match any known optical patterns. the cause of the optical signal issue was not determined since product was not returned, inconclusive data logs review, and insufficient clinical details.
6891 | 1220648-2025-30949 | 22691194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22691194&pc=OZD | 04/10/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 04/10/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087213 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, a 'placement signal not reliable' alarm was triggered. the alarms were disabled, and support continued without interruption. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs and console was returned. on the returned product found the front panel optical connector contaminated, unable to remove the debris upon cleaning (likely unrelated to the reported failure mode). the root cause for the optical signal issue was a lamp malfunction. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. please refer to investigation checklist for indicated "(b)(4) addresses s/w change for usb release."
6892 | 1220648-2025-30950 | 22691379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22691379&pc=OZD | 03/03/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/03/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596370 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, a 'placement signal not reliable' alarm appeared. however, both the motor current and flow remained within range, and support continued without interruption. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. impella cp was returned for evaluation. upon visual inspection, a kink in the catheter just past the red handle kink protector was observed. pump did not pass laser continuity testing and light was observed emitting from the kink in the catheter. pump was connected to the console and alarm reproduced when connected to console. data logs were reviewed and lt log of case shows a sudden change in optical 5s parameters consistent with a fiber break. clinical details noted that a alarm was noted on support and no ps waveform was observed on aic. pump remained in patient for an additional 2 days before explant. the root cause of the optical signal issue was due to a damaged fiber in the catheter right past the red handle kink protector.
6893 | 1220648-2025-30951 | 22691523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22691523&pc=OZD | 03/03/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/03/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592821 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a 5.5 pump to support a patient. it was believed that the optical sensor on the implanted impella 5.5 pump was failing. the placement signal readings were both elevated and dampened. due to these irregularities, the placement signal was disabled. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned for investigation. the data logs show a 5-day gradual upward trend in the placement signal, followed by a '1048 - placement signal not reliable' alarm and a spike to 3000, which persisted until the end of the case run. the root cause of the optical signal issue was not determined since product was not returned and no conclusive evidence from data logs on the failure pattern. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. please refer to investigation checklist for indicated "(b)(4) had similar optical failure complaint related to subcomponent lot # 1803573."
6894 | 1220648-2025-30952 | 22692381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22692381&pc=OZD | 03/05/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/05/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a 5.5 pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal issues were observed. an echocardiogram was performed, and the impella appeared to be in the appropriate position. support was continued. the investigating engineer determined that the issue likely originated from the console. no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs and console have been returned. this console was tested after arriving. there were no issues found in the analog output, which confirms that there is no issue with the optical bench. a small reflective material was found on the front panel optical connector. this is unrelated to the reported failure mode. the root cause of the oscillating contrast was most likely due to a lamp malfunction. a review of the device history record (dhr) for the suspected product and historical complaint data was conducted. please refer to device investigation checklist for indicated " qn (b)(4) addresses additional optical system tests to increase detection of nonconformities experienced at the subassembly level (pn 0042-4009). " and "22-11001: defective cordset 6000-0171."
6895 | 1220648-2025-30953 | 22692534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22692534&pc=PYX | 03/10/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 03/10/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600775 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an impella rp flex became inaccurate during support. although the pump's motor current remained stable, the observed flow and placement signal readings were inconsistent with the expected support level. the device was zeroed, but the issue persisted. despite this, support was able to continue using the implanted pump. no patient harm resulted from the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. data logs and pump have been returned. the clinical stated that the patient was in suction due to excessive fluid removal from dialysis on (b)(6) 2025. the data logs were reviewed and there was an increase in the ps and increase in flows before the flow calculation went out. during this time, cvp was negative and suction alarms were occurring. based on the clinical, data log analysis and returned product testing, the root cause of the placement signal issue is most likely due to the patients condition as they were in suction and had a lack of volume.
6896 | 1220648-2025-30954 | 22692656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22692656&pc=OZD | 02/26/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 02/26/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581884 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. the pump lost its optical signal but remained operational until weaning and explant. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. clinical details reported that at the end of the case, there was a sudden spike in ps measurement values until they were no longer measurable. the values also did not match up to other monitoring devices and pump position was confirmed. data log review confirmed that the alarm triggered at the end of the case. during the majority of the case, there were abnormal high values being read by the raw optical signal. as reported, at the end of the case, there was a sudden spike in the calculated values matching up with the abnormally high values in the raw optical signal that had been filtered out up to that point. there were no changes in optical signal metrics at this time. returned product was investigated and there were no visual abnormalities. upon review of data logs and the returned pump, no problem was found.
6897 | 1220648-2025-30955 | 22692720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22692720&pc=OZD | 02/01/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 02/01/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607173 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed. during support, "placement signal unreliable" alarms were triggered. the alarm was disabled, and support continued uninterrupted. there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned. the data logs show the pump is first connected within specification. the signal completely improves once the pump is turned off towards the end and then turned back on. due to lack of product returned, the true root cause of the optical signal issue cannot be determined.
6898 | 1220648-2025-30956 | 22692757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22692757&pc=OZD | 02/04/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 02/04/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508027 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was placed. during support, an "impella position unknown" alarm appeared in conjunction with altered signals. the issue resolved without the need for intervention. support continued, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. no product was returned. the data logs show the pump is first connected within specification. the data logs show motor current spikes with speed deviations after which there were impella position unknown and impella position in aorta alarms. there were also suction alarms. clinical noted patient was febrile. the root cause of the optical signal issue was most likely due to biomaterial on the sensor face as evidenced by motor current spikes.
6899 | 1220648-2025-30957 | 22692766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22692766&pc=OZD | 02/10/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 02/10/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581014 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, "placement signal unreliable" alarms were triggered. the pump was repositioned, but the issue persisted. support continued without interruption. no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the returned product was inspected and there were no abnormalities. the parameters on the returned product were within specification. the data logs from the case had no abnormalities. upon review of the imc logs, it was revealed that the pressure on the pump was applied before the pump was connected to the automated impella controller (aic). the root cause of the optical signal issue (placement signal low) was likely an offset error as imc logs reveal pressure was applied to the pump before connecting it to the aic, resulting in a calibration error that caused large deviations.
6900 | 1220648-2025-30958 | 22692855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22692855&pc=OZD | 02/06/2025 | 08/01/2025 |  | 08/01/2025 |  |  | 02/06/2025 | 08/01/2025 |  |  |  |  | 08/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551807 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was placed to support a patient with acute myocardial infarction and cardiogenic shock. during support, 'placement signal unreliable' alarms were triggered. the alarm was silenced, and support continued without interruption. no known patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue has been completed. the pump was not returned for investigation. data log was only provided till the first 3 days. as per the clinical, the alarm was reported on 6th day of pump use. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided.
6901 | 1220648-2025-30876 | 22673850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22673850&pc=OZD | 07/17/2025 | 08/01/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693543 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 57 KG | Non Hispanic | Asian | Death | Yes | Yes | Positioning Failure | Pumping Stopped | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Through communication with the abiomed field representative, it was discovered that a different complaint event occurred with this device (pump sn (b)(6)). according to the field representative, "the patient had a stroke after the second pump was put in." the pump stop event is associated with the first pump that was implanted. please see manufacturer report number for impella 5.5 pump one 1220648-2025-45710. upon implanting the second pump, the old graft (which was on the innominate artery), was removed and a new one was placed in the same spot. within 12 hours, the patient had a stroke, and thus, was removed from the transplant list. the patient was withdrawn from care and four days later expired. | The user facility reported an impella 5.5 had positional issues. while repositioning, the pump's cable was unplugged to relieve the torque. when plugged back in the pump failed to re-start. the team made the decision to remove and replace the pump. the patients outcome is critical. |  |  |  |  | As noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the following fields were corrected based on the additional information received by the field representative, that the initial manufacturer device report 1220648-2025-30876 reported a malfunction that should have been against the first pump implanted, which has been reported in manufacturer device report 1220648-2025-45710. this report is for the second pump that was implanted and the patient experienced a stroke and expired. b.1 product problem was erroneously selected. b.2 death and date of death was inadvertently omitted. b.5 corrective statement and patient narrative were updated to reflect the correct information. h.1 type of report was erroneously selected. h.6 health effect - clinical code, health effect - impact code, and medical device problem code were all erroneously entered. h.10 instructions for use were updated to reflect the new information provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6902 | 1220648-2025-30877 | 22674210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674210&pc=OZD | 07/17/2025 | 09/16/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694960 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella cp implant and during support, there was a high purge pressure alarm. the purge solution was changed to bicarb and support continued with the alarms continuing. the patient's procedure was complete and the pump was explanted. patient survived. | [5 empty] | The investigation of high purge pressure (hpp) has been completed. the impella device was received for investigation. data log review showed purge pressure trend up over 5 minutes while flows decreased before triggering hpp alarm. after hpp was triggered, abnormal purge ticks were noted and no motor current (mc) spike was seen. the pump was returned and inspected. no biomaterial was found on the impeller blades nor in the purge line. the pump was run for 15 minutes and the hpp was unable to be recreated. the cause of the high purge pressure was unable to be determined due to inconclusive evidence per log review and issue did not reproduce with return product. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and.
6903 | 1220648-2025-30878 | 22674232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674232&pc=OZD | 07/17/2025 | 09/18/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697932 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient began bleeding from their femoral access site during impella cp support. the heparin drip was held to manage the condition. later that same day, the family made the decision to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time values, heparin was stopped to achieve hemostasis as per the clinical description. b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30878.
6904 | 1220648-2025-30879 | 22674338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674338&pc=OZD | 07/17/2025 | 09/03/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165517 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic)'s plug for the white cable was broken. the console was just transported back to them. this issue did not occur during patient support. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - cabling issue has been completed. data analysis: not informative for this investigation. the imc logs did not show any anomalies. device analysis: the console ic3283 was sent to field service (fs) for further investigation. it was determined that the external optical connector dust cover was broken. additionally, the console was tested with the pump and the purge cassette, and no problems were observed. root cause: no issues were found related to the reported failure mode of the aic - cabling issue. the broken dust cover was most likely caused by user-related issues. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
6905 | 1220648-2025-30880 | 22674395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674395&pc=OZD | 07/17/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693658 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient with an ejection fraction of 10% was taken to the operating room for insertion of impella 5.5. the 5.5 was successfully inserted. it was reported that during support awaiting left ventricular assist device they removed central lines and gave the patient broad spectrum antibiotics until bacteria could be identified. additionally, the patient had a surgical pocket wash out. the patient continues on support. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock & the instructions for use for the cp with smartassist states the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6906 | 1220648-2025-30881 | 22674505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674505&pc=OZD | 07/18/2025 | 09/10/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660007 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 44 KG | Non Hispanic | White | Malfunction | No | Yes | Incomplete or Inadequate Priming | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump being placed to support the mitral valve surgery/replacement. the pump was primed but the priming failed so the setup was aborted and a new pump was used. no patient harm was reported. | [5 empty] | The investigation of priming issue has been completed. the pump and the purge cassette were returned for investigation. according to the clinical report, the doctor was not able to see purge fluid exiting the pump despite all troubleshooting efforts. data log shows multiple repeated case start steps and manual priming steps during the case. no visual abnormalities were noted on the returned product. the pump was successfully primed during testing, and all purge parameters were found to be within specification range. purge fluid was observed coming out of the pump during testing. there was no issue found during the case. the device functioned/operated to specification. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
6907 | 1220648-2025-30882 | 22674712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674712&pc=OZD | 07/19/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/19/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673795 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, there were unexplained incidences of flat motor current, flat ao, and flat left ventricle waveform with out any change of impella position. there were occasional suction alarms. at the end of the case, staff were able to settle on p5 without issue but above p5 would experience the same setting issues. the patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6908 | 1220648-2025-30883 | 22674766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674766&pc=OZD | 07/17/2025 | 11/25/2025 |  | 07/31/2025 |  |  | 07/20/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1306675 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 pump to support the patient with multiple cardiac and systemic comorbidities. automated impella controller (aic) unexpectedly shut itself down today. upon automatic reboot "impella defective" alarm present as well as "unexpected controller shutdown". white cable unplugged and replugged, impella support reinitiated at p-6. "unexpected controller shutdown" alarm still present. new aic obtained and device was transferred to new aic per abiomed recommendation. no alarms present after switch. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. data analysis: ¿ after 3 days and 18 hours of support with the pump (sn: (b)(6), the console unexpectedly shut down on the complaint date, july 20, 2025, due to an ¿io-graphics¿ process crash. ¿ following an automatic reboot, the console displayed ¿impella defective (error reading eeprom) ¿ alarm,¿ event #3 and ¿unexpected controller shutdown ¿ alarm.¿ event # 603. this confirms the reported failure mode. device analysis: this console was tested after arriving at fs. the unexpected controller shutdown could not be reproduced during a 23-hour test of the aic with the test cassette and pump at p-level p9. upon power cycling the console 10 times, each time powering on with the pump connected, the console was able to recognize the pump connection and read the pump sn # without any issues with reading the calibration data. therefore, the impella defective alarm could not be reproduced. root cause: ¿ the root cause of the unexpected controller shutdown could not be determined due to the inability to reproduce it. ¿ there was no ¿controller failure¿ alarm in the logs. it is most likely that the unexpected controller shutdown was misreported as a ¿controller failure¿. ¿ no log evidence was found to confirm an ¿optical signal issue¿. it is most likely that the ¿impella defective¿ alarm was misreported as an ¿optical signal issue¿. the cause of the ¿impella defective¿ alarm was not determined due to the inability to reproduce it. d3 address line 1 updated. d10 concomitant medical products and therapy dates updated.
6909 | 1220648-2025-30884 | 22674971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674971&pc=OZD | 07/20/2025 | 09/16/2025 |  | 07/31/2025 |  |  | 07/20/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701095 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed suspected hemolysis during impella cp support. it was noted that the patient had an elevated lactate dehydrogenase at the time of the concern. intermittent suction with shallow positioning was observed and volume was subsequently given. the condition was believed to have been successfully managed. it was additionally reported that the access site began to ooze and a stable hematoma was found in the left groin. an additional suture and manual pressure were applied to manage the condition. support continued uninterrupted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation for the reported access site bleeding and hemolysis has been completed. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely anticoagulation management related since the act was high (233) at the time of bleeding.the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
6910 | 1220648-2025-30885 | 22674993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22674993&pc=OZD | 07/20/2025 | 09/12/2025 |  | 07/31/2025 |  |  | 07/20/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609923 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient began to bleed from their groin site. manual pressure was held to achieve hemostasis and a pressure bandage was applied. the following day, the patient continued to bleed and vascular surgery was consulted to assess the right leg. the angle match was adjusted and heparin was put on hold until the groin site improved. the patient was moved the next day, at which point the reperfusion sheath was accidentally pulled out. the patient had not distal pulse at that point in time and the decision was made to remove the pump. an intra-aortic balloon pump was placed for ongoing support. the patient remained in stable condition. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia and access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding is determined to be use issue because reperfusion sheath was accidentally pully out during transport of the patient causing the bleed and the managed with manual pressure and angle match. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30885.
6911 | 1220648-2025-30886 | 22675215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22675215&pc=OZD | 07/20/2025 | 09/25/2025 |  | 07/31/2025 |  |  | 07/20/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528829 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support ongoing when after 3 days of support the pump stopped and troubleshooting failed to restart the pump for continued support. the pump was explanted and patient was noted to be in stable condition. | [5 empty] | The investigation of pump stop has been completed. the impella device was not returned for evaluation. due to a lack of product returned and no data logs, the root cause of the pump stop cannot be determined. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6912 | 1220648-2025-30887 | 22675272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22675272&pc=OZD | 07/17/2025 | 09/16/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686597 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was supporting a patient admitted in with cardiogenic shock and cardiomyopathy. the patient had arrhythmia during support while awaiting a transplant. the ventricular tachycardia was medically managed. on day 15 of the support, the transplant occurred and the team observed the red plug of the pump to be bleeding under the surgical drape. the patient survived. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of product damage (hole in catheter) and vf/vt/af/at has been completed. impella 5.5 returned for evaluation. product damge (hole in catheter): upon visual inspection of returned pump, catheter was cut completely and pump and cannula side of catheter were not returned. returned section of catheter was inspected and no damage to catheter was observed. blood was present in the red handle when opened. the cause of the product damage (hole in catheter) could not be definitively determined as insufficient product was returned for analysis and limited clinical details were provided. vf/vt: clinic the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
6913 | 1220648-2025-30888 | 22675293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22675293&pc=OZD | 07/20/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/20/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582157 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had ventricular tachycardia. an echocardiogram was called in and the impella migrated back to the inlet being just beyond the aortic valve. cardiologist repositioned the pump and ectopy resolved. additionally, the patient had oozing at the access site after the patient had a lot of movement. the site was redressed and angle matching was performed. three units of blood products were given to the patient. the oozing continued therefore the pump was explanted and intra-aortic balloon pump was implanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6914 | 1220648-2025-30889 | 22675403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22675403&pc=OZD | 07/17/2025 | 09/18/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700986 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 103 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient decompensated during a kidney transplant and suffered non-st-segment elevation myocardial infarction in the post-anesthesia care unit going into cardiogenic shock. the patient was emergently placed on venoarterial extracorporeal membrane oxygenation (va-ecmo) and taken to catheterization lab to revascularize the left anterior descending artery. the impella 5.5 implant for orthotopic heart transplantation versus left ventricular assist device was evaluated. it was reported that the right axillary approach with graft attached was successful. a 0.018 wire was placed in the left ventricle and an impella 5.5 was inserted into the impella sheath but did not advance more than halfway thru the artery. multiple wires were used and exchanged without success. interventional cardiologist was consulted and a balloon was used to widen the artery but was also unsuccessful. a gore dryseal 22 french was inserted through the graft into the artery but again would not advance past where the impella 5.5 had. after multiple methods and attempts, a cardiothoracic surgeon decided to abort the implant and have the patient remain on va-ecmo. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since 5.5 inserted into impella sheath but did not advance more than halfway thru the artery. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6915 | 1220648-2025-30890 | 22675665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22675665&pc=OZD | 07/18/2025 | 09/17/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591678 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 68 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in after being transferred from community to the tertiary center. an impella 14 x 25 centimeter sheath was inserted. the impella cp was inserted and advanced over the wire with no issues. the decision was made to leave the dislodged portion of previous pump until the patient was stabilized and take to surgery. while exchanging the peel away from the repositioning sheath, the impella cp was watched closely under fluoroscopy. while on fluoroscopy, during repositioning, an odd shape of the catheter was noticed and a kink was located between the motor and bend of the cannula. flows were diminished and the decision was made to leave across the valve as it was the last impella on-site. the patient was sent to the intensive care unit (icu) with a possible plan of extraction of the dislodged impella if the patient stabilized. the patient declined in the icu and passed away on support. | [5 empty] | The impella device was not received from the customer as it was left inside the patient when they passed away. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported low pump flow and product damage (cannula kink) has been completed. data logs confirmed the failure mode & clinical narrative. logs showed after pump started & run for about 30 minutes without issue, flow suddenly dropped to 0.5 l/min that triggered auto suction response & suction alarms. this event remained to the rest of the case. product was not returned for review. based on clinical description & data review, the root cause of low flow was most likely due to kinked cannula. the root cause of the kinked cannula was unable to be determined as limited clinical details were provided and no product was returned for review.
6916 | 1220648-2025-30891 | 22675987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22675987&pc=OZD | 07/17/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601550 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella 5.5 placed in a patient with non-st-segment elevation myocardial infarction after the intra-aortic balloon pump was explanted and a 4-vessel coronary artery bypass graft was planned at the second, tertiary hospital. the patient suffered from ventricular tachycardia that was treated by defibrillation on the second day of impella 5.5 support. two units of red blood cells were given and albumin. two amiodarone boluses were done. the patient was cardioverted twice for rapid ventricular response. the pump remained on despite this until care was withdrawn and the patient expired after just under four days. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6917 | 1220648-2025-30892 | 22676237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22676237&pc=OZD | 07/17/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672961 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that after the impella cp was placed, there was moderate oozing noted from the groin site despite angle-matching. activated clotting time was noted to be over 400. pressure was held for 20 minutes and perclose was tied down firmly with minimal bleeding after. in the iintensive care unit, systemic glucose tolerance test was on hold for activated clotting time out of range. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6918 | 1220648-2025-30893 | 22676561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22676561&pc=OZD | 07/17/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697029 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 118 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, ventricular tachycardia occurred. the patient was coded and shocked multiple times. the decision was made to place venoarterial extracorporeal membrane oxygenation with impella 5.5 support. however, the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6919 | 1220648-2025-30894 | 22676848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22676848&pc=OZD | 07/17/2025 | 09/15/2025 |  | 07/31/2025 |  |  | 07/17/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664022 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Ectopic Heartbeat | Heart Block | [6 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with shortness of breath was taken to the cardiac catheterization laboratory (ccl) for anterior st-segment elevation myocardial infarction per electrocardiogram. ectopy was noted upon introduction of the pulmonary artery catheter for right heart catheterization. the patient was transferred. however, the patient immediately decompensated. complete heart block followed by ventricular tachycardia (vt) was noted. the patient was emergently taken back to lab and defibrillated again. the decision was made to place the impella cp in the setting of acute myocardial infarction shock. during support, it was reported there were rhythm changes including trigeminy, bigeminy, andvt runs. electrolytes were stable and amiodarone bolus given. the patient was mildly hypotensive. levophed was added. however, the patient did not respond. the patient was supported with vasopressin and 250 albumin bolus was given as well. it was also noted of concerns for hemolysis with lactate dehydrogenase in the 1800s and hemoglobin drop to 8.6. after some manual examination to patient¿s abdomen and groin site it became oozy and the impella position in aorta alarm noted. a stat echocardiogram showed the pigtail still across. heparin was paused and femstop was applied. one unit of packed red blood cells was given. the patient was taken urgently to the ccl where the device was successfully repositioned under fluoroscopy guidance. higher p-levels were trialed to assess for any position change but remained appropriate. the patient was changed back to p-4 to mitigate risk of further hemolysis. the impella cp was successfully weaned and explanted. | [5 empty] | The investigation of access site bleeding, positioning issues, vf/vt/af/at, and hemolysis has been completed. the impella device was not returned for evaluation. access site bleed: based on clinical description, the cause of access site bleeding was most likely use related while physician was examining the patient's groin. positioning issue: the cause of positioning issue was most likely use related since pump pulled out while physician was examining the patient's groin. hemolysis: data logs confirmed pump was in improper position corresponded with the time hemolysis and pump in aorta were reported per clinical description that triggered alarms impella position in aorta. pump then was in proper position & no issues were found on the logs to the rest of the case. the root cause of hemolysis was most likely due to pump was not in the proper position since hemolysis was improved after positioning of the pump. vt/vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. (b)(6). | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6920 | 1220648-2025-30895 | 22677074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677074&pc=OZD | 07/18/2025 | 09/16/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553733 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 66 KG | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being prepped for an impella cp implant and during the preparation, an alarm reading purge system open continued to be active. efforts were made to de air the system and cassettes were changed without success. once de airing the line, the automated impella controller (aic) advanced directly to the placement screen, bypassing the calibration screen. the aic was replaced to no avail. the impella cp was aborted leading to prep and implant of an additional impella cp. there was no patient harm. | [5 empty] | The investigation of priming issue has been completed. the pump and two purge cassettes were returned for investigation. according to the clinical report, the troubleshooting method did not resolve the purge system open alarm, and the pump was replaced. no physical abnormalities were observed on the returned products. the pump was successfully primed using both returned purge cassettes without issue. purge values were within specification limits during testing. data log shows a purge system open alarm during case start, which resolved a few minutes after multiple case start attempts were reported. full logs of the case start were not available for review. there was no issue found during the case. the device functioned/operated to specification. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6921 | 1220648-2025-30896 | 22677086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677086&pc=OZD | 07/21/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/21/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693831 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Failure to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed for the st-segment elevation myocardial infarction patient. the pump was placed but there was low flow and kinking observed. the pump was removed from the left ventricle (lv) and then flushed and re-delivered to the lv. team had to use a new .018 wire, a v18, as the original wire had kinked. this was known to be against recommendations but the pump was placed again to the lv. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6922 | 1220648-2025-30897 | 22677231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677231&pc=OZD | 07/18/2025 | 09/15/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661860 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Malposition of Device | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal alarm that the pump was positioned deep. the doctor did not want to adjust the depth and was left the pump as is. care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. data review: data logs confirmed pump was in ventricle near the end of the case. pump then was stopped manually and disconnected from the console. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-30897.
6923 | 1220648-2025-30898 | 22677266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677266&pc=OZD | 07/18/2025 | 09/17/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1457787 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 66 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being prepped for an impella cp implant and during the preparation, an alarm reading purge system open continued to be active. efforts were made to de air the system and cassettes were changed without success. once de airing the line, the automated impella controller (aic) was advanced directly to the placement screen, bypassing the calibration screen. the aic was replaced to no avail. the impella cp was aborted leading to prep and implant of an additional impella cp. | [5 empty] | The investigation for the reported aic-purge issue-other has been completed. the logs indicate that a "purge system open (#407) alarm" was triggered shortly after the calibration was completed, and the alarm remained uncleared until the aic was shut down. initially, the case start was reverted to step 6 since the pump had not yet been started. subsequently, the pump was disconnected and then reconnected. the purge cassette was also removed and reinserted; however, the alarm persisted. eventually, the case start wizard was canceled and exited, at which point the alarm was cleared simultaneously with the exit of the wizard. the aic was then shut down. the root cause for the purge system open #407 alarm could not be determined as aic was not returned. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6924 | 1220648-2025-30899 | 22677360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677360&pc=OZD | 07/17/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  | 02/28/2026 |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471282 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation that resolved without treatment. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 and a5, ethnicity, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6925 | 1220648-2025-30900 | 22677450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677450&pc=OZD | 07/19/2025 | 09/24/2025 |  | 07/31/2025 |  |  | 07/19/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635967 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male | 80 KG |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a high purge pressure alarm. no kinks or blockages were visible in the purge line. lysis was done and the impella ran without issues. the patient survived. | [5 empty] | The investigation of high purge pressure (hpp) has been completed. the impella device was not returned for evaluation. data log review showed purge pressure build up within 1 hour with corresponding purge flow decrease. after 24 hours pp high alarms ceased and purge metrics returned to an acceptable range and returned to baseline within 72 hours. the cause of the high purge pressure was most likely biomaterial buildup due to purge pressures increasing and purge flows decreasing over the course of 1 hour. purge issues resolved after lysis protocol was executed. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6926 | 1220648-2025-30901 | 22677655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677655&pc=OZD | 07/19/2025 | 12/03/2025 |  | 07/31/2025 |  |  | 07/19/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885741 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 120 KG | Non Hispanic | White | Malfunction | No | Yes | Temperature Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during support of an implanted impella 5.5 the automated impella controller (aic) became warm to the touch. the controller alarmed for high temperature but remained on for support. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation was completed. data analysis: console logs from the complaint date revealed the ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) that self-resolved in 10 seconds. there was no preceding battery temperature alarm #48, which occurs when battery temperature is between 50degc and 60degc. battery and power data embedded in lt logs do not show any temperature rise or elevated battery temperatures, as lt log data is recorded every 60s. device analysis: the console ic12204 was returned to field service for further investigation. the reported complaint could not be reproduced. the console¿s batteries and power management system were tested, and no issues were observed. root cause: the root cause of the battery temperature high alarm (#49, due to battery temperature over 60degc) was due to a momentary communication glitch between the batteries and pbm, where a single sample from battery data (in this case value of battery temperature) was corrupt. h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation.
6927 | 1220648-2025-30902 | 22677656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677656&pc=OZD | 07/18/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682719 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 53 KG |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp could not be implanted due to the start up sequence not initiating. a second impella cp was subsequently implanted for planned support. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6928 | 1220648-2025-30903 | 22677673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677673&pc=OZD | 07/12/2025 | 09/09/2025 |  | 07/31/2025 |  |  | 07/12/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686413 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had oozing at the access site. pressure dressing was applied and bivalirudin was on hold. additionally, the pump flow became saturated. the patient decompensated and the impella 5.5 was explanted and replaced with extracorporeal membrane oxygenation. upon explant, there was a thrombus noted on the outside of inflow. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of saturated flow, thrombus, and access site bleeding has been completed. the impella device was not returned for evaluation. saturated flow: data logs show on last day of support, a small motor current spike and speed deviation with s slight drop in pump flow. 8 hours after first motor current spike, a second larger motor current spike with speed deviation and fully saturated flows. clinical details noted saturated flows when patient was decompensating prior to pump explant. additionally, thrombus was present on pump on explant. the cause of the saturated flow was most likely due to biomaterial based on returned data log analysis and clinical details. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. access site bleeding: the cause of the access site bleeding could not be definitively determined as limited clinical details were provided.
6929 | 1220648-2025-30904 | 22677725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677725&pc=OZD | 07/19/2025 | 09/12/2025 |  | 07/31/2025 |  |  | 07/19/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687335 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had oozing at the access site. angle matching was performed and one unit of blood products was transfused. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
6930 | 1220648-2025-30905 | 22677945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677945&pc=OZD | 07/18/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689414 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and had a hematoma at the access site. pressure was held and two units of blood products were given. additionally, the patient had a retroperitoneal hematoma and was transfused two units of blood. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6931 | 1220648-2025-30906 | 22677975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22677975&pc=OZD | 07/19/2025 | 10/02/2025 |  | 07/31/2025 |  |  | 07/19/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684350 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 89 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Gastrointestinal Hemorrhage | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a large hematoma on the lateral thorax/axilla. the hematoma was evacuated in the operating room and the patient was transfused five units of packed red blood cells and two units of fresh frozen plasma. the purge was switched from heparin to sodium bicarb. the following day, the patient continued to bleed and a gastrointestinal bleed was identified. seven units of packed red blood cells were transfused. the patient went into pulseless electrical activity, and the family decided to withdraw care. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was most likely use issue wrong composition of purge was administered to the patient most likely leading to the bleed. vascular damage peripheral vessel: patient had gi bleed. the root cause of vascular damage gi bleed cannot be determined since the product was not returned and insufficient clinical details were provided.
6932 | 1220648-2025-30907 | 22678070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22678070&pc=OZD | 07/18/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692823 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 65 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 developed a gastrointestinal bleed. heparin was discontinued. impella 5.5 support continues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6933 | 1220648-2025-30908 | 22678510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22678510&pc=OZD | 07/21/2025 | 07/31/2025 |  | 07/31/2025 |  |  | 07/21/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593580 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon the patient arriving at icu after impella cp placement, hematoma and bleeding noted at insertion site. manual pressure was held for the duration of 1 hour. heparin and angiomax drip were held and one unit of prbc was administered. the bleeding was resolved. later weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6934 | 1220648-2025-30909 | 22679337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22679337&pc=OZD | 07/18/2025 | 09/11/2025 |  | 07/31/2025 |  |  | 07/18/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673977 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 80 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and had bleeding from the access site due to excessive movement. albumin was given and a femstop was applied with a sandbag. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement since patient was moving excessively and bleeding from the impella site.
6935 | 1220648-2025-30910 | 22679447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22679447&pc=OZD | 07/19/2025 | 10/01/2025 |  | 07/31/2025 |  |  | 07/19/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1705772 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA 5.5 SERIAL NUMBER (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal error alarm appeared on the automated impella controller (aic) during impella 5.5 support. the alarm self resolved. ultimately, care was withdrawn and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | A.2, a.4, and a.5 were left blank as patient information is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console and data logs were received for investigation. the root cause for the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. d9 device returned to manufacture date added.
6936 | 1220648-2025-45705 | 22679867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22679867&pc=OZD | 06/28/2025 | 07/30/2025 |  | 07/31/2025 |  |  | 06/28/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665426 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Life Threatening | Death | Required Intervention | Hospitalization | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient was on aggrastat for percutaneous coronary intervention and therefore, hemostasis took a little longer to achieve. a jackson-pratt bulb was placed to assess for bleeding and a unit of packed red blood cells (prbc) were administered. the patient was transferred to the unit on support and a time thereafter, the patient was administered two additional units of prbcs with platelets, cryoprecipitate, and desmopressin which achieved hemostasis with only 120ml of drain output. the patient, previously placed on extracorporeal membrane oxygenation (ecmo) and continuous renal replacement therapy, continued on impella mcs. approximately five days later, the patient was slowing improving and was taken off ecmo. the patient was still on dobutamine but all other inopressors were off. two days later, it was noted the patient remained intubated and off sedation with inconsistent responses. foley catheter was removed; patient remained on crrt. it was further noted a nasal gastric tube was placed and caused significant nosebleed requiring two units of prbcs for a hemoglobin of 6.7. dobutamine was attempted to be weaned; however, the patient was hypotensive. aggrastat was changed out for brilinta and heparin was on hold due to the nosebleed. the following day, the patient started coughing and brady down and then ventricular tachycardia arrested. cardiopulmonary resuscitation was administered for two minutes with 100mg of lidocaine and defibrillator x200 joules. the patient was converted back to normal sinus rhythm. later this day, the physician encountered difficulty with the left ventricular (lv) placement signal and the device¿s position. the adjust lv signal was reperformed multiple times by the physician making his own adjustments. the care team informed of possible issues related to the adjustments and need to monitor patient hemodynamics and imaging for position. two days later, a temporary pacer catheter was placed due to episodes of bradycardia and extended pauses in rhythm. the patient experienced a short run of ventricular tachycardia with no intervention required. however, it was noted the patient would develop atrial fibrillation with rapid ventricular response. days following, the patient was noted to be doing well. the patient was extubated, sitting up and worked on strengthening activities. however later that night, the patient aspirated after vomiting and required emergent reintubation. the patient would continue impella 5.5 mcs and now a couple of days later, the patient developed a ¿massive¿ gastrointestinal bleed (mallory-weiss tear). heparin was stopped. the gastrointestinal bleed caused the patient to go into cardiac arrest, and the decision was made by the family to withdraw care. the patient subsequently expired. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not returned and therefore, an evaluation/analysis of the device was not possible. device history review was completed, and the pump passed all post-sterile inspection checks. adverse events: ventricular tachycardia, bradycardia, epistaxis, gastrointestinal bleed, and access site adverse event (bleeding), in order to make a root cause determination, adequate clinical details would have to be provided. the root cause of these adverse events was unable to be determined due to insufficient clinical details. pertaining to the device malfunction, event logs were received and analyzed. pertaining to optical sensor issue, the long term event log was analyzed and showed no major issues with 5s optical signal. all variations were in relevance to p-level. left ventricle offset applied for a diff of > 40 mmhg (from -25 to ~ +25 mmhg). the placement signal observed to decline consistently, and several placement signal low alarms were observed, and later, the placement signal monitoring was disabled. real time event log showed no major abnormalities with 5s parameters. the placement signal was starting to downtrend after an applied left ventricular offset. the cause of optical signal issue was not determined due to inconclusive evidence per log analysis, clinical and without return product for analysis. regarding the device in wrong position, the reported information noted the position was confirmed at 5 centimeters deep in left ventricle. however, the left ventricle outflow tract was noted to be narrow, and the device was lying on septum. the positioning issue coincided with early ventricular fibrillation/ventricular tachycardia, and bradycardia. however, the cause of position issue was not determined due to inconclusive evidence per clinical and since product was not returned for analysis. h.6 codes were added accordingly based on the above-noted. e.1 zip code extension was inadvertently omitted from the initial report submitted and was added.
6937 | 1220648-2025-45709 | 22680502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22680502&pc=OZD | 07/27/2025 | 07/29/2025 |  | 07/31/2025 |  |  | 07/27/2025 | 07/31/2025 |  |  |  |  | 07/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698038 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced severe limb ischemia in the left leg where the impella cp was inserted. a vascular surgeon placed a 6 french distal reperfusion sheath and fasciotomy as performed. the patient was then sent to the icu in critical condition. it was additionally reported that the impella was found to be slightly deep at 5.7 cm, however, the physician did not want to reposition as the patient experienced long runs of ventricular tachycardia when the pump had been repositioned in the cardiac catheterization laboratory. on (b)(6) 2025 the decision was made to escalate the patient to an impella 5.5 for increased support. the patient was successfully escalated to the impella 5.5 and was noted to be in stable condition following implant of the impella 5.5, however the following day the decision was made to withdraw care. the patient¿s care was withdrawn, and the patient expired. there is not enough evidence to exclude the impella cp device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of the ischemia and ventricular tachycardia was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks.
6938 | 1220648-2025-30840 | 22663111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663111&pc=OZD | 07/06/2025 | 09/15/2025 |  | 07/30/2025 |  |  | 07/06/2025 | 07/30/2025 |  |  | 07/31/2025 |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024383463 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump experienced low purge pressure coinciding with a fluid leak at the purge filter during patient support. a purge bypass option was discussed with the md. however the decision was ultimately made to explant the pump. patient survived. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. a.5.: unknown. | The investigation of purge leak ¿ pump has been completed. pump was returned for investigation. data log review was not required for this case run. the returned pump was primed, and the purge leak was observed from the sidearm filter. upon close examination, the sidearm was found cracked where the leak was observed. the cracks are most likely consistent with chemical interaction. the cause of the purge leak was sidearm filter damage. instructions for use: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30840.
6939 | 1220648-2025-30841 | 22663164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663164&pc=OZD | 07/08/2025 | 09/23/2025 |  | 07/30/2025 |  |  | 07/08/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687763 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella implant and during the implant, the peel away sheath could not be placed correctly due to the patient anatomy. there was no patient harm. the implant was aborted. patient survived. | [5 empty] | The investigation of difficulty inserting/removing introducer has been completed. difficulty inserting introducer: clinical details report that the introducer insertion procedure was aborted due to the patient's anatomy. there was no harm to the patient. returned product was investigated and the 14 fr x 13 cm introducer had both a sheath kink and damage to the tip. based on the clinical details provided that the patient's anatomy was difficult and the damage noted on returned product, the cause of the difficulty inserting introducer was determined to be patient condition. introducer damage - mechanical: based on the findings with returned product, the failure mode (fm) introducer damage was added. clinical details reported that the patient anatomy prevented the user from being able to insert the introducer. returned product had damage/deformation to the tip. based on the clinical details provided that the patient's anatomy was difficult and the damage noted on returned product, the cause of the introducer damage was determined to be a sheath tip split. product damage (sheath kink): based on the findings with returned product, the fm product damage (sheath kink) was added. clinical details reported that the patient anatomy prevented the user from being able to insert the introducer. returned product had a kink in the sheath. based on the clinical details provided that the patient's anatomy was difficult and the damage noted on returned product, the cause of the sheath kink was determined to be patient condition. d9 device returned to manufacturer date added d6a implant date was erroneously entered on the initial manufacturer device report number 1220648-2025-30841. h6 medical device problem code 2682 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30841. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿.
6940 | 1220648-2025-30842 | 22663205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663205&pc=OZD | 07/08/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/08/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684324 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient developed hematoma at the impella access site. the next day, the hematoma appeared to be getting larger. the patient will be taken to get drain placed at hematoma site. bival glucose tolerance test was stopped. the pump was explanted three days later. | [5 empty] | The investigation for the access site bleed and thrombocytopenia has been completed. the device was not returned for evaluation. without sufficient clinical details nor the device return, the root cause of the access site bleed and the thrombocytopenia could not be determined. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation)."
6941 | 1220648-2025-30843 | 22663301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663301&pc=OZD | 02/03/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/09/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 75 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 75-year-old patient had bleeding at the access site. the first attempt to stop the bleeding was through the pull-up stitching technique with gauzes, but it was unsuccessful. then, we used local tranexamic acid at the site of the repositioning sheath and partially locked the knots of the two proglides. nevertheless, external manual compression was the only way to limit the bleeding. a femstop was applied. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ missing details follow this example: a.1- a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
6942 | 1220648-2025-30844 | 22663476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663476&pc=OZD | 05/27/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/16/2025 | 07/30/2025 |  |  | 01/31/2025 |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371073 |  | 004413 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  | Asian | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team was notified by post market of an expired pump being used. this was discovered in prep for pqss, "during our review of this month's pqss, we identified a usage reported for the impella 2.5 associated with case number 01468673 and serial number (b)(6) for the account of veiva scientific india private limi in india for the month of may." no known patient harm or injury. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6943 | 1220648-2025-30845 | 22663565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663565&pc=OZD | 11/05/2024 | 07/30/2025 |  | 07/30/2025 |  |  | 07/11/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509882 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a 5.5 pump for support in (b)(6) 2024 for the patient admitted in with cardiomyopathy and other systemic comorbidities with plan to admit for a coronary artery bypass graft . the 5.5 supported for 2 days and was successfully weaned and explanted. later in (b)(6) 2025 the impact study reported upon the adverse event of anemia. the anemia prompted transfusion. the thought was that the bleed and anemia could have been from perioperative bleeding or graft site anastamosis. the relationship is deemed "possible" per the impact study. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6944 | 1220648-2025-30761 | 22663699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663699&pc=OZD | 06/20/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/09/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female | 124 KG | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella 5.5 pump placed to support a coronary artery bypass graft. when the support was ongoing and on day two, the pump was being repositioned and the team in error unplugged the impella 5.5 to relieve the cord torque and in doing this the pump stopped and could not be restarted. the automated impella controller was exchanged and multiple attempts to troubleshoot were made. the patient's outcome is stable. | [5 empty] | The investigation has been completed. the console logs were consistent with what was reported in the complaint. the console was returned and tested. the reported pump stop was unable to be reproduced while testing the console with a known good purge cassette kit and pump on an engineering lab bench. the reported pump stop was a normal response to the pump being disconnected while it was running. the root cause of the reported pump stop was use related. however, the console had issues detecting the pump following this disconnection. the pump detection issue was likely due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a software mitigation in place. the reported pump stop was an expected response to the user disconnecting the pump.
6945 | 1220648-2025-30846 | 22663712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663712&pc=OZD | 07/12/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/12/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665933 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impela 5.5 support had ventricular tachycardia run, echocardiogram was done at bedside and impella was repositioned. it was further reported that the patient was heparin-induced thrombocytopenia positive and was started on argatroban. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
6946 | 1220648-2025-30847 | 22663783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663783&pc=OZD | 07/15/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/15/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682728 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 66 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella cp placed via the femoral to support with acute myocardial infarction complicated by cardiogenic shock. the patient had an ooze that prompted the team to tighten the perclose. the pump remained on for support. the patient survived the procedure and is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6947 | 1220648-2025-30848 | 22663906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663906&pc=OZD | 07/15/2025 | 09/10/2025 |  | 07/30/2025 |  |  | 07/15/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that had a cp pump placed to support the patient admitted in with cardiogenic shock and cardiomyopathy. during the support there were signs of hemolysis the team tried to troubleshoot away and treat. the decision was made to withdraw all care and patient expired after 63 hours of support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation has been completed. hemolysis: complaint device not returned for analyzing. data log reviewed, no mc spike outside of p-level changes observed. few quick resolved impella positioning in ventricle and placement low alarms seen of first day of support witch resolved after impella repositioning and pulled back impella to 3.4 cm from av. no suction occurred. on 7/14 downward trending ps observed. clinical mentioned, brown color urine and evidence of hemolysis observed as ldh 3040, all lfts elevated. cr up to 2.6, serum k 6.1. on 7/16 urine was yellow indicating hemolysis resolved and reported that the repositioning resolved the hemolysis on 7/15. the cause of the hemolysis was most likely improper positioning of the device since repositioning resolved the hemolysis. positioning issues: fm positioning issues was added because it was reported that repositioning resolved the hemolysis injury. the cause of the positioning issue was not determined due to insufficient clinical details.
6948 | 1220648-2025-30849 | 22663913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663913&pc=OZD | 07/16/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/16/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681696 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient endured supraventricular tachycardia during impella cp support. the patient was cardioverted twice and vasopressors were increased in response to the condition. support continued uninterrupted following the intervention. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6949 | 1220648-2025-30850 | 22663999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22663999&pc=OZD | 07/15/2025 | 09/10/2025 |  | 07/30/2025 |  |  | 07/15/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690942 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, the patient had ventricular tachycardia. the impella was found deep and was pulled back. no further ectopy was noted. additionally, high purge pressure was noted. tissue plasma activator was added to the purge. the purge cassette was replaced and the issue resolved. | [5 empty] | The investigation of vf/vt/af/at, positioning issues, and high purge pressure (hpp) has been completed. the impella device was not returned for evaluation. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. positioning issues: the cause of the positioning issues was not determined since product was not returned and insufficient clinical details. high purge pressure: data log review showed hpp and purge blocked alarms occurring after 6 hours after case start. purge pressure increase and flow decrease with no motor current (mc) spike indicate biomaterial buildup. no bag change occurred when purge pressures began to climb. after 7 hours of alarms, a bag change procedure resolved the purge alarms and both purge pressure and flows begin to return to normal. the cause of the high purge pressure was most likely biomaterial buildup due to data log review indicating biomaterial build up the resolved with a bag change around the time clinical details state tissue plasminogen activator (tpa) was added to the purge. d6b explant date added. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
6950 | 1220648-2025-30851 | 22664050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664050&pc=OZD | 07/17/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/17/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667330 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and had a gastrointestinal bleed. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6951 | 1220648-2025-30852 | 22664168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664168&pc=OZD | 07/15/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/23/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025597573 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on extracorporeal membrane oxygenation with ejection fraction of 10%, was placed on impella cp. the impella was inserted from the left subclavian and extracorporeal membrane oxygenation with impella (ecpella) was administered. during support it was reported that the patient had a cerebral hemorrhage and that due to a poor neurological prognosis of the cerebral hemorrhage, it was decided not to undergo additional treatment, and was treated for multiple organ failure. the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6952 | 1220648-2025-30853 | 22664186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664186&pc=OZD | 07/17/2025 | 09/08/2025 |  | 07/30/2025 |  |  | 07/17/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025602622 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female | 165 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted as the pump was unable to cross due to the patient's anatomy. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the pump could not be inserted due to patient's anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6953 | 1220648-2025-30855 | 22664384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664384&pc=OZD | 07/13/2025 | 09/10/2025 |  | 07/30/2025 |  |  | 07/13/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696166 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Thrombosis/Thrombus | Gastrointestinal Hemorrhage | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient developed hemolysis. the pump migrated too deep and was repositioned. staff monitored patient's labs. heparin was eventually switched to bival which helped the patient's hemolysis. additionally, the patient had an extensive thrombus in transverse arch into descending. heparin was switched to bival and the staff monitored. furthermore, the patient had a gastrointestinal bleed. later care was withdrawn and the patient expired. | [5 empty] | The investigation of thrombus, hemolysis, positioning issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. hemolysis: the cause of the hemolysis was most likely due to improper position based on clinical details noting pump was deep. positioning issues: the cause of the positioning issues was most likely due to patient movement as clinical details noted that patient was extremely agitated and was difficult to restrain. vascular damage - peripheral vessel: no additional details provided and no intervention noted. the cause of the vascular damage could not be definitively determined as limited clinical details were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6954 | 1220648-2025-30854 | 22664390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664390&pc=OZD | 07/17/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/17/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653457 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 and during support the anticontamination sleeve broke. tegaderm was placed. the patient survived and was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6955 | 1220648-2025-30856 | 22664468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664468&pc=OZD | 06/23/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/17/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661724 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured a right axillary hematoma from the impella site with a "possible" relationship to the impella device used: ¿baseline h/h unknown. nadir h/h= 6.8/20.5 (on (b)(6) 2025 at 0334). ct chest a and p on (b)(6) 2025: pleural effusion. evolving hematoma 4x cm at impella site. washout on (b)(6) 2025 w/o gross evidence of infx. axillary dacron graft trimmed to short length and sewed over. graft and wound c/s sent. wound vac placed at 125 mm hg suction.¿ additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right brachial thrombus with a "possible" relationship to the impella device used: ¿pt had right brachial thrombus found intra op during impella removal. thrombectomy performed. ulnar signal at conclusion of case. per vascular no additional anticoagulation is needed.¿ the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
6956 | 1220648-2025-30857 | 22664644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664644&pc=OZD | 07/15/2025 | 11/18/2025 |  | 07/30/2025 |  |  | 07/15/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261099 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had a "system check failed" error and to change consoles. this issue occurred during start up. there was no patient harm reported. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation has been completed. data analysis: the console (ic4664) logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to this communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. d9 updated. h3 updated.
6957 | 1220648-2025-30858 | 22664658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22664658&pc=OZD | 07/17/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/17/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680976 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 88 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Disseminated Intravascular Coagulation (DIC) | Hematoma | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and developed a hematoma and disseminated intravascular coagulation. the patient was transfused with packed red blood cells, platelets, and cryoprecipitate. the patient was taken to surgery for a hematoma evacuation and became hypoxic during transport. additionally, the patient had ventricular tachycardia. amiodarone was given. impella 5.5 support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6958 | 1220648-2025-30859 | 22665121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22665121&pc=OZD | 07/11/2025 | 09/16/2025 |  | 07/30/2025 |  |  | 07/15/2025 | 07/30/2025 |  |  | 01/31/2026 |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025529037 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump experienced high purge pressure issues throughout the course of patient support. the purge cassette was replaced, but the issue remained. the pump current continued to be monitored as a result of the noted issue. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressure has been completed. no data logs were returned for analysis. thrombus fm added as evidenced by returned product. the returned pump was inspected and biomaterial was observed in the purge gap and by impeller. the pump was run in the lab after clearing and high purge pressure did not reproduce. there were no other pump abnormalities. the root cause of the high purge pressure was most likely biomaterial as evidenced by the thrombus blocking the purge gap on returned product and issue not reproducing after clearing. the root cause of the thrombus was unable to be determined due to insufficient clinical details. corrections to the initial mfr report: d9 device return date has been added. h3 device evaluation updated to yes. h6 clinical code 4440 added.
6959 | 1220648-2025-30860 | 22665230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22665230&pc=OZD | 07/17/2025 | 09/24/2025 |  | 07/30/2025 |  |  | 07/17/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675144 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and a perforation occurred. it was unknown if the perforation was in the right or left ventricle. it was also unknown what caused the perforation. according to the physicians it could happen during insertion of temporary pacemaker electrode transcatheter aortic valve implantation. , multiple guidewires exchanged. the patient received a transfusion. patient survived | [5 empty] | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
6960 | 1220648-2025-30861 | 22665603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22665603&pc=OZD | 07/18/2025 | 12/02/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1452267 |  | 0042-0010-EU |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that during prepping for impella cp placement and when flushing the pump, no liquid came out and there was an error message. the purge cassette was changed but the same problem occurred. the cassette was exchanged again and the console, but the issue was not resolved. another new console and a new purge cassette were used and this is worked successfully. there was no patient harm. | [5 empty] | D4 serial number and unique identifier (udi) number was erroneously entered on the initial manufacturer device report number 1220648-2025-30861. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-30861. h8 was erroneously selected on the initial manufacturer device report number 1220648-2025-30861. | The investigation for the reported issue has been completed. analysis: on the complaint date of 7/18, uninitiated pump 608100 is connected to (b)(4). even before the pump is connected, there are purge fluid not detected errors received by the user at which point the user reboots (b)(4). during this time, there is no purge pressure despite the purge flow ticks incrementing, which is what triggered the error to the user. the console is booted back up with the purge cassette connected but no pump is connected. on cases with (b)(4) after the complaint case, the console successfully purged and the customer cleaned the purge assembly. pump 608100 is moved over to (b)(4) still with cassette 61929 and got the same purge fluid not detected error. new cassette 92989 is inserted but the purge fluid not detected error remained. analysis and testing of both purge did not indicate any malfunction, and showed no evidence of contamination. root cause: the cause of the purge fluid not detected errors was most likely use-related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The medical safety officer reviewed and determined the following: review of the automated impella controller (aic) at bedside demonstrated that "heavily clogged with glucose that the conveyor wheel could not run." cleaning these conveyor wheels resulted in the operation of the aics. these acts were performed prior to start of support and did not contribute to the outcome of the patient's demise. the demise outcome is captured under manufacture device report 1220648-2025-31630 for impella cp pump.
6961 | 1220648-2025-30862 | 22666039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22666039&pc=OZD | 07/18/2025 | 12/03/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1452267 |  | 0042-0010-EU |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that during prepping for impella cp placement and when flushing the pump, no liquid came out and there was an error message. the purge cassette was changed but the same problem occurred. the cassette was exchanged again and the console, but the issue was not resolved. another new console and a new purge cassette were used and this is worked successfully. the detailed examination afterwards revealed that both consoles were heavily stuck to the wheel of the purge cassette. after intensive cleaning and retesting, the error could not be reproduced. the fault was therefore rectified. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported issue has been completed. analysis: on the complaint date of 7/18, uninitiated pump 608100 is connected to (b)(6). even before the pump is connected, there are purge fluid not detected errors received by the user at which point the user reboots (b)(6). during this time, there is no purge pressure despite the purge flow ticks incrementing, which is what triggered the error to the user. the console is booted back up with the purge cassette connected but no pump is connected. this was all performed with purge cassette sn(b)(6). pump 608100 is moved over to (b)(6) still with cassette (b)(6) and got the same purge fluid not detected error with very low purge pressure. new cassette (b)(6) is inserted but the purge fluid not detected error remained. there was an alleged third console used which did not malfunction but its serial number could not be determined. analysis and testing of both purge cassettes did not indicate any malfunction and showed no evidence of contamination. root cause: the cause of the purge fluid not detected errors was most likely use-related. | D.4. serial number and unique identifier (udi) number was erroneously entered on the initial manufacturer device report number 1220648-2025-30862. h.5. was erroneously selected on the initial manufacturer device report number 1220648-2025-30862. h.6. was erroneously selected on the initial manufacturer device report number 1220648-2025-30862. | The medical safety officer reviewed and determined the following: review of the automated impella controller (aic) at bedside demonstrated that "heavily clogged with glucose that the conveyor wheel could not run." cleaning these conveyor wheels resulted in the operation of the aics. these acts were performed prior to start of support and did not contribute to the outcome of the patient's demise. the demise outcome is captured under manufacture device report 1220648-2025-31630 for impella cp pump.
6962 | 1220648-2025-30863 | 22666477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22666477&pc=OZD | 07/18/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella support and during the explant, patient arterial waveform became diminished on the same side of 5.5 implant. it was then discovered a thrombosis had traveled down the patient's left arm and had obstructed flow. vascular surgery was consulted. vascular surgeon then restored flow to left arm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
6963 | 1220648-2025-30864 | 22666584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22666584&pc=OZD | 07/18/2025 | 09/05/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller displayed a controller error alarm when it was turned on. a second controller was used to successfully support the patient. no harm to the patient was reported. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: when console ic11400 was booted up on 7/18, there was an immediate controller error (defective optical sensor lamp) alarm issued. the 5s triggered by the bad lamp had a voltage of 0.47v and 0.48v which is the value for a lamp outputting no light. device analysis: the failure mode was not reproduced during bench top testing. 5s testing showed the light outputting less than 1v and the lamp was visibly outputting light. there was no evidence for why the failure mode occurred on one boot up in the field. root cause: the cause of the controller error was an optical lamp outputting no light. the cause of the optical lamp not outputting light intermittently was most likely a defective lamp. d9 device returned to manufacturer date added. | Patient information was not provided. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6964 | 1220648-2025-30865 | 22666631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22666631&pc=OZD | 07/18/2025 | 09/26/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627416 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 62 KG | Hispanic |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a loss of placement signal. the patient's procedure was completed successfully. the patient was weaned from the pump and survived. | [5 empty] | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. the cause of the optical signal issue issue was most likely a damaged sensor due to placement signal (ps) railing to 3000mmhg, %s parameter signature in the data logs, and clinical information stating shockwave was used in the pci procedure. the distance between the shockwave and sensor at the time of failure is not known. | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6965 | 1220648-2025-30866 | 22666744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22666744&pc=OZD | 07/18/2025 | 09/16/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075036 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) displayed a controller error alarm during impella 5.5 support. the aic was swapped and support continued without issue. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. based on the device and data analysis the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" was due to a firmware issue (optical bench fw anomaly). h8 usage of device was erroneously selected on the initial manufacturer device report number 1220648-2025-30866. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6966 | 1220648-2025-30867 | 22667051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667051&pc=OZD | 07/18/2025 | 09/16/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673556 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient admitted in with third degree heart block, seizures, no pulse, and ongoing cardiopulmonary resuscitation had an impella cp implanted. however, the pump flow was not appropriate and the medical team made the decision to withdraw care. the patient expired. | [5 empty] | The investigation for the reported low pump flow has been completed. clinical reports despite proper positioning and volume, the pump would not achieve good flow levels. the product was not returned for investigation. data log review shows minor variation in placement signal at the time the pump was started, as well a #14 suction alarm, but no motor current spikes outside p-level changes. after the suction alarm occurred, flow levels varied significantly and decreased. the cause of the low pump flow was not determined due to lack of product returned and insufficient clinical details. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
6967 | 1220648-2025-30868 | 22667130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667130&pc=OZD | 07/18/2025 | 09/11/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568033 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 63 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted as the pump could not pass the patient's anatomy. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since the impella was unable to pass through the vasculature. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
6968 | 1220648-2025-30869 | 22667291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667291&pc=PYX | 06/09/2025 | 08/06/2025 |  | 07/30/2025 |  |  | 06/09/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026655490 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella rp flex support and during support, the patient had hemolysis. patient's urine looked hemolyzed with lactate dehydrogenase (ldh) 1800. the pump position was confirmed on echo however labs were still hemolyzing. patient survived. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6969 | 1220648-2025-30870 | 22667296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667296&pc=OZD | 07/18/2025 | 09/16/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682711 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Male | 61 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and had oozing at the access site and hematuria. forward tension sutures with angle matching were applied, heparin was withheld, and a femstop was placed to obtain hemostasis. the patient was later successfully weaned from the pump and survived. | [5 empty] | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. data review: data logs showed pump ran without issue during support. placement signal ps trending down throughout. product was not returned for review. the root cause of access site bleeding was most likely anticoagulation management related. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6970 | 1220648-2025-30871 | 22667369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667369&pc=OZD | 07/18/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672006 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 133 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and had oozing at the axillary pocket that was repaired with stitches and surgicel. additionally, the patient had ectopy that was treated with lidocaine. the patient was later successfully weaned from the pump and survived. additional information was received. this occurred intra procedure and the patient did not have bleeding in the axillary upon closure and discharge to the intensive care unit. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6971 | 1220648-2025-30872 | 22667630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667630&pc=OZD | 05/25/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 05/25/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655960 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 97 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient endured ventricular tachycardia during pump support. later, care was withdrawn and the patient expired. | [5 empty] | To determine the true root cause of the ventricular tachycardia the clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the ventricular tachycardia could not determined.
6972 | 1220648-2025-30873 | 22667699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667699&pc=OZD | 07/18/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657734 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient underwent coronary artery bypass graft, aortic valve replacement, maze. the patient required cardiac massage due to pre-bypass ventricular fibrillation event. intra-aortic balloon pump was placed to right femoral. due to bi-ventricular failure and unable to separate from cardio-pulmonary bypass, the patient underwent right axillary impella 5.5 placement with al1 and straight wire under transesophageal echocardiogram guidance. there was a left ventricle perforation noticed by surgeon and primarily repaired with single stitch. the patient subsequently unable to separate from cardio-pulmonary bypass with impella and expired on the table. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
6973 | 1220648-2025-30874 | 22667805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22667805&pc=OZD | 07/18/2025 | 09/17/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  |  |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700777 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 66 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to patient anatomy. the patient was implanted with an impella cp instead. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because despite multiple implantation with 018 wire and then changing to an 035 advantage glide, the 5.5 was aborted due to patient anatomy. d6a and d6b were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30874. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
6974 | 1220648-2025-30875 | 22668158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22668158&pc=OZD | 07/18/2025 | 07/30/2025 |  | 07/30/2025 |  |  | 07/18/2025 | 07/30/2025 |  |  | 10/31/2025 |  | 07/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410032 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 104 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp had little to no flow down the leg. the plan was to remain on impella over the weekend, and the staff was monitoring the distal pulses, but the family decided to withdraw care. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation."
6975 | 1220648-2025-30796 | 22641858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22641858&pc=OZD | 07/15/2025 | 09/24/2025 |  | 07/29/2025 |  |  | 07/15/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667454 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a cardiogenic shock patient who the day prior had two valves replaced. the pump was supporting when due to low flows and a reported cannula kink, the pump was removed and replaced. no harm was noted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (cannula kink) and low pump flow has been completed. the impella device was returned for evaluation low pump flow: return product analysis showed cannula kink. no other abnormalities were observed. no data logs were returned. the root cause of low pump flow was most likely cannula kink but not clear how the kink had occurred. product damage (cannula kink): the root cause of product damage (cannula kink) cannot be determined since the insufficient clinical details were provided.
6976 | 1220648-2025-30797 | 22641965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22641965&pc=OZD | 07/08/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/08/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654188 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient had a 5.5 pump placed to support with biventricular failure. the patient was put on the transplant list. the pump did support for 25 days until transplant and pump removal. the support did have 1 malfunction of a purge cassette (pc) leak that was remedied by pc change. no purge pressure (pp) or purge flow (pf) issues from the leak. additionally there was a injury of becoming heparin-induced thrombocytopenic during the support. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
6977 | 1220648-2025-30798 | 22641986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22641986&pc=OZD | 07/12/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/14/2025 | 07/29/2025 |  |  | 05/31/2025 |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350062 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male | 85 KG | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during the implant, there was a massive rupture of one of big arteries. vascular surgery advice was taken but the procedure could not be achieved due to massive instability and ongoing resuscitation. additionally, there was an issue with the pump upon start up. the pump did not start in auto mode but directly on a p-level. the console was restarted but still this did not resolve the issue. the catheter also went from p5 to p0 without interaction. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6978 | 1220648-2025-30799 | 22642029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642029&pc=OZD | 07/10/2025 | 09/05/2025 |  | 07/29/2025 |  |  | 07/10/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688164 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, the pump had suction alarms. albumin bolus was given to the patient. additionally, the pump was accidentally pulled back across new aortic valve replacement. the decision was made to replace the pump. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow and positioning issues has been completed. the impella device was not returned for evaluation. low pump flow: complaint device not retuned for investigation. data log reviewed: no motor current (mc) spike outside p-level changes seen. many suction occurred during impella support. on first day on support 7/8 impella low flow at p9 and 8 with placement spikes observed, purge spike seen during support. impella positioning in aorta at 7/14 with correlate with clinical detail "surgeon pulled pump back across new aortic valve replacement accidentally ". from clinical details patient had suction alarms resolved with albumin bolus, another suctions event when patient sat up in chair, decided reduce flows to p7 when patient sit on chair and when in bed can go back to p9. the cause of the pump flow and suction issues most likely due to patient condition related since flow back to range after suction alarms resolved with albumin bolus and p-level changes. positioning issues: the cause of positioning issue most likely use related due to improper technique since the surgeon pulled pump back across new aortic valve replacement accidentally and decided to place a new pump.
6979 | 1220648-2025-30801 | 22642099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642099&pc=OZD | 07/10/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/10/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671820 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that impella cp was placed in the usual manner, post percutaneous coronary intervention, swan was placed in left ventricle, temp pacer placed at the beginning of procedure in the right ventricle. thrombectomy and pci performed on right coronary artery (rca), position verified with echocardiogram (echo) and flow confirmed down the leg with past impella sheath. while attempting to get patient out of the room the patient went into refractory ventricular fibrillation and was shocked 20 times, cardiopulmonary resuscitation performed for 5 min return of spontaneous circulation achieved, regained hemodynamic stability. site was clean and dry free of hematoma or bleeding. during support decision was made to return patient to cath lab for rp flex insertion due to hemodynamic instability. patient continued to cough up blood and clot from endotracheal tube. the patient was heparinized. the impella was successfully weaned and explanted however, it was noted that the site was initially preclosed, suture broke and additional perclose was deployed suture broke, one additional attempt made to do dry closure but wire would not stay positioned past impella, so manual pressure held for 25 minutes site clean dry and free of hematoma or bleeding. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
6980 | 1220648-2025-30800 | 22642101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642101&pc=OZD | 07/18/2025 | 09/16/2025 |  | 07/29/2025 |  |  | 07/18/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644192 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and became hypotensive and developed bleeding from around the sheath and the hemostatic valve. the patient had a pre-existing dissection that was unrelated to the impella. the pump was explanted and the patient survived. | [5 empty] | The investigation of vascular damage - peripheral vessel and vf/vt/af/at has been completed. vascular damage - peripheral vessel: the product was returned for investigation, and blood was observed in the hemostatic valve, but there were no other abnormalities on the sheath or pump. the cause of the peripheral vascular damage was most likely use related, as the vessel dissection occurred when the needle stick was performed, implying the needle stick was likely performed improperly. ¿vf/vt/af/at: the root cause of the hypotension was unable to be determined due to insufficient clinical details. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
6981 | 1220648-2025-30802 | 22642120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642120&pc=OZD | 07/11/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/11/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694702 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that some difficulty was encountered with the preclose delivery system getting trapped in the artery coinciding with impella cp implant. vascular surgery successfully dislodged the preclose with no noted damage to the artery and without surgical intervention. support was able to proceed with the impella cp pump. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following: cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
6982 | 1220648-2025-30803 | 22642237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642237&pc=OZD | 07/18/2025 | 12/03/2025 |  | 07/29/2025 |  |  | 07/18/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic)'s dust cover over the white plug connection site broke. this issue did not have patient involvement. | [5 empty] | The investigation has been completed. data review: console logs did not contain any data relevant to the complaint. device analysis: console was tested after arriving in field service. the reported broken dust cover was confirmed during inspection of the console. conclusion: the root cause of the broken pump connector cover assembly was likely use related. h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation. | The reported issue had no patient involvement. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
6983 | 1220648-2025-30804 | 22642270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642270&pc=OZD | 07/11/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/11/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658623 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the impella 5.5 had frequent position unknown alarms and placement signal was disabled. echocardiogram showed left ventricle measurement at 5cm at 42.5cm catheter position. nurse thought that the team slightly repositioned during echocardiogram. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
6984 | 1220648-2025-30805 | 22642326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642326&pc=OZD | 07/13/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/13/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684346 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient needed care escalation from the previously placed intra-aortic balloon pump. after the impella 5.5 device was placed, the pump's optical signal was inappropriately giving flow/position information. placement signal was disable and the impella 5.5 device remained in use, supporting the patient. now two days later, the patient had a code stroke called due to confusion and facial drooping. heparin drip was stopped and computed tomography was completed. the result of the computer tomography scan was negative. however, it was noted the patient was still lethargic and acting "off." and while the patient was confused, the patient pulled in the swan-ganz catheter into the right ventricle. heparin drip was restarted the next morning. the patient was noted to be in stable condition. however, the following day the patient was decompensating. the patient intubated, sedated, ventilated and placed on venous arterial extracorporeal membrane oxygenation. the next two days the patient was moved to status 1 on the transplant list and the next day the patient was successfully bridged to transplant. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6985 | 1220648-2025-30806 | 22642573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642573&pc=OZD | 07/14/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/14/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555016 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced limb ischemia and subsequently expired. the patient was admitted with increased shortness of breath and palpitations found to be in atrial fibrillation with rapid ventricular response. the patient was started on medication by mouth for rate control. the following day the patient began to have more shortness of breath and required intubation and continued to decompensate requiring multiple vasoactive drips. the patient was taken to catheterization lab for an impella cp placement, which was successfully completed. shock call was made, and the decision to transfer the patient to advanced heart failure management facility. the next day on arrival, pulses are absent bilaterally and arterial ultrasound to assess for flow to lower extremities were ordered. extracorporeal membrane oxygenation (ecmo) was recommended, and the patient's family declined the ecmo therapy. the family opted for comfort measures only, and the patient subsequently expired. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
6986 | 1220648-2025-30807 | 22642708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642708&pc=OZD | 07/14/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/14/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627447 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, thrombus was noted around catheter in descending while doing transesophageal echocardiogram assessment. the physician was made aware and planned to start the patient on systemic anticoagulation. the patient was reported to be in stable condition following the event. the patient continued on impella mcs for another seven days before being successfully weaned and the device explanted with a survived outcome. | [5 empty] | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
6987 | 1220648-2025-30808 | 22642938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22642938&pc=OZD | 07/14/2025 | 09/11/2025 |  | 07/29/2025 |  |  | 07/14/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682518 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed on a patient for persistent hypotension and cardiogenic shock (cgs) secondary to right heart failure. discussion was had about the decreased flows, physician declined to investigate further. peel away sheath removed and repositioning sheath advanced. forward tension sutures applied. during support it was reported impella site clean, dry and intact however other sites were oozing significantly as well as increased drainage from sternum. 2 units fresh frozen plasma given and 2 units of plasma given, oozing slowed down as well as drainage from wound vac. auto turned off, suction alarms continued until on p4 flowing at 1.8l, echo performed shows it measuring at 3.4. left ventricle apex hypertrophic, discussion had about flows being lower than expected and potentially shallowing impella to see if flows improve if suction alarms continue. multiple suction events occurred that corrected with reducing p-level and stopped pulling fluid on continuous renal replacement therapy. they continued to have issues with suction and unable to escalate p-level without suction. impella position was 3.0 cm from the annulus in-flow in the mid left ventricle space. he has had ongoing occult blood loss, we did discuss labs measuring hemolysis in the setting of anuria. his white blood cells remained significantly elevated on broad spectrum antibiotics. his sternal wound had grown candida. shock picture was mixed right sided cardiogenic shock with concurrent septic shock. there had also been a steady decline in platelets. impella dressing changed for oozing. the patient was actively decompensating in the setting of max inotrope and pressor support. the patient continued to demonstrate severe occult blood loss, severe acidosis, severe right ventricle dysfunction as such the patient progressed to a slow chemical code progressing to pulseless electrical activity (pea) and advanced cardiovascular life support algorithms. physician placed impella at p-2 and pulled into the descending aorta prior to pulseless electrical activity, he felt the impella may have contributed to arrhythmia and possibly hemolysis. they did not request to save the device for return for internal review. ultimately resuscitation efforts were seen as futile and the patient was allowed to expire with family at the bedside. the impella was unplugged and turned off by staff. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of low pump flow and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. data log shows multiple suction alarms throughout the case. pump flow was not able to exceed 2 l/min at p-5. placement signal was fluctuating during the case. there were no motor current spikes or positioning issues reported during the case run. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details. the cause of the vascular damage issue was most likely patient condition related since multiple site were bleeding at same time.
6988 | 1220648-2025-30811 | 22643610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22643610&pc=OZD | 07/17/2025 | 12/04/2025 |  | 07/29/2025 |  |  | 07/17/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0042-0040-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 85 KG |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a controller error alarm appeared following a pump stop during impella 5.5 support. the automated impella controller was swapped in response to the alarm. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Analysis: 5.5 pump 578763 was initially connected to ic2221 on (b)(6). the pump was run for 35 days on this console before persistent pump stop alarms started occurring (#13 motor current high) and "impella stopped. controller failure" alarms (#138, pump speed deviation). multiple attempts to restart the pump on this console were unsuccessful. at this point, the pump was moved over to ic5535 and the pump stops continued similar to on ic2221. investigation into pump 578763 found that the electrical lines were open on the pump which would have caused the pump stops, regardless of the console used (see investigation (b)(4) for more detail). root cause: there was no problem found with either console as the pump stops with subsequent controller alarms persisted across both consoles and was found to be caused by the pump's open electrical lines. d10 concomitant medical products and therapy dates updated.
6989 | 1220648-2025-30810 | 22644039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644039&pc=OZD | 07/16/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677407 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp in a younger patient who was admitted with syncope. the patient was in a st-segment elevation myocardial infarction and found to have coronary artery disease. during support, the patient had a coughing episode with subsequent suctioning alarms. the pump could not go above p-3 without suction and 1.1 liters per minute of flow. an echocardiogram at bedside was done and attempted repositioning as the device looked rotated towards the mitral. it was unable to be flipped and the patient remained onpp-2. during echocardiogram, a left ventricle thrombus was discovered. the patient had the impella cp removed and switched to an intra-aortic balloon pump due to left ventricle thrombus and flow issues. the patient was noted in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
6990 | 1220648-2025-30812 | 22644117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644117&pc=OZD | 07/18/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/18/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677016 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient endured atrial fibrillation, worsening thrombocytopenia, and an impella site hematoma during impella 5.5 support. the thrombocytopenia had an unknown/undetermined relationship with the impella device. heparin products were discontinued and the patient was on bivalrudin. the atrial fibrillation was believed to have a possible relationship with the impella device. the patient was monitored for fourteen days to determine the presence of an atrial fibrillation burden. finally, the hematoma at the axillary site was reported to have a possible relationship with the impella device. the hematoma was expanding and the patient was brought to the operating room for an evacuation. blood products were given and a generalized surface bleed from the pectoralis muscle noticably improved. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6991 | 1220648-2025-30813 | 22644175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644175&pc=OZD | 07/14/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/14/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696971 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 118 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and after the implant, the patient went into ventricular fibrillation. the patient was shocked three times and several rounds of chest compressions were preformed along with amiodarone given. upon the fourth shock, the patient was converted to normal sinus rhythm and impella flows were increased to 3.6 liters per minute. additionally, the pump was noted to be shallow. the doctor decided to leave as is and support was continued. the patient survived and is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
6992 | 1220648-2025-30814 | 22644344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644344&pc=OZD | 07/11/2025 | 09/23/2025 |  | 07/29/2025 |  |  | 07/11/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687231 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a thrombus and low pump flow. the pump was removed. | [5 empty] | The investigation of thrombus and low pump flow has been completed. data review: on (b)(6) 2025 at 19:58, a concurrent drop in motor current and pump flow was observed. pump was turned off and on to troubleshoot but was unsuccessful in resolving the low pump flow. no motor current spikes outside of the restart were observed. low flow: pump was returned for analysis. visual inspection found biomaterial on if cage. no other issues were found on the pump. the cause of the low pump flow was most likely biomaterial ingestion per the returned product and clinical report of a clot found at the inlet. thrombus: the cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30814. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
6993 | 1220648-2025-30815 | 22644447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644447&pc=OZD | 07/15/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/18/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574473 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was successfully implanted in a patient however, the sleeve became torn and later, the patient expired. | [5 empty] | A2, a3, a4, and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
6994 | 1220648-2025-30816 | 22644458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644458&pc=OZD | 07/16/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519323 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient lost pulses in bilateral lower extremities during impella cp support. when pulses became absent, the decision was made to remove the impella cp and escalate to impella 5.5 support. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6995 | 1220648-2025-30817 | 22644753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644753&pc=OZD | 07/18/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/18/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640082 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced suction events. the pump was confirmed to be in the correct position and no biomaterial was visualized. the pump was explanted. | [5 empty] | A2, a4, and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
6996 | 1220648-2025-30818 | 22644830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644830&pc=OZD | 07/18/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/18/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691896 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 64 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella 5.5 implant the patient experienced some bleeding from their mouth. the patient was transfused two units of blood while platelet levels began to trend downward. the following day, a ct scan was performed due to acute changes in the patient's neuro status and a intraparenchymal hematoma was discovered. heparin was discontinued as further actions were being discussed. the patient's family was called in response to the hemorrhagic stroke and the decision was made to withdraw care. the patient passed away shortly after. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
6997 | 1220648-2025-30819 | 22644971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22644971&pc=OZD | 07/16/2025 | 08/13/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664180 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 75 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Disability | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device followed by an impella rp flex device for bipella mechanical circulatory support and following bipella implant, the patient began to develop hemolysis. four units of blood were given to treat the condition. the following day, the patient was diagnosed with ischemia in both extremities. lactic remains elevated, contributing to limb ischemia. the impella cp pump was removed, a fasciotomy was performed, and the patient was escalated to impella 5.5 support, which was placed with some difficulty. support ramped up and position was stable. later that same day, the physician believed that the right leg was dead. additional intervention was not performed however. also, it was noted the belly kept getting bigger, a lot of blood seen around the liver on transesophageal echocardiogram. it was noted at that said time, there was no plan for any intervention. due to the patient's prognosis, care was withdrawn and the patient expired. | The complainant reported that following an impella cp implant, the patient began to develop hemolysis. four units of blood were given to treat the condition. the following day, the patient was diagnosed with ischemia in bilateral limbs. the pump was removed. a fasciotomy was performed and the patient was escalated to impella 5.5 support. later that same day, the physician believed that the right leg was dead. additional intervention was not performed. the procedural outcome was the patient survived the surgery. |  |  |  |  | Given the details of the event, there is not enough evidence to exclude the impella cp as an associated factor in the patient's demise outcome. hemolysis along with limb ischemia present that progressed and along with vascular surgical repair related to the impella cp device removal. refer manufacturer report number 1220648-2025-45796 for the report on the impella rp flex. the event was revisited, and the demise outcome determined to be associated with the impella cp. consequently, updates were made to the following section(s): b2: to include disability or permanent damage/death/date of death. b5: to include bipella support and demise outcome. h1: change to death type of reportable event. h6: health effect clinical code(s)/impact code(s)/device problem code(s)/component code(s) repopulated based on event with a demise outcome association. h10: related report number. note: h6 investigation coding as submitted with initial report remains unchanged. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
6998 | 1220648-2025-30820 | 22645086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22645086&pc=OZD | 07/16/2025 | 09/03/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  | 02/28/2026 |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474108 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had bleeding from the access site. two units of packed red blood cells were transfused and the patient was treated with levosimendan. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation has been completed. no product was returned. the root cause of access site bleeding was most likely anticoagulation management since the bleeding stopped after stopping the heparin. | A4 and a5, race, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
6999 | 1220648-2025-30821 | 22645099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22645099&pc=OZD | 11/24/2024 | 07/29/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544098 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient was implanted with an impella cp and experienced access site bleeding. perclose sutures were used and the pump was explanted. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7000 | 1220648-2025-30822 | 22645122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22645122&pc=OZD | 07/16/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664189 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient with a st-segment elevation myocardial infarction admitted to the community and transferred to the tertiary center. the impella cp pump was placed but there was a gastrointestinal (gi) bleed during the support. the gi bleed revealed with bloody stools. blood was given and anticoagulation assessed. the pump support remained on and was explanted after 6 days of use. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
7001 | 1220648-2025-30823 | 22645349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22645349&pc=OZD | 07/16/2025 | 09/12/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689320 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 105 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the doctor did not place a perclose and the patient was highly anticoagulated. bleeding occurred at the access site. heparin was withheld. additionally, there was suction alarms and albumin was given. support continued until weaning which was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of bleeding issue most likely was anticoagulation management due to high activated clotting time (act) and resolved with holding heparin. low pump flow: data log reviewed, no motor current (mc) spike, no placement signal issues observed. quick resolved position alarms in aorta and ventricle which not impact the reported issue, purge spike seen. suctions observed on 7/16 to18 which not impacted pump flows and correlated with clinical details, 7/16 albumin also given for low volume and suction alarms on console. the cause of suction alarms most likely was patient condition related since it resolved after blood and medications managed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7002 | 1220648-2025-30824 | 22645470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22645470&pc=OZD | 07/16/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600621 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 68 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited lower extremities had been dusky since admission, and the patient had been complaining of discomfort. ultrasound and doppler confirmed blood flow, however, blood culture results were positive. ultimately, care was withdrawn and the patient expired. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7003 | 1220648-2025-30825 | 22645602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22645602&pc=OZD | 07/14/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/14/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565608 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient endured tachycardic arrythmia/atrial fibrillation during impella cp support. the patient was converted and received amiodarone to manage the condition. in addition, it was noted that the patient had some trace bleeding from their access site. the site was redressed and support continued uninterrupted. the patients outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7004 | 1220648-2025-30826 | 22645649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22645649&pc=OZD | 07/16/2025 | 09/19/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699967 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support the acute myocardial infarction and cardiogenic shock patient. the pump was placed but there was persistent oozing at the groin. the team adjusted closure with perclose, a c-clamp, and pump repositioning. the pump remained in use despite this but, soon there was hemolysis observed and the pump noted to be alarming for being out of left ventricle positioning. the pump was explanted on day of insertion and patient noted to be stable. | [5 empty] | The investigation of access site bleeding, positioning issues, and hemolysis has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was most likely due to pump was not in the proper position since the clinical team confirmed the pump in improper position caused hemolysis. g2 report source; foreign was missing from manufacturer device report 1220648-2025-30826. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7005 | 1220648-2025-30827 | 22646078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22646078&pc=OZD | 07/16/2025 | 09/09/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574741 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Ventricular Fibrillation | Asystole | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the initial implant attempt of an impella cp pump, the patient went asystolic and coded. the patient received cardiopulmonary resuscitation for over an hour and received several shocks for ventricular fibrillation. it was believed by the physician that a possible aortic aneurysm was preventing advancement of the pump. the patient eventually passed away after return of spontaneous circulation could not be achieved. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation of delivery issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to be patient condition because of possible aortic aneurysm preventing the passage of pump into heart valve. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. b2 outcome, requiring intervention was inadvertently not selected on the initial manufacturer device report number 1220648-2025-30827.
7006 | 1220648-2025-30828 | 22646477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22646477&pc=OZD | 05/30/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661702 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a study patient implanted with an impella 5.5 endured non-sustained ventricular tachycardia. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7007 | 1220648-2025-30829 | 22646482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22646482&pc=OZD | 07/17/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/17/2025 | 07/29/2025 |  |  | 09/30/2025 |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401773 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Material Split, Cut or Torn | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was placed in a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was placed but on the day after implant, hematuria was noted and hemolysis was suspected. this was observed as the patient was transferred from community to tertiary center. continuous renal replacement therapy was began for the hemolysis at the tertiary center. the pump remained on with planned explant and escalation to an impella 5.5. two days later there was a product malfunction with purge pressure/pump flow and the red plug was seen to be leaking blood. the team attempted to troubleshoot and left the pump on for support. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7008 | 1220648-2025-30830 | 22646548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22646548&pc=OZD | 05/30/2025 | 08/29/2025 |  | 07/29/2025 |  |  | 07/17/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1598659 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) had controller errors. no further information was provided. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. the second console (ic9097) confirmed the occurrences of the reported failure mode. the lt log shows two occurrences of a controller error, event #1045 caused by an invalid crc on data packets from the optical pressure-measuring board. during each occurrence, rt logs confirmed that all signals received from the optical bench (snr, contrast, lamp, gain) are represented as -16384 values, and the ps spiked to 3000 due to the crc error. the console ic9097 was not returned for this investigation. the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. d4 serial number, lot number and udi number was erroneously entered on the initial manufacturer device report number and updated with the correct information.
7009 | 1220648-2025-30831 | 22646678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22646678&pc=OZD | 07/17/2025 | 09/09/2025 |  | 07/29/2025 |  |  | 07/17/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026647355 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced low pump flow and pump stop. the automated impella controller was replaced without resolution. the pump was replaced to resolve the issue. the patient's outcome is stable. | [5 empty] | The investigation of pump stop and low pump flow has been completed. the pump was returned for investigation. no physical damage was noted on the returned product. the cannula was inspected, and biomaterial was found on the impeller. the pump was then tested in a bucket of water, showing low pump flow and suction alarm. the biomaterial was able to be removed since it had moved further into the outflow cage. after biomaterial removal, the pump ran as expected in the bucket of water. data log: the log confirms a pump stop with high motor current during the initial ramp-up process of pump start. motor current gradually rose and reached around 938 at the time of pump stop. pump flow was approximately 5l/min. no logs were available on the backup console. clinical details suggest low pump flow and pump stop on both utilized consoles. the cause of the pump stop issue was most likely biomaterial ingestion, based on returned product investigation. the cause of the low pump flow issue was most likely biomaterial ingestion, based on returned product investigation. thrombus failure mode was as investigated failure mode based on returned product investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7010 | 1220648-2025-30832 | 22646835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22646835&pc=OZD | 07/17/2025 | 09/23/2025 |  | 07/29/2025 |  |  | 07/17/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696199 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had an access site bleed. stitches were required to obtain hemostasis. impella support continues. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because bleeding complication was managed by placing an additional suture at the site. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7011 | 1220648-2025-30833 | 22646995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22646995&pc=OZD | 07/17/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/17/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487395 |  | 1000100 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella 5.5 implant, the purge pressure began to climb coinciding with a rising left ventricle signal metric line. an echocardoigram was performed to verify the positioning of the device and it was discovered that the entire left ventricle had been clotted off. the decision was made to end the supported procedure and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7012 | 1220648-2025-30834 | 22647085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22647085&pc=OZD | 07/16/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673893 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and after the implant, the patient had multiple episodes of ventricular tachycardia/ventricular fibrillation with multiple shocks. additionally, the pump was sitting a little deep and the device was repositioned however, the doctor elected to leave a little deep as waveforms and flows are appropriate. the patient continued to deteriorate requiring maxed out epi, levo, and vaso with borderline map. remained very acidotic and note oxygenating well. map dropped into the 40s to 50s with suction at all reduced p levels. threads pulse noted so initiated cardiopulmonary resuscitation. after multiple rounds the family elected to stop life saving measures and the patient passed shortly after. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7013 | 1220648-2025-30835 | 22647599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22647599&pc=OZD | 07/18/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/18/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682461 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 64 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced optical signal loss and "impella in aorta" alarms. the pump was repositioned to resolve the alarm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7014 | 1220648-2025-30836 | 22647679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22647679&pc=OZD | 07/14/2025 | 07/29/2025 |  | 07/29/2025 |  |  | 07/14/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696828 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient experienced some oozing from their impella site during impella cp support. the dressing was changed to manage the condition. in addition, it was documented that the patient developed hematuria prompting a lowering of the support level until the patient's urine cleared. support continued uninterrupted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7015 | 1220648-2025-30837 | 22648859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22648859&pc=OZD | 07/16/2025 | 09/12/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685027 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 54 KG | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp (after placement of antegrade perfusion sheath) femorally to support the critical patient. the patient had a renal history and was on extracorporeal membrane oxygenation prior to impella cp placement. the impella cp was supporting when on the second day there was concerns of hemolysis and low flow that were treated by blood products and continuous renal replacement therapy. however, eventually care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of low pump flow and hemolysis has been completed. the impella device was not returned for evaluation. the pump was not returned for investigation. data log shows the placement signal trend and motor current trend while running on a steady p-level. the flow fluctuated throughout the case with several suction alarms in the middle of the case while running on p1. no motor current spike or positioning issue was reported during the case run. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details. the cause of the hemolysis issue was most likely patient condition related, based on data log review and provided clinical details.
7016 | 1220648-2025-30838 | 22649097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22649097&pc=OZD | 07/16/2025 | 09/09/2025 |  | 07/29/2025 |  |  | 07/16/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690304 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 being placed to support a cardiomyopathy patient. the physician could not pass the impella 5.5 through the patients small axillary artery. the impella 5.5 was aborted and an impella cp was placed in the axillary without issue. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery is determined to be patient condition because 5.5 was unable to pass through vasculature due to small axillary artery.
7017 | 1220648-2025-30839 | 22649285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22649285&pc=OZD | 07/17/2025 | 09/08/2025 |  | 07/29/2025 |  |  | 07/17/2025 | 07/29/2025 |  |  |  |  | 07/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694517 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 61 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp, the pump would not pass an eccentric piece of calcium in the patient's abdominal aorta. the implant was aborted and the patient survived. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. root cause of delivery issue was most likely patient condition since the pump could not be inserted due to eccentric piece of calcium that was in the abdominal aorta. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7018 | 1220648-2025-30738 | 22628668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22628668&pc=OZD | 07/12/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/12/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636787 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a thrombus appeared to be present in the left ventricle during impella 5.5 support. systemic heparin was initiated upon discovery of the thrombus. support continued uninterrupted. weaning was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7019 | 1220648-2025-30743 | 22628857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22628857&pc=OZD | 07/15/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690301 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 74 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had a possible cerebrovascular accident. the staff were unsure what type of stroke it was. care was withdrawn and patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7020 | 1220648-2025-30748 | 22628977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22628977&pc=OZD | 07/04/2025 | 08/26/2025 |  | 07/28/2025 |  |  | 07/04/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689297 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Cardiac Arrest | Thrombocytopenia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The type of reportable event was changed from death to serious injury as per medical safety officer review, heparin-induced thrombocytopenia is a serious injury type of reportable event. the demise outcome is associated to the impella rp flex device captured under manufacturer device report 1220648-2025-30933. | The user facility reported that the patient had a history of chronic kidney disease, diabetes, and severe multi-vessel coronary artery disease. the patient was brought to the hospital and continued to decline. a decision was made for bipella placement. while the impella cp was being inserted, the patient went bradycardic and lost blood pressure (bp). code was called and cardiopulmonary resuscitation was started. the impella cp was quickly advanced. the wire was removed and the pump was started. shortly thereafter, the team was able to get a return of spontaneous circulation. in the intensive care unit, the patient was heparin-induced thrombocytopenia positive. the patient was switched to bivalirudin. ultimately, the family decided to withdraw care and the patient expired. |  |  |  |  | The investigation of thrombocytopenia and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the type of reportable event was changed from death to serious injury as per medical safety officer (mso) review, heparin-induced thrombocytopenia is a serious injury type of reportable event. the demise outcome is associated to the impella rp flex device captured under manufacturer device report 1220648-2025-30933. the following fields were updated based on the mso review of manufacturer device report. b2. death and date of death was erroneously entered. h1. type of reportable event was erroneously selected. h6. health effect - impact code 1802 was erroneously entered. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
7021 | 1220648-2025-30749 | 22629089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22629089&pc=OZD | 07/11/2025 | 09/24/2025 |  | 07/28/2025 |  |  | 07/11/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673576 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | Component Missing | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and after implant the pump did not start. consequently, the impella cp device was explanted and successfully replaced. it was further reported the 14f peel-away sheath was not in the impella cp package as expected, and consequently, the team secured and used a long sheath from a stand-alone package. there was no report or indication of harm to the patient. | [5 empty] | Some patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of pump did not start has been completed. the impella device was not returned for evaluation. the pump was not returned. the cause of the pump unable to start was not determined since product was not returned, data logs showing the pump attempting to be started were not recovered, and insufficient clinical details. b5 erroneously included a packaging issue in the narrative on the initial manufacturer device report number 1220648-2025-30749. g1 reporting contact first and last name was erroneously entered on the initial manufacturer device report number 1220648-2025-30749. h6 medical device problem code 2306 and component code 4747 was erroneously entered on the initial manufacturer device report number 1220648-2025-30749.
7022 | 1220648-2025-30750 | 22629241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22629241&pc=OZD | 07/11/2025 | 09/26/2025 |  | 07/28/2025 |  |  | 07/11/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672256 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day there was bleeding from access site. due to the high-volume administration, high anticoagulation and the bleeding in the groin, a total of four units of blood were administered. after optimizing the pump's angle of entry and the anticoagulation, the bleeding was stopped. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. access site bleeding - major: the product was returned for investigation and no kinks or abnormalities were observed. the root cause of the access site bleeding issue was most likely use related since optimizing angle of entry and optimizing anticoagulation achieved hemostasis.
7023 | 1220648-2025-30752 | 22629460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22629460&pc=OZD | 07/12/2025 | 09/08/2025 |  | 07/28/2025 |  |  | 07/12/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp was bleeding from the extracorporeal membrane oxygenation site and heparin was held. it was further reported that impella cp was pulled back while in operating room and it was unable to be repositioned. the cp was exchanged. | [5 empty] | The investigation of positioning issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issue was most likely use related as the pump was pulled out and unable to be repositioned per clinical. vascular damage peripheral: the cause of the vascular damage peripheral vessel was not determined due to lack of clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings and cautions: cautions: "physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta." caution: "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance."
7024 | 1220648-2025-30753 | 22629600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22629600&pc=OZD | 07/14/2025 | 11/30/2025 |  | 07/28/2025 |  |  | 07/14/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116498 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was implanted with an impella 5.5 device for mechanical circulatory support (mcs) as elective placement. the patient was deferred for orthotopic heart transplant/durable left ventricular assist device candidacy and given options for revascularization. approximately four days into the support, the automated impella controller (aic) had an unexpected controller shutdown. aic automatically restarted, resumed to previous settings, and impella mcs continued. there was no harm to the patient as a result of this event. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed. | Data analysis: the console logs confirm the occurrence of an unexpected reboot on the date of the complaint. there are no alarms or watchdog triggers immediately preceding the event to provide insight into the cause. there is an rtlogger overflow but this occurs ~31 minutes before the unexpected reboot. there are multiple instances of #1036 calculator placement signal errors but no accompanying alarms. device analysis: the console was opened and a visual inspection revealed no clues as to the cause of the unexpected reboot. a test pump was connected and ran at p-level 4 for ~24 hours with no occurrence of spontaneous reboots and no occurrence of noteworthy alarms. 5s parameters were obtained and returned a low snr value of ~500. the optical connector was cleaned but resulted in only marginal improvement of the snr value. the optical bench ccd output was evaluated and found to be notably distorted. root cause: the root cause of the spontaneous, unexpected reboot is unknown due to a lack of correlating entries in the data logs and an inability to reproduce the issue in the lab. h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation.
7025 | 1220648-2025-30754 | 22629657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22629657&pc=OZD | 07/13/2025 | 09/05/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626230 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and two days into the support, the patient had oozing at the access site. one unit of blood products were administered, and heparin was withheld. the patient was noted to be in stable condition following the event; impella mcs continued. the next day it was noted the patient remained sedated/vented in bed; urinary output was increasing with good color. a fixed dose of heparin drip was infusing, and access site oozing had resolved. subsequently, the device was successfully weaned and explanted with a survived outcome. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be anticoagulation management because the bleeding was resolved by fixing the dosage of heparin. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer - ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7026 | 1220648-2025-30755 | 22629866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22629866&pc=OZD | 07/13/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667382 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 85 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient has rhabdomyolysis and was on impella cp support. during support the patient was noted to have some hemolysis as there is some blood in the urine and oozing at the insertion site. care team suspects a combination of rhabdomyolysis and impella cp. continuous renal replacement therapy was started that day and the care team escalated the patient to impella 5.5 the following day. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7027 | 1220648-2025-30756 | 22630418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22630418&pc=OZD | 07/13/2025 | 09/08/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661953 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp to support the patient admitted in with arrhythmia, cardiogenetic, and cardiomyopathy. the pump was placed however there was access site bleeding which prompted blood products to be given and when the pump alarmed for out of position, the pump was repositioned. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of bleeding issue cannot be determined due to insufficient clinical details. positioning issues: the cause of positioning issue most likely was due patient movement occurred due preparing patient for transport and physician repositioned impella from 90cm to 87cm using repositioning guide.
7028 | 1220648-2025-30758 | 22630640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22630640&pc=OZD | 07/15/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683070 |  | 0042-0010-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after connecting the impella cp pump to an automated impella controller (aic), an error message appeared indicating that there was no placement signal. the pump was removed and replaced in response to the failure. there was no noted patient injury. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7029 | 1220648-2025-30759 | 22630840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22630840&pc=OZD | 09/27/2024 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539706 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Thrombocytopenia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Abiomed¿s long-term outcome and quality indicator (loqi) impella registry reported that a patient was implanted with impella 5.5 for high-risk surgery on (b)(6) 2024. it was reported that the "patient endured thrombocytopenia with a "possible" relationship to the impella device used: ¿patient developed worsening blood loss anemia after impella implant. nadir hb: 7.2 on (b)(6) post-impella implant. transfused with a total of 2 packed red blood cells (prbc), and 4 platelets during impella support from on (b)(6). plavix, lovenox on hold. resume as able." additionally, the patient developed worsening of acute kidney injury (aki) with a "possible" relationship to the impella device used: ¿aki requiring continuous renal replacement therapy (crrt). creatinine trended upwards prior to impella implant. initially, crrt recommended but family refused dialysis. the support was for just over 3 days when the patient had care withdrawn and patient expired. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7030 | 1220648-2025-30760 | 22630843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22630843&pc=OZD | 07/16/2025 | 12/10/2025 |  | 07/28/2025 |  |  | 07/16/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657953 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 68 KG |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | Followup | [7 empty] | Additional information was received stating that the pump was explanted and the patient was coagulopathic and too unstable to go into surgery for impella 5.5. the patient was having bleeding issues prior to the cp implant due to complications of cardiogenic shock. the patient expired 2-3 days later from a subdural hemorrhage. it is unknown if the impella cp caused or contributed to the subdural hemorrhage. | The complainant reported that after an automated impella controller (aic) to aic transfer, an implanted impella cp pump failed to resume support. despite attempting on four different aics, the consoles would not recognize the impella cp device. the pump was subsequently removed and plans were made to escalate to impella 5.5 support, but the patient is currently too unstable for surgery. the patient was reported to be in critical condition. |  |  |  |  | Clinical details state that during console transfer upon arrival to a second site, the pump was not detected. another console, the original console, and two more consoles were attempted and none resolved the issue. the pump was explanted. lt log review showed no alarms or abnormalities related to pump detection. imc log review showed the pump successfully connecting to (b)(6) at initial plugin. on (b)(6) 2025, the pump was unplugged and the logs ended shortly after with no more connection attempts seen on (b)(6). the logs for the other consoles were not provided. the pump was returned and evaluated. there was a large separation/gap observed on the patient cable plug and fragments from the broken ferrule that could be seen through the pin-side of the patient cable plug. when electrical testing the communication lines, there were open lines found at the connections with pin 1, 5, and 11. there were no abnormalities with the electrical connection to pin 6. ct/xrays of the patient cable plug showed visual narrowing and/or clear separation at the connections to pin 1, 5, and 11. the cause of the pump detection issues was open electrical lines in the patient cable plug based on returned product analysis. the cause of the electrical opens is unknown due to a lack of clinical details. the cause of the injury was not determined due to insufficient clinical details. | H6 health impact clinical code and impact code updated with additional information. a4 patient weight was inadvertently omitted on the initial manufacturer device report number. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7031 | 1220648-2025-30762 | 22630897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22630897&pc=OZD | 07/14/2025 | 10/07/2025 |  | 07/28/2025 |  |  | 07/14/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693897 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a placement signal not reliable alarm appeared. as the flows remained within the expected range, the pump was able to continue to support the patient throughout the scheduled primary percutaneous coronary intervention until planned wean and explant. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue has been completed. the impella device was returned for evaluation. the cause of the optical signal issue related to the pump was not determined since the issue did not reproduce with the returned pump.
7032 | 1220648-2025-30763 | 22630950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22630950&pc=OZD | 07/16/2025 | 12/03/2025 |  | 07/28/2025 |  |  | 07/16/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302879 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | False Alarm | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preventative maintenance, the speaker of an automated impella controller (aic) began making elecrical/static sounds. the aic was repaired using a new power battery manager (pbm), however a controller error alarm then appeared. there was no patient involvement. | [5 empty] | The investigation has been completed. data analysis: the console logs show a controller error "pbm voltage out-of-spec #501" which was triggered immediately after the "impella stopped" alarm (#18) on the reported date. this sequence matches the reported complaint. during pm section 13 (short circuit simulation), the procedure intentionally triggers the "impella stopped alarm", which sometimes is followed by a controller error alarm. these alarm also are used to verify audio alarm functionality. this behavior is expected during pm procedure. there were no logs associated static noise coming from the speaker which was likely resolved via the pbm replacement. device analysis: console (b)(4) is a service loaner that had just undergone its annual preventive maintenance (pm) and was tested on an engineering lab bench. the static noise coming from the speaker was not reproduced given it was likely resolved via the pbm replacement. the controller error alarm is expected during section 13 of the pm procedure. there is no problem found with this device with respect to that issue. no issues were observed with the aic related to the reported false buzzer alarm. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7033 | 1220648-2025-30765 | 22630994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22630994&pc=OZD | 07/16/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/16/2025 | 07/28/2025 |  |  | 02/28/2026 |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025467727 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a patient arrest outside in community and given cardiopulmonary resuscitation (cpr). the patient was defibrillated and intubated and return of spontaneous circulation was achieved. the patient was taken the cardiac catheterization lab and percutaneous coronary intervention (pci) was performed and an intra-aortic balloon pump (iabp) was placed. unable to evaluate neuro status. pupils pinpoint and fixed. the patient continued to decompensate and the decision was made to escalate the iabp to impella cp. the patient was transferred to the tertiary center on support. while on support the patient suffered a gastrointestinal bleed and had blood products infused. the family made choice to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. b2: date of death unknown. d6b: explant date unknown.
7034 | 1220648-2025-30764 | 22631064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631064&pc=OZD | 06/01/2024 | 07/28/2025 |  | 07/28/2025 |  |  | 07/16/2025 | 07/28/2025 |  |  | 02/28/2026 |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025474136 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received an impact study stating the patient endured a stroke after explant. there was a possible relationship between the impella device and the impella procedure and the stroke. the study stated, "readmit to icu for slurred speech. imaging ct head shows: subacute left cerebellar changes; partially occlusive thrombus within the basilar tip extending into the proximal left sca." the patient outcome is unknown. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7035 | 1220648-2025-30767 | 22631189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631189&pc=OZD | 07/10/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/10/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680920 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 116 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that shortly following removal of the 14 french sheath, the access site began to pulsate blood. in addition, the patient's leg was cool and doppler pulses could not be found. after twenty minutes of manual pressure, femostop was placed on the site and pulses returned. the patient survived and was in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7036 | 1220648-2025-30766 | 22631303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631303&pc=OZD | 06/30/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/16/2025 | 07/28/2025 |  |  | 06/30/2025 |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370649 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant was placing a cp for the support of the patient transferred from community to tertiary center for care of the st-elevation myocardial infarction. the 1st cp failed to be recognized by the console. the pump was swapped out and 2nd pump was detected and support initiated. there patient was reported as stable and no harm was reported. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed.
7037 | 1220648-2025-30768 | 22631397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631397&pc=OZD | 07/13/2025 | 09/09/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682107 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 94 KG | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Increase in Pressure | Detachment of Device or Device Component |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had purge system blocked/purge pressure high alarms. no kinks or blockages were observed. the purge cassette changed with no improvement. tissue plasminogen activator protocol was done. additionally, the care team repositioned impella cp for positional alarms, and tuohy-borst was broke. the catheter was secured by using a stat lock at 97cm from 90cm. it was further reported that the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of high purge pressure, positioning issues, and product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. data log shows a 3-minute rise in purge pressure followed by a high purge pressure alarm. the purge pressure resolved with a gradual decreasing trend after 4 hours of case run. no motor current spike was reported during the time of purge pressure rise. the cause of the high purge pressure issue was most likely biomaterial buildup, based on data log review and given clinical details. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the product damage (sleeve damage) issue was not determined since the product was not returned and sufficient clinical information, including clinical images, was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7038 | 1220648-2025-30769 | 22631403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631403&pc=OZD | 07/15/2025 | 09/10/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698101 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | Other | [22 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a moderate amount of bleeding was seen from the right groin / impella sheath. in addition, the patient developed hemolysis. femostop was applied and an additional suture was added to achieve hemostasis. blood products were also given to treat the bleed and hemolysis. the patient's family made the decision to forgo escalation of support and the patient was made comfort care. the patient passed away following withdrawal of care. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: access bleed: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation has been completed. pump was not returned for investigation. the data log was not provided or could not be retrieved from the remote server. the cause of the access site bleeding was most likely attributed to high patient act values reported after transfer to the icu. the patient was on antiplatelet therapy at time of bleeding. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided including data log.
7039 | 1220648-2025-30770 | 22631654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631654&pc=OZD | 07/10/2025 | 09/08/2025 |  | 07/28/2025 |  |  | 07/10/2025 | 07/28/2025 |  |  | 09/30/2025 |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446031 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after an impella cp was placed, oozing and hematoma was noted to the left groin. pink/red urine was also noted. in the ntensive care unit, the patient was still oozing a significant amount. the patient had received three blood thinners. two units for blood products were administered. it was further reported that on day two of impella support, the impella cp was pulled back into the aorta and unable to cross the valve again. the impella cp was removed. | [5 empty] | The investigation of access site bleeding, hemolysis, and positioning issues has been completed. the pump was not returned for investigation. data log shows a trending placement signal with motor current fluctuations while on a steady p-level. no motor current spikes or positioning issues were recorded during the case run. access site bleeding: the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. hemolysis: the cause of the hemolysis issue was most likely patient condition related, based on data log review and provided clinical details. positioning issues: the cause of the positioning issue was most likely use related since it was pulled out in aorta during use. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7040 | 1220648-2025-30771 | 22631692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631692&pc=OZD | 07/15/2025 | 09/23/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668830 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient had an attempted impella cp implant and during the implant, the impella wire was already in the left ventricle, and the impella had just been threaded onto the wire. the pump was already in the vascular system, but it had not yet been placed in the left ventricle. chest compressions caused the wire to slip back into the aorta, and the impella could not be positioned in the left ventricle. the implant was aborted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. chest compressions was causing the wire to slip back into the aorta and impella cannot be positioned in left ventricle and insertion was aborted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7041 | 1220648-2025-30772 | 22631804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631804&pc=OZD | 07/12/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/12/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676228 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 88 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after one day of support, the patient developed hemolysis. the staff was aware of the situation and planned to monitor the condition via fluid status, lab draws, and positioning checks. support continued for one more day, at which point the family chose to place the patient on comfort care. the patient subsequently expired following that decision. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller. displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7042 | 1220648-2025-30773 | 22631901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631901&pc=OZD | 07/15/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689522 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a cp pump placed for the stemi patient. prior medical history and underlying conditions not known for the patient due to language barrier. the pump was supporting without heparin in the purge and there was a persistent ooze at the groin site. this was treated by infusion of blood products. the pump remains on despite the harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7043 | 1220648-2025-30776 | 22631977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631977&pc=OZD | 07/13/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692947 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and approximately eight days after start of the support, the patient experienced sustained ventricular tachycardia and hypotension. amiodarone bolus was administered with impella 5.5 support level increased back to p-8. some suspicion of afterload agent contribution was noted. a point-of-care ultrasound was completed to spot check the impella 5.5 position and optimal placement was noted. the patient continued on support for two additional days before being successfully weaned for recovery, and the device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7044 | 1220648-2025-30774 | 22631997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22631997&pc=OZD | 07/12/2025 | 09/09/2025 |  | 07/28/2025 |  |  | 07/12/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676994 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 81 KG |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient presented with congestive heart failure exacerbation. given poor right ventricle function, the patient was not a ventricular assist device candidate. a decision for impella 5.5 placement was made. the impella 5.5 passed through the graft and into the artery. the pump was positioned under transesophageal echocardiogram. during support, there were suction alarms twice that resolved on their own. it was also reported that the patient had a gastrointestinal bleed. the patient had sustained suction alarm while laying on their side. the impella 5.5 was successfully weaned and explanted as a bridge to transplant. the patient survived. | [5 empty] | The investigation of vascular damage - peripheral vessel and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: the cause of the low pump flow was patient condition related due to repositioning efforts not resolving the issue and patient volume a noted concern by the clinical team. vascular damage - peripheral vessel: the cause of the vascular damage - peripheral vessel was not determined since product was not returned and insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7045 | 1220648-2025-30775 | 22632019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22632019&pc=OZD | 07/15/2025 | 09/10/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683274 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had oozing around repo sheath even with forward tension sutures. the doctor decided to place figure of 8 as well as administer lido with epi. oozing stopped. additionally, the patient had red urine with hemolyzed labs. the doctor declined to reposition due to the position looking acceptable on echo. furthermore, staff was unable to doppler pulses on either leg. duplex was performed. support continued. | [5 empty] | The investigation was completed. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. the alarms then were resolved & no other issues were found on the logs to the rest of the case. product was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was most likely due to pump was not in the proper position since hemolysis was managed after repositioning of the pump. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: acces site bleeding: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ limb ischemia: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7046 | 1220648-2025-30777 | 22632062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22632062&pc=OZD | 07/13/2025 | 09/19/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1373815 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller (aic) had an alarm indicating a controller error. the aorta placement signal and left ventricle placement signal were not displayed. the alarm was muted and the left ventricle metrics appeared. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) and optical signal issue has been completed. based on the device analysis and data log the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc.
7047 | 1220648-2025-30778 | 22632165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22632165&pc=OZD | 07/14/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/14/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611278 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 70 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support the patient with planned high-risk percutaneous coronary intervention. the pump was explanted after 24 hours and the closure was complicated. perclose x2 deployed in right femoral artery (rfa) without good hemostasis- access obtained on left femoral artery and destination sheath placed through 7fr sheath. balloon inflated via up and over technique and then covered stent deployed after balloon inflations. it was discussed with vascular the need for femoral arterial cut down if unable to deploy covered stent. successful covered stent deployment. manual pressure applied to rfa site with good hemostasis noted and no oozing/bleeding. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings and cautions: cautions: "physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta." caution: "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance."
7048 | 1220648-2025-30780 | 22632486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22632486&pc=OZD | 07/13/2025 | 09/09/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589942 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the insertion of an impella cp device was unsuccessful. "extreme" difficulty crossing the aortic valve was encountered due to lack of cardiac activity. "numerous" of troubleshooting scenarios were tried and were unsuccessful. cardiac massage was performed. the guidewire crossed. multiple catheters were unable to advance. an "extensive" clot was subsequently noted in aorta and left atrium. the procedure was aborted. the chest was packed and the patient was returned to the unit. the patient was noted to have survived the event. however, the outcome following the event was unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and thrombus has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue was most likely patient condition since the pump could not be inserted due to extensive clot noted in aorta with extreme difficulty crossing aortic valve due to lack of cardiac activity. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7049 | 1220648-2025-30779 | 22632750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22632750&pc=OZD | 07/15/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had an attempted impella implant and during the implant, the 14fr peel away sheath was fully occluded in the vessel. the sheath was explanted and implant was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7050 | 1220648-2025-30781 | 22632830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22632830&pc=OZD | 07/15/2025 | 09/08/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674855 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella cp implant and during the implant, the patient's right femoral artery was accessed and there were issues with the angle of the stick which led to multiple issues with prolapsing the wire and complications advancing catheters. perclose failed and the pump was explanted and pressure was held. the pump was implanted using the left femoral artery. while prepping the patient for transport in unit, a right femoral hematoma was noted. manual pressure was held, four units of blood products were given and 3 units of albumin. support continued. | [5 empty] | The investigation of delivery issues and access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding : the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. delivery issue : the root cause of the delivery issue was most likely use related since there were issues with the angle of the stick which led to multiple issues with prolapsing the wire and complications advancing catheters. the pump was implanted using the left femoral artery successfully. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: access site bleeding: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ delivery issues: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7051 | 1220648-2025-30782 | 22632902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22632902&pc=OZD | 07/13/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682507 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 101 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and two days after the start of support, the patient would undergo coronary artery bypass graft surgery x 3. following it was noted tat the pump was shallow and pointed towards the papillary. additionally, pump flow was unable to flow above p-4. the pump was explanted thereafter with a survived outcome. the patient's status following the event was unknown. however, there was no report or indication of any adverse events. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.